## Sheet1
Microwave Heating of Nanotubes Project
5/12/2015 P.R. Cluxton
Calibration Check of IR Sensor
Using Silicone Oil Bath with Black Painted Metal Can Target
Using the following instruments:
Fluke model 54 - 2B, S/N 25940064WS with Type T thermocouple
ThermoProbe model TL1-R, S/N 1-12730
Omega IR Sensor model OS-MINI 802-D-MA, S/N OM003360, Emissivity set to default 0.95
Computer with National Instruments Interface and Mwave program
Time of Day | Fluke 54 | ThermoProbe | Omega IR | Computer
hh:mm:ss | degrees C | degrees C | degrees C | degrees C
12:43 | 22.7 | 22.59 | 24 | 24.08
12:56 | 36.3 | 36.11 | 38.3 | 38.47
1:10 | 48 | 47.88 | 51.2 | 51.34
1:22 | 58.9 | 58.86 | 62.8 | 62.92
2:07 | 69.2 | 69.4 | 74.9 | 75
2:20 | 89.4 | 89.22 | 96.3 | 96.29
2:32 | 108.7 | 108.68 | 117.3 | 117.24
2:43 | 126.4 | 126.36 | 137.3 | 137.33
2:52 | 141.9 | 141.67 | 153.7 | 153.85
3:03 | 149 | 148.78 | 162.7 | 162.87
OmegaIR: | used to measure the temperature of the sample
ThermoProbe: | used to monitor the temperatur of the waveguide as reference

## Sheet2
Microwave Heating of Nanotubes Project
5/12/2015 P.R. Cluxton
Calibration Check of IR Sensor
Using Silicone Oil Bath with Teflon Cell and Cover Glass with Black Painted Metal Can Target
Using the following instruments:
Fluke model 54 - 2B, S/N 25940064WS with Type T thermocouple
ThermoProbe model TL1-R, S/N 1-12730
Omega IR Sensor model OS-MINI 802-D-MA, S/N OM003360, Emissivity set to default 0.95
Computer with National Instruments Interface and Mwave program
Time of Day | Fluke 54 | ThermoProbe | Omega IR | Computer
hh:mm:ss | degrees C | degrees C | degrees C | degrees C
5:27 | 53.5 | 53.4 | 56.4 | 56.3
5:38 | 72 | 71.85 | 68.3 | 68.55
5:53 | 89.6 | 89.58 | 90 | 90
6:07 | 112.8 | 112.7 | 111.4 | 111.63
6:20 | 132.3 | 132.3 | 132.2 | 132.57
6:32 | 150 | 150.03 | 148.4 | 148.4
5/13/2015
7:15 | 23.3 | 23.17 | 24.3 | 24.43
OmegaIR: | used to measure the temperature of the sample
ThermoProbe: | used to monitor the temperatur of the waveguide as reference