## Input Output
1) Paste emission spectrum from Ocean Optics spectrometer in column I | [6 empty] | Absorbance | Wavelength | Absorbance
2) Modify column D in the Germicidal Calculator tab to calculate radiometer | [5 empty] | 400 |  | 200 | 1.164
sensitivity relative to the weighted average wavelength |  |  |  |  | Peak | 399 |  | 201 | 0.979
3) Paste water absorbance spectrum (200 - 300 nm) in column L |  |  |  |  | 269.7393 | 398 |  | 202 | 0.813
4) Enter radiometer factor for weighted average wavelength into radiometer |  |  |  |  | Rad. Factor: | 397 |  | 203 | 0.684
5) Enter irradiance reading for center of petri dish (mW/cm2) |  |  |  |  | 0.0013121 | 396 |  | 204 | 0.574
6) Enter petri factor | [5 empty] | 395 |  | 205 | 0.484
7) Enter reflection factor for weighted average wavelength | [5 empty] | 394 |  | 206 | 0.407
8) Enter sample depth and distance from sample to beam below. | [5 empty] | 393 |  | 207 | 0.341
9) Enter desired UV doses | [5 empty] | 392 |  | 208 | 0.297
10) Outputs are in blue | [5 empty] | 391 |  | 209 | 0.251
[6 empty] | 390 |  | 210 | 0.222
MP UV Dose Calculations | [5 empty] | 389 |  | 211 | 0.202
Input: | Irradiance readingat 270 nm = |  | 0.473 | mW/cm2 |  | 388 |  | 212 | 0.178
| Petri factor = |  | 0.89111 |  |  | 387 |  | 213 | 0.165
| Sample depth in the petri dish = |  | 0.6 | cm |  | 386 |  | 214 | 0.152
|  | Distance from sample to beam = | 4 | cm |  | 385 |  | 215 | 0.139
| Reflection Factor = |  | 0.97721 | at 270 nm |  | 384 |  | 216 | 0.131
Output: | Divergence Factor = |  | 0.86957 |  |  | 383 |  | 217 | 0.125
| UVT (254) = |  | 92.2571 |  |  | 382 |  | 218 | 0.119
| Incident irradiance = |  | 0.34412 | mW/cm2 |  | 381 |  | 219 | 0.113
| Average irradiance = |  | 0.33744 | mW/cm2 |  | 380 |  | 220 | 0.108
|  | Germicidal irradiance = | 0.32092 | mW/cm2 |  | 379 |  | 221 | 0.104
|  | Action spectra irradiance = | 0.30835 | mW/cm2 |  | 378 |  | 222 | 0.1
[6 empty] | 377 |  | 223 | 0.097
[6 empty] | 376 |  | 224 | 0.092
| UV Dose Required |  |  |  |  | 375 |  | 225 | 0.089
|  | Exposure Time |  |  |  | 374 |  | 226 | 0.085
| (mJ/cm2) | (sec) | min | sec |  | 373 |  | 227 | 0.081
| 5 | 14.8174 | 0 | 15 |  | 372 |  | 228 | 0.077
| 10 | 29.6347 | 0 | 30 |  | 371 |  | 229 | 0.074
| 15 | 44.4521 | 0 | 45 |  | 370 |  | 230 | 0.07
| 20 | 59.2695 | 0 | 60 |  | 369 |  | 231 | 0.067
| 30 | 88.9042 | 1 | 29 |  | 368 |  | 232 | 0.064
[6 empty] | 367 |  | 233 | 0.06
[6 empty] | 366 |  | 234 | 0.057
[6 empty] | 365 |  | 235 | 0.055
[6 empty] | 364 |  | 236 | 0.052
[6 empty] | 363 |  | 237 | 0.05
[6 empty] | 362 |  | 238 | 0.049
[6 empty] | 361 |  | 239 | 0.046
[6 empty] | 360 |  | 240 | 0.046
[6 empty] | 359 |  | 241 | 0.044
[6 empty] | 358 |  | 242 | 0.042
[6 empty] | 357 |  | 243 | 0.042
[6 empty] | 356 |  | 244 | 0.041
[6 empty] | 355 |  | 245 | 0.04
[6 empty] | 354 |  | 246 | 0.04
[6 empty] | 353 |  | 247 | 0.039
[6 empty] | 352 |  | 248 | 0.039
[6 empty] | 351 |  | 249 | 0.038
[6 empty] | 350 |  | 250 | 0.037
[6 empty] | 349 |  | 251 | 0.036
[6 empty] | 348 |  | 252 | 0.036
[6 empty] | 347 |  | 253 | 0.035
[6 empty] | 346 |  | 254 | 0.035
[6 empty] | 345 |  | 255 | 0.034
[6 empty] | 344 |  | 256 | 0.034
[6 empty] | 343 |  | 257 | 0.032
[6 empty] | 342 |  | 258 | 0.033
[6 empty] | 341 |  | 259 | 0.032
[6 empty] | 340 |  | 260 | 0.032
[6 empty] | 339 |  | 261 | 0.033
[6 empty] | 338 |  | 262 | 0.03
[6 empty] | 337 |  | 263 | 0.033
[6 empty] | 336 |  | 264 | 0.031
[6 empty] | 335 |  | 265 | 0.03
[6 empty] | 334 |  | 266 | 0.031
[6 empty] | 333 |  | 267 | 0.031
[6 empty] | 332 |  | 268 | 0.03
[6 empty] | 331 |  | 269 | 0.03
[6 empty] | 330 |  | 270 | 0.029
[6 empty] | 329 |  | 271 | 0.028
[6 empty] | 328 |  | 272 | 0.029
[6 empty] | 327 |  | 273 | 0.029
[6 empty] | 326 |  | 274 | 0.027
[6 empty] | 325 |  | 275 | 0.027
[6 empty] | 324 |  | 276 | 0.026
[6 empty] | 323 |  | 277 | 0.026
[6 empty] | 322 |  | 278 | 0.025
[6 empty] | 321 |  | 279 | 0.025
[6 empty] | 320 |  | 280 | 0.024
[6 empty] | 319 |  | 281 | 0.023
[6 empty] | 318 |  | 282 | 0.023
[6 empty] | 317 |  | 283 | 0.022
[6 empty] | 316 |  | 284 | 0.022
[6 empty] | 315 |  | 285 | 0.02
[6 empty] | 314 |  | 286 | 0.019
[6 empty] | 313 |  | 287 | 0.017
[6 empty] | 312 |  | 288 | 0.017
[6 empty] | 311 |  | 289 | 0.016
[6 empty] | 310 |  | 290 | 0.016
[6 empty] | 309 |  | 291 | 0.014
[6 empty] | 308 |  | 292 | 0.014
[6 empty] | 307 |  | 293 | 0.013
[6 empty] | 306 |  | 294 | 0.012
[6 empty] | 305 |  | 295 | 0.011
[6 empty] | 304 |  | 296 | 0.011
[6 empty] | 303 |  | 297 | 0.011
[6 empty] | 302 |  | 298 | 0.011
[6 empty] | 301 |  | 299 | 0.01
[6 empty] | 300 |  | 300 | 0.01
[6 empty] | 299 |  | 301
[6 empty] | 298 |  | 302
[6 empty] | 297 |  | 303
[6 empty] | 296 |  | 304
[6 empty] | 295 |  | 305
[6 empty] | 294 |  | 306
[6 empty] | 293 |  | 307
[6 empty] | 292 |  | 308
[6 empty] | 291 |  | 309
[6 empty] | 290 |  | 310
[6 empty] | 289 |  | 311
[6 empty] | 288 |  | 312
[6 empty] | 287 |  | 313
[6 empty] | 286 |  | 314
[6 empty] | 285 |  | 315
[6 empty] | 284 |  | 316
[6 empty] | 283 |  | 317
[6 empty] | 282 |  | 318
[6 empty] | 281 |  | 319
[6 empty] | 280 |  | 320
[6 empty] | 279 |  | 321
[6 empty] | 278 |  | 322
[6 empty] | 277 |  | 323
[6 empty] | 276 |  | 324
[6 empty] | 275 |  | 325
[6 empty] | 274 |  | 326
[6 empty] | 273 |  | 327
[6 empty] | 272 |  | 328
[6 empty] | 271 |  | 329
[6 empty] | 270 |  | 330
[6 empty] | 269 |  | 331
[6 empty] | 268 |  | 332
[6 empty] | 267 |  | 333
[6 empty] | 266 |  | 334
[6 empty] | 265 |  | 335
[6 empty] | 264 |  | 336
[6 empty] | 263 |  | 337
[6 empty] | 262 |  | 338
[6 empty] | 261 |  | 339
[6 empty] | 260 |  | 340
[6 empty] | 259 |  | 341
[6 empty] | 258 |  | 342
[6 empty] | 257 |  | 343
[6 empty] | 256 |  | 344
[6 empty] | 255 |  | 345
[6 empty] | 254 |  | 346
[6 empty] | 253 |  | 347
[6 empty] | 252 |  | 348
[6 empty] | 251 |  | 349
[6 empty] | 250 |  | 350
[6 empty] | 249 |  | 351
[6 empty] | 248 |  | 352
[6 empty] | 247 |  | 353
[6 empty] | 246 |  | 354
[6 empty] | 245 |  | 355
[6 empty] | 244 |  | 356
[6 empty] | 243 |  | 357
[6 empty] | 242 |  | 358
[6 empty] | 241 |  | 359
[6 empty] | 240 |  | 360
[6 empty] | 239 |  | 361
[6 empty] | 238 |  | 362
[6 empty] | 237 |  | 363
[6 empty] | 236 |  | 364
[6 empty] | 235 |  | 365
[6 empty] | 234 |  | 366
[6 empty] | 233 |  | 367
[6 empty] | 232 |  | 368
[6 empty] | 231 |  | 369
[6 empty] | 230 |  | 370
[6 empty] | 229 |  | 371
[6 empty] | 228 |  | 372
[6 empty] | 227 |  | 373
[6 empty] | 226 |  | 374
[6 empty] | 225 |  | 375
[6 empty] | 224 |  | 376
[6 empty] | 223 |  | 377
[6 empty] | 222 |  | 378
[6 empty] | 221 |  | 379
[6 empty] | 220 |  | 380
[6 empty] | 219 |  | 381
[6 empty] | 218 |  | 382
[6 empty] | 217 |  | 383
[6 empty] | 216 |  | 384
[6 empty] | 215 |  | 385
[6 empty] | 214 |  | 386
[6 empty] | 213 |  | 387
[6 empty] | 212 |  | 388
[6 empty] | 211 |  | 389
[6 empty] | 210 |  | 390
[6 empty] | 209 |  | 391
[6 empty] | 208 |  | 392
[6 empty] | 207 |  | 393
[6 empty] | 206 |  | 394
[6 empty] | 205 |  | 395
[6 empty] | 204 |  | 396
[6 empty] | 203 |  | 397
[6 empty] | 202 |  | 398
[6 empty] | 201 |  | 399
[6 empty] | 200 |  | 400

## Petri
[8 empty] | UV Dose Calculations for a Low Pressure UV Lamp | [14 empty] | Calculations for carrying out a Dose-Reponse Curve
[8 empty] | Jim Bolton - Bolton Photosciences Inc., 92 Main St., Ayr, ON, Canada N0B 1E0 | [14 empty] | for a low pressure UV lamp
[8 empty] | Tel: 519-741-6283; Fax: 519-632-8941; Email: jbolton@boltonuv.com | [14 empty] | Programmed by Jim Bolton (519-741-6283)
Steps:
1. Draw a 0.5 cm x 0.5 cm grid and place the center of the grid at the center of the collimated beam
2. Measure the UV irradiance (with a radiometer) every 0.5 cm in the x and y directions and place into columns C and H below.
3. Place the Petri dish "inner" radius into Cell Q27.
4. The Petri factor is calculated in Cell Q20
UV Distribution across the Petri Dish for a low pressure UV Lamp
|  | Meter |  |  |  |  | Meter
x | y | Reading |  |  | x | y | Reading
|  | mW/cm2 | Ratio |  |  |  | mW/cm2 | Ratio
|  | y |  |  |  |  | x
0 | -6 |  | 0 |  | -6 | 0 |  | 0
0 | -5.5 |  | 0 |  | -5.5 | 0 |  | 0
0 | -5 |  | 0 |  | -5 | 0 |  | 0
0 | -4.5 |  | 0 |  | -4.5 | 0 |  | 0 | [6 empty] | Petri factor = | 0.89111
0 | -4 |  | 0 |  | -4 | 0 |  | 0
0 | -3.5 |  | 0 |  | -3.5 | 0 |  | 0
0 | -3 |  | 0 |  | -3 | 0 |  | 0
0 | -2.5 | 2.38 | 0.50746 |  | -2.5 | 0 | 3.05 | 0.64482
0 | -2 | 2.74 | 0.58422 |  | -2 | 0 | 3.56 | 0.75264
0 | -1.5 | 3.31 | 0.70576 |  | -1.5 | 0 | 3.96 | 0.83721
0 | -1 | 3.83 | 0.81663 |  | -1 | 0 | 4.39 | 0.92812 | [6 empty] | Petri Dish Radius = | 1.75 | cm
0 | -0.5 | 4.23 | 0.90192 |  | -0.5 | 0 | 4.65 | 0.98309
0 | 0 | 4.69 | 1 |  | 0 | 0 | 4.73 | 1
0 | 0.5 | 4.91 | 1.0469 |  | 0.5 | 0 | 4.55 | 0.96195
0 | 1 | 4.89 | 1.0426 |  | 1 | 0 | 4.31 | 0.91121
0 | 1.5 | 4.69 | 1 |  | 1.5 | 0 | 3.97 | 0.83932
0 | 2 | 4.34 | 0.92537 |  | 2 | 0 | 3.56 | 0.75264
0 | 2.5 | 3.9 | 0.83156 |  | 2.5 | 0 | 3.14 | 0.66385
0 | 3 |  | 0 |  | 3 | 0 |  | 0
0 | 3.5 |  | 0 |  | 3.5 | 0 |  | 0
0 | 4 |  | 0 |  | 4 | 0 |  | 0
0 | 4.5 |  | 0 |  | 4.5 | 0 |  | 0
0 | 5 |  | 0 |  | 5 | 0 |  | 0
0 | 5.5 |  | 0 |  | 5.5 | 0 |  | 0
0 | 6 |  | 0 |  | 6 | 0 |  | 0
[21 empty] | 3.5 | 4 | 4.5 | 5
x y |  | -6 | -5.5 | 5 | -4.5 | -4 | -3.5 | -3 | -2.5 | -2 | -1.5 | -1 | -0.5 | 0 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 |  |  |  |  | 5.5 | 6
-6 | [27 empty] | 0 | 0
-5.5 | [27 empty] | 0 | 0
-5 | [27 empty] | 0 | 0
-4.5 | [27 empty] | 0 | 0
-4 | [27 empty] | 0 | 0
-3.5 | [27 empty] | 0 | 0
-3 | [27 empty] | 0 | 0
-2.5 | [27 empty] | 0 | 0
-2 | [27 empty] | 0 | 0
-1.5 | [12 empty] | 0.75509 | 0.83721 | 0.87648 | [12 empty] | 2.4688 | 3
-1 | [11 empty] | 0.75793 | 0.83709 | 0.92812 | 0.97165 | 0.9677 | [11 empty] | 4.4625 | 5
-0.5 | [10 empty] | 0.69382 | 0.80282 | 0.88666 | 0.98309 | 1.0292 | 1.025 | 0.98309 | [10 empty] | 6.4037 | 7
0 | [10 empty] | 0.70576 | 0.81663 | 0.90192 | 1 | 1.0469 | 1.0426 | 1 | [10 empty] | 6.5139 | 7
0.5 | [10 empty] | 0.6789 | 0.78555 | 0.8676 | 0.96195 | 1.0071 | 1.003 | 0.96195 | [10 empty] | 6.266 | 7
1 | [11 empty] | 0.74412 | 0.82183 | 0.91121 | 0.95395 | 0.95006 | [11 empty] | 4.3812 | 5
1.5 | [12 empty] | 0.757 | 0.83932 | 0.87869 | [12 empty] | 2.475 | 3
2 | [27 empty] | 0 | 0
2.5 | [27 empty] | 0 | 0
3 | [27 empty] | 0 | 0
3.5 | [27 empty] | 0 | 0
4 | [27 empty] | 0 | 0
4.5 | [27 empty] | 0 | 0
5 | [27 empty] | 0 | 0
5.5 | [27 empty] | 0 | 0
6 | [27 empty] | 0 | 0
[28 empty] | 32.971 | 37
[27 empty] | Average Petri factor = | 0.89111

## Germicidal Calculator
Petri Factor = | 0.89111 |  | Incident Irradiance = | 0.473
Sample Depth = | 0.6 | cm | Average Irradiance = | 0.33744
Sample to beam = | 4 | cm
Divergence Factor = | 0.86957 |  | TRLE | 1
Reflection Factor = | 0.97721
Sensor Factor = | 0.96078
Wavelength (nm) | Rated Lamp Energy (RLE) over 200-300 nm | Radiometer Calibration Scan Data (Linden) | Radiometer Sensitivity (fr) (relative to 270 nm) | RLE x radioeter sensitivity | L(l) RLE/TRLE (Same as column B) | Ii | Water Absorbance | Water Factor | Average Irradiance Calc | Absorbance DNA | Relative Absorbance | Germicidal Irraidance | Action spectra |  | Action Spectra Irradiance
200 | -2.758e-05 | 0.0015617 | 1.2226 | -3.3719e-05 | -2.758e-05 | -9.4909e-06 | 1.164 | 0.49731 | -4.7199e-06 | 0.83 | 1.431 | -6.7544e-06 | 5.2451 | 4.196 | -1.9805e-05
201 | 8.2059e-05 | 0.0016241 | 1.2715 | 0.00010434 | 8.2059e-05 | 2.8238e-05 | 0.979 | 0.54817 | 1.5479e-05 | 0.82 | 1.4138 | 2.1884e-05 | 5.6644 | 4.5314 | 7.0142e-05
202 | 0 | 0.0016865 | 1.3203 | 0 | 0 | 0 | 0.813 | 0.60075 | 0 | 0.81 | 1.3966 | 0 | 6.0802 | 4.864 | 0
203 | 6.1452e-05 | 0.0017625 | 1.3798 | 8.479e-05 | 6.1452e-05 | 2.1147e-05 | 0.684 | 0.64691 | 1.368e-05 | 0.8 | 1.3793 | 1.8869e-05 | 6.4892 | 5.1912 | 7.1015e-05
204 | 3.9579e-05 | 0.0018384 | 1.4392 | 5.6963e-05 | 3.9579e-05 | 1.362e-05 | 0.574 | 0.69043 | 9.4036e-06 | 0.79 | 1.3621 | 1.2808e-05 | 6.8878 | 5.5101 | 5.1815e-05
205 | 5.9811e-05 | 0.0018895 | 1.4793 | 8.8476e-05 | 5.9811e-05 | 2.0582e-05 | 0.484 | 0.72922 | 1.5009e-05 | 0.776 | 1.3379 | 2.0081e-05 | 7.2728 | 5.818 | 8.7322e-05
206 | -1.4742e-05 | 0.0019407 | 1.5193 | -2.2398e-05 | -1.4742e-05 | -5.073e-06 | 0.407 | 0.7649 | -3.8804e-06 | 0.762 | 1.3138 | -5.098e-06 | 7.6406 | 6.1123 | -2.3718e-05
207 | 7.4209e-05 | 0.0019779 | 1.5484 | 0.00011491 | 7.4209e-05 | 2.5537e-05 | 0.341 | 0.79746 | 2.0365e-05 | 0.748 | 1.2897 | 2.6263e-05 | 7.9879 | 6.3901 | 0.00013013
208 | 1.6366e-05 | 0.0020151 | 1.5776 | 2.5818e-05 | 1.6366e-05 | 5.6318e-06 | 0.297 | 0.82024 | 4.6194e-06 | 0.735 | 1.2672 | 5.8539e-06 | 8.3111 | 6.6487 | 3.0713e-05
209 | 1.3927e-05 | 0.0020608 | 1.6133 | 2.2469e-05 | 1.3927e-05 | 4.7926e-06 | 0.251 | 0.84503 | 4.0499e-06 | 0.71 | 1.2241 | 4.9577e-06 | 8.607 | 6.8854 | 2.7885e-05
210 | 1.4343e-05 | 0.0021064 | 1.649 | 2.3651e-05 | 1.4343e-05 | 4.9356e-06 | 0.222 | 0.86119 | 4.2505e-06 | 0.685 | 1.181 | 5.02e-06 | 8.872 | 7.0974 | 3.0167e-05
211 | 1.7217e-05 | 0.0021453 | 1.6795 | 2.8915e-05 | 1.7217e-05 | 5.9245e-06 | 0.202 | 0.87259 | 5.1697e-06 | 0.66 | 1.1379 | 5.8827e-06 | 9.1027 | 7.2819 | 3.7645e-05
212 | 2.3511e-05 | 0.0021842 | 1.7099 | 4.0201e-05 | 2.3511e-05 | 8.0904e-06 | 0.178 | 0.88653 | 7.1724e-06 | 0.645 | 1.1121 | 7.9762e-06 | 9.2957 | 7.4363 | 5.3336e-05
213 | 1.5581e-06 | 0.0022142 | 1.7334 | 2.7009e-06 | 1.5581e-06 | 5.3619e-07 | 0.165 | 0.89421 | 4.7946e-07 | 0.624 | 1.0759 | 5.1584e-07 | 9.4476 | 7.5579 | 3.6237e-06
214 | 2.3207e-05 | 0.0022442 | 1.7569 | 4.0773e-05 | 2.3207e-05 | 7.986e-06 | 0.152 | 0.90198 | 7.2032e-06 | 0.603 | 1.0397 | 7.4889e-06 | 9.5549 | 7.6437 | 5.5059e-05
215 | 2.888e-05 | 0.0022661 | 1.774 | 5.1234e-05 | 2.888e-05 | 9.938e-06 | 0.139 | 0.90984 | 9.042e-06 | 0.582 | 1.0034 | 9.0732e-06 | 9.6143 | 7.6912 | 6.9544e-05
216 | 7.9646e-06 | 0.0022879 | 1.7912 | 1.4266e-05 | 7.9646e-06 | 2.7408e-06 | 0.131 | 0.91473 | 2.507e-06 | 0.56 | 0.96552 | 2.4206e-06 | 9.6222 | 7.6975 | 1.9298e-05
217 | 3.4393e-05 | 0.0023127 | 1.8106 | 6.2271e-05 | 3.4393e-05 | 1.1835e-05 | 0.125 | 0.91842 | 1.087e-05 | 0.54 | 0.93103 | 1.012e-05 | 9.5753 | 7.66 | 8.3262e-05
218 | 1.2723e-05 | 0.0023375 | 1.83 | 2.3284e-05 | 1.2723e-05 | 4.3784e-06 | 0.119 | 0.92212 | 4.0374e-06 | 0.52 | 0.89655 | 3.6197e-06 | 9.4701 | 7.5759 | 3.0587e-05
219 | 3.8929e-05 | 0.002371 | 1.8562 | 7.2259e-05 | 3.8929e-05 | 1.3396e-05 | 0.113 | 0.92585 | 1.2403e-05 | 0.5 | 0.86207 | 1.0692e-05 | 9.3033 | 7.4424 | 9.2306e-05
220 | 1.4687e-05 | 0.0024045 | 1.8824 | 2.7646e-05 | 1.4687e-05 | 5.054e-06 | 0.108 | 0.92897 | 4.695e-06 | 0.48 | 0.82759 | 3.8855e-06 | 9.0713 | 7.2568 | 3.407e-05
221 | 1.5545e-05 | 0.0023973 | 1.8767 | 2.9173e-05 | 1.5545e-05 | 5.3492e-06 | 0.104 | 0.93148 | 4.9827e-06 | 0.453 | 0.78103 | 3.8917e-06 | 8.773 | 7.0182 | 3.497e-05
222 | 2.4083e-05 | 0.00239 | 1.8711 | 4.5061e-05 | 2.4083e-05 | 8.2875e-06 | 0.1 | 0.934 | 7.7405e-06 | 0.426 | 0.73448 | 5.6853e-06 | 8.4169 | 6.7333 | 5.2119e-05
223 | 2.694e-05 | 0.0023994 | 1.8784 | 5.0604e-05 | 2.694e-05 | 9.2704e-06 | 0.097 | 0.93589 | 8.6761e-06 | 0.399 | 0.68793 | 5.9686e-06 | 8.0134 | 6.4105 | 5.5618e-05
224 | 3.4996e-05 | 0.0024088 | 1.8858 | 6.5996e-05 | 3.4996e-05 | 1.2043e-05 | 0.092 | 0.93906 | 1.1309e-05 | 0.37 | 0.63793 | 7.2143e-06 | 7.5731 | 6.0583 | 6.8513e-05
225 | 1.2547e-05 | 0.0024061 | 1.8837 | 2.3635e-05 | 1.2547e-05 | 4.3178e-06 | 0.089 | 0.94097 | 4.0629e-06 | 0.35 | 0.60345 | 2.4517e-06 | 7.1066 | 5.6851 | 2.3098e-05
226 | 1.5937e-05 | 0.0024035 | 1.8816 | 2.9987e-05 | 1.5937e-05 | 5.4843e-06 | 0.085 | 0.94352 | 5.1745e-06 | 0.337 | 0.58103 | 3.0066e-06 | 6.6246 | 5.2995 | 2.7422e-05
227 | 1.1652e-05 | 0.0024126 | 1.8887 | 2.2008e-05 | 1.1652e-05 | 4.0097e-06 | 0.081 | 0.94608 | 3.7935e-06 | 0.324 | 0.55862 | 2.1191e-06 | 6.1376 | 4.91 | 1.8626e-05
228 | 3.1821e-05 | 0.0024217 | 1.8959 | 6.0329e-05 | 3.1821e-05 | 1.095e-05 | 0.077 | 0.94865 | 1.0388e-05 | 0.31 | 0.53448 | 5.5521e-06 | 5.6563 | 4.5249 | 4.7004e-05
229 | 2.6789e-05 | 0.0024115 | 1.8879 | 5.0574e-05 | 2.6789e-05 | 9.2184e-06 | 0.074 | 0.95058 | 8.7628e-06 | 0.305 | 0.52586 | 4.608e-06 | 5.1912 | 4.1528 | 3.6391e-05
230 | 2.0933e-05 | 0.0024013 | 1.8799 | 3.9352e-05 | 2.0933e-05 | 7.2035e-06 | 0.07 | 0.95317 | 6.8661e-06 | 0.3 | 0.51724 | 3.5514e-06 | 4.7529 | 3.8022 | 2.6106e-05
231 | 3.227e-05 | 0.0024014 | 1.88 | 6.0667e-05 | 3.227e-05 | 1.1105e-05 | 0.067 | 0.95511 | 1.0606e-05 | 0.3 | 0.51724 | 5.4859e-06 | 4.3499 | 3.4798 | 3.6908e-05
232 | 3.2641e-05 | 0.0024016 | 1.8801 | 6.1369e-05 | 3.2641e-05 | 1.1232e-05 | 0.064 | 0.95707 | 1.075e-05 | 0.3 | 0.51724 | 5.5604e-06 | 3.9822 | 3.1856 | 3.4246e-05
233 | 2.9937e-05 | 0.0023984 | 1.8777 | 5.6211e-05 | 2.9937e-05 | 1.0302e-05 | 0.06 | 0.95968 | 9.8864e-06 | 0.305 | 0.52586 | 5.1989e-06 | 3.6475 | 2.9179 | 2.8848e-05
234 | 3.2295e-05 | 0.0023953 | 1.8752 | 6.056e-05 | 3.2295e-05 | 1.1113e-05 | 0.057 | 0.96164 | 1.0687e-05 | 0.31 | 0.53448 | 5.712e-06 | 3.3438 | 2.6749 | 2.8587e-05
235 | 2.5154e-05 | 0.0023887 | 1.87 | 4.7039e-05 | 2.5154e-05 | 8.656e-06 | 0.055 | 0.96295 | 8.3353e-06 | 0.32 | 0.55172 | 4.5988e-06 | 3.0688 | 2.4549 | 2.0463e-05
236 | 4.1371e-05 | 0.0023821 | 1.8649 | 7.7153e-05 | 4.1371e-05 | 1.4237e-05 | 0.052 | 0.96492 | 1.3737e-05 | 0.33 | 0.56897 | 7.816e-06 | 2.8203 | 2.2562 | 3.0994e-05
237 | 3.0269e-05 | 0.0023659 | 1.8522 | 5.6063e-05 | 3.0269e-05 | 1.0416e-05 | 0.05 | 0.96624 | 1.0064e-05 | 0.35 | 0.60345 | 6.0734e-06 | 2.5962 | 2.0769 | 2.0903e-05
238 | 3.7412e-05 | 0.0023496 | 1.8394 | 6.8817e-05 | 3.7412e-05 | 1.2874e-05 | 0.049 | 0.9669 | 1.2448e-05 | 0.37 | 0.63793 | 7.9411e-06 | 2.3944 | 1.9155 | 2.3844e-05
239 | 4.3187e-05 | 0.0023429 | 1.8341 | 7.9212e-05 | 4.3187e-05 | 1.4862e-05 | 0.046 | 0.96889 | 1.4399e-05 | 0.39 | 0.67241 | 9.6822e-06 | 2.2126 | 1.77 | 2.5487e-05
240 | 6.097e-05 | 0.0023361 | 1.8289 | 0.00011151 | 6.097e-05 | 2.0981e-05 | 0.046 | 0.96889 | 2.0328e-05 | 0.41 | 0.7069 | 1.437e-05 | 2.0487 | 1.6389 | 3.3316e-05
241 | 8.9086e-05 | 0.0023177 | 1.8144 | 0.00016164 | 8.9086e-05 | 3.0656e-05 | 0.044 | 0.97021 | 2.9743e-05 | 0.433 | 0.74655 | 2.2205e-05 | 1.9005 | 1.5203 | 4.5219e-05
242 | 0.00010147 | 0.0022992 | 1.8 | 0.00018265 | 0.00010147 | 3.4919e-05 | 0.042 | 0.97154 | 3.3925e-05 | 0.456 | 0.78621 | 2.6672e-05 | 1.7658 | 1.4126 | 4.7923e-05
243 | 0.00015695 | 0.0022817 | 1.7862 | 0.00028035 | 0.00015695 | 5.4009e-05 | 0.042 | 0.97154 | 5.2472e-05 | 0.479 | 0.82586 | 4.3334e-05 | 1.6425 | 1.3139 | 6.8944e-05
244 | 0.00022861 | 0.0022641 | 1.7725 | 0.00040521 | 0.00022861 | 7.867e-05 | 0.041 | 0.97221 | 7.6483e-05 | 0.5 | 0.86207 | 6.5934e-05 | 1.5283 | 1.2226 | 9.351e-05
245 | 0.00039469 | 0.0022321 | 1.7474 | 0.00068969 | 0.00039469 | 0.00013582 | 0.04 | 0.97287 | 0.00013213 | 0.52 | 0.89655 | 0.00011847 | 1.4218 | 1.1374 | 0.00015029
246 | 0.00066453 | 0.0022001 | 1.7224 | 0.0011446 | 0.00066453 | 0.00022868 | 0.04 | 0.97287 | 0.00022247 | 0.54 | 0.93103 | 0.00020713 | 1.3237 | 1.0589 | 0.00023558
247 | 0.0011113 | 0.0021777 | 1.7048 | 0.0018946 | 0.0011113 | 0.00038241 | 0.039 | 0.97354 | 0.00037229 | 0.56 | 0.96552 | 0.00035946 | 1.2353 | 0.98824 | 0.00036792
248 | 0.0020253 | 0.0021553 | 1.6873 | 0.0034172 | 0.0020253 | 0.00069694 | 0.039 | 0.97354 | 0.0006785 | 0.59 | 1.0172 | 0.0006902 | 1.1582 | 0.92652 | 0.00062864
249 | 0.0032798 | 0.0021371 | 1.6731 | 0.0054874 | 0.0032798 | 0.0011287 | 0.038 | 0.9742 | 0.0010995 | 0.61 | 1.0517 | 0.0011564 | 1.0936 | 0.87486 | 0.00096194
250 | 0.0051951 | 0.002119 | 1.6589 | 0.008618 | 0.0051951 | 0.0017877 | 0.037 | 0.97487 | 0.0017428 | 0.63 | 1.0862 | 0.001893 | 1.043 | 0.83437 | 0.0014541
251 | 0.0078413 | 0.0020712 | 1.6215 | 0.012714 | 0.0078413 | 0.0026983 | 0.036 | 0.97554 | 0.0026323 | 0.65 | 1.1207 | 0.00295 | 1.0077 | 0.80614 | 0.002122
252 | 0.011258 | 0.0020234 | 1.5841 | 0.017834 | 0.011258 | 0.0038742 | 0.036 | 0.97554 | 0.0037794 | 0.665 | 1.1466 | 0.0043333 | 0.98917 | 0.79131 | 0.0029907
253 | 0.015153 | 0.0019932 | 1.5604 | 0.023645 | 0.015153 | 0.0052144 | 0.035 | 0.97621 | 0.0050904 | 0.68 | 1.1724 | 0.005968 | 0.98876 | 0.79098 | 0.0040264
254 | 0.019697 | 0.001963 | 1.5368 | 0.030269 | 0.019697 | 0.006778 | 0.035 | 0.97621 | 0.0066167 | 0.695 | 1.1983 | 0.0079287 | 1.0078 | 0.80625 | 0.0053347
255 | 0.024038 | 0.0019167 | 1.5005 | 0.036069 | 0.024038 | 0.008272 | 0.034 | 0.97688 | 0.0080807 | 0.71 | 1.2241 | 0.0098919 | 1.0459 | 0.83667 | 0.0067609
256 | 0.028166 | 0.0018703 | 1.4642 | 0.04124 | 0.028166 | 0.0096923 | 0.034 | 0.97688 | 0.0094682 | 0.72 | 1.2414 | 0.011754 | 1.0991 | 0.87928 | 0.0083251
257 | 0.031509 | 0.0018276 | 1.4308 | 0.045084 | 0.031509 | 0.010843 | 0.032 | 0.97822 | 0.010607 | 0.725 | 1.25 | 0.013258 | 1.1635 | 0.9308 | 0.0098728
258 | 0.033587 | 0.001785 | 1.3974 | 0.046935 | 0.033587 | 0.011558 | 0.033 | 0.97755 | 0.011298 | 0.73 | 1.2586 | 0.01422 | 1.2351 | 0.98802 | 0.011163
259 | 0.034579 | 0.0017649 | 1.3817 | 0.047778 | 0.034579 | 0.011899 | 0.032 | 0.97822 | 0.01164 | 0.735 | 1.2672 | 0.014751 | 1.3095 | 1.0476 | 0.012194
260 | 0.034571 | 0.0017448 | 1.366 | 0.047223 | 0.034571 | 0.011896 | 0.032 | 0.97822 | 0.011637 | 0.725 | 1.25 | 0.014547 | 1.382 | 1.1056 | 0.012866
261 | 0.032975 | 0.0016698 | 1.3072 | 0.043106 | 0.032975 | 0.011347 | 0.033 | 0.97755 | 0.011092 | 0.7125 | 1.2284 | 0.013627 | 1.4478 | 1.1582 | 0.012847
262 | 0.03043 | 0.0015947 | 1.2485 | 0.037991 | 0.03043 | 0.010471 | 0.03 | 0.97956 | 0.010257 | 0.7 | 1.2069 | 0.01238 | 1.5018 | 1.2014 | 0.012324
263 | 0.027517 | 0.001563 | 1.2237 | 0.033671 | 0.027517 | 0.0094691 | 0.033 | 0.97755 | 0.0092565 | 0.6925 | 1.194 | 0.011052 | 1.5392 | 1.2313 | 0.011398
264 | 0.02421 | 0.0015314 | 1.1988 | 0.029024 | 0.02421 | 0.0083312 | 0.031 | 0.97889 | 0.0081553 | 0.685 | 1.181 | 0.0096317 | 1.555 | 1.244 | 0.010145
265 | 0.021232 | 0.0014956 | 1.1708 | 0.024859 | 0.021232 | 0.0073062 | 0.03 | 0.97956 | 0.0071569 | 0.675 | 1.1638 | 0.0083291 | 1.5461 | 1.2368 | 0.0088519
266 | 0.018781 | 0.0014598 | 1.1428 | 0.021464 | 0.018781 | 0.006463 | 0.031 | 0.97889 | 0.0063265 | 0.662 | 1.1414 | 0.007221 | 1.5163 | 1.213 | 0.0076741
267 | 0.016955 | 0.0014427 | 1.1295 | 0.01915 | 0.016955 | 0.0058346 | 0.031 | 0.97889 | 0.0057114 | 0.649 | 1.119 | 0.0063909 | 1.4713 | 1.177 | 0.0067222
268 | 0.01591 | 0.0014257 | 1.1161 | 0.017758 | 0.01591 | 0.005475 | 0.03 | 0.97956 | 0.0053631 | 0.635 | 1.0948 | 0.0058717 | 1.4166 | 1.1332 | 0.0060777
269 | 0.015858 | 0.0013689 | 1.0717 | 0.016994 | 0.015858 | 0.0054569 | 0.03 | 0.97956 | 0.0053454 | 0.622 | 1.0724 | 0.0057324 | 1.3579 | 1.0863 | 0.0058068
270 | 0.01751 | 0.0013121 | 1.0272 | 0.017986 | 0.01751 | 0.0060254 | 0.029 | 0.98023 | 0.0059063 | 0.6 | 1.0345 | 0.0061099 | 1.3009 | 1.0407 | 0.0061466
271 | 0.0207 | 0.0012774 | 1 | 0.0207 | 0.0207 | 0.0071234 | 0.028 | 0.98091 | 0.0069873 | 0.58 | 1 | 0.0069873 | 1.25 | 1 | 0.0069873
272 | 0.025512 | 0.0012426 | 0.97279 | 0.024817 | 0.025512 | 0.008779 | 0.029 | 0.98023 | 0.0086054 | 0.56 | 0.96552 | 0.0083087 | 1.2054 | 0.96431 | 0.0082983
273 | 0.031729 | 0.0012081 | 0.94577 | 0.030009 | 0.031729 | 0.010919 | 0.029 | 0.98023 | 0.010703 | 0.54 | 0.93103 | 0.0099647 | 1.1661 | 0.93282 | 0.0099838
274 | 0.038228 | 0.0011736 | 0.91874 | 0.035122 | 0.038228 | 0.013155 | 0.027 | 0.98158 | 0.012913 | 0.52 | 0.89655 | 0.011577 | 1.131 | 0.90474 | 0.011683
275 | 0.043585 | 0.0011338 | 0.8876 | 0.038686 | 0.043585 | 0.014998 | 0.027 | 0.98158 | 0.014722 | 0.5 | 0.86207 | 0.012691 | 1.0991 | 0.87926 | 0.012944
276 | 0.047032 | 0.001094 | 0.85646 | 0.040281 | 0.047032 | 0.016184 | 0.026 | 0.98225 | 0.015897 | 0.48 | 0.82759 | 0.013156 | 1.0695 | 0.85558 | 0.013601
277 | 0.047403 | 0.0010564 | 0.82704 | 0.039204 | 0.047403 | 0.016312 | 0.026 | 0.98225 | 0.016023 | 0.46 | 0.7931 | 0.012708 | 1.0412 | 0.83291 | 0.013345
278 | 0.045205 | 0.0010188 | 0.79762 | 0.036056 | 0.045205 | 0.015556 | 0.025 | 0.98293 | 0.01529 | 0.44 | 0.75862 | 0.011599 | 1.0131 | 0.81044 | 0.012392
279 | 0.040919 | 0.00097406 | 0.76256 | 0.031203 | 0.040919 | 0.014081 | 0.025 | 0.98293 | 0.01384 | 0.42 | 0.72414 | 0.010022 | 0.98425 | 0.78737 | 0.010898
280 | 0.035087 | 0.00092926 | 0.72749 | 0.025526 | 0.035087 | 0.012074 | 0.024 | 0.9836 | 0.011876 | 0.4 | 0.68966 | 0.0081905 | 0.95366 | 0.7629 | 0.0090604
281 | 0.029244 | 0.00087526 | 0.68521 | 0.020038 | 0.029244 | 0.010063 | 0.023 | 0.98428 | 0.009905 | 0.376 | 0.64828 | 0.0064212 | 0.92032 | 0.73623 | 0.0072924
282 | 0.02372 | 0.00082126 | 0.64294 | 0.015251 | 0.02372 | 0.0081626 | 0.023 | 0.98428 | 0.0080343 | 0.352 | 0.6069 | 0.004876 | 0.88346 | 0.70674 | 0.0056782
283 | 0.018964 | 0.00078303 | 0.613 | 0.011625 | 0.018964 | 0.0065259 | 0.022 | 0.98496 | 0.0064277 | 0.328 | 0.56552 | 0.003635 | 0.84318 | 0.67452 | 0.0043357
284 | 0.015216 | 0.00074479 | 0.58307 | 0.0088722 | 0.015216 | 0.0052362 | 0.022 | 0.98496 | 0.0051574 | 0.305 | 0.52586 | 0.0027121 | 0.79984 | 0.63985 | 0.0033
285 | 0.012067 | 0.0007104 | 0.55615 | 0.0067111 | 0.012067 | 0.0041525 | 0.02 | 0.98631 | 0.0040957 | 0.28 | 0.48276 | 0.0019772 | 0.75377 | 0.603 | 0.0024697
286 | 0.0094768 | 0.000676 | 0.52922 | 0.0050153 | 0.0094768 | 0.0032611 | 0.019 | 0.98699 | 0.0032187 | 0.255 | 0.43966 | 0.0014151 | 0.70531 | 0.56423 | 0.0018161
287 | 0.0076324 | 0.00065196 | 0.5104 | 0.0038955 | 0.0076324 | 0.0026264 | 0.017 | 0.98835 | 0.0025958 | 0.23 | 0.39655 | 0.0010294 | 0.65481 | 0.52383 | 0.0013598
288 | 0.0061958 | 0.00062791 | 0.49157 | 0.0030457 | 0.0061958 | 0.0021321 | 0.017 | 0.98835 | 0.0021072 | 0.21 | 0.36207 | 0.00076297 | 0.6026 | 0.48207 | 0.0010158
289 | 0.0050671 | 0.00059241 | 0.46378 | 0.00235 | 0.0050671 | 0.0017437 | 0.016 | 0.98903 | 0.0017245 | 0.19 | 0.32759 | 0.00056493 | 0.54903 | 0.43921 | 0.00075743
290 | 0.0041261 | 0.0005569 | 0.43598 | 0.0017989 | 0.0041261 | 0.0014198 | 0.016 | 0.98903 | 0.0014043 | 0.17 | 0.2931 | 0.0004116 | 0.49444 | 0.39554 | 0.00055544
291 | 0.0033899 | 0.00051359 | 0.40207 | 0.001363 | 0.0033899 | 0.0011665 | 0.014 | 0.99039 | 0.0011553 | 0.15 | 0.25862 | 0.00029879 | 0.43916 | 0.35132 | 0.00040589
292 | 0.0027997 | 0.00047027 | 0.36816 | 0.0010307 | 0.0027997 | 0.00096343 | 0.014 | 0.99039 | 0.00095417 | 0.14 | 0.24138 | 0.00023032 | 0.38354 | 0.30682 | 0.00029276
293 | 0.0023665 | 0.0004505 | 0.35268 | 0.0008346 | 0.0023665 | 0.00081434 | 0.013 | 0.99107 | 0.00080707 | 0.125 | 0.21552 | 0.00017394 | 0.32792 | 0.26233 | 0.00021172
294 | 0.0019621 | 0.00043072 | 0.3372 | 0.00066162 | 0.0019621 | 0.0006752 | 0.012 | 0.99176 | 0.00066963 | 0.11 | 0.18966 | 0.000127 | 0.27265 | 0.21811 | 0.00014605
295 | 0.0016064 | 0.00040247 | 0.31508 | 0.00050616 | 0.0016064 | 0.0005528 | 0.011 | 0.99244 | 0.00054862 | 0.095 | 0.16379 | 8.9861e-05 | 0.21805 | 0.17444 | 9.57e-05
296 | 0.0013166 | 0.00037422 | 0.29296 | 0.00038572 | 0.0013166 | 0.00045307 | 0.011 | 0.99244 | 0.00044965 | 0.08 | 0.13793 | 6.202e-05 | 0.16448 | 0.13158 | 5.9165e-05
297 | 0.0010788 | 0.00034845 | 0.27279 | 0.00029429 | 0.0010788 | 0.00037124 | 0.011 | 0.99244 | 0.00036843 | 0.071 | 0.12241 | 4.5101e-05 | 0.11227 | 0.089816 | 3.3091e-05
298 | 0.00088304 | 0.00032269 | 0.25262 | 0.00022307 | 0.00088304 | 0.00030387 | 0.011 | 0.99244 | 0.00030157 | 0.062 | 0.1069 | 3.2237e-05 | 0.06177 | 0.049414 | 1.4902e-05
299 | 0.00074076 | 0.00030325 | 0.23741 | 0.00017586 | 0.00074076 | 0.00025491 | 0.01 | 0.99312 | 0.00025316 | 0.053 | 0.091379 | 2.3133e-05 | 0.013312 | 0.01065 | 2.696e-06
300 | 0.00061422 | 0.00028382 | 0.2222 | 0.00013648 | 0.00061422 | 0.00021137 | 0.01 | 0.99312 | 0.00020991 | 0.045 | 0.077586 | 1.6286e-05 | 0 | 0 | 0
TRLE: | 1 |  |  | 1.0408 |  | 0.34412 |  | Average Irradiance | 0.33744 |  | 0 | 0.32092 | Action | #VALUE! | 0.30835
|  |  | Lamp Correction Factor, fp | 0.96078
|  |  | Incident IR= | 0.34412

## Emission spectrum 3 rows
|  | Interpolation over germicidal range (200 to 300) |  |  |  | Calculating rated lamp energy (RLE)
Spectrometer Measurement |  | x | x1, x2 | y1, y2 | y |  | interpolated | RLE
200.11 | -9.7292e-05 | 200.11 | 200.11 | -9.7292e-05 | -9.7292e-05 | 200.11 | -9.7292e-05 | -2.758e-05 | -0.019469
200.23 | 0.00038555 |  |  |  |  | 201 | 0.00028947 | 8.2059e-05 | 0.058184
200.35 | -0.00017188 | 201 | 200.94 | 7.2952e-05 | 0.00028947 | 202 | 0 | 0 | 0
200.46 | 0.00017025 |  | 201.06 | 0.00050599 |  | 203 | 0.00021678 | 6.1452e-05 | 0.044006
200.58 | -0.00022495 | 202 | 202 | 0 | 0 | 204 | 0.00013962 | 3.9579e-05 | 0.028482
200.7 | 0 |  | 202.12 | 4.9932e-05 |  | 205 | 0.00021099 | 5.9811e-05 | 0.043252
200.82 | 0.00027609 | 203 | 202.95 | 6.2399e-05 | 0.00021678 | 206 | -5.2004e-05 | -1.4742e-05 | -0.010713
200.94 | 7.2952e-05 |  | 203.06 | 0.00040203 |  | 207 | 0.00026178 | 7.4209e-05 | 0.054188
201.06 | 0.00050599 | 204 | 203.89 | 8.7352e-05 | 0.00013962 | 208 | 5.7732e-05 | 1.6366e-05 | 0.012008
201.17 | 0 |  | 204.01 | 0.00014437 |  | 209 | 4.9129e-05 | 1.3927e-05 | 0.010268
201.29 | -1.7745e-05 | 205 | 204.95 | 0.00016152 | 0.00021099 | 210 | 5.0595e-05 | 1.4343e-05 | 0.010625
201.41 | 0.00024617 |  | 205.07 | 0.00028024 |  | 211 | 6.0733e-05 | 1.7217e-05 | 0.012815
201.53 | -0.00012196 | 206 | 205.89 | 6.2922e-05 | -5.2004e-05 | 212 | 8.2935e-05 | 2.3511e-05 | 0.017582
201.65 | 0.00013808 |  | 206.01 | -6.2452e-05 |  | 213 | 5.4965e-06 | 1.5581e-06 | 0.0011708
201.77 | 0.0002224 | 207 | 206.95 | 0.00025955 | 0.00026178 | 214 | 8.1865e-05 | 2.3207e-05 | 0.017519
201.88 | -6.7818e-05 |  | 207.07 | 0.00025781 |  | 215 | 0.00010187 | 2.888e-05 | 0.021903
202 | 0 | 208 | 207.9 | 0.00030191 | 5.7732e-05 | 216 | 2.8096e-05 | 7.9646e-06 | 0.0060687
202.12 | 4.9932e-05 |  | 208.01 | 3.3314e-05 |  | 217 | 0.00012132 | 3.4393e-05 | 0.026327
202.24 | 0 | 209 | 208.95 | 8.4222e-05 | 4.9129e-05 | 218 | 4.4883e-05 | 1.2723e-05 | 0.0097845
202.36 | 1.6337e-05 |  | 209.07 | 0 |  | 219 | 0.00013732 | 3.8929e-05 | 0.030074
202.47 | 0.00045315 | 210 | 209.89 | -5.0097e-05 | 5.0595e-05 | 220 | 5.1809e-05 | 1.4687e-05 | 0.011398
202.59 | 0.00028859 |  | 210.01 | 5.9749e-05 |  | 221 | 5.4836e-05 | 1.5545e-05 | 0.012119
202.71 | 0.00034945 | 211 | 210.95 | 5.6942e-05 | 6.0733e-05 | 222 | 8.4956e-05 | 2.4083e-05 | 0.01886
202.83 | 7.8692e-05 |  | 211.07 | 6.604e-05 |  | 223 | 9.5032e-05 | 2.694e-05 | 0.021192
202.95 | 6.2399e-05 | 212 | 211.89 | 9.0546e-05 | 8.2935e-05 | 224 | 0.00012345 | 3.4996e-05 | 0.027653
203.06 | 0.00040203 |  | 212.01 | 0.00010805 |  | 225 | 4.4262e-05 | 1.2547e-05 | 0.009959
203.18 | 0.00026056 | 213 | 212.95 | 1.7206e-05 | 5.4965e-06 | 226 | 5.622e-05 | 1.5937e-05 | 0.012706
203.3 | 0.00013676 |  | 213.06 | -8.5549e-06 |  | 227 | 4.1103e-05 | 1.1652e-05 | 0.0093305
203.42 | 4.5194e-05 | 214 | 214 | 8.1865e-05 | 8.1865e-05 | 228 | 0.00011225 | 3.1821e-05 | 0.025593
203.54 | 0.00014934 |  | 214.12 | 0.00010586 |  | 229 | 9.4499e-05 | 2.6789e-05 | 0.02164
203.65 | -1.4808e-05 | 215 | 214.94 | 8.646e-05 | 0.00010187 | 230 | 7.3843e-05 | 2.0933e-05 | 0.016984
203.77 | 0.00016149 |  | 215.06 | 0.00011729 |  | 231 | 0.00011383 | 3.227e-05 | 0.026296
203.89 | 8.7352e-05 | 216 | 215.99 | 2.2508e-05 | 2.8096e-05 | 232 | 0.00011514 | 3.2641e-05 | 0.026714
204.01 | 0.00014437 |  | 216.11 | 8.956e-05 |  | 233 | 0.0001056 | 2.9937e-05 | 0.024606
204.13 | 0.00040095 | 217 | 216.93 | 5.7631e-05 | 0.00012132 | 234 | 0.00011392 | 3.2295e-05 | 0.026658
204.24 | 0.00026977 |  | 217.04 | 0.00015772 |  | 235 | 8.8733e-05 | 2.5154e-05 | 0.020852
204.36 | 1.4077e-05 | 218 | 217.98 | 5.5234e-05 | 4.4883e-05 | 236 | 0.00014594 | 4.1371e-05 | 0.034442
204.48 | 0.00041858 |  | 218.1 | -6.8718e-06 |  | 237 | 0.00010678 | 3.0269e-05 | 0.025306
204.6 | 0.00041487 | 219 | 218.91 | 0.00015276 | 0.00013732 | 238 | 0.00013197 | 3.7412e-05 | 0.03141
204.72 | 0.00038371 |  | 219.03 | 0.00013218 |  | 239 | 0.00015235 | 4.3187e-05 | 0.036411
204.83 | 0.0002173 | 220 | 219.97 | 1.9063e-05 | 5.1809e-05 | 240 | 0.00021508 | 6.097e-05 | 0.051618
204.95 | 0.00016152 |  | 220.08 | 0.00013913 |  | 241 | 0.00031426 | 8.9086e-05 | 0.075736
205.07 | 0.00028024 | 221 | 220.9 | 5.5048e-05 | 5.4836e-05 | 242 | 0.00035796 | 0.00010147 | 0.086626
205.19 | 0.00013229 |  | 221.02 | 5.4793e-05 |  | 243 | 0.00055365 | 0.00015695 | 0.13454
205.3 | 0.00043278 | 222 | 221.95 | 8.7804e-05 | 8.4956e-05 | 244 | 0.00080645 | 0.00022861 | 0.19677
205.42 | 0.00026005 |  | 222.06 | 8.1538e-05 |  | 245 | 0.0013923 | 0.00039469 | 0.34111
205.54 | 0.00032236 | 223 | 223 | 9.5032e-05 | 9.5032e-05 | 246 | 0.0023442 | 0.00066453 | 0.57667
205.66 | 5.1144e-05 |  | 223.11 | 0.00017799 |  | 247 | 0.0039202 | 0.0011113 | 0.96828
205.78 | 0.00012683 | 224 | 223.93 | 2.1484e-05 | 0.00012345 | 248 | 0.0071444 | 0.0020253 | 1.7718
205.89 | 6.2922e-05 |  | 224.04 | 0.00018172 |  | 249 | 0.01157 | 0.0032798 | 2.8809
206.01 | -6.2452e-05 | 225 | 224.98 | 3.5937e-05 | 4.4262e-05 | 250 | 0.018326 | 0.0051951 | 4.5815
206.13 | -9.9191e-05 |  | 225.09 | 8.1725e-05 |  | 251 | 0.027661 | 0.0078413 | 6.9428
206.25 | 0.00016002 | 226 | 225.91 | 0.00013292 | 5.622e-05 | 252 | 0.039714 | 0.011258 | 10.008
206.37 | 0.00017108 |  | 226.02 | 3.9175e-05 |  | 253 | 0.053454 | 0.015153 | 13.5238
206.48 | 9.7051e-05 | 227 | 226.95 | 7.0463e-05 | 4.1103e-05 | 254 | 0.069482 | 0.019697 | 17.6484
206.6 | -4.819e-05 |  | 227.07 | 0 |  | 255 | 0.084797 | 0.024038 | 21.6232
206.72 | 2.3925e-05 | 228 | 228 | 0.00011225 | 0.00011225 | 256 | 0.099356 | 0.028166 | 25.4352
206.84 | 2.3758e-05 |  | 228.11 | 1.3402e-05 |  | 257 | 0.11115 | 0.031509 | 28.566
206.95 | 0.00025955 | 229 | 228.92 | 6.8996e-05 | 9.4499e-05 | 258 | 0.11848 | 0.033587 | 30.5681
207.07 | 0.00025781 |  | 229.04 | 0.00010725 |  | 259 | 0.12198 | 0.034579 | 31.5928
207.19 | -2.3283e-05 | 230 | 229.97 | 9.0758e-05 | 7.3843e-05 | 260 | 0.12195 | 0.034571 | 31.7075
207.31 | 5.7819e-05 |  | 230.08 | 2.8738e-05 |  | 261 | 0.11632 | 0.032975 | 30.36
207.42 | 3.4457e-05 | 231 | 230.89 | 5.9649e-05 | 0.00011383 | 262 | 0.10734 | 0.03043 | 28.124
207.54 | -5.7043e-05 |  | 231.01 | 0.00011876 |  | 263 | 0.097068 | 0.027517 | 25.529
207.66 | 0.00029465 | 232 | 231.94 | 0.00012997 | 0.00011514 | 264 | 0.085403 | 0.02421 | 22.5465
207.78 | -6.755e-05 |  | 232.05 | 0.00010279 |  | 265 | 0.074897 | 0.021232 | 19.8476
207.9 | 0.00030191 | 233 | 232.98 | 9.229e-05 | 0.0001056 | 266 | 0.066253 | 0.018781 | 17.6232
208.01 | 3.3314e-05 |  | 233.09 | 0.00016552 |  | 267 | 0.059811 | 0.016955 | 15.9696
208.13 | 0 | 234 | 233.9 | 5.7442e-05 | 0.00011392 | 268 | 0.056125 | 0.01591 | 15.0415
208.25 | 5.4777e-05 |  | 234.02 | 0.00012522 |  | 269 | 0.055939 | 0.015858 | 15.0476
208.37 | 0.00023939 | 235 | 234.94 | 8.3651e-05 | 8.8733e-05 | 270 | 0.061767 | 0.01751 | 16.677
208.48 | -1.0808e-05 |  | 235.06 | 9.3815e-05 |  | 271 | 0.073022 | 0.0207 | 19.789
208.6 | -5.3675e-05 | 236 | 235.98 | 0.00015278 | 0.00014594 | 272 | 0.089994 | 0.025512 | 24.4784
208.72 | 0.00021328 |  | 236.1 | 0.00011174 |  | 273 | 0.11193 | 0.031729 | 30.5564
208.84 | -2.1191e-05 | 237 | 236.91 | 4.9879e-05 | 0.00010678 | 274 | 0.13485 | 0.038228 | 36.9499
208.95 | 8.4222e-05 |  | 237.02 | 0.00011942 |  | 275 | 0.15375 | 0.043585 | 42.2808
209.07 | 0 | 238 | 237.94 | 0.00013698 | 0.00013197 | 276 | 0.16591 | 0.047032 | 45.7907
209.19 | 0.00039507 |  | 238.06 | 0.00012697 |  | 277 | 0.16722 | 0.047403 | 46.3189
209.31 | 8.267e-05 | 239 | 238.98 | 0.00014438 | 0.00015235 | 278 | 0.15946 | 0.045205 | 44.3306
209.42 | -2.0541e-05 |  | 239.1 | 0.00019218 |  | 279 | 0.14434 | 0.040919 | 40.2719
209.54 | -1.0208e-05 | 240 | 239.9 | 0.00021536 | 0.00021508 | 280 | 0.12377 | 0.035087 | 34.6565
209.66 | 0.00015216 |  | 240.02 | 0.00021502 |  | 281 | 0.10316 | 0.029244 | 28.9877
209.78 | 0.00027219 | 241 | 240.94 | 0.00033493 | 0.00031426 | 282 | 0.083675 | 0.02372 | 23.5964
209.89 | -5.0097e-05 |  | 241.05 | 0.00029703 |  | 283 | 0.066898 | 0.018964 | 18.9321
210.01 | 5.9749e-05 | 242 | 241.97 | 0.00035341 | 0.00035796 | 284 | 0.053677 | 0.015216 | 15.2442
210.13 | 0.00013858 |  | 242.09 | 0.00037161 |  | 285 | 0.042568 | 0.012067 | 12.1318
210.25 | 4.919e-05 | 243 | 242.89 | 0.00045241 | 0.00055365 | 286 | 0.03343 | 0.0094768 | 9.561
210.36 | 0.00025423 |  | 243.01 | 0.00056285 |  | 287 | 0.026924 | 0.0076324 | 7.7272
210.48 | 0.00022355 | 244 | 243.93 | 0.0008048 | 0.00080645 | 288 | 0.021856 | 0.0061958 | 6.2946
210.6 | 0.00014492 |  | 244.04 | 0.00080739 |  | 289 | 0.017875 | 0.0050671 | 5.1657
210.72 | 5.7618e-05 | 245 | 244.96 | 0.0013804 | 0.0013923 | 290 | 0.014555 | 0.0041261 | 4.221
210.83 | 0 |  | 245.07 | 0.0014131 |  | 291 | 0.011958 | 0.0033899 | 3.4799
210.95 | 5.6942e-05 | 246 | 245.99 | 0.0023238 | 0.0023442 | 292 | 0.0098762 | 0.0027997 | 2.8838
211.07 | 6.604e-05 |  | 246.11 | 0.0025686 |  | 293 | 0.0083479 | 0.0023665 | 2.4459
211.19 | 0.00013131 | 247 | 246.91 | 0.0037521 | 0.0039202 | 294 | 0.0069215 | 0.0019621 | 2.0349
211.3 | 3.7299e-05 |  | 247.02 | 0.0039575 |  | 295 | 0.0056668 | 0.0016064 | 1.6717
211.42 | 4.6345e-05 | 248 | 247.94 | 0.0068126 | 0.0071444 | 296 | 0.0046445 | 0.0013166 | 1.3748
211.54 | 5.529e-05 |  | 248.05 | 0.0074209 |  | 297 | 0.0038056 | 0.0010788 | 1.1303
211.66 | 0.0002657 | 249 | 248.97 | 0.011397 | 0.01157 | 298 | 0.003115 | 0.00088304 | 0.92827
211.77 | -2.7324e-05 |  | 249.08 | 0.012031 |  | 299 | 0.0026131 | 0.00074076 | 0.78132
211.89 | 9.0546e-05 | 250 | 250 | 0.018326 | 0.018326 | 300 | 0.0021667 | 0.00061422 | 0.65002
212.01 | 0.00010805 |  | 250.11 | 0.019338 |  |  | 3.5276 | 1 | 951.5273
212.13 | 8.0585e-05 | 251 | 250.91 | 0.026772 | 0.027661 |  |  | Weighted Avg: | 269.7393
212.24 | 9.793e-05 |  | 251.03 | 0.027957
212.36 | 2.6555e-05 | 252 | 251.94 | 0.038966 | 0.039714
212.48 | 5.2805e-05 |  | 252.05 | 0.040338
212.59 | -0.00010502 | 253 | 252.97 | 0.052839 | 0.053454
212.71 | 7.8319e-05 |  | 253.08 | 0.055093
212.83 | 0.00013844 | 254 | 253.99 | 0.069377 | 0.069482
212.95 | 1.7206e-05 |  | 254.11 | 0.070635
213.06 | -8.5549e-06 | 255 | 254.9 | 0.083532 | 0.084797
213.18 | 7.6564e-05 |  | 255.02 | 0.08505
213.3 | 0.00017763 | 256 | 255.93 | 0.098573 | 0.099356
213.42 | 7.5694e-05 |  | 256.04 | 0.099804
213.53 | 7.5271e-05 | 257 | 256.95 | 0.11076 | 0.11115
213.65 | 6.6546e-05 |  | 257.07 | 0.1117
213.77 | 9.102e-05 | 258 | 257.98 | 0.11837 | 0.11848
213.88 | 2.469e-05 |  | 258.09 | 0.11898
214 | 8.1865e-05 | 259 | 259 | 0.12198 | 0.12198
214.12 | 0.00010586 |  | 259.11 | 0.12285
214.24 | 8.1008e-05 | 260 | 259.91 | 0.12214 | 0.12195
214.35 | 0.00013704 |  | 260.02 | 0.12191
214.47 | 1.6039e-05 | 261 | 260.93 | 0.11692 | 0.11632
214.59 | -7.9792e-06 |  | 261.04 | 0.11598
214.7 | 0.00028581 | 262 | 261.95 | 0.10788 | 0.10734
214.82 | 7.8999e-05 |  | 262.06 | 0.1067
214.94 | 8.646e-05 | 263 | 262.97 | 0.097678 | 0.097068
215.06 | 0.00011729 |  | 263.08 | 0.095443
215.17 | 0.00010112 | 264 | 263.99 | 0.085483 | 0.085403
215.29 | 0.0001006 |  | 264.1 | 0.084608
215.41 | 3.8501e-05 | 265 | 264.89 | 0.075585 | 0.074897
215.52 | 4.596e-05 |  | 265.01 | 0.074834
215.64 | 1.5239e-05 | 266 | 265.91 | 0.066957 | 0.066253
215.76 | 0.00017436 |  | 266.02 | 0.066096
215.87 | 7.5413e-06 | 267 | 266.93 | 0.059894 | 0.059811
215.99 | 2.2508e-05 |  | 267.04 | 0.059764
216.11 | 8.956e-05 | 268 | 267.94 | 0.05589 | 0.056125
216.23 | 4.4552e-05 |  | 268.06 | 0.05636
216.34 | 0.00025117 | 269 | 268.96 | 0.055662 | 0.055939
216.46 | 0.00013229 |  | 269.07 | 0.056424
216.58 | 0.00025593 | 270 | 269.97 | 0.061586 | 0.061767
216.69 | 0.00012368 |  | 270.08 | 0.062249
216.81 | -4.3437e-05 | 271 | 270.99 | 0.07287 | 0.073022
216.93 | 5.7631e-05 |  | 271.1 | 0.074544
217.04 | 0.00015772 | 272 | 272 | 0.089994 | 0.089994
217.16 | -4.9942e-05 |  | 272.11 | 0.092399
217.28 | 0.00016331 | 273 | 272.9 | 0.10941 | 0.11193
217.4 | 0.00011306 |  | 273.01 | 0.11218
217.51 | 7.0334e-05 | 274 | 273.91 | 0.1328 | 0.13485
217.63 | 0.00012601 |  | 274.02 | 0.13531
217.75 | 8.3632e-05 | 275 | 274.92 | 0.15209 | 0.15375
217.86 | 9.0177e-05 |  | 275.03 | 0.15437
217.98 | 5.5234e-05 | 276 | 275.93 | 0.16587 | 0.16591
218.1 | -6.8718e-06 |  | 276.04 | 0.16593
218.21 | 0.00018465 | 277 | 276.94 | 0.16708 | 0.16722
218.33 | 0.00016334 |  | 277.05 | 0.16733
218.45 | -2.0318e-05 | 278 | 277.95 | 0.16009 | 0.15946
218.56 | 0.00013477 |  | 278.06 | 0.15871
218.68 | 2.0122e-05 | 279 | 278.96 | 0.14518 | 0.14434
218.8 | 2.0023e-05 |  | 279.07 | 0.14288
218.91 | 0.00015276 | 280 | 279.96 | 0.12432 | 0.12377
219.03 | 0.00013218 |  | 280.08 | 0.12268
219.15 | 3.2886e-05 | 281 | 280.97 | 0.10377 | 0.10316
219.26 | 0.00014398 |  | 281.08 | 0.10153
219.38 | 0.00012373 | 282 | 281.98 | 0.083994 | 0.083675
219.5 | 1.2961e-05 |  | 282.09 | 0.082241
219.62 | 0 | 283 | 282.98 | 0.067146 | 0.066898
219.73 | 0.00017322 |  | 283.09 | 0.065781
219.85 | 0.00014686 | 284 | 283.99 | 0.053797 | 0.053677
219.97 | 1.9063e-05 |  | 284.1 | 0.052473
220.08 | 0.00013913 | 285 | 284.99 | 0.042662 | 0.042568
220.2 | 0.00015736 |  | 285.1 | 0.041624
220.32 | 0.00013782 | 286 | 285.99 | 0.033488 | 0.03343
220.43 | 6.8575e-05 |  | 286.1 | 0.032852
220.55 | 6.2043e-05 | 287 | 286.99 | 0.02696 | 0.026924
220.67 | 6.7925e-05 |  | 287.1 | 0.026563
220.78 | 0.00017208 | 288 | 287.99 | 0.021906 | 0.021856
220.9 | 5.5048e-05 |  | 288.11 | 0.021309
221.02 | 5.4793e-05 | 289 | 288.99 | 0.017914 | 0.017875
221.13 | 6.0583e-05 |  | 289.11 | 0.01744
221.25 | 0.00018085 | 290 | 289.99 | 0.014585 | 0.014555
221.37 | -6.0001e-06 |  | 290.11 | 0.014225
221.48 | 3.5827e-05 | 291 | 290.99 | 0.011987 | 0.011958
221.6 | 2.9709e-05 |  | 291.1 | 0.011671
221.71 | 7.6872e-05 | 292 | 291.99 | 0.0098772 | 0.0098762
221.83 | 0.00018828 |  | 292.1 | 0.0098658
221.95 | 8.7804e-05 | 293 | 292.99 | 0.0083706 | 0.0083479
222.06 | 8.1538e-05 |  | 293.1 | 0.0081205
222.18 | 4.6363e-05 | 294 | 293.99 | 0.0069446 | 0.0069215
222.3 | 1.7299e-05 |  | 294.1 | 0.0066907
222.41 | 0.00012049 | 295 | 294.98 | 0.0056702 | 0.0056668
222.53 | 3.9956e-05 |  | 295.09 | 0.0056517
222.65 | 6.2463e-05 | 296 | 295.98 | 0.0046395 | 0.0046445
222.76 | 0.00012426 |  | 296.09 | 0.004667
222.88 | 1.6858e-05 | 297 | 296.97 | 0.0037824 | 0.0038056
223 | 9.5032e-05 |  | 297.08 | 0.0038674
223.11 | 0.00017799 | 298 | 297.96 | 0.0031437 | 0.003115
223.23 | 5.5351e-05 |  | 298.07 | 0.0030648
223.35 | 0.00018175 | 299 | 298.96 | 0.0026279 | 0.0026131
223.46 | 9.3163e-05 |  | 299.07 | 0.0025872
223.58 | 0.00010906 | 300 | 299.95 | 0.0021785 | 0.0021667
223.7 | 1.0851e-05 |  | 300.06 | 0.0021526
223.81 | 9.1774e-05
223.93 | 2.1484e-05
224.04 | 0.00018172
224.16 | 0.00010635
224.28 | 4.7626e-05
224.39 | -5.266e-06
224.51 | 7.3355e-05
224.63 | 4.6914e-05
224.74 | 0.00011409
224.86 | 8.2556e-05
224.98 | 3.5937e-05
225.09 | 8.1725e-05
225.21 | 0.00011689
225.32 | 4.0446e-05
225.44 | 9.5569e-05
225.56 | 1.0006e-05
225.67 | 7.9609e-05
225.79 | 9.4032e-05
225.91 | 0.00013292
226.02 | 3.9175e-05
226.14 | 7.7934e-05
226.25 | 0.0001066
226.37 | 9.6409e-05
226.49 | 1.918e-05
226.6 | 5.7241e-05
226.72 | -5.6947e-05
226.83 | 3.305e-05
226.95 | 7.0463e-05
227.07 | 0
227.18 | 8.3713e-05
227.3 | 9.7189e-05
227.42 | 0.00010591
227.53 | 5.0396e-05
227.65 | 6.3808e-05
227.76 | 4.5347e-05
227.88 | 7.6708e-05
228 | 0.00011225
228.11 | 1.3402e-05
228.23 | -4.4453e-06
228.34 | 0.00010172
228.46 | 2.6399e-05
228.58 | 7.8794e-05
228.69 | 5.2267e-05
228.81 | 0.00010834
228.92 | 6.8996e-05
229.04 | 0.00010725
229.16 | 0.0001494
229.27 | 7.2198e-05
229.39 | 7.6053e-05
229.5 | 9.2494e-05
229.62 | 8.7874e-05
229.74 | 4.165e-05
229.85 | 0.00010363
229.97 | 9.0758e-05
230.08 | 2.8738e-05
230.2 | 9.3984e-05
230.32 | 7.7289e-05
230.43 | 0.0001498
230.55 | 2.0153e-05
230.66 | 0.00010034
230.78 | 0.00011586
230.89 | 5.9649e-05
231.01 | 0.00011876
231.13 | 8.2751e-05
231.24 | 0.00010984
231.36 | 0.00010547
231.47 | 0.00014002
231.59 | 0.00013167
231.71 | 4.6272e-05
231.82 | 6.1426e-05
231.94 | 0.00012997
232.05 | 0.00010279
232.17 | 8.7203e-05
232.28 | 8.6858e-05
232.4 | 9.7805e-05
232.52 | 6.7454e-05
232.63 | 0.00011946
232.75 | 0.00011529
232.86 | 0.00011116
232.98 | 9.229e-05
233.09 | 0.00016552
233.21 | 8.0625e-05
233.33 | 2.1912e-05
233.44 | 6.9154e-05
233.56 | 0.00011244
233.67 | 9.0376e-05
233.79 | 0.00013689
233.9 | 5.7442e-05
234.02 | 0.00012522
234.13 | 5.7056e-05
234.25 | 0.00010308
234.37 | 8.1481e-05
234.48 | 5.2966e-05
234.6 | 6.6868e-05
234.71 | 0.0001298
234.83 | 5.9446e-05
234.94 | 8.3651e-05
235.06 | 9.3815e-05
235.17 | 0.00014896
235.29 | 5.5271e-05
235.41 | 0.0001481
235.52 | 8.9281e-05
235.64 | 0.00015751
235.75 | 8.8775e-05
235.87 | 0.00011917
235.98 | 0.00015278
236.1 | 0.00011174
236.21 | 0.00018237
236.33 | 8.4204e-05
236.44 | 0.00012764
236.56 | 0.00012395
236.68 | 8.3542e-05
236.79 | 7.9995e-05
236.91 | 4.9879e-05
237.02 | 0.00011942
237.14 | 0.00012245
237.25 | 0.00012879
237.37 | 5.2718e-05
237.48 | 0.00018737
237.6 | 0.00013122
237.71 | 0.00013425
237.83 | 0.00012419
237.94 | 0.00013698
238.06 | 0.00012697
238.17 | 0.00019498
238.29 | 0.00019787
238.4 | 0.00017484
238.52 | 0.00018419
238.64 | 8.709e-05
238.75 | 0.00024795
238.87 | 0.00018322
238.98 | 0.00014438
239.1 | 0.00019218
239.21 | 0.00020463
239.33 | 0.00022665
239.44 | 0.00025492
239.56 | 0.00015908
239.67 | 0.00020649
239.79 | 0.00020616
239.9 | 0.00021536
240.02 | 0.00021502
240.13 | 0.00022419
240.25 | 0.0002081
240.36 | 0.00027079
240.48 | 0.00031133
240.59 | 0.00026069
240.71 | 0.00026983
240.82 | 0.00031339
240.94 | 0.00033493
241.05 | 0.00029703
241.17 | 0.00025297
241.28 | 0.00036506
241.4 | 0.00037092
241.51 | 0.00037988
241.63 | 0.00037325
241.74 | 0.00033865
241.86 | 0.0003538
241.97 | 0.00035341
242.09 | 0.00037161
242.2 | 0.00047024
242.32 | 0.00046362
242.43 | 0.00044155
242.55 | 0.00046272
242.66 | 0.00041609
242.78 | 0.00052356
242.89 | 0.00045241
243.01 | 0.00056285
243.12 | 0.00059011
243.24 | 0.0005068
243.35 | 0.0005893
243.47 | 0.00076682
243.58 | 0.00069888
243.7 | 0.00074748
243.81 | 0.00075632
243.93 | 0.0008048
244.04 | 0.00080739
244.16 | 0.0010116
244.27 | 0.00090119
244.38 | 0.001029
244.5 | 0.001108
244.61 | 0.0011412
244.73 | 0.0011803
244.84 | 0.0012652
244.96 | 0.0013804
245.07 | 0.0014131
245.19 | 0.0015828
245.3 | 0.0016857
245.42 | 0.0018037
245.53 | 0.0017694
245.65 | 0.0019903
245.76 | 0.0020832
245.88 | 0.0021912
245.99 | 0.0023238
246.11 | 0.0025686
246.22 | 0.0027134
246.33 | 0.0028732
246.45 | 0.00313
246.56 | 0.0033347
246.68 | 0.0034426
246.79 | 0.0036046
246.91 | 0.0037521
247.02 | 0.0039575
247.14 | 0.0044627
247.25 | 0.0046989
247.37 | 0.0049533
247.48 | 0.0052747
247.59 | 0.0055634
247.71 | 0.0059485
247.82 | 0.0064811
247.94 | 0.0068126
248.05 | 0.0074209
248.17 | 0.0077068
248.28 | 0.0081443
248.4 | 0.0085334
248.51 | 0.0090459
248.62 | 0.009649
248.74 | 0.010101
248.85 | 0.010785
248.97 | 0.011397
249.08 | 0.012031
249.2 | 0.012716
249.31 | 0.013499
249.43 | 0.014237
249.54 | 0.014914
249.65 | 0.01588
249.77 | 0.016312
249.88 | 0.017257
250 | 0.018326
250.11 | 0.019338
250.23 | 0.02006
250.34 | 0.021207
250.45 | 0.022041
250.57 | 0.02366
250.68 | 0.024331
250.8 | 0.025531
250.91 | 0.026772
251.03 | 0.027957
251.14 | 0.029324
251.25 | 0.030494
251.37 | 0.031807
251.48 | 0.033018
251.6 | 0.03404
251.71 | 0.035693
251.82 | 0.037395
251.94 | 0.038966
252.05 | 0.040338
252.17 | 0.041752
252.28 | 0.04325
252.4 | 0.04487
252.51 | 0.046405
252.62 | 0.047875
252.74 | 0.04971
252.85 | 0.051462
252.97 | 0.052839
253.08 | 0.055093
253.19 | 0.056478
253.31 | 0.058207
253.42 | 0.059971
253.54 | 0.061775
253.65 | 0.063944
253.76 | 0.065655
253.88 | 0.067014
253.99 | 0.069377
254.11 | 0.070635
254.22 | 0.073022
254.33 | 0.074605
254.45 | 0.076249
254.56 | 0.07819
254.68 | 0.080095
254.79 | 0.081949
254.9 | 0.083532
255.02 | 0.08505
255.13 | 0.086718
255.25 | 0.088321
255.36 | 0.090414
255.47 | 0.092003
255.59 | 0.093452
255.7 | 0.095789
255.81 | 0.096936
255.93 | 0.098573
256.04 | 0.099804
256.16 | 0.10099
256.27 | 0.10309
256.38 | 0.10394
256.5 | 0.1055
256.61 | 0.1068
256.72 | 0.1081
256.84 | 0.1091
256.95 | 0.11076
257.07 | 0.1117
257.18 | 0.1125
257.29 | 0.11421
257.41 | 0.11469
257.52 | 0.11501
257.63 | 0.11617
257.75 | 0.1176
257.86 | 0.1177
257.98 | 0.11837
258.09 | 0.11898
258.2 | 0.11989
258.32 | 0.12023
258.43 | 0.12096
258.54 | 0.12137
258.66 | 0.12132
258.77 | 0.12134
258.88 | 0.12247
259 | 0.12198
259.11 | 0.12285
259.23 | 0.12257
259.34 | 0.12264
259.45 | 0.12288
259.57 | 0.12265
259.68 | 0.12219
259.79 | 0.12156
259.91 | 0.12214
260.02 | 0.12191
260.13 | 0.12086
260.25 | 0.12033
260.36 | 0.12013
260.47 | 0.11911
260.59 | 0.11849
260.7 | 0.1179
260.81 | 0.11822
260.93 | 0.11692
261.04 | 0.11598
261.15 | 0.11505
261.27 | 0.11375
261.38 | 0.11308
261.49 | 0.11232
261.61 | 0.11178
261.72 | 0.11006
261.83 | 0.10958
261.95 | 0.10788
262.06 | 0.1067
262.17 | 0.10535
262.29 | 0.10467
262.4 | 0.10332
262.51 | 0.10192
262.63 | 0.10096
262.74 | 0.099923
262.85 | 0.09841
262.97 | 0.097678
263.08 | 0.095443
263.19 | 0.094259
263.31 | 0.093526
263.42 | 0.091535
263.53 | 0.091331
263.65 | 0.089392
263.76 | 0.088664
263.87 | 0.087164
263.99 | 0.085483
264.1 | 0.084608
264.21 | 0.083783
264.33 | 0.081697
264.44 | 0.080984
264.55 | 0.079406
264.67 | 0.078496
264.78 | 0.077281
264.89 | 0.075585
265.01 | 0.074834
265.12 | 0.073821
265.23 | 0.072662
265.34 | 0.071724
265.46 | 0.070642
265.57 | 0.069676
265.68 | 0.068738
265.8 | 0.068018
265.91 | 0.066957
266.02 | 0.066096
266.14 | 0.065328
266.25 | 0.064547
266.36 | 0.063294
266.47 | 0.062542
266.59 | 0.061988
266.7 | 0.061533
266.81 | 0.060711
266.93 | 0.059894
267.04 | 0.059764
267.15 | 0.058985
267.26 | 0.058656
267.38 | 0.058065
267.49 | 0.057454
267.6 | 0.057388
267.72 | 0.056699
267.83 | 0.056576
267.94 | 0.05589
268.06 | 0.05636
268.17 | 0.055572
268.28 | 0.055735
268.39 | 0.055299
268.51 | 0.05535
268.62 | 0.055231
268.73 | 0.055346
268.84 | 0.055921
268.96 | 0.055662
269.07 | 0.056424
269.18 | 0.056969
269.3 | 0.057348
269.41 | 0.057654
269.52 | 0.058294
269.63 | 0.059179
269.75 | 0.06002
269.86 | 0.060533
269.97 | 0.061586
270.08 | 0.062249
270.2 | 0.063333
270.31 | 0.06425
270.42 | 0.065134
270.54 | 0.067074
270.65 | 0.068027
270.76 | 0.070016
270.87 | 0.071218
270.99 | 0.07287
271.1 | 0.074544
271.21 | 0.076128
271.32 | 0.077912
271.44 | 0.079625
271.55 | 0.081722
271.66 | 0.083561
271.77 | 0.085666
271.89 | 0.08764
272 | 0.089994
272.11 | 0.092399
272.22 | 0.094457
272.34 | 0.096681
272.45 | 0.099225
272.56 | 0.10184
272.67 | 0.10417
272.79 | 0.10698
272.9 | 0.10941
273.01 | 0.11218
273.12 | 0.11446
273.24 | 0.11772
273.35 | 0.11941
273.46 | 0.12232
273.57 | 0.12552
273.69 | 0.12825
273.8 | 0.13062
273.91 | 0.1328
274.02 | 0.13531
274.13 | 0.13791
274.25 | 0.13934
274.36 | 0.14224
274.47 | 0.14425
274.58 | 0.14712
274.7 | 0.14853
274.81 | 0.15083
274.92 | 0.15209
275.03 | 0.15437
275.15 | 0.15587
275.26 | 0.15811
275.37 | 0.15913
275.48 | 0.16097
275.59 | 0.16206
275.71 | 0.16308
275.82 | 0.1645
275.93 | 0.16587
276.04 | 0.16593
276.16 | 0.16622
276.27 | 0.16811
276.38 | 0.16757
276.49 | 0.16757
276.6 | 0.16814
276.72 | 0.16799
276.83 | 0.16715
276.94 | 0.16708
277.05 | 0.16733
277.16 | 0.16678
277.28 | 0.1661
277.39 | 0.16579
277.5 | 0.16488
277.61 | 0.16351
277.73 | 0.16307
277.84 | 0.16086
277.95 | 0.16009
278.06 | 0.15871
278.17 | 0.15694
278.29 | 0.15526
278.4 | 0.15384
278.51 | 0.15233
278.62 | 0.15031
278.73 | 0.14836
278.85 | 0.14642
278.96 | 0.14518
279.07 | 0.14288
279.18 | 0.14048
279.29 | 0.13792
279.4 | 0.13645
279.52 | 0.13391
279.63 | 0.13099
279.74 | 0.12889
279.85 | 0.12725
279.96 | 0.12432
280.08 | 0.12268
280.19 | 0.11995
280.3 | 0.1178
280.41 | 0.11477
280.52 | 0.11252
280.64 | 0.11068
280.75 | 0.10841
280.86 | 0.10589
280.97 | 0.10377
281.08 | 0.10153
281.19 | 0.098587
281.31 | 0.097021
281.42 | 0.094339
281.53 | 0.092099
281.64 | 0.090278
281.75 | 0.088527
281.86 | 0.086477
281.98 | 0.083994
282.09 | 0.082241
282.2 | 0.080332
282.31 | 0.078194
282.42 | 0.076635
282.53 | 0.074482
282.65 | 0.072564
282.76 | 0.0706
282.87 | 0.068847
282.98 | 0.067146
283.09 | 0.065781
283.2 | 0.064061
283.32 | 0.062463
283.43 | 0.060842
283.54 | 0.059842
283.65 | 0.057571
283.76 | 0.056224
283.87 | 0.055102
283.99 | 0.053797
284.1 | 0.052473
284.21 | 0.050878
284.32 | 0.04963
284.43 | 0.048332
284.54 | 0.047132
284.65 | 0.045792
284.77 | 0.044662
284.88 | 0.043798
284.99 | 0.042662
285.1 | 0.041624
285.21 | 0.040219
285.32 | 0.039491
285.43 | 0.038408
285.55 | 0.037353
285.66 | 0.036111
285.77 | 0.03541
285.88 | 0.034442
285.99 | 0.033488
286.1 | 0.032852
286.21 | 0.031797
286.33 | 0.031422
286.44 | 0.030517
286.55 | 0.029932
286.66 | 0.028901
286.77 | 0.028219
286.88 | 0.02767
286.99 | 0.02696
287.1 | 0.026563
287.22 | 0.025755
287.33 | 0.025252
287.44 | 0.024496
287.55 | 0.023979
287.66 | 0.02352
287.77 | 0.022791
287.88 | 0.022265
287.99 | 0.021906
288.11 | 0.021309
288.22 | 0.020896
288.33 | 0.020232
288.44 | 0.019855
288.55 | 0.019359
288.66 | 0.019085
288.77 | 0.018596
288.88 | 0.018139
288.99 | 0.017914
289.11 | 0.01744
289.22 | 0.016935
289.33 | 0.016678
289.44 | 0.016298
289.55 | 0.016108
289.66 | 0.015365
289.77 | 0.015271
289.88 | 0.014713
289.99 | 0.014585
290.11 | 0.014225
290.22 | 0.013923
290.33 | 0.013407
290.44 | 0.013245
290.55 | 0.01309
290.66 | 0.012789
290.77 | 0.012595
290.88 | 0.012405
290.99 | 0.011987
291.1 | 0.011671
291.21 | 0.01166
291.33 | 0.011369
291.44 | 0.010915
291.55 | 0.010756
291.66 | 0.010602
291.77 | 0.010308
291.88 | 0.010346
291.99 | 0.0098772
292.1 | 0.0098658
292.21 | 0.0096992
292.32 | 0.0095027
292.43 | 0.0091567
292.55 | 0.0089203
292.66 | 0.0087423
292.77 | 0.0088224
292.88 | 0.0084237
292.99 | 0.0083706
293.1 | 0.0081205
293.21 | 0.0078884
293.32 | 0.0077411
293.43 | 0.0075627
293.54 | 0.0075894
293.65 | 0.0072241
293.76 | 0.0071511
293.87 | 0.0070644
293.99 | 0.0069446
294.1 | 0.0066907
294.21 | 0.0066052
294.32 | 0.0064999
294.43 | 0.0064568
294.54 | 0.00609
294.65 | 0.0060508
294.76 | 0.0059885
294.87 | 0.0058224
294.98 | 0.0056702
295.09 | 0.0056517
295.2 | 0.0054494
295.31 | 0.0053674
295.42 | 0.00519
295.53 | 0.0051883
295.64 | 0.0049882
295.76 | 0.0049566
295.87 | 0.0049191
295.98 | 0.0046395
296.09 | 0.004667
296.2 | 0.0044579
296.31 | 0.0044419
296.42 | 0.0042892
296.53 | 0.0042698
296.64 | 0.0041103
296.75 | 0.0041334
296.86 | 0.0039952
296.97 | 0.0037824
297.08 | 0.0038674
297.19 | 0.0037563
297.3 | 0.0036752
297.41 | 0.0035408
297.52 | 0.0034405
297.63 | 0.0033604
297.74 | 0.0033511
297.85 | 0.0032287
297.96 | 0.0031437
298.07 | 0.0030648
298.19 | 0.003127
298.3 | 0.0029195
298.41 | 0.0029134
298.52 | 0.0028977
298.63 | 0.0027209
298.74 | 0.0026239
298.85 | 0.0027112
298.96 | 0.0026279
299.07 | 0.0025872
299.18 | 0.0024635
299.29 | 0.0025055
299.4 | 0.0024282
299.51 | 0.002346
299.62 | 0.0022559
299.73 | 0.0022789
299.84 | 0.0021989
299.95 | 0.0021785
300.06 | 0.0021526
31.289
0

## Emission spectrum 2 rows
|  | Interpolation over germicidal range (200 to 300) |  |  |  | Calculating rated lamp energy (RLE)
Spectrometer Measurement |  | x | x1, x2 | y1, y2 | y |  | interpolated | RLE
200.11 | -2.53E-04 | 200.11 | 200.11 | -2.53E-04 | -2.53E-04 | 200.11 | -2.53E-04 | -8.4658e-05 | -0.050628
200.23 | 9.64e-05 |  |  |  |  | 201 | -0.00012755 | -4.268e-05 | -0.025638
200.35 | 0.000134 | 201 | 200.94 | -0.000237 | -0.00012755 | 202 | -0.000134 | -4.4839e-05 | -0.027068
200.46 | -0.00017 |  | 201.06 | -1.81e-05 |  | 203 | 0 | 0 | 0
200.58 | 0.000262 | 202 | 202 | -0.000134 | -0.000134 | 204 | 9.2717e-05 | 3.1025e-05 | 0.018914
200.7 | 9.29e-05 |  | 202.12 | -6.66e-05 |  | 205 | -0.00020308 | -6.7955e-05 | -0.041632
200.82 | 0.000313 | 203 | 202.95 | 0 | 0 | 206 | 0.00012492 | 4.1799e-05 | 0.025733
200.94 | -0.000237 |  | 203.06 | 0 |  | 207 | -0.00014805 | -4.9541e-05 | -0.030647
201.06 | -1.81e-05 | 204 | 203.89 | 0.00016 | 9.2717e-05 | 208 | 3.4455e-05 | 1.1529e-05 | 0.0071665
201.17 | 0 |  | 204.01 | 8.66e-05 |  | 209 | -5.525e-05 | -1.8488e-05 | -0.011547
201.29 | 3.55e-05 | 205 | 204.95 | -0.00031 | -0.00020308 | 210 | 5.89e-05 | 1.9709e-05 | 0.012369
201.41 | 0.000158 |  | 205.07 | -5.34e-05 |  | 211 | 6.8892e-05 | 2.3052e-05 | 0.014536
201.53 | 0.000122 | 206 | 205.89 | -0.000151 | 0.00012492 | 212 | 8.0321e-06 | 2.6877e-06 | 0.0017028
201.65 | 0.000501 |  | 206.01 | 0.00015 |  | 213 | 1.89e-05 | 6.3242e-06 | 0.0040257
201.77 | 0.000171 | 207 | 206.95 | -0.000165 | -0.00014805 | 214 | 6.55e-05 | 2.1917e-05 | 0.014017
201.88 | 0.00022 |  | 207.07 | 0.000105 |  | 215 | 0.0001765 | 5.906e-05 | 0.037948
202 | -0.000134 | 208 | 207.9 | 0.000157 | 3.4455e-05 | 216 | -6.875e-05 | -2.3005e-05 | -0.01485
202.12 | -6.66e-05 |  | 208.01 | 2.22e-05 |  | 217 | 7.8545e-06 | 2.6283e-06 | 0.0017044
202.24 | 0.000379 | 209 | 208.95 | -0.000147 | -5.525e-05 | 218 | 0.00011363 | 3.8024e-05 | 0.024772
202.36 | -0.000147 |  | 209.07 | 7.32e-05 |  | 219 | 2.1525e-05 | 7.2026e-06 | 0.004714
202.47 | 0.000194 | 210 | 209.89 | 5.01e-05 | 5.89e-05 | 220 | -7.0764e-05 | -2.3679e-05 | -0.015568
202.59 | 9.62e-05 |  | 210.01 | 5.97e-05 |  | 221 | 1.015e-05 | 3.3964e-06 | 0.0022431
202.71 | 0 | 211 | 210.95 | -5.69e-05 | 6.8892e-05 | 222 | 0.00014295 | 4.7835e-05 | 0.031736
202.83 | 0.000142 |  | 211.07 | 0.000245 |  | 223 | -7.27e-05 | -2.4327e-05 | -0.016212
202.95 | 0 | 212 | 211.89 | 0 | 8.0321e-06 | 224 | 5.0136e-05 | 1.6776e-05 | 0.011231
203.06 | 0 |  | 212.01 | 7.2e-05 |  | 225 | 6.7173e-05 | 2.2477e-05 | 0.015114
203.18 | 0.000169 | 213 | 212.95 | 7.74e-05 | 1.89e-05 | 226 | 1.6109e-05 | 5.3904e-06 | 0.0036407
203.3 | 1.52e-05 |  | 213.06 | -5.13e-05 |  | 227 | 9.105e-05 | 3.0467e-05 | 0.020668
203.42 | 0.000362 | 214 | 214 | 6.55e-05 | 6.55e-05 | 228 | 0.000144 | 4.8185e-05 | 0.032832
203.54 | -0.000164 |  | 214.12 | 4.89e-05 |  | 229 | 2.8593e-05 | 9.5678e-06 | 0.0065479
203.65 | -4.44e-05 | 215 | 214.94 | 0.000244 | 0.0001765 | 230 | 7.7473e-05 | 2.5924e-05 | 0.017819
203.77 | 0.000161 |  | 215.06 | 0.000109 |  | 231 | 3.235e-05 | 1.0825e-05 | 0.0074728
203.89 | 0.00016 | 216 | 215.99 | -7.5e-05 | -6.875e-05 | 232 | 7.4773e-05 | 2.502e-05 | 0.017347
204.01 | 8.66e-05 |  | 216.11 | 0 |  | 233 | 0.00010725 | 3.5889e-05 | 0.02499
204.13 | 1.43e-05 | 217 | 216.93 | 2.16e-05 | 7.8545e-06 | 234 | 3.76e-05 | 1.2582e-05 | 0.0087984
204.24 | 0.000298 |  | 217.04 | 0 |  | 235 | 4.355e-05 | 1.4573e-05 | 0.010234
204.36 | 0.000183 | 218 | 217.98 | 0.000124 | 0.00011363 | 236 | 6.8983e-05 | 2.3083e-05 | 0.01628
204.48 | -0.000209 |  | 218.1 | 6.18e-05 |  | 237 | 2.3573e-05 | 7.8878e-06 | 0.0055867
204.6 | 0.000124 | 219 | 218.91 | 4.65e-05 | 2.1525e-05 | 238 | 0.0001138 | 3.8079e-05 | 0.027084
204.72 | -0.000137 |  | 219.03 | 1.32e-05 |  | 239 | 7.8667e-05 | 2.6323e-05 | 0.018801
204.83 | 8.15e-05 | 220 | 219.97 | -0.000121 | -7.0764e-05 | 240 | 0.00013517 | 4.5229e-05 | 0.03244
204.95 | -0.00031 |  | 220.08 | 6.32e-05 |  | 241 | 0.00015491 | 5.1835e-05 | 0.037333
205.07 | -5.34e-05 | 221 | 220.9 | -3.06e-05 | 1.015e-05 | 242 | 0.00025525 | 8.5411e-05 | 0.06177
205.19 | -0.000132 |  | 221.02 | 1.83e-05 |  | 243 | 0.00033717 | 0.00011282 | 0.081932
205.3 | 2.62e-05 | 222 | 221.95 | 0.000199 | 0.00014295 | 244 | 0.00040145 | 0.00013433 | 0.097955
205.42 | 0.000273 |  | 222.06 | 7.57e-05 |  | 245 | 0.00066873 | 0.00022377 | 0.16384
205.54 | 2.58e-05 | 223 | 223 | -7.27e-05 | -7.27e-05 | 246 | 0.0011 | 0.00036808 | 0.2706
205.66 | -0.000102 |  | 223.11 | 8.34e-05 |  | 247 | 0.0018745 | 0.00062725 | 0.46301
205.78 | 0 | 224 | 223.93 | 5.37e-05 | 5.0136e-05 | 248 | 0.0031782 | 0.0010635 | 0.78819
205.89 | -0.000151 |  | 224.04 | 4.81e-05 |  | 249 | 0.0053691 | 0.0017966 | 1.3369
206.01 | 0.00015 | 225 | 224.98 | 8.21e-05 | 6.7173e-05 | 250 | 0.00847 | 0.0028342 | 2.1175
206.13 | -2.48e-05 |  | 225.09 | 0 |  | 251 | 0.01325 | 0.0044337 | 3.3258
206.25 | -2.46e-05 | 226 | 225.91 | 8.86e-05 | 1.6109e-05 | 252 | 0.019482 | 0.0065189 | 4.9094
206.37 | 7.33e-05 |  | 226.02 | 0 |  | 253 | 0.0274 | 0.0091685 | 6.9322
206.48 | -0.000158 | 227 | 226.95 | 8.93e-05 | 9.105e-05 | 254 | 0.036592 | 0.012244 | 9.2943
206.6 | 0.000145 |  | 227.07 | 9.35e-05 |  | 255 | 0.046433 | 0.015537 | 11.8405
206.72 | -4.79e-05 | 228 | 228 | 0.000144 | 0.000144 | 256 | 0.057236 | 0.019152 | 14.6525
206.84 | 0.000131 |  | 228.11 | 9.38e-05 |  | 257 | 0.066483 | 0.022246 | 17.0862
206.95 | -0.000165 | 229 | 228.92 | -8.62e-06 | 2.8593e-05 | 258 | 0.074927 | 0.025072 | 19.3312
207.07 | 0.000105 |  | 229.04 | 4.72e-05 |  | 259 | 0.0814 | 0.027238 | 21.0826
207.19 | 0.000198 | 230 | 229.97 | 6.19e-05 | 7.7473e-05 | 260 | 0.084636 | 0.028321 | 22.0055
207.31 | 0.000139 |  | 230.08 | 0.000119 |  | 261 | 0.085373 | 0.028567 | 22.2823
207.42 | 8.04e-05 | 231 | 230.89 | 8.35e-05 | 3.235e-05 | 262 | 0.083245 | 0.027855 | 21.8103
207.54 | -0.000103 |  | 231.01 | 2.77e-05 |  | 263 | 0.078682 | 0.026328 | 20.6933
207.66 | 0.000181 | 232 | 231.94 | 2.29e-05 | 7.4773e-05 | 264 | 0.072527 | 0.024269 | 19.1472
207.78 | 2.25e-05 |  | 232.05 | 0.000118 |  | 265 | 0.066433 | 0.02223 | 17.6048
207.9 | 0.000157 | 233 | 232.98 | 0.000118 | 0.00010725 | 266 | 0.060273 | 0.020168 | 16.0325
208.01 | 2.22e-05 |  | 233.09 | 5.89e-05 |  | 267 | 0.0555 | 0.018571 | 14.8185
208.13 | 0.000232 | 234 | 233.9 | 6.46e-05 | 3.76e-05 | 268 | 0.05265 | 0.017618 | 14.1102
208.25 | 0.000164 |  | 234.02 | 3.22e-05 |  | 269 | 0.053109 | 0.017771 | 14.2863
208.37 | -4.35e-05 | 235 | 234.94 | 5.93e-05 | 4.355e-05 | 270 | 0.057718 | 0.019313 | 15.5839
208.48 | 0.000151 |  | 235.06 | 2.78e-05 |  | 271 | 0.068245 | 0.022836 | 18.4945
208.6 | -2.15e-05 | 236 | 235.98 | 6.45e-05 | 6.8983e-05 | 272 | 0.0841 | 0.028141 | 22.8752
208.72 | 4.27e-05 |  | 236.1 | 9.14e-05 |  | 273 | 0.10473 | 0.035043 | 28.5905
208.84 | 5.3e-05 | 237 | 236.91 | 6.98e-05 | 2.3573e-05 | 274 | 0.12682 | 0.042435 | 34.7482
208.95 | -0.000147 |  | 237.02 | 1.33e-05 |  | 275 | 0.14745 | 0.049341 | 40.55
209.07 | 7.32e-05 | 238 | 237.94 | 9.46e-05 | 0.0001138 | 276 | 0.16 | 0.053539 | 44.16
209.19 | 5.2e-05 |  | 238.06 | 0.000133 |  | 277 | 0.163 | 0.054542 | 45.151
209.31 | 0.000165 | 239 | 238.98 | 6.74e-05 | 7.8667e-05 | 278 | 0.15609 | 0.052231 | 43.3933
209.42 | -0.000123 |  | 239.1 | 0.000135 |  | 279 | 0.14127 | 0.047272 | 39.4151
209.54 | -2.04e-05 | 240 | 239.9 | 0.000146 | 0.00013517 | 280 | 0.122 | 0.040823 | 34.16
209.66 | -7.1e-05 |  | 240.02 | 0.000133 |  | 281 | 0.10245 | 0.034283 | 28.7897
209.78 | -5.04e-05 | 241 | 240.94 | 0.00015 | 0.00015491 | 282 | 0.083664 | 0.027995 | 23.5931
209.89 | 5.01e-05 |  | 241.05 | 0.000159 |  | 283 | 0.066909 | 0.022389 | 18.9353
210.01 | 5.97e-05 | 242 | 241.97 | 0.000267 | 0.00025525 | 284 | 0.053073 | 0.017759 | 15.0727
210.13 | 3.96e-05 |  | 242.09 | 0.00022 |  | 285 | 0.041882 | 0.014014 | 11.9363
210.25 | -1.97e-05 | 243 | 242.89 | 0.000295 | 0.00033717 | 286 | 0.033218 | 0.011115 | 9.5004
210.36 | 5.87e-05 |  | 243.01 | 0.000341 |  | 287 | 0.026645 | 0.008916 | 7.6472
210.48 | -0.000107 | 244 | 243.93 | 0.000355 | 0.00040145 | 288 | 0.021275 | 0.007119 | 6.1272
210.6 | 0.000116 |  | 244.04 | 0.000428 |  | 289 | 0.017458 | 0.0058418 | 5.0455
210.72 | 9.6e-05 | 245 | 244.96 | 0.00064 | 0.00066873 | 290 | 0.014083 | 0.0047125 | 4.0842
210.83 | 0.000277 |  | 245.07 | 0.000719 |  | 291 | 0.011773 | 0.0039393 | 3.4259
210.95 | -5.69e-05 | 246 | 245.99 | 0.00109 | 0.0011 | 292 | 0.0097073 | 0.0032482 | 2.8345
211.07 | 0.000245 |  | 246.11 | 0.00121 |  | 293 | 0.0080627 | 0.0026979 | 2.3624
211.19 | 0.000113 | 247 | 246.91 | 0.00176 | 0.0018745 | 294 | 0.0068009 | 0.0022757 | 1.9995
211.3 | 3.73e-05 |  | 247.02 | 0.0019 |  | 295 | 0.0055782 | 0.0018666 | 1.6456
211.42 | 0.000102 | 248 | 247.94 | 0.00314 | 0.0031782 | 296 | 0.0044918 | 0.001503 | 1.3296
211.54 | -9.22e-06 |  | 248.05 | 0.00321 |  | 297 | 0.00372 | 0.0012448 | 1.1048
211.66 | -1.83e-05 | 249 | 248.97 | 0.00529 | 0.0053691 | 298 | 0.0030482 | 0.00102 | 0.90836
211.77 | 9.11e-05 |  | 249.08 | 0.00558 |  | 299 | 0.0024755 | 0.00082833 | 0.74016
211.89 | 0 | 250 | 250 | 0.00847 | 0.00847 | 300 | 0.0020536 | 0.00068718 | 0.61609
212.01 | 7.2e-05 |  | 250.11 | 0.00894 |  |  | 2.9885 | 0.99992 | 811.7295
212.13 | 3.58e-05 | 251 | 250.91 | 0.0128 | 0.01325 |  |  | Weighted Avg: | 271.618
212.24 | 0.000169 |  | 251.03 | 0.0134
212.36 | -7.97e-05 | 252 | 251.94 | 0.0191 | 0.019482
212.48 | 9.68e-05 |  | 252.05 | 0.0198
212.59 | 3.5e-05 | 253 | 252.97 | 0.0271 | 0.0274
212.71 | -3.48e-05 |  | 253.08 | 0.0282
212.83 | 0.00013 | 254 | 253.99 | 0.0365 | 0.036592
212.95 | 7.74e-05 |  | 254.11 | 0.0376
213.06 | -5.13e-05 | 255 | 254.9 | 0.0456 | 0.046433
213.18 | 5.1e-05 |  | 255.02 | 0.0466
213.3 | 0.000211 | 256 | 255.93 | 0.0566 | 0.057236
213.42 | 1.68e-05 |  | 256.04 | 0.0576
213.53 | 5.85e-05 | 257 | 256.95 | 0.0659 | 0.066483
213.65 | 1.66e-05 |  | 257.07 | 0.0673
213.77 | -3.31e-05 | 258 | 257.98 | 0.0748 | 0.074927
213.88 | 0.000107 |  | 258.09 | 0.0755
214 | 6.55e-05 | 259 | 259 | 0.0814 | 0.0814
214.12 | 4.89e-05 |  | 259.11 | 0.0815
214.24 | -7.29e-05 | 260 | 259.91 | 0.0848 | 0.084636
214.35 | 4.84e-05 |  | 260.02 | 0.0846
214.47 | -8.02e-06 | 261 | 260.93 | 0.0855 | 0.085373
214.59 | -3.19e-05 |  | 261.04 | 0.0853
214.7 | 8.73e-05 | 262 | 261.95 | 0.0832 | 0.083245
214.82 | 0.000111 |  | 262.06 | 0.0833
214.94 | 0.000244 | 263 | 262.97 | 0.0789 | 0.078682
215.06 | 0.000109 |  | 263.08 | 0.0781
215.17 | 0.000187 | 264 | 263.99 | 0.0726 | 0.072527
215.29 | 5.42e-05 |  | 264.1 | 0.0718
215.41 | 7.7e-06 | 265 | 264.89 | 0.0668 | 0.066433
215.52 | 0.000115 |  | 265.01 | 0.0664
215.64 | 3.81e-05 | 266 | 265.91 | 0.0606 | 0.060273
215.76 | -1.52e-05 |  | 266.02 | 0.0602
215.87 | 1.51e-05 | 267 | 266.93 | 0.0555 | 0.0555
215.99 | -7.5e-05 |  | 267.04 | 0.0555
216.11 | 0 | 268 | 267.94 | 0.0528 | 0.05265
216.23 | -5.94e-05 |  | 268.06 | 0.0525
216.34 | 9.6e-05 | 269 | 268.96 | 0.053 | 0.053109
216.46 | 3.67e-05 |  | 269.07 | 0.0533
216.58 | -2.19e-05 | 270 | 269.97 | 0.0575 | 0.057718
216.69 | 2.18e-05 |  | 270.08 | 0.0583
216.81 | 6.52e-05 | 271 | 270.99 | 0.0681 | 0.068245
216.93 | 2.16e-05 |  | 271.1 | 0.0697
217.04 | 0 | 272 | 272 | 0.0841 | 0.0841
217.16 | 0.000128 |  | 272.11 | 0.0861
217.28 | 7.1e-05 | 273 | 272.9 | 0.102 | 0.10473
217.4 | -1.41e-05 |  | 273.01 | 0.105
217.51 | 7.74e-05 | 274 | 273.91 | 0.126 | 0.12682
217.63 | 8.4e-05 |  | 274.02 | 0.127
217.75 | 7.67e-05 | 275 | 274.92 | 0.146 | 0.14745
217.86 | 0.000201 |  | 275.03 | 0.148
217.98 | 0.000124 | 276 | 275.93 | 0.16 | 0.16
218.1 | 6.18e-05 |  | 276.04 | 0.16
218.21 | 2.05e-05 | 277 | 276.94 | 0.163 | 0.163
218.33 | 6.81e-05 |  | 277.05 | 0.163
218.45 | -0.000122 | 278 | 277.95 | 0.157 | 0.15609
218.56 | 7.41e-05 |  | 278.06 | 0.155
218.68 | 9.39e-05 | 279 | 278.96 | 0.142 | 0.14127
218.8 | 0.000113 |  | 279.07 | 0.14
218.91 | 4.65e-05 | 280 | 279.96 | 0.123 | 0.122
219.03 | 1.32e-05 |  | 280.08 | 0.12
219.15 | 0.000112 | 281 | 280.97 | 0.103 | 0.10245
219.26 | 0.000105 |  | 281.08 | 0.101
219.38 | 0.00013 | 282 | 281.98 | 0.0841 | 0.083664
219.5 | 1.94e-05 |  | 282.09 | 0.0817
219.62 | 5.8e-05 | 283 | 282.98 | 0.0672 | 0.066909
219.73 | 7.7e-05 |  | 283.09 | 0.0656
219.85 | 5.11e-05 | 284 | 283.99 | 0.0532 | 0.053073
219.97 | -0.000121 |  | 284.1 | 0.0518
220.08 | 6.32e-05 | 285 | 284.99 | 0.042 | 0.041882
220.2 | 8.18e-05 |  | 285.1 | 0.0407
220.32 | 0.000169 | 286 | 285.99 | 0.0333 | 0.033218
220.43 | 0.000118 |  | 286.1 | 0.0324
220.55 | 4.96e-05 | 287 | 286.99 | 0.0267 | 0.026645
220.67 | -6.18e-05 |  | 287.1 | 0.0261
220.78 | -1.84e-05 | 288 | 287.99 | 0.0213 | 0.021275
220.9 | -3.06e-05 |  | 288.11 | 0.021
221.02 | 1.83e-05 | 289 | 288.99 | 0.0175 | 0.017458
221.13 | 9.69e-05 |  | 289.11 | 0.017
221.25 | 0 | 290 | 289.99 | 0.0141 | 0.014083
221.37 | 0.000216 |  | 290.11 | 0.0139
221.48 | 3.58e-05 | 291 | 290.99 | 0.0118 | 0.011773
221.6 | 0.000125 |  | 291.1 | 0.0115
221.71 | -1.77e-05 | 292 | 291.99 | 0.00973 | 0.0097073
221.83 | 7.06e-05 |  | 292.1 | 0.00948
221.95 | 0.000199 | 293 | 292.99 | 0.00809 | 0.0080627
222.06 | 7.57e-05 |  | 293.1 | 0.00779
222.18 | 5.8e-06 | 294 | 293.99 | 0.00683 | 0.0068009
222.3 | 0.000104 |  | 294.1 | 0.00651
222.41 | 4.02e-05 | 295 | 294.98 | 0.00562 | 0.0055782
222.53 | 7.99e-05 |  | 295.09 | 0.00539
222.65 | 5.68e-06 | 296 | 295.98 | 0.00449 | 0.0044918
222.76 | 5.08e-05 |  | 296.09 | 0.0045
222.88 | 0.000129 | 297 | 296.97 | 0.00375 | 0.00372
223 | -7.27e-05 |  | 297.08 | 0.00364
223.11 | 8.34e-05 | 298 | 297.96 | 0.00311 | 0.0030482
223.23 | 8.86e-05 |  | 298.07 | 0.00294
223.35 | 0.00011 | 299 | 298.96 | 0.00249 | 0.0024755
223.46 | -5.48e-06 |  | 299.07 | 0.00245
223.58 | -7.09e-05 | 300 | 299.95 | 0.00204 | 0.0020536
223.7 | 3.26e-05 |  | 300.06 | 0.00207
223.81 | 5.4e-05
223.93 | 5.37e-05
224.04 | 4.81e-05
224.16 | 1.06e-05
224.28 | 0.000122
224.39 | 7.9e-05
224.51 | 9.96e-05
224.63 | 4.69e-05
224.74 | 0.000114
224.86 | 7.74e-05
224.98 | 8.21e-05
225.09 | 0
225.21 | 1.52e-05
225.32 | 2.53e-05
225.44 | 0.000126
225.56 | 3.5e-05
225.67 | 4.48e-05
225.79 | 0.000114
225.91 | 8.86e-05
226.02 | 0
226.14 | 9.74e-05
226.25 | 7.75e-05
226.37 | 1.93e-05
226.49 | 6.23e-05
226.6 | 5.25e-05
226.72 | 3.32e-05
226.83 | 0.000146
226.95 | 8.93e-05
227.07 | 9.35e-05
227.18 | 0.000126
227.3 | 6.02e-05
227.42 | -0.000106
227.53 | 0.00011
227.65 | -5.47e-05
227.76 | 9.52e-05
227.88 | 3.61e-05
228 | 0.000144
228.11 | 9.38e-05
228.23 | 4.45e-05
228.34 | 6.63e-05
228.46 | 4.84e-05
228.58 | -8.75e-06
228.69 | 3.92e-05
228.81 | 8.23e-05
228.92 | -8.62e-06
229.04 | 4.72e-05
229.16 | 1.28e-05
229.27 | 0.000106
229.39 | 3.8e-05
229.5 | 7.15e-05
229.62 | 6.28e-05
229.74 | 8.33e-05
229.85 | 4.56e-05
229.97 | 6.19e-05
230.08 | 0.000119
230.2 | 8.17e-06
230.32 | 9.36e-05
230.43 | 4.45e-05
230.55 | 0.000117
230.66 | -4.01e-06
230.78 | 3.6e-05
230.89 | 8.35e-05
231.01 | 2.77e-05
231.13 | 7.49e-05
231.24 | 9.42e-05
231.36 | 8.98e-05
231.47 | 7.78e-05
231.59 | 5.81e-05
231.71 | 8.1e-05
231.82 | 5.37e-05
231.94 | 2.29e-05
232.05 | 0.000118
232.17 | 0
232.28 | -4.15e-05
232.4 | -3.39e-05
232.52 | 7.12e-05
232.63 | 5.97e-05
232.75 | 5.21e-05
232.86 | 6.3e-05
232.98 | 0.000118
233.09 | 5.89e-05
233.21 | 7.7e-05
233.33 | 7.3e-05
233.44 | 1.46e-05
233.56 | 1.81e-05
233.67 | 7.59e-05
233.79 | 5.4e-05
233.9 | 6.46e-05
234.02 | 3.22e-05
234.13 | 7.13e-05
234.25 | 8.53e-05
234.37 | 4.25e-05
234.48 | 6.71e-05
234.6 | 0.000109
234.71 | 1.05e-05
234.83 | 6.64e-05
234.94 | 5.93e-05
235.06 | 2.78e-05
235.17 | 0
235.29 | 0.000121
235.41 | 4.13e-05
235.52 | 7.21e-05
235.64 | 0.000134
235.75 | 7.85e-05
235.87 | 2.38e-05
235.98 | 6.45e-05
236.1 | 9.14e-05
236.21 | 5.07e-05
236.33 | 0.000108
236.44 | 0.000104
236.56 | 3.35e-05
236.68 | 6.68e-05
236.79 | 7e-05
236.91 | 6.98e-05
237.02 | 1.33e-05
237.14 | 6.62e-05
237.25 | 3.63e-05
237.37 | 0.000122
237.48 | 8.88e-05
237.6 | -1.64e-05
237.71 | 0.00019
237.83 | 1.96e-05
237.94 | 9.46e-05
238.06 | 0.000133
238.17 | 6.5e-05
238.29 | 0.000159
238.4 | 3.24e-06
238.52 | 0.000168
238.64 | 9.68e-05
238.75 | 9.98e-05
238.87 | 0.000209
238.98 | 6.74e-05
239.1 | 0.000135
239.21 | 8.95e-05
239.33 | 0.000102
239.44 | 5.74e-05
239.56 | 9.23e-05
239.67 | 0.000133
239.79 | 0.000102
239.9 | 0.000146
240.02 | 0.000133
240.13 | 0.000164
240.25 | 0.000151
240.36 | 0.000189
240.48 | 0.000148
240.59 | 0.000144
240.71 | 0.000122
240.82 | 0.000191
240.94 | 0.00015
241.05 | 0.000159
241.17 | 0.000119
241.28 | 0.000147
241.4 | 0.000221
241.51 | 0.000112
241.63 | 0.000106
241.74 | 0.000261
241.86 | 0.00022
241.97 | 0.000267
242.09 | 0.00022
242.2 | 0.000235
242.32 | 0.000229
242.43 | 0.000207
242.55 | 0.00025
242.66 | 0.00025
242.78 | 0.00028
242.89 | 0.000295
243.01 | 0.000341
243.12 | 0.00027
243.24 | 0.000319
243.35 | 0.000273
243.47 | 0.000316
243.58 | 0.000356
243.7 | 0.000435
243.81 | 0.000386
243.93 | 0.000355
244.04 | 0.000428
244.16 | 0.000477
244.27 | 0.000486
244.38 | 0.000504
244.5 | 0.000501
244.61 | 0.000595
244.73 | 0.00057
244.84 | 0.000662
244.96 | 0.00064
245.07 | 0.000719
245.19 | 0.000712
245.3 | 0.000727
245.42 | 0.000763
245.53 | 0.000836
245.65 | 0.000933
245.76 | 0.000908
245.88 | 0.00102
245.99 | 0.00109
246.11 | 0.00121
246.22 | 0.00119
246.33 | 0.00139
246.45 | 0.00138
246.56 | 0.00146
246.68 | 0.00164
246.79 | 0.00165
246.91 | 0.00176
247.02 | 0.0019
247.14 | 0.00202
247.25 | 0.00221
247.37 | 0.00219
247.48 | 0.00238
247.59 | 0.00252
247.71 | 0.00278
247.82 | 0.00302
247.94 | 0.00314
248.05 | 0.00321
248.17 | 0.00337
248.28 | 0.00374
248.4 | 0.00396
248.51 | 0.00407
248.62 | 0.00455
248.74 | 0.0047
248.85 | 0.00495
248.97 | 0.00529
249.08 | 0.00558
249.2 | 0.00592
249.31 | 0.00613
249.43 | 0.0063
249.54 | 0.00693
249.65 | 0.0071
249.77 | 0.0077
249.88 | 0.00787
250 | 0.00847
250.11 | 0.00894
250.23 | 0.00951
250.34 | 0.0101
250.45 | 0.0104
250.57 | 0.0109
250.68 | 0.0113
250.8 | 0.0121
250.91 | 0.0128
251.03 | 0.0134
251.14 | 0.0141
251.25 | 0.0148
251.37 | 0.0154
251.48 | 0.0159
251.6 | 0.0168
251.71 | 0.0175
251.82 | 0.0185
251.94 | 0.0191
252.05 | 0.0198
252.17 | 0.0206
252.28 | 0.0216
252.4 | 0.0222
252.51 | 0.0232
252.62 | 0.0243
252.74 | 0.0249
252.85 | 0.026
252.97 | 0.0271
253.08 | 0.0282
253.19 | 0.0292
253.31 | 0.0299
253.42 | 0.0311
253.54 | 0.0322
253.65 | 0.033
253.76 | 0.0345
253.88 | 0.0354
253.99 | 0.0365
254.11 | 0.0376
254.22 | 0.039
254.33 | 0.0398
254.45 | 0.041
254.56 | 0.0422
254.68 | 0.0435
254.79 | 0.0444
254.9 | 0.0456
255.02 | 0.0466
255.13 | 0.0482
255.25 | 0.0492
255.36 | 0.0504
255.47 | 0.0515
255.59 | 0.0531
255.7 | 0.054
255.81 | 0.0553
255.93 | 0.0566
256.04 | 0.0576
256.16 | 0.0591
256.27 | 0.0598
256.38 | 0.0611
256.5 | 0.0618
256.61 | 0.0635
256.72 | 0.0643
256.84 | 0.0654
256.95 | 0.0659
257.07 | 0.0673
257.18 | 0.0684
257.29 | 0.0687
257.41 | 0.0698
257.52 | 0.0707
257.63 | 0.0723
257.75 | 0.0729
257.86 | 0.0737
257.98 | 0.0748
258.09 | 0.0755
258.2 | 0.0761
258.32 | 0.0773
258.43 | 0.0778
258.54 | 0.079
258.66 | 0.0792
258.77 | 0.0804
258.88 | 0.0808
259 | 0.0814
259.11 | 0.0815
259.23 | 0.0823
259.34 | 0.083
259.45 | 0.0824
259.57 | 0.0844
259.68 | 0.084
259.79 | 0.0838
259.91 | 0.0848
260.02 | 0.0846
260.13 | 0.085
260.25 | 0.0851
260.36 | 0.0851
260.47 | 0.0853
260.59 | 0.0852
260.7 | 0.0852
260.81 | 0.0858
260.93 | 0.0855
261.04 | 0.0853
261.15 | 0.0854
261.27 | 0.0848
261.38 | 0.0853
261.49 | 0.0849
261.61 | 0.0844
261.72 | 0.0842
261.83 | 0.0842
261.95 | 0.0832
262.06 | 0.0833
262.17 | 0.0829
262.29 | 0.0821
262.4 | 0.0817
262.51 | 0.0815
262.63 | 0.0803
262.74 | 0.0799
262.85 | 0.0796
262.97 | 0.0789
263.08 | 0.0781
263.19 | 0.0767
263.31 | 0.0774
263.42 | 0.0761
263.53 | 0.0754
263.65 | 0.0748
263.76 | 0.0742
263.87 | 0.0734
263.99 | 0.0726
264.1 | 0.0718
264.21 | 0.071
264.33 | 0.0706
264.44 | 0.0697
264.55 | 0.0686
264.67 | 0.0677
264.78 | 0.0676
264.89 | 0.0668
265.01 | 0.0664
265.12 | 0.0653
265.23 | 0.0643
265.34 | 0.064
265.46 | 0.0632
265.57 | 0.0624
265.68 | 0.0618
265.8 | 0.0612
265.91 | 0.0606
266.02 | 0.0602
266.14 | 0.0591
266.25 | 0.0591
266.36 | 0.0585
266.47 | 0.0574
266.59 | 0.057
266.7 | 0.0565
266.81 | 0.0559
266.93 | 0.0555
267.04 | 0.0555
267.15 | 0.055
267.26 | 0.0542
267.38 | 0.0536
267.49 | 0.0534
267.6 | 0.0535
267.72 | 0.0531
267.83 | 0.0528
267.94 | 0.0528
268.06 | 0.0525
268.17 | 0.0522
268.28 | 0.0523
268.39 | 0.0522
268.51 | 0.0524
268.62 | 0.0521
268.73 | 0.0526
268.84 | 0.0527
268.96 | 0.053
269.07 | 0.0533
269.18 | 0.0537
269.3 | 0.054
269.41 | 0.0547
269.52 | 0.055
269.63 | 0.0553
269.75 | 0.0558
269.86 | 0.0569
269.97 | 0.0575
270.08 | 0.0583
270.2 | 0.0594
270.31 | 0.0605
270.42 | 0.0615
270.54 | 0.0628
270.65 | 0.0636
270.76 | 0.0648
270.87 | 0.0664
270.99 | 0.0681
271.1 | 0.0697
271.21 | 0.071
271.32 | 0.0725
271.44 | 0.0743
271.55 | 0.0761
271.66 | 0.0778
271.77 | 0.0797
271.89 | 0.0821
272 | 0.0841
272.11 | 0.0861
272.22 | 0.0884
272.34 | 0.0908
272.45 | 0.093
272.56 | 0.095
272.67 | 0.0971
272.79 | 0.0996
272.9 | 0.102
273.01 | 0.105
273.12 | 0.108
273.24 | 0.11
273.35 | 0.113
273.46 | 0.115
273.57 | 0.118
273.69 | 0.12
273.8 | 0.123
273.91 | 0.126
274.02 | 0.127
274.13 | 0.13
274.25 | 0.133
274.36 | 0.135
274.47 | 0.137
274.58 | 0.139
274.7 | 0.142
274.81 | 0.143
274.92 | 0.146
275.03 | 0.148
275.15 | 0.149
275.26 | 0.151
275.37 | 0.152
275.48 | 0.154
275.59 | 0.156
275.71 | 0.157
275.82 | 0.158
275.93 | 0.16
276.04 | 0.16
276.16 | 0.161
276.27 | 0.162
276.38 | 0.162
276.49 | 0.163
276.6 | 0.163
276.72 | 0.163
276.83 | 0.162
276.94 | 0.163
277.05 | 0.163
277.16 | 0.162
277.28 | 0.163
277.39 | 0.161
277.5 | 0.161
277.61 | 0.16
277.73 | 0.159
277.84 | 0.158
277.95 | 0.157
278.06 | 0.155
278.17 | 0.154
278.29 | 0.153
278.4 | 0.151
278.51 | 0.15
278.62 | 0.147
278.73 | 0.146
278.85 | 0.144
278.96 | 0.142
279.07 | 0.14
279.18 | 0.138
279.29 | 0.137
279.4 | 0.135
279.52 | 0.133
279.63 | 0.13
279.74 | 0.128
279.85 | 0.125
279.96 | 0.123
280.08 | 0.12
280.19 | 0.119
280.3 | 0.116
280.41 | 0.114
280.52 | 0.112
280.64 | 0.11
280.75 | 0.108
280.86 | 0.105
280.97 | 0.103
281.08 | 0.101
281.19 | 0.0984
281.31 | 0.0959
281.42 | 0.094
281.53 | 0.0919
281.64 | 0.0902
281.75 | 0.0882
281.86 | 0.0859
281.98 | 0.0841
282.09 | 0.0817
282.2 | 0.0804
282.31 | 0.078
282.42 | 0.0765
282.53 | 0.0745
282.65 | 0.072
282.76 | 0.0705
282.87 | 0.0685
282.98 | 0.0672
283.09 | 0.0656
283.2 | 0.0638
283.32 | 0.0622
283.43 | 0.0603
283.54 | 0.0589
283.65 | 0.0575
283.76 | 0.056
283.87 | 0.0548
283.99 | 0.0532
284.1 | 0.0518
284.21 | 0.051
284.32 | 0.0492
284.43 | 0.0479
284.54 | 0.047
284.65 | 0.0456
284.77 | 0.0444
284.88 | 0.0432
284.99 | 0.042
285.1 | 0.0407
285.21 | 0.0401
285.32 | 0.0389
285.43 | 0.0379
285.55 | 0.0369
285.66 | 0.0361
285.77 | 0.0351
285.88 | 0.034
285.99 | 0.0333
286.1 | 0.0324
286.21 | 0.0321
286.33 | 0.0308
286.44 | 0.0299
286.55 | 0.0295
286.66 | 0.0286
286.77 | 0.0279
286.88 | 0.0273
286.99 | 0.0267
287.1 | 0.0261
287.22 | 0.0254
287.33 | 0.0247
287.44 | 0.0238
287.55 | 0.0236
287.66 | 0.0229
287.77 | 0.0224
287.88 | 0.022
287.99 | 0.0213
288.11 | 0.021
288.22 | 0.0203
288.33 | 0.0202
288.44 | 0.0194
288.55 | 0.0191
288.66 | 0.0186
288.77 | 0.0181
288.88 | 0.0178
288.99 | 0.0175
289.11 | 0.017
289.22 | 0.0166
289.33 | 0.0164
289.44 | 0.0159
289.55 | 0.0156
289.66 | 0.0152
289.77 | 0.0149
289.88 | 0.0144
289.99 | 0.0141
290.11 | 0.0139
290.22 | 0.0136
290.33 | 0.0134
290.44 | 0.0131
290.55 | 0.0127
290.66 | 0.0125
290.77 | 0.0124
290.88 | 0.012
290.99 | 0.0118
291.1 | 0.0115
291.21 | 0.0111
291.33 | 0.0109
291.44 | 0.0109
291.55 | 0.0106
291.66 | 0.0104
291.77 | 0.0101
291.88 | 0.0101
291.99 | 0.00973
292.1 | 0.00948
292.21 | 0.00944
292.32 | 0.00925
292.43 | 0.00898
292.55 | 0.0088
292.66 | 0.00856
292.77 | 0.00833
292.88 | 0.00813
292.99 | 0.00809
293.1 | 0.00779
293.21 | 0.0077
293.32 | 0.00764
293.43 | 0.00738
293.54 | 0.00721
293.65 | 0.00713
293.76 | 0.00703
293.87 | 0.00687
293.99 | 0.00683
294.1 | 0.00651
294.21 | 0.00649
294.32 | 0.00629
294.43 | 0.00597
294.54 | 0.00594
294.65 | 0.00586
294.76 | 0.00576
294.87 | 0.00563
294.98 | 0.00562
295.09 | 0.00539
295.2 | 0.00523
295.31 | 0.00519
295.42 | 0.00508
295.53 | 0.00492
295.64 | 0.00487
295.76 | 0.00477
295.87 | 0.00461
295.98 | 0.00449
296.09 | 0.0045
296.2 | 0.00438
296.31 | 0.00423
296.42 | 0.00417
296.53 | 0.00407
296.64 | 0.00405
296.75 | 0.00396
296.86 | 0.00384
296.97 | 0.00375
297.08 | 0.00364
297.19 | 0.00353
297.3 | 0.00345
297.41 | 0.00338
297.52 | 0.00332
297.63 | 0.00329
297.74 | 0.00329
297.85 | 0.00311
297.96 | 0.00311
298.07 | 0.00294
298.19 | 0.00296
298.3 | 0.00298
298.41 | 0.00285
298.52 | 0.00271
298.63 | 0.0027
298.74 | 0.00262
298.85 | 0.00256
298.96 | 0.00249
299.07 | 0.00245
299.18 | 0.00238
299.29 | 0.00242
299.4 | 0.00228
299.51 | 0.0023
299.62 | 0.00222
299.73 | 0.0022
299.84 | 0.00203
299.95 | 0.00204
300.06 | 0.00207
26.5506
0