## Input Output
1) Paste emission spectrum from Ocean Optics spectrometer in column I | [6 empty] | wavelength | Spectrum |  | Absorbance | Wavelength | Absorbance
2) Modify column D in the Germicidal Calculator tab to calculate radiometer | [6 empty] | 200.11 | -9.7292e-05 | 400 |  | 200 | 1.164
sensitivity relative to the weighted average wavelength | [5 empty] | Peak | 200.23 | 9.6387e-05 | 399 |  | 201 | 0.979
3) Paste water absorbance spectrum (200 - 300 nm) in column L | [5 empty] | 261.0103 | 200.35 | -7.639e-05 | 398 |  | 202 | 0.813
4) Enter radiometer factor for weighted average wavelength into radiometer | [5 empty] | Rad. Factor: | 200.46 | 0 | 397 |  | 203 | 0.684
5) Enter irradiance reading for center of petri dish (mW/cm2) | [5 empty] | 0.0016698 | 200.58 | 0.00016871 | 396 |  | 204 | 0.574
6) Enter petri factor | [6 empty] | 200.7 | 0.00018574 | 395 |  | 205 | 0.484
7) Enter reflection factor for weighted average wavelength | [6 empty] | 200.82 | -0.00012884 | 394 |  | 206 | 0.407
8) Enter sample depth and distance from sample to beam below. | [6 empty] | 200.94 | -3.6476e-05 | 393 |  | 207 | 0.341
9) Enter desired UV doses | [6 empty] | 201.06 | 0.00032528 | 392 |  | 208 | 0.297
10) Outputs are in blue | [6 empty] | 201.17 | 0.00010744 | 391 |  | 209 | 0.251
[7 empty] | 201.29 | 0.00014196 | 390 |  | 210 | 0.222
MP UV Dose Calculations | [6 empty] | 201.41 | 0.00029892 | 389 |  | 211 | 0.202
Input: | Irradiance reading at 261 nm = |  |  | 0.1944 | mW/cm2 |  | 201.53 | -5.2269e-05 | 388 |  | 212 | 0.178
| Petri factor = |  |  | 0.87353 |  |  | 201.65 | 0.0001726 | 387 |  | 213 | 0.165
| Sample depth in the petri dish = |  |  | 0.6 | cm |  | 201.77 | 0.00037636 | 386 |  | 214 | 0.152
|  |  | Distance from sample to beam = | 4 | cm |  | 201.88 | -3.3909e-05 | 385 |  | 215 | 0.139
| Reflection Factor = |  |  | 0.9769 | at 261 nm |  | 202 | -1.6798e-05 | 384 |  | 216 | 0.131
Output: | Divergence Factor = |  |  | 0.86957 |  |  | 202.12 | -1.6644e-05 | 383 |  | 217 | 0.125
| UVT (254) = |  |  | 92.2571 |  |  | 202.24 | 4.9475e-05 | 382 |  | 218 | 0.119
| Incident irradiance = |  |  | 0.1443 | mW/cm2 |  | 202.36 | 3.2674e-05 | 381 |  | 219 | 0.113
| Average irradiance = |  |  | 0.14115 | mW/cm2 |  | 202.47 | 0.00045315 | 380 |  | 220 | 0.108
|  |  | Germicidal Irradiance = | 0.13443 | mW/cm2 |  | 202.59 | 0.00017636 | 379 |  | 221 | 0.104
|  |  | Action spectra irradiance = | 0.14888 | mW/cm2 |  | 202.71 | 0.00028592 | 378 |  | 222 | 0.1
[7 empty] | 202.83 | 6.2954e-05 | 377 |  | 223 | 0.097
| UV Dose Required | Average Irradiance |  | Exposure Time |  |  | 202.95 | -0.0001404 | 376 |  | 224 | 0.092
| (mJ/cm2) |  | (sec) | min | sec |  | 203.06 | 9.2777e-05 | 375 |  | 225 | 0.089
| 5 |  | 35.4234 | 0 | 36 |  | 203.18 | 0.00015327 | 374 |  | 226 | 0.085
| 10 |  | 70.8467 | 1 | 11 |  | 203.3 | -7.5976e-05 | 373 |  | 227 | 0.081
| 15 |  | 106.2701 | 1 | 47 |  | 203.42 | 6.0259e-05 | 372 |  | 228 | 0.077
| 20 |  | 141.6934 | 2 | 22 |  | 203.54 | 0.00010454 | 371 |  | 229 | 0.074
| 30 |  | 212.5401 | 3 | 33 |  | 203.65 | 0.00013327 | 370 |  | 230 | 0.07
[7 empty] | 203.77 | 8.8088e-05 | 369 |  | 231 | 0.067
[7 empty] | 203.89 | -1.4559e-05 | 368 |  | 232 | 0.064
[7 empty] | 204.01 | -2.8874e-05 | 367 |  | 233 | 0.06
[7 empty] | 204.13 | 4.2958e-05 | 366 |  | 234 | 0.057
[7 empty] | 204.24 | 7.0993e-05 | 365 |  | 235 | 0.055
[7 empty] | 204.36 | -0.00011262 | 364 |  | 236 | 0.052
[7 empty] | 204.48 | 0.00015348 | 363 |  | 237 | 0.05
[7 empty] | 204.6 | 0.00023509 | 362 |  | 238 | 0.049
[7 empty] | 204.72 | 0.00012333 | 361 |  | 239 | 0.046
[7 empty] | 204.83 | 0.00013581 | 360 |  | 240 | 0.046
[7 empty] | 204.95 | 9.4221e-05 | 359 |  | 241 | 0.044
[7 empty] | 205.07 | 0.00022686 | 358 |  | 242 | 0.042
[7 empty] | 205.19 | 7.9374e-05 | 357 |  | 243 | 0.042
[7 empty] | 205.3 | 6.5573e-05 | 356 |  | 244 | 0.041
[7 empty] | 205.42 | 0.00015603 | 355 |  | 245 | 0.04
[7 empty] | 205.54 | 9.026e-05 | 354 |  | 246 | 0.04
[7 empty] | 205.66 | 3.8358e-05 | 353 |  | 247 | 0.039
[7 empty] | 205.78 | 3.805e-05 | 352 |  | 248 | 0.039
[7 empty] | 205.89 | 5.0338e-05 | 351 |  | 249 | 0.038
[7 empty] | 206.01 | 8.7433e-05 | 350 |  | 250 | 0.037
[7 empty] | 206.13 | -7.4393e-05 | 349 |  | 251 | 0.036
[7 empty] | 206.25 | 0.00016002 | 348 |  | 252 | 0.036
[7 empty] | 206.37 | -2.4441e-05 | 347 |  | 253 | 0.035
[7 empty] | 206.48 | 0.00014558 | 346 |  | 254 | 0.035
[7 empty] | 206.6 | -0.00019276 | 345 |  | 255 | 0.034
[7 empty] | 206.72 | -7.1776e-05 | 344 |  | 256 | 0.034
[7 empty] | 206.84 | 0.00014255 | 343 |  | 257 | 0.032
[7 empty] | 206.95 | 0.00025955 | 342 |  | 258 | 0.033
[7 empty] | 207.07 | 4.6874e-05 | 341 |  | 259 | 0.032
[7 empty] | 207.19 | 0.0001979 | 340 |  | 260 | 0.032
[7 empty] | 207.31 | 0.00013877 | 339 |  | 261 | 0.033
[7 empty] | 207.42 | -2.2971e-05 | 338 |  | 262 | 0.03
[7 empty] | 207.54 | -7.986e-05 | 337 |  | 263 | 0.033
[7 empty] | 207.66 | 0.00027199 | 336 |  | 264 | 0.031
[7 empty] | 207.78 | -2.2517e-05 | 335 |  | 265 | 0.03
[7 empty] | 207.9 | 0.00014537 | 334 |  | 266 | 0.031
[7 empty] | 208.01 | 9.9942e-05 | 333 |  | 267 | 0.031
[7 empty] | 208.13 | 7.7203e-05 | 332 |  | 268 | 0.03
[7 empty] | 208.25 | 0.00015337 | 331 |  | 269 | 0.03
[7 empty] | 208.37 | 9.7932e-05 | 330 |  | 270 | 0.029
[7 empty] | 208.48 | 5.4039e-05 | 329 |  | 271 | 0.028
[7 empty] | 208.6 | -8.5879e-05 | 328 |  | 272 | 0.029
[7 empty] | 208.72 | 0.00011731 | 327 |  | 273 | 0.029
[7 empty] | 208.84 | 7.4167e-05 | 326 |  | 274 | 0.027
[7 empty] | 208.95 | 0.00010528 | 325 |  | 275 | 0.027
[7 empty] | 209.07 | -3.1384e-05 | 324 |  | 276 | 0.026
[7 empty] | 209.19 | 0.00015595 | 323 |  | 277 | 0.026
[7 empty] | 209.31 | 0.00017567 | 322 |  | 278 | 0.025
[7 empty] | 209.42 | -0.00012325 | 321 |  | 279 | 0.025
[7 empty] | 209.54 | 0 | 320 |  | 280 | 0.024
[7 empty] | 209.66 | 2.0288e-05 | 319 |  | 281 | 0.023
[7 empty] | 209.78 | 0.00015121 | 318 |  | 282 | 0.023
[7 empty] | 209.89 | 6.0117e-05 | 317 |  | 283 | 0.022
[7 empty] | 210.01 | 4.9791e-05 | 316 |  | 284 | 0.022
[7 empty] | 210.13 | 0.00020787 | 315 |  | 285 | 0.02
[7 empty] | 210.25 | -2.9514e-05 | 314 |  | 286 | 0.019
[7 empty] | 210.36 | 0.00013689 | 313 |  | 287 | 0.017
[7 empty] | 210.48 | 0.00017495 | 312 |  | 288 | 0.017
[7 empty] | 210.6 | 0.00013525 | 311 |  | 289 | 0.016
[7 empty] | 210.72 | 9.603e-05 | 310 |  | 290 | 0.016
[7 empty] | 210.83 | -9.5455e-06 | 309 |  | 291 | 0.014
[7 empty] | 210.95 | -9.4903e-05 | 308 |  | 292 | 0.014
[7 empty] | 211.07 | 0.00013208 | 307 |  | 293 | 0.013
[7 empty] | 211.19 | 1.8758e-05 | 306 |  | 294 | 0.012
[7 empty] | 211.3 | 0.00013987 | 305 |  | 295 | 0.011
[7 empty] | 211.42 | 8.342e-05 | 304 |  | 296 | 0.011
[7 empty] | 211.54 | 0.0001843 | 303 |  | 297 | 0.011
[7 empty] | 211.66 | 0.00027486 | 302 |  | 298 | 0.011
[7 empty] | 211.77 | -0.00017305 | 301 |  | 299 | 0.01
[7 empty] | 211.89 | 7.2437e-05 | 300 |  | 300 | 0.01
[7 empty] | 212.01 | 0.00013506 | 299 |  | 301
[7 empty] | 212.13 | 2.6862e-05 | 298 |  | 302
[7 empty] | 212.24 | 0.00015135 | 297 |  | 303
[7 empty] | 212.36 | -1.7703e-05 | 296 |  | 304
[7 empty] | 212.48 | 2.6402e-05 | 295 |  | 305
[7 empty] | 212.59 | -8.7513e-06 | 294 |  | 306
[7 empty] | 212.71 | 8.7022e-06 | 293 |  | 307
[7 empty] | 212.83 | 8.6522e-05 | 292 |  | 308
[7 empty] | 212.95 | 6.0221e-05 | 291 |  | 309
[7 empty] | 213.06 | -7.6994e-05 | 290 |  | 310
[7 empty] | 213.18 | -0.00011059 | 289 |  | 311
[7 empty] | 213.3 | 0.00017763 | 288 |  | 312
[7 empty] | 213.42 | 0.00015139 | 287 |  | 313
[7 empty] | 213.53 | -8.3635e-06 | 286 |  | 314
[7 empty] | 213.65 | 8.3182e-05 | 285 |  | 315
[7 empty] | 213.77 | 5.7922e-05 | 284 |  | 316
[7 empty] | 213.88 | 4.938e-05 | 283 |  | 317
[7 empty] | 214 | 4.0932e-05 | 282 |  | 318
[7 empty] | 214.12 | 7.3285e-05 | 281 |  | 319
[7 empty] | 214.24 | -3.2403e-05 | 280 |  | 320
[7 empty] | 214.35 | 0.00015316 | 279 |  | 321
[7 empty] | 214.47 | 7.2177e-05 | 278 |  | 322
[7 empty] | 214.59 | -0.00010373 | 277 |  | 323
[7 empty] | 214.7 | 7.1453e-05 | 276 |  | 324
[7 empty] | 214.82 | 1.58e-05 | 275 |  | 325
[7 empty] | 214.94 | 7.86e-06 | 274 |  | 326
[7 empty] | 215.06 | 0.00014074 | 273 |  | 327
[7 empty] | 215.17 | 7.0006e-05 | 272 |  | 328
[7 empty] | 215.29 | 8.5126e-05 | 271 |  | 329
[7 empty] | 215.41 | 6.1602e-05 | 270 |  | 330
[7 empty] | 215.52 | 6.894e-05 | 269 |  | 331
[7 empty] | 215.64 | -7.6194e-05 | 268 |  | 332
[7 empty] | 215.76 | 0.00013646 | 267 |  | 333
[7 empty] | 215.87 | 0.0001282 | 266 |  | 334
[7 empty] | 215.99 | 6.7525e-05 | 265 |  | 335
[7 empty] | 216.11 | -7.4633e-06 | 264 |  | 336
[7 empty] | 216.23 | 1.4851e-05 | 263 |  | 337
[7 empty] | 216.34 | 0.00014775 | 262 |  | 338
[7 empty] | 216.46 | 0.00014699 | 261 |  | 339
[7 empty] | 216.58 | 0.00026324 | 260 |  | 340
[7 empty] | 216.69 | 0.00013823 | 259 |  | 341
[7 empty] | 216.81 | 0.00010859 | 258 |  | 342
[7 empty] | 216.93 | -7.2039e-05 | 257 |  | 343
[7 empty] | 217.04 | 7.8858e-05 | 256 |  | 344
[7 empty] | 217.16 | -2.8538e-05 | 255 |  | 345
[7 empty] | 217.28 | 9.9407e-05 | 254 |  | 346
[7 empty] | 217.4 | 7.773e-05 | 253 |  | 347
[7 empty] | 217.51 | 0.00019694 | 252 |  | 348
[7 empty] | 217.63 | 0.00012601 | 251 |  | 349
[7 empty] | 217.75 | 0.00014636 | 250 |  | 350
[7 empty] | 217.86 | 7.6303e-05 | 249 |  | 351
[7 empty] | 217.98 | 5.5234e-05 | 248 |  | 352
[7 empty] | 218.1 | 2.0615e-05 | 247 |  | 353
[7 empty] | 218.21 | 0.0001231 | 246 |  | 354
[7 empty] | 218.33 | 4.764e-05 | 245 |  | 355
[7 empty] | 218.45 | -4.0637e-05 | 244 |  | 356
[7 empty] | 218.56 | 0.00012803 | 243 |  | 357
[7 empty] | 218.68 | 7.378e-05 | 242 |  | 358
[7 empty] | 218.8 | 3.3371e-05 | 241 |  | 359
[7 empty] | 218.91 | -5.9777e-05 | 240 |  | 360
[7 empty] | 219.03 | 8.5919e-05 | 239 |  | 361
[7 empty] | 219.15 | 6.5772e-06 | 238 |  | 362
[7 empty] | 219.26 | 0.00018979 | 237 |  | 363
[7 empty] | 219.38 | 5.2099e-05 | 236 |  | 364
[7 empty] | 219.5 | 7.1284e-05 | 235 |  | 365
[7 empty] | 219.62 | 2.5794e-05 | 234 |  | 366
[7 empty] | 219.73 | 0.00011548 | 233 |  | 367
[7 empty] | 219.85 | 6.3853e-05 | 232 |  | 368
[7 empty] | 219.97 | 1.2708e-05 | 231 |  | 369
[7 empty] | 220.08 | 3.7944e-05 | 230 |  | 370
[7 empty] | 220.2 | 5.0356e-05 | 229 |  | 371
[7 empty] | 220.32 | 0.00025058 | 228 |  | 372
[7 empty] | 220.43 | 0.00013715 | 227 |  | 373
[7 empty] | 220.55 | -8.0656e-05 | 226 |  | 374
[7 empty] | 220.67 | -1.235e-05 | 225 |  | 375
[7 empty] | 220.78 | 0.00011677 | 224 |  | 376
[7 empty] | 220.9 | 9.7863e-05 | 223 |  | 377
[7 empty] | 221.02 | 9.741e-05 | 222 |  | 378
[7 empty] | 221.13 | 6.0583e-06 | 221 |  | 379
[7 empty] | 221.25 | 0.0001266 | 220 |  | 380
[7 empty] | 221.37 | 0 | 219 |  | 381
[7 empty] | 221.48 | 9.5539e-05 | 218 |  | 382
[7 empty] | 221.6 | 5.9418e-06 | 217 |  | 383
[7 empty] | 221.71 | 0.00010644 | 216 |  | 384
[7 empty] | 221.83 | 8.8259e-05 | 215 |  | 385
[7 empty] | 221.95 | 2.3414e-05 | 214 |  | 386
[7 empty] | 222.06 | 0.00012231 | 213 |  | 387
[7 empty] | 222.18 | 9.2727e-05 | 212 |  | 388
[7 empty] | 222.3 | 4.6131e-05 | 211 |  | 389
[7 empty] | 222.41 | 9.1801e-05 | 210 |  | 390
[7 empty] | 222.53 | 4.5664e-05 | 209 |  | 391
[7 empty] | 222.65 | 4.5427e-05 | 208 |  | 392
[7 empty] | 222.76 | 9.6018e-05 | 207 |  | 393
[7 empty] | 222.88 | -3.9335e-05 | 206 |  | 394
[7 empty] | 223 | 7.2672e-05 | 205 |  | 395
[7 empty] | 223.11 | 0.00017243 | 204 |  | 396
[7 empty] | 223.23 | 0.00011624 | 203 |  | 397
[7 empty] | 223.35 | 0.0001432 | 202 |  | 398
[7 empty] | 223.46 | 3.8361e-05 | 201 |  | 399
[7 empty] | 223.58 | 8.7246e-05 | 200 |  | 400
[7 empty] | 223.7 | 4.8828e-05
[7 empty] | 223.81 | 2.1594e-05
[7 empty] | 223.93 | 5.9081e-05
[7 empty] | 224.04 | 2.1378e-05
[7 empty] | 224.16 | 6.9129e-05
[7 empty] | 224.28 | 0.00011113
[7 empty] | 224.39 | 4.7394e-05
[7 empty] | 224.51 | 6.2875e-05
[7 empty] | 224.63 | 5.7339e-05
[7 empty] | 224.74 | 8.2978e-05
[7 empty] | 224.86 | 9.2876e-05
[7 empty] | 224.98 | -5.1339e-06
[7 empty] | 225.09 | 2.0431e-05
[7 empty] | 225.21 | 0.00013213
[7 empty] | 225.32 | 7.5836e-05
[7 empty] | 225.44 | 8.0479e-05
[7 empty] | 225.56 | 7.5043e-05
[7 empty] | 225.67 | 2.4878e-05
[7 empty] | 225.79 | 6.4338e-05
[7 empty] | 225.91 | 0
[7 empty] | 226.02 | 0.00010283
[7 empty] | 226.14 | 4.8709e-05
[7 empty] | 226.25 | 8.2375e-05
[7 empty] | 226.37 | 7.2307e-05
[7 empty] | 226.49 | 1.918e-05
[7 empty] | 226.6 | 8.5862e-05
[7 empty] | 226.72 | 9.4912e-06
[7 empty] | 226.83 | 0.00013692
[7 empty] | 226.95 | 0.00012214
[7 empty] | 227.07 | -3.2716e-05
[7 empty] | 227.18 | 5.1158e-05
[7 empty] | 227.3 | 9.2561e-05
[7 empty] | 227.42 | 0.00010591
[7 empty] | 227.53 | 6.8721e-05
[7 empty] | 227.65 | 1.3673e-05
[7 empty] | 227.76 | 3.6278e-05
[7 empty] | 227.88 | 9.9269e-05
[7 empty] | 228 | 0.00010776
[7 empty] | 228.11 | 9.3817e-05
[7 empty] | 228.23 | 7.5569e-05
[7 empty] | 228.34 | 6.1919e-05
[7 empty] | 228.46 | -1.7599e-05
[7 empty] | 228.58 | 4.3774e-05
[7 empty] | 228.69 | 3.92e-05
[7 empty] | 228.81 | 8.6671e-05
[7 empty] | 228.92 | 0.0001423
[7 empty] | 229.04 | 9.8673e-05
[7 empty] | 229.16 | 0.00014513
[7 empty] | 229.27 | 6.3704e-05
[7 empty] | 229.39 | 4.2251e-05
[7 empty] | 229.5 | 4.6247e-05
[7 empty] | 229.62 | 0.0001088
[7 empty] | 229.74 | 9.5794e-05
[7 empty] | 229.85 | 7.8756e-05
[7 empty] | 229.97 | 5.7755e-05
[7 empty] | 230.08 | 7.3898e-05
[7 empty] | 230.2 | 5.7208e-05
[7 empty] | 230.32 | 9.7628e-05
[7 empty] | 230.43 | 0.00011337
[7 empty] | 230.55 | 8.4641e-05
[7 empty] | 230.66 | 0.00012442
[7 empty] | 230.78 | 7.5909e-05
[7 empty] | 230.89 | 6.7602e-05
[7 empty] | 231.01 | 7.5212e-05
[7 empty] | 231.13 | 3.9405e-05
[7 empty] | 231.24 | 0.00015692
[7 empty] | 231.36 | 0.00011719
[7 empty] | 231.47 | 0.0001089
[7 empty] | 231.59 | 0.00011618
[7 empty] | 231.71 | 1.928e-05
[7 empty] | 231.82 | 6.5265e-05
[7 empty] | 231.94 | 0.0001338
[7 empty] | 232.05 | 0.0001066
[7 empty] | 232.17 | 9.8577e-05
[7 empty] | 232.28 | 5.287e-05
[7 empty] | 232.4 | 0.00011285
[7 empty] | 232.52 | 7.4949e-05
[7 empty] | 232.63 | 9.333e-05
[7 empty] | 232.75 | 0.00013016
[7 empty] | 232.86 | 0.00010745
[7 empty] | 232.98 | 0.00011444
[7 empty] | 233.09 | 0.00012874
[7 empty] | 233.21 | 0.00011361
[7 empty] | 233.33 | 8.0345e-05
[7 empty] | 233.44 | 9.4632e-05
[7 empty] | 233.56 | 9.0675e-05
[7 empty] | 233.67 | 3.615e-05
[7 empty] | 233.79 | 0.00015491
[7 empty] | 233.9 | 0.0001077
[7 empty] | 234.02 | 0.00016458
[7 empty] | 234.13 | 7.8452e-05
[7 empty] | 234.25 | 0.00011375
[7 empty] | 234.37 | 0.00010628
[7 empty] | 234.48 | 7.7684e-05
[7 empty] | 234.6 | 0.00011966
[7 empty] | 234.71 | 0.00011577
[7 empty] | 234.83 | 0.00011889
[7 empty] | 234.94 | 9.7593e-05
[7 empty] | 235.06 | 0.00013204
[7 empty] | 235.17 | 0.00014549
[7 empty] | 235.29 | 0.00013127
[7 empty] | 235.41 | 8.9548e-05
[7 empty] | 235.52 | 0.00012362
[7 empty] | 235.64 | 0.00017463
[7 empty] | 235.75 | 5.1216e-05
[7 empty] | 235.87 | 0.00010895
[7 empty] | 235.98 | 0.00017315
[7 empty] | 236.1 | 0.00012867
[7 empty] | 236.21 | 0.00014522
[7 empty] | 236.33 | 0.00013136
[7 empty] | 236.44 | 0.00015452
[7 empty] | 236.56 | 9.3801e-05
[7 empty] | 236.68 | 0.00014035
[7 empty] | 236.79 | 0.00013999
[7 empty] | 236.91 | 0.00011306
[7 empty] | 237.02 | 0.00013932
[7 empty] | 237.14 | 0.00014562
[7 empty] | 237.25 | 0.00019483
[7 empty] | 237.37 | 0.00013839
[7 empty] | 237.48 | 0.00020052
[7 empty] | 237.6 | 9.8415e-05
[7 empty] | 237.71 | 0.00018992
[7 empty] | 237.83 | 0.0001634
[7 empty] | 237.94 | 0.00014351
[7 empty] | 238.06 | 0.00019208
[7 empty] | 238.17 | 0.00021773
[7 empty] | 238.29 | 0.00025625
[7 empty] | 238.4 | 0.00020074
[7 empty] | 238.52 | 0.00018096
[7 empty] | 238.64 | 0.00016128
[7 empty] | 238.75 | 0.00027049
[7 empty] | 238.87 | 0.00016714
[7 empty] | 238.98 | 0.00019251
[7 empty] | 239.1 | 0.00028826
[7 empty] | 239.21 | 0.00026539
[7 empty] | 239.33 | 0.00028411
[7 empty] | 239.44 | 0.00028678
[7 empty] | 239.56 | 0.00025771
[7 empty] | 239.67 | 0.00023825
[7 empty] | 239.79 | 0.00032669
[7 empty] | 239.9 | 0.0002882
[7 empty] | 240.02 | 0.00031621
[7 empty] | 240.13 | 0.00036312
[7 empty] | 240.25 | 0.00039097
[7 empty] | 240.36 | 0.00031172
[7 empty] | 240.48 | 0.00038366
[7 empty] | 240.59 | 0.00035491
[7 empty] | 240.71 | 0.00035454
[7 empty] | 240.82 | 0.00041054
[7 empty] | 240.94 | 0.00042571
[7 empty] | 241.05 | 0.00045024
[7 empty] | 241.17 | 0.00040288
[7 empty] | 241.28 | 0.00042746
[7 empty] | 241.4 | 0.00048001
[7 empty] | 241.51 | 0.00047952
[7 empty] | 241.63 | 0.00046656
[7 empty] | 241.74 | 0.00047535
[7 empty] | 241.86 | 0.00045932
[7 empty] | 241.97 | 0.00050222
[7 empty] | 242.09 | 0.00054812
[7 empty] | 242.2 | 0.00063111
[7 empty] | 242.32 | 0.00067071
[7 empty] | 242.43 | 0.00062682
[7 empty] | 242.55 | 0.00063855
[7 empty] | 242.66 | 0.00068115
[7 empty] | 242.78 | 0.00070834
[7 empty] | 242.89 | 0.0007294
[7 empty] | 243.01 | 0.00072278
[7 empty] | 243.12 | 0.00074379
[7 empty] | 243.24 | 0.00081088
[7 empty] | 243.35 | 0.00081949
[7 empty] | 243.47 | 0.00088952
[7 empty] | 243.58 | 0.00099927
[7 empty] | 243.7 | 0.0010201
[7 empty] | 243.81 | 0.0010442
[7 empty] | 243.93 | 0.0011659
[7 empty] | 244.04 | 0.0012203
[7 empty] | 244.16 | 0.0012806
[7 empty] | 244.27 | 0.0013472
[7 empty] | 244.38 | 0.0013222
[7 empty] | 244.5 | 0.0014773
[7 empty] | 244.61 | 0.0015654
[7 empty] | 244.73 | 0.0017354
[7 empty] | 244.84 | 0.0018658
[7 empty] | 244.96 | 0.0019258
[7 empty] | 245.07 | 0.0020252
[7 empty] | 245.19 | 0.0021155
[7 empty] | 245.3 | 0.0022243
[7 empty] | 245.42 | 0.0025337
[7 empty] | 245.53 | 0.0025629
[7 empty] | 245.65 | 0.0027136
[7 empty] | 245.76 | 0.0029486
[7 empty] | 245.88 | 0.0032413
[7 empty] | 245.99 | 0.0033341
[7 empty] | 246.11 | 0.0036028
[7 empty] | 246.22 | 0.0039322
[7 empty] | 246.33 | 0.0040371
[7 empty] | 246.45 | 0.0043087
[7 empty] | 246.56 | 0.0045371
[7 empty] | 246.68 | 0.005014
[7 empty] | 246.79 | 0.0050967
[7 empty] | 246.91 | 0.0057734
[7 empty] | 247.02 | 0.0059877
[7 empty] | 247.14 | 0.0062781
[7 empty] | 247.25 | 0.0067503
[7 empty] | 247.37 | 0.0072015
[7 empty] | 247.48 | 0.0075444
[7 empty] | 247.59 | 0.0080606
[7 empty] | 247.71 | 0.0087123
[7 empty] | 247.82 | 0.0092721
[7 empty] | 247.94 | 0.0098616
[7 empty] | 248.05 | 0.010507
[7 empty] | 248.17 | 0.011188
[7 empty] | 248.28 | 0.011586
[7 empty] | 248.4 | 0.012446
[7 empty] | 248.51 | 0.013061
[7 empty] | 248.62 | 0.01384
[7 empty] | 248.74 | 0.01488
[7 empty] | 248.85 | 0.015691
[7 empty] | 248.97 | 0.016594
[7 empty] | 249.08 | 0.017376
[7 empty] | 249.2 | 0.018113
[7 empty] | 249.31 | 0.019229
[7 empty] | 249.43 | 0.020539
[7 empty] | 249.54 | 0.021448
[7 empty] | 249.65 | 0.022537
[7 empty] | 249.77 | 0.023966
[7 empty] | 249.88 | 0.024803
[7 empty] | 250 | 0.026192
[7 empty] | 250.11 | 0.027944
[7 empty] | 250.23 | 0.029016
[7 empty] | 250.34 | 0.030628
[7 empty] | 250.45 | 0.031939
[7 empty] | 250.57 | 0.033786
[7 empty] | 250.68 | 0.035624
[7 empty] | 250.8 | 0.037025
[7 empty] | 250.91 | 0.038503
[7 empty] | 251.03 | 0.040366
[7 empty] | 251.14 | 0.042365
[7 empty] | 251.25 | 0.043874
[7 empty] | 251.37 | 0.046128
[7 empty] | 251.48 | 0.048312
[7 empty] | 251.6 | 0.049839
[7 empty] | 251.71 | 0.052416
[7 empty] | 251.82 | 0.05412
[7 empty] | 251.94 | 0.055935
[7 empty] | 252.05 | 0.058398
[7 empty] | 252.17 | 0.060754
[7 empty] | 252.28 | 0.063041
[7 empty] | 252.4 | 0.06571
[7 empty] | 252.51 | 0.067953
[7 empty] | 252.62 | 0.070457
[7 empty] | 252.74 | 0.072625
[7 empty] | 252.85 | 0.075499
[7 empty] | 252.97 | 0.077577
[7 empty] | 253.08 | 0.080559
[7 empty] | 253.19 | 0.083081
[7 empty] | 253.31 | 0.085609
[7 empty] | 253.42 | 0.088201
[7 empty] | 253.54 | 0.090689
[7 empty] | 253.65 | 0.09347
[7 empty] | 253.76 | 0.095753
[7 empty] | 253.88 | 0.098468
[7 empty] | 253.99 | 0.10191
[7 empty] | 254.11 | 0.10425
[7 empty] | 254.22 | 0.10712
[7 empty] | 254.33 | 0.10998
[7 empty] | 254.45 | 0.11212
[7 empty] | 254.56 | 0.11457
[7 empty] | 254.68 | 0.11757
[7 empty] | 254.79 | 0.11936
[7 empty] | 254.9 | 0.1233
[7 empty] | 255.02 | 0.12595
[7 empty] | 255.13 | 0.12801
[7 empty] | 255.25 | 0.13054
[7 empty] | 255.36 | 0.13366
[7 empty] | 255.47 | 0.13546
[7 empty] | 255.59 | 0.13867
[7 empty] | 255.7 | 0.14073
[7 empty] | 255.81 | 0.14359
[7 empty] | 255.93 | 0.14568
[7 empty] | 256.04 | 0.14825
[7 empty] | 256.16 | 0.15026
[7 empty] | 256.27 | 0.15191
[7 empty] | 256.38 | 0.15511
[7 empty] | 256.5 | 0.15743
[7 empty] | 256.61 | 0.159
[7 empty] | 256.72 | 0.16112
[7 empty] | 256.84 | 0.16305
[7 empty] | 256.95 | 0.16439
[7 empty] | 257.07 | 0.16686
[7 empty] | 257.18 | 0.16762
[7 empty] | 257.29 | 0.16975
[7 empty] | 257.41 | 0.17105
[7 empty] | 257.52 | 0.17376
[7 empty] | 257.63 | 0.1747
[7 empty] | 257.75 | 0.17626
[7 empty] | 257.86 | 0.17739
[7 empty] | 257.98 | 0.17832
[7 empty] | 258.09 | 0.17927
[7 empty] | 258.2 | 0.18106
[7 empty] | 258.32 | 0.18128
[7 empty] | 258.43 | 0.18262
[7 empty] | 258.54 | 0.18271
[7 empty] | 258.66 | 0.18312
[7 empty] | 258.77 | 0.18362
[7 empty] | 258.88 | 0.18474
[7 empty] | 259 | 0.18478
[7 empty] | 259.11 | 0.18626
[7 empty] | 259.23 | 0.18428
[7 empty] | 259.34 | 0.18592
[7 empty] | 259.45 | 0.18567
[7 empty] | 259.57 | 0.18628
[7 empty] | 259.68 | 0.1856
[7 empty] | 259.79 | 0.18594
[7 empty] | 259.91 | 0.18511
[7 empty] | 260.02 | 0.18468
[7 empty] | 260.13 | 0.18467
[7 empty] | 260.25 | 0.18483
[7 empty] | 260.36 | 0.18299
[7 empty] | 260.47 | 0.1831
[7 empty] | 260.59 | 0.18176
[7 empty] | 260.7 | 0.18246
[7 empty] | 260.81 | 0.18025
[7 empty] | 260.93 | 0.17946
[7 empty] | 261.04 | 0.17942
[7 empty] | 261.15 | 0.17761
[7 empty] | 261.27 | 0.17711
[7 empty] | 261.38 | 0.17572
[7 empty] | 261.49 | 0.1749
[7 empty] | 261.61 | 0.17353
[7 empty] | 261.72 | 0.17183
[7 empty] | 261.83 | 0.17072
[7 empty] | 261.95 | 0.1681
[7 empty] | 262.06 | 0.16701
[7 empty] | 262.17 | 0.1646
[7 empty] | 262.29 | 0.16477
[7 empty] | 262.4 | 0.16193
[7 empty] | 262.51 | 0.15998
[7 empty] | 262.63 | 0.15838
[7 empty] | 262.74 | 0.15654
[7 empty] | 262.85 | 0.15432
[7 empty] | 262.97 | 0.15234
[7 empty] | 263.08 | 0.15027
[7 empty] | 263.19 | 0.14837
[7 empty] | 263.31 | 0.14678
[7 empty] | 263.42 | 0.14454
[7 empty] | 263.53 | 0.14233
[7 empty] | 263.65 | 0.14043
[7 empty] | 263.76 | 0.13761
[7 empty] | 263.87 | 0.13539
[7 empty] | 263.99 | 0.13401
[7 empty] | 264.1 | 0.1318
[7 empty] | 264.21 | 0.12966
[7 empty] | 264.33 | 0.12749
[7 empty] | 264.44 | 0.12565
[7 empty] | 264.55 | 0.12336
[7 empty] | 264.67 | 0.12186
[7 empty] | 264.78 | 0.11902
[7 empty] | 264.89 | 0.11701
[7 empty] | 265.01 | 0.11543
[7 empty] | 265.12 | 0.11271
[7 empty] | 265.23 | 0.1103
[7 empty] | 265.34 | 0.10851
[7 empty] | 265.46 | 0.10669
[7 empty] | 265.57 | 0.10428
[7 empty] | 265.68 | 0.10287
[7 empty] | 265.8 | 0.10059
[7 empty] | 265.91 | 0.098744
[7 empty] | 266.02 | 0.09615
[7 empty] | 266.14 | 0.094909
[7 empty] | 266.25 | 0.092631
[7 empty] | 266.36 | 0.090241
[7 empty] | 266.47 | 0.089031
[7 empty] | 266.59 | 0.086862
[7 empty] | 266.7 | 0.085261
[7 empty] | 266.81 | 0.083911
[7 empty] | 266.93 | 0.081736
[7 empty] | 267.04 | 0.080115
[7 empty] | 267.15 | 0.078413
[7 empty] | 267.26 | 0.076626
[7 empty] | 267.38 | 0.075008
[7 empty] | 267.49 | 0.0738
[7 empty] | 267.6 | 0.07119
[7 empty] | 267.72 | 0.070116
[7 empty] | 267.83 | 0.068639
[7 empty] | 267.94 | 0.067709
[7 empty] | 268.06 | 0.065639
[7 empty] | 268.17 | 0.064095
[7 empty] | 268.28 | 0.062701
[7 empty] | 268.39 | 0.061385
[7 empty] | 268.51 | 0.059977
[7 empty] | 268.62 | 0.05841
[7 empty] | 268.73 | 0.057077
[7 empty] | 268.84 | 0.055737
[7 empty] | 268.96 | 0.054028
[7 empty] | 269.07 | 0.053377
[7 empty] | 269.18 | 0.052022
[7 empty] | 269.3 | 0.051151
[7 empty] | 269.41 | 0.049798
[7 empty] | 269.52 | 0.048667
[7 empty] | 269.63 | 0.047358
[7 empty] | 269.75 | 0.046573
[7 empty] | 269.86 | 0.045019
[7 empty] | 269.97 | 0.044353
[7 empty] | 270.08 | 0.043203
[7 empty] | 270.2 | 0.041923
[7 empty] | 270.31 | 0.041372
[7 empty] | 270.42 | 0.040416
[7 empty] | 270.54 | 0.039253
[7 empty] | 270.65 | 0.038421
[7 empty] | 270.76 | 0.037266
[7 empty] | 270.87 | 0.036829
[7 empty] | 270.99 | 0.035584
[7 empty] | 271.1 | 0.035069
[7 empty] | 271.21 | 0.034429
[7 empty] | 271.32 | 0.033188
[7 empty] | 271.44 | 0.032332
[7 empty] | 271.55 | 0.031588
[7 empty] | 271.66 | 0.031063
[7 empty] | 271.77 | 0.030218
[7 empty] | 271.89 | 0.029477
[7 empty] | 272 | 0.029075
[7 empty] | 272.11 | 0.028475
[7 empty] | 272.22 | 0.027489
[7 empty] | 272.34 | 0.02688
[7 empty] | 272.45 | 0.026263
[7 empty] | 272.56 | 0.02546
[7 empty] | 272.67 | 0.025087
[7 empty] | 272.79 | 0.024175
[7 empty] | 272.9 | 0.023958
[7 empty] | 273.01 | 0.02339
[7 empty] | 273.12 | 0.022774
[7 empty] | 273.24 | 0.022115
[7 empty] | 273.35 | 0.021804
[7 empty] | 273.46 | 0.021289
[7 empty] | 273.57 | 0.020914
[7 empty] | 273.69 | 0.020106
[7 empty] | 273.8 | 0.019721
[7 empty] | 273.91 | 0.019413
[7 empty] | 274.02 | 0.018576
[7 empty] | 274.13 | 0.018467
[7 empty] | 274.25 | 0.017701
[7 empty] | 274.36 | 0.017351
[7 empty] | 274.47 | 0.016961
[7 empty] | 274.58 | 0.016714
[7 empty] | 274.7 | 0.01607
[7 empty] | 274.81 | 0.016046
[7 empty] | 274.92 | 0.01548
[7 empty] | 275.03 | 0.015056
[7 empty] | 275.15 | 0.014748
[7 empty] | 275.26 | 0.014331
[7 empty] | 275.37 | 0.014122
[7 empty] | 275.48 | 0.013485
[7 empty] | 275.59 | 0.013356
[7 empty] | 275.71 | 0.013334
[7 empty] | 275.82 | 0.012715
[7 empty] | 275.93 | 0.012498
[7 empty] | 276.04 | 0.012312
[7 empty] | 276.16 | 0.011938
[7 empty] | 276.27 | 0.011818
[7 empty] | 276.38 | 0.011535
[7 empty] | 276.49 | 0.011202
[7 empty] | 276.6 | 0.010895
[7 empty] | 276.72 | 0.010727
[7 empty] | 276.83 | 0.01064
[7 empty] | 276.94 | 0.010113
[7 empty] | 277.05 | 0.0098154
[7 empty] | 277.16 | 0.0096069
[7 empty] | 277.28 | 0.0093905
[7 empty] | 277.39 | 0.0093076
[7 empty] | 277.5 | 0.0091436
[7 empty] | 277.61 | 0.0087873
[7 empty] | 277.73 | 0.0087257
[7 empty] | 277.84 | 0.0085798
[7 empty] | 277.95 | 0.008438
[7 empty] | 278.06 | 0.0081874
[7 empty] | 278.17 | 0.0079347
[7 empty] | 278.29 | 0.0078586
[7 empty] | 278.4 | 0.0076503
[7 empty] | 278.51 | 0.0074622
[7 empty] | 278.62 | 0.00718
[7 empty] | 278.73 | 0.0070909
[7 empty] | 278.85 | 0.0070297
[7 empty] | 278.96 | 0.0068266
[7 empty] | 279.07 | 0.0066642
[7 empty] | 279.18 | 0.0065477
[7 empty] | 279.29 | 0.0064055
[7 empty] | 279.4 | 0.0061985
[7 empty] | 279.52 | 0.0061045
[7 empty] | 279.63 | 0.005899
[7 empty] | 279.74 | 0.0057754
[7 empty] | 279.85 | 0.0056064
[7 empty] | 279.96 | 0.0056048
[7 empty] | 280.08 | 0.005534
[7 empty] | 280.19 | 0.0052862
[7 empty] | 280.3 | 0.0052074
[7 empty] | 280.41 | 0.0051862
[7 empty] | 280.52 | 0.0049795
[7 empty] | 280.64 | 0.0049493
[7 empty] | 280.75 | 0.0047415
[7 empty] | 280.86 | 0.0046775
[7 empty] | 280.97 | 0.00456
[7 empty] | 281.08 | 0.0045642
[7 empty] | 281.19 | 0.0042699
[7 empty] | 281.31 | 0.0044269
[7 empty] | 281.42 | 0.0043366
[7 empty] | 281.53 | 0.0039913
[7 empty] | 281.64 | 0.0041228
[7 empty] | 281.75 | 0.0038818
[7 empty] | 281.86 | 0.0039105
[7 empty] | 281.98 | 0.0038929
[7 empty] | 282.09 | 0.0037861
[7 empty] | 282.2 | 0.0037282
[7 empty] | 282.31 | 0.0035364
[7 empty] | 282.42 | 0.00352
[7 empty] | 282.53 | 0.0033823
[7 empty] | 282.65 | 0.0034365
[7 empty] | 282.76 | 0.0032488
[7 empty] | 282.87 | 0.0032939
[7 empty] | 282.98 | 0.0031998
[7 empty] | 283.09 | 0.003157
[7 empty] | 283.2 | 0.0030714
[7 empty] | 283.32 | 0.0029885
[7 empty] | 283.43 | 0.0029704
[7 empty] | 283.54 | 0.0029521
[7 empty] | 283.65 | 0.0029042
[7 empty] | 283.76 | 0.0027898
[7 empty] | 283.87 | 0.0028273
[7 empty] | 283.99 | 0.0027435
[7 empty] | 284.1 | 0.0027092
[7 empty] | 284.21 | 0.0026676
[7 empty] | 284.32 | 0.0025544
[7 empty] | 284.43 | 0.0025743
[7 empty] | 284.54 | 0.0024101
[7 empty] | 284.65 | 0.0024306
[7 empty] | 284.77 | 0.0024511
[7 empty] | 284.88 | 0.0023616
[7 empty] | 284.99 | 0.002267
[7 empty] | 285.1 | 0.0022047
[7 empty] | 285.21 | 0.0021932
[7 empty] | 285.32 | 0.0021407
[7 empty] | 285.43 | 0.0021175
[7 empty] | 285.55 | 0.0020606
[7 empty] | 285.66 | 0.0020283
[7 empty] | 285.77 | 0.0020297
[7 empty] | 285.88 | 0.0019082
[7 empty] | 285.99 | 0.0019552
[7 empty] | 286.1 | 0.0018343
[7 empty] | 286.21 | 0.0018835
[7 empty] | 286.33 | 0.0018632
[7 empty] | 286.44 | 0.0018241
[7 empty] | 286.55 | 0.001792
[7 empty] | 286.66 | 0.0018147
[7 empty] | 286.77 | 0.0017638
[7 empty] | 286.88 | 0.0017249
[7 empty] | 286.99 | 0.0016486
[7 empty] | 287.1 | 0.0016149
[7 empty] | 287.22 | 0.0016495
[7 empty] | 287.33 | 0.0016536
[7 empty] | 287.44 | 0.0014823
[7 empty] | 287.55 | 0.0015358
[7 empty] | 287.66 | 0.0014817
[7 empty] | 287.77 | 0.0015118
[7 empty] | 287.88 | 0.0015045
[7 empty] | 287.99 | 0.0015345
[7 empty] | 288.11 | 0.0013905
[7 empty] | 288.22 | 0.0013949
[7 empty] | 288.33 | 0.001397
[7 empty] | 288.44 | 0.0014406
[7 empty] | 288.55 | 0.0013804
[7 empty] | 288.66 | 0.0012927
[7 empty] | 288.77 | 0.0013062
[7 empty] | 288.88 | 0.0012899
[7 empty] | 288.99 | 0.0012279
[7 empty] | 289.11 | 0.0012392
[7 empty] | 289.22 | 0.0012205
[7 empty] | 289.33 | 0.0011634
[7 empty] | 289.44 | 0.0012086
[7 empty] | 289.55 | 0.0011611
[7 empty] | 289.66 | 0.0011181
[7 empty] | 289.77 | 0.001125
[7 empty] | 289.88 | 0.0010465
[7 empty] | 289.99 | 0.0011521
[7 empty] | 290.11 | 0.0010447
[7 empty] | 290.22 | 0.0010806
[7 empty] | 290.33 | 0.0010294
[7 empty] | 290.44 | 0.0010208
[7 empty] | 290.55 | 0.0009634
[7 empty] | 290.66 | 0.0010257
[7 empty] | 290.77 | 0.001017
[7 empty] | 290.88 | 0.00097746
[7 empty] | 290.99 | 0.00095565
[7 empty] | 291.1 | 0.00092956
[7 empty] | 291.21 | 0.00097589
[7 empty] | 291.33 | 0.00090817
[7 empty] | 291.44 | 0.00088011
[7 empty] | 291.55 | 0.00087399
[7 empty] | 291.66 | 0.00090057
[7 empty] | 291.77 | 0.00090296
[7 empty] | 291.88 | 0.00083401
[7 empty] | 291.99 | 0.0008431
[7 empty] | 292.1 | 0.00079392
[7 empty] | 292.21 | 0.00079876
[7 empty] | 292.32 | 0.0008465
[7 empty] | 292.43 | 0.00083184
[7 empty] | 292.55 | 0.00074033
[7 empty] | 292.66 | 0.00077515
[7 empty] | 292.77 | 0.00074159
[7 empty] | 292.88 | 0.00075497
[7 empty] | 292.99 | 0.00074708
[7 empty] | 293.1 | 0.00074343
[7 empty] | 293.21 | 0.00071232
[7 empty] | 293.32 | 0.00071927
[7 empty] | 293.43 | 0.00073459
[7 empty] | 293.54 | 0.00068479
[7 empty] | 293.65 | 0.00067922
[7 empty] | 293.76 | 0.0006695
[7 empty] | 293.87 | 0.00067866
[7 empty] | 293.99 | 0.0006857
[7 empty] | 294.1 | 0.0006595
[7 empty] | 294.21 | 0.00057744
[7 empty] | 294.32 | 0.00059712
[7 empty] | 294.43 | 0.00063327
[7 empty] | 294.54 | 0.00058256
[7 empty] | 294.65 | 0.00059807
[7 empty] | 294.76 | 0.00056406
[7 empty] | 294.87 | 0.00059396
[7 empty] | 294.98 | 0.00057236
[7 empty] | 295.09 | 0.00061636
[7 empty] | 295.2 | 0.0005662
[7 empty] | 295.31 | 0.0005651
[7 empty] | 295.42 | 0.00055585
[7 empty] | 295.53 | 0.00055887
[7 empty] | 295.64 | 0.00049497
[7 empty] | 295.76 | 0.00053454
[7 empty] | 295.87 | 0.00057194
[7 empty] | 295.98 | 0.00052649
[7 empty] | 296.09 | 0.00048522
[7 empty] | 296.2 | 0.00046026
[7 empty] | 296.31 | 0.00050358
[7 empty] | 296.42 | 0.00050462
[7 empty] | 296.53 | 0.00046376
[7 empty] | 296.64 | 0.00048685
[7 empty] | 296.75 | 0.00042628
[7 empty] | 296.86 | 0.00046932
[7 empty] | 296.97 | 0.00045072
[7 empty] | 297.08 | 0.00043214
[7 empty] | 297.19 | 0.00048092
[7 empty] | 297.3 | 0.0004268
[7 empty] | 297.41 | 0.00043002
[7 empty] | 297.52 | 0.00044114
[7 empty] | 297.63 | 0.00040506
[7 empty] | 297.74 | 0.00040637
[7 empty] | 297.85 | 0.00044304
[7 empty] | 297.96 | 0.00044239
[7 empty] | 298.07 | 0.00039679
[7 empty] | 298.19 | 0.00042552
[7 empty] | 298.3 | 0.00041318
[7 empty] | 298.41 | 0.00039507
[7 empty] | 298.52 | 0.00039258
[7 empty] | 298.63 | 0.00039591
[7 empty] | 298.74 | 0.00039146
[7 empty] | 298.85 | 0.00040445
[7 empty] | 298.96 | 0.00039612
[7 empty] | 299.07 | 0.00037042
[7 empty] | 299.18 | 0.00034669
[7 empty] | 299.29 | 0.0003615
[7 empty] | 299.4 | 0.00042037
[7 empty] | 299.51 | 0.00040249
[7 empty] | 299.62 | 0.0003272
[7 empty] | 299.73 | 0.00038013
[7 empty] | 299.84 | 0.00032803
[7 empty] | 299.95 | 0.00029707
[7 empty] | 300.06 | 0.00033848

## Petri
[8 empty] | UV Dose Calculations for a Low Pressure UV Lamp | [14 empty] | Calculations for carrying out a Dose-Reponse Curve
[8 empty] | Jim Bolton - Bolton Photosciences Inc., 92 Main St., Ayr, ON, Canada N0B 1E0 | [14 empty] | for a low pressure UV lamp
[8 empty] | Tel: 519-741-6283; Fax: 519-632-8941; Email: jbolton@boltonuv.com | [14 empty] | Programmed by Jim Bolton (519-741-6283)
Steps:
1. Draw a 0.5 cm x 0.5 cm grid and place the center of the grid at the center of the collimated beam
2. Measure the UV irradiance (with a radiometer) every 0.5 cm in the x and y directions and place into columns C and H below.
3. Place the Petri dish "inner" radius into Cell Q27.
4. The Petri factor is calculated in Cell Q20
UV Distribution across the Petri Dish for a low pressure UV Lamp
|  | Meter |  |  |  |  | Meter
x | y | Reading |  |  | x | y | Reading
|  | mW/cm2 | Ratio |  |  |  | mW/cm2 | Ratio
|  | y |  |  |  |  | x
0 | -6 |  | 0 |  | -6 | 0 |  | 0
0 | -5.5 |  | 0 |  | -5.5 | 0 |  | 0
0 | -5 |  | 0 |  | -5 | 0 |  | 0
0 | -4.5 |  | 0 |  | -4.5 | 0 |  | 0 | [6 empty] | Petri factor = | 0.87353
0 | -4 |  | 0 |  | -4 | 0 |  | 0
0 | -3.5 |  | 0 |  | -3.5 | 0 |  | 0
0 | -3 |  | 0 |  | -3 | 0 |  | 0
0 | -2.5 | 0.94 | 0.48354 |  | -2.5 | 0 | 1.192 | 0.61254
0 | -2 | 1.079 | 0.55504 |  | -2 | 0 | 1.388 | 0.71326
0 | -1.5 | 1.319 | 0.6785 |  | -1.5 | 0 | 1.583 | 0.81346
0 | -1 | 1.56 | 0.80247 |  | -1 | 0 | 1.777 | 0.91316 | [6 empty] | Petri Dish Radius = | 1.75 | cm
0 | -0.5 | 1.753 | 0.90175 |  | -0.5 | 0 | 1.9 | 0.97636
0 | 0 | 1.944 | 1 |  | 0 | 0 | 1.946 | 1
0 | 0.5 | 2.02 | 1.0391 |  | 0.5 | 0 | 1.914 | 0.98356
0 | 1 | 1.953 | 1.0046 |  | 1 | 0 | 1.81 | 0.93011
0 | 1.5 | 1.811 | 0.93158 |  | 1.5 | 0 | 1.56 | 0.80164
0 | 2 | 1.635 | 0.84105 |  | 2 | 0 | 1.332 | 0.68448
0 | 2.5 | 1.414 | 0.72737 |  | 2.5 | 0 | 1.146 | 0.5889
0 | 3 |  | 0 |  | 3 | 0 |  | 0
0 | 3.5 |  | 0 |  | 3.5 | 0 |  | 0
0 | 4 |  | 0 |  | 4 | 0 |  | 0
0 | 4.5 |  | 0 |  | 4.5 | 0 |  | 0
0 | 5 |  | 0 |  | 5 | 0 |  | 0
0 | 5.5 |  | 0 |  | 5.5 | 0 |  | 0
0 | 6 |  | 0 |  | 6 | 0 |  | 0
[21 empty] | 3.5 | 4 | 4.5 | 5
x y |  | -6 | -5.5 | 5 | -4.5 | -4 | -3.5 | -3 | -2.5 | -2 | -1.5 | -1 | -0.5 | 0 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 |  |  |  |  | 5.5 | 6
-6 | [27 empty] | 0 | 0
-5.5 | [27 empty] | 0 | 0
-5 | [27 empty] | 0 | 0
-4.5 | [27 empty] | 0 | 0
-4 | [27 empty] | 0 | 0
-3.5 | [27 empty] | 0 | 0
-3 | [27 empty] | 0 | 0
-2.5 | [27 empty] | 0 | 0
-2 | [27 empty] | 0 | 0
-1.5 | [12 empty] | 0.73354 | 0.81346 | 0.84527 | [12 empty] | 2.3923 | 3
-1 | [11 empty] | 0.73278 | 0.82344 | 0.91316 | 0.94885 | 0.91738 | [11 empty] | 4.3356 | 5
-0.5 | [10 empty] | 0.66246 | 0.7835 | 0.88043 | 0.97636 | 1.0145 | 0.98088 | 0.90956 | [10 empty] | 6.2077 | 7
0 | [10 empty] | 0.6785 | 0.80247 | 0.90175 | 1 | 1.0391 | 1.0046 | 0.93158 | [10 empty] | 6.358 | 7
0.5 | [10 empty] | 0.66734 | 0.78927 | 0.88692 | 0.98356 | 1.022 | 0.98811 | 0.91627 | [10 empty] | 6.2535 | 7
1 | [11 empty] | 0.74639 | 0.83873 | 0.93011 | 0.96648 | 0.93442 | [11 empty] | 4.4161 | 5
1.5 | [12 empty] | 0.72288 | 0.80164 | 0.83298 | [12 empty] | 2.3575 | 3
2 | [27 empty] | 0 | 0
2.5 | [27 empty] | 0 | 0
3 | [27 empty] | 0 | 0
3.5 | [27 empty] | 0 | 0
4 | [27 empty] | 0 | 0
4.5 | [27 empty] | 0 | 0
5 | [27 empty] | 0 | 0
5.5 | [27 empty] | 0 | 0
6 | [27 empty] | 0 | 0
[28 empty] | 32.3207 | 37
[27 empty] | Average Petri factor = | 0.87353

## Germicidal Calculator
Petri Factor = | 0.87353 |  | Incident Irradiance = | 0.1944
Sample Depth = | 0.6 | cm | Average Irradiance = | 0.14115
Sample to beam = | 4 | cm
Divergence Factor = | 0.86957 |  | TRLE | 1
Reflection Factor = | 0.9769
Sensor Factor = | 1.0003
Wavelength (nm) | Rated Lamp Energy (RLE) over 200-300 nm | Radiometer Calibration Scan Data (Linden) | Radiometer Sensitivity (fr) (relative to 261 nm) | RLE x radiometer sensitivity | L(l) RLE/TRLE (Same as column B) | Ii | Water Absorbance | Water Factor | Average Irradiance Calc | Absorbance DNA | Relative Absorbance | Germicidal Irradiance | Action Spectra |  | Action Spectrum Absorbance
200 | -3.714e-05 | 0.0015617 | 0.93525 | -3.4735e-05 | -3.714e-05 | -5.3594e-06 | 1.164 | 0.49731 | -2.6653e-06 | 0.83 | 1.1649 | -3.1048e-06 | 23.5339 | 21.1878 | -5.6472e-05
201 | 5.5124e-05 | 0.0016241 | 0.97264 | 5.3616e-05 | 5.5124e-05 | 7.9546e-06 | 0.979 | 0.54817 | 4.3604e-06 | 0.82 | 1.1509 | 5.0183e-06 | 22.9553 | 20.6669 | 9.0116e-05
202 | -6.4124e-06 | 0.0016865 | 1.01 | -6.4768e-06 | -6.4124e-06 | -9.2534e-07 | 0.813 | 0.60075 | -5.559e-07 | 0.81 | 1.1368 | -6.3197e-07 | 22.3766 | 20.1459 | -1.1199e-05
203 | -1.3136e-05 | 0.0017625 | 1.0555 | -1.3865e-05 | -1.3136e-05 | -1.8955e-06 | 0.684 | 0.64691 | -1.2262e-06 | 0.8 | 1.1228 | -1.3768e-06 | 21.7979 | 19.6248 | -2.4065e-05
204 | -1.0567e-05 | 0.0018384 | 1.101 | -1.1634e-05 | -1.0567e-05 | -1.5248e-06 | 0.574 | 0.69043 | -1.0528e-06 | 0.79 | 1.1088 | -1.1673e-06 | 21.2192 | 19.1038 | -2.0113e-05
205 | 5.7065e-05 | 0.0018895 | 1.1316 | 6.4575e-05 | 5.7065e-05 | 8.2347e-06 | 0.484 | 0.72922 | 6.0049e-06 | 0.776 | 1.0891 | 6.5401e-06 | 20.6406 | 18.5829 | 0.00011159
206 | 3.2196e-05 | 0.0019407 | 1.1622 | 3.742e-05 | 3.2196e-05 | 4.6461e-06 | 0.407 | 0.7649 | 3.5538e-06 | 0.762 | 1.0695 | 3.8007e-06 | 20.0619 | 18.0619 | 6.4188e-05
207 | 9.6786e-05 | 0.0019779 | 1.1845 | 0.00011465 | 9.6786e-05 | 1.3967e-05 | 0.341 | 0.79746 | 1.1138e-05 | 0.748 | 1.0498 | 1.1693e-05 | 19.4832 | 17.5409 | 0.00019537
208 | 3.9728e-05 | 0.0020151 | 1.2068 | 4.7945e-05 | 3.9728e-05 | 5.7329e-06 | 0.297 | 0.82024 | 4.7024e-06 | 0.735 | 1.0316 | 4.8509e-06 | 18.9045 | 17.0199 | 8.0034e-05
209 | 1.8452e-05 | 0.0020608 | 1.2341 | 2.2772e-05 | 1.8452e-05 | 2.6627e-06 | 0.251 | 0.84503 | 2.2501e-06 | 0.71 | 0.99649 | 2.2422e-06 | 18.3258 | 16.4989 | 3.7123e-05
210 | 1.9336e-05 | 0.0021064 | 1.2615 | 2.4391e-05 | 1.9336e-05 | 2.7902e-06 | 0.222 | 0.86119 | 2.4029e-06 | 0.685 | 0.9614 | 2.3102e-06 | 17.7472 | 15.978 | 3.8393e-05
211 | -1.2473e-07 | 0.0021453 | 1.2848 | -1.6025e-07 | -1.2473e-07 | -1.7999e-08 | 0.202 | 0.87259 | -1.5706e-08 | 0.66 | 0.92632 | -1.4549e-08 | 17.1685 | 15.457 | -2.4277e-07
212 | 2.7168e-05 | 0.0021842 | 1.3081 | 3.5537e-05 | 2.7168e-05 | 3.9204e-06 | 0.178 | 0.88653 | 3.4756e-06 | 0.645 | 0.90526 | 3.1463e-06 | 16.5898 | 14.9359 | 5.1911e-05
213 | -8.2053e-07 | 0.0022142 | 1.326 | -1.088e-06 | -8.2053e-07 | -1.1841e-07 | 0.165 | 0.89421 | -1.0588e-07 | 0.624 | 0.87579 | -9.2728e-08 | 16.0111 | 14.4149 | -1.5262e-06
214 | 1.5625e-05 | 0.0022442 | 1.344 | 2.1001e-05 | 1.5625e-05 | 2.2548e-06 | 0.152 | 0.90198 | 2.0338e-06 | 0.603 | 0.84632 | 1.7212e-06 | 15.4325 | 13.894 | 2.8257e-05
215 | 2.8363e-05 | 0.0022661 | 1.3571 | 3.8492e-05 | 2.8363e-05 | 4.0929e-06 | 0.139 | 0.90984 | 3.7239e-06 | 0.582 | 0.81684 | 3.0418e-06 | 14.8538 | 13.373 | 4.98e-05
216 | 2.3391e-05 | 0.0022879 | 1.3702 | 3.2051e-05 | 2.3391e-05 | 3.3755e-06 | 0.131 | 0.91473 | 3.0876e-06 | 0.56 | 0.78596 | 2.4268e-06 | 14.2751 | 12.852 | 3.9682e-05
217 | 9.1565e-06 | 0.0023127 | 1.385 | 1.2682e-05 | 9.1565e-06 | 1.3213e-06 | 0.125 | 0.91842 | 1.2135e-06 | 0.54 | 0.75789 | 9.1972e-07 | 13.6964 | 12.331 | 1.4964e-05
218 | 1.8882e-05 | 0.0023375 | 1.3999 | 2.6433e-05 | 1.8882e-05 | 2.7248e-06 | 0.119 | 0.92212 | 2.5126e-06 | 0.52 | 0.72982 | 1.8338e-06 | 13.1177 | 11.81 | 2.9674e-05
219 | 1.8894e-05 | 0.002371 | 1.4199 | 2.6829e-05 | 1.8894e-05 | 2.7265e-06 | 0.113 | 0.92585 | 2.5243e-06 | 0.5 | 0.70175 | 1.7715e-06 | 12.5391 | 11.2891 | 2.8497e-05
220 | 7.4785e-06 | 0.0024045 | 1.44 | 1.0769e-05 | 7.4785e-06 | 1.0792e-06 | 0.108 | 0.92897 | 1.0025e-06 | 0.48 | 0.67368 | 6.7538e-07 | 11.9604 | 10.7681 | 1.0795e-05
221 | 3.7214e-05 | 0.0023973 | 1.4357 | 5.3427e-05 | 3.7214e-05 | 5.3701e-06 | 0.104 | 0.93148 | 5.0021e-06 | 0.453 | 0.63579 | 3.1803e-06 | 11.3552 | 10.2232 | 5.1138e-05
222 | 2.6098e-05 | 0.00239 | 1.4313 | 3.7355e-05 | 2.6098e-05 | 3.7661e-06 | 0.1 | 0.934 | 3.5175e-06 | 0.426 | 0.59789 | 2.1031e-06 | 10.7501 | 9.6784 | 3.4044e-05
223 | 2.7742e-05 | 0.0023994 | 1.437 | 3.9864e-05 | 2.7742e-05 | 4.0032e-06 | 0.097 | 0.93589 | 3.7466e-06 | 0.399 | 0.56 | 2.0981e-06 | 10.1449 | 9.1335 | 3.422e-05
224 | 1.3394e-05 | 0.0024088 | 1.4426 | 1.9323e-05 | 1.3394e-05 | 1.9329e-06 | 0.092 | 0.93906 | 1.8151e-06 | 0.37 | 0.5193 | 9.4257e-07 | 9.5398 | 8.5887 | 1.5589e-05
225 | -1.8542e-07 | 0.0024061 | 1.441 | -2.6719e-07 | -1.8542e-07 | -2.6757e-08 | 0.089 | 0.94097 | -2.5178e-08 | 0.35 | 0.49123 | -1.2368e-08 | 8.9346 | 8.0439 | -2.0253e-07
226 | 3.2117e-05 | 0.0024035 | 1.4394 | 4.6229e-05 | 3.2117e-05 | 4.6346e-06 | 0.085 | 0.94352 | 4.3728e-06 | 0.337 | 0.47298 | 2.0683e-06 | 8.3294 | 7.4991 | 3.2792e-05
227 | 2.1994e-05 | 0.0024126 | 1.4449 | 3.1779e-05 | 2.1994e-05 | 3.1739e-06 | 0.081 | 0.94608 | 3.0027e-06 | 0.324 | 0.45474 | 1.3655e-06 | 7.7243 | 6.9542 | 2.0882e-05
228 | 4.1136e-05 | 0.0024217 | 1.4503 | 5.9661e-05 | 4.1136e-05 | 5.9361e-06 | 0.077 | 0.94865 | 5.6313e-06 | 0.31 | 0.43509 | 2.4501e-06 | 7.1191 | 6.4094 | 3.6093e-05
229 | 4.3219e-05 | 0.0024115 | 1.4442 | 6.2416e-05 | 4.3219e-05 | 6.2366e-06 | 0.074 | 0.95058 | 5.9284e-06 | 0.305 | 0.42807 | 2.5378e-06 | 6.514 | 5.8646 | 3.4768e-05
230 | 2.3728e-05 | 0.0024013 | 1.4381 | 3.4122e-05 | 2.3728e-05 | 3.424e-06 | 0.07 | 0.95317 | 3.2637e-06 | 0.3 | 0.42105 | 1.3742e-06 | 5.9088 | 5.3198 | 1.7362e-05
231 | 2.8469e-05 | 0.0024014 | 1.4382 | 4.0943e-05 | 2.8469e-05 | 4.1082e-06 | 0.067 | 0.95511 | 3.9238e-06 | 0.3 | 0.42105 | 1.6521e-06 | 5.4152 | 4.8753 | 1.913e-05
232 | 4.5413e-05 | 0.0024016 | 1.4382 | 6.5315e-05 | 4.5413e-05 | 6.5533e-06 | 0.064 | 0.95707 | 6.2719e-06 | 0.3 | 0.42105 | 2.6408e-06 | 4.9216 | 4.4309 | 2.779e-05
233 | 4.4679e-05 | 0.0023984 | 1.4364 | 6.4175e-05 | 4.4679e-05 | 6.4473e-06 | 0.06 | 0.95968 | 6.1873e-06 | 0.305 | 0.42807 | 2.6486e-06 | 4.428 | 3.9865 | 2.4666e-05
234 | 5.9208e-05 | 0.0023953 | 1.4345 | 8.4932e-05 | 5.9208e-05 | 8.5439e-06 | 0.057 | 0.96164 | 8.2161e-06 | 0.31 | 0.43509 | 3.5747e-06 | 3.9343 | 3.5421 | 2.9102e-05
235 | 4.383e-05 | 0.0023887 | 1.4305 | 6.2701e-05 | 4.383e-05 | 6.3248e-06 | 0.055 | 0.96295 | 6.0905e-06 | 0.32 | 0.44912 | 2.7354e-06 | 3.4407 | 3.0977 | 1.8866e-05
236 | 6.3268e-05 | 0.0023821 | 1.4266 | 9.0259e-05 | 6.3268e-05 | 9.1298e-06 | 0.052 | 0.96492 | 8.8096e-06 | 0.33 | 0.46316 | 4.0802e-06 | 2.9471 | 2.6533 | 2.3374e-05
237 | 5.1361e-05 | 0.0023659 | 1.4169 | 7.2772e-05 | 5.1361e-05 | 7.4116e-06 | 0.05 | 0.96624 | 7.1614e-06 | 0.35 | 0.49123 | 3.5179e-06 | 2.4535 | 2.2089 | 1.5819e-05
238 | 6.4054e-05 | 0.0023496 | 1.4071 | 9.0131e-05 | 6.4054e-05 | 9.2432e-06 | 0.049 | 0.9669 | 8.9373e-06 | 0.37 | 0.5193 | 4.6411e-06 | 1.9599 | 1.7645 | 1.577e-05
239 | 7.958e-05 | 0.0023429 | 1.4031 | 0.00011166 | 7.958e-05 | 1.1484e-05 | 0.046 | 0.96889 | 1.1126e-05 | 0.39 | 0.54737 | 6.0903e-06 | 1.4662 | 1.3201 | 1.4688e-05
240 | 0.00011893 | 0.0023361 | 1.3991 | 0.00016639 | 0.00011893 | 1.7162e-05 | 0.046 | 0.96889 | 1.6628e-05 | 0.41 | 0.57544 | 9.5682e-06 | 0.9726 | 0.87564 | 1.456e-05
241 | 0.00016762 | 0.0023177 | 1.388 | 0.00023265 | 0.00016762 | 2.4188e-05 | 0.044 | 0.97021 | 2.3467e-05 | 0.433 | 0.60772 | 1.4262e-05 | 0.97456 | 0.87741 | 2.059e-05
242 | 0.0001961 | 0.0022992 | 1.3769 | 0.00027002 | 0.0001961 | 2.8297e-05 | 0.042 | 0.97154 | 2.7492e-05 | 0.456 | 0.64 | 1.7595e-05 | 0.97652 | 0.87917 | 2.417e-05
243 | 0.00027612 | 0.0022817 | 1.3664 | 0.0003773 | 0.00027612 | 3.9846e-05 | 0.042 | 0.97154 | 3.8712e-05 | 0.479 | 0.67228 | 2.6025e-05 | 0.97847 | 0.88093 | 3.4102e-05
244 | 0.00045828 | 0.0022641 | 1.3559 | 0.0006214 | 0.00045828 | 6.6132e-05 | 0.041 | 0.97221 | 6.4294e-05 | 0.5 | 0.70175 | 4.5118e-05 | 0.98043 | 0.88269 | 5.6752e-05
245 | 0.00074895 | 0.0022321 | 1.3368 | 0.0010012 | 0.00074895 | 0.00010808 | 0.04 | 0.97287 | 0.00010514 | 0.52 | 0.72982 | 7.6737e-05 | 0.98239 | 0.88445 | 9.2995e-05
246 | 0.0012813 | 0.0022001 | 1.3176 | 0.0016882 | 0.0012813 | 0.0001849 | 0.04 | 0.97287 | 0.00017988 | 0.54 | 0.75789 | 0.00013633 | 0.98435 | 0.88621 | 0.00015941
247 | 0.0022709 | 0.0021777 | 1.3042 | 0.0029616 | 0.0022709 | 0.00032769 | 0.039 | 0.97354 | 0.00031902 | 0.56 | 0.78596 | 0.00025074 | 0.9863 | 0.88798 | 0.00028328
248 | 0.0038989 | 0.0021553 | 1.2907 | 0.0050325 | 0.0038989 | 0.00056263 | 0.039 | 0.97354 | 0.00054774 | 0.59 | 0.82807 | 0.00045357 | 0.98826 | 0.88974 | 0.00048735
249 | 0.006416 | 0.0021371 | 1.2799 | 0.0082116 | 0.006416 | 0.00092585 | 0.038 | 0.9742 | 0.00090196 | 0.61 | 0.85614 | 0.00077221 | 0.99022 | 0.8915 | 0.0008041
250 | 0.0099985 | 0.002119 | 1.269 | 0.012688 | 0.0099985 | 0.0014428 | 0.037 | 0.97487 | 0.0014066 | 0.63 | 0.88421 | 0.0012437 | 0.99217 | 0.89326 | 0.0012564
251 | 0.015231 | 0.0020712 | 1.2404 | 0.018893 | 0.015231 | 0.002198 | 0.036 | 0.97554 | 0.0021442 | 0.65 | 0.91228 | 0.0019561 | 0.99413 | 0.89502 | 0.0019191
252 | 0.021865 | 0.0020234 | 1.2118 | 0.026496 | 0.021865 | 0.0031553 | 0.036 | 0.97554 | 0.0030781 | 0.665 | 0.93333 | 0.0028729 | 0.99609 | 0.89678 | 0.0027604
253 | 0.029925 | 0.0019932 | 1.1937 | 0.035721 | 0.029925 | 0.0043182 | 0.035 | 0.97621 | 0.0042155 | 0.68 | 0.95439 | 0.0040232 | 0.99804 | 0.89855 | 0.0037878
254 | 0.038977 | 0.001963 | 1.1756 | 0.045822 | 0.038977 | 0.0056246 | 0.035 | 0.97621 | 0.0054908 | 0.695 | 0.97544 | 0.0053559 | 1 | 0.90031 | 0.0049434
255 | 0.047911 | 0.0019167 | 1.1478 | 0.054994 | 0.047911 | 0.0069138 | 0.034 | 0.97688 | 0.0067539 | 0.71 | 0.99649 | 0.0067302 | 1.0092 | 0.90863 | 0.0061368
256 | 0.056236 | 0.0018703 | 1.1201 | 0.062989 | 0.056236 | 0.008115 | 0.034 | 0.97688 | 0.0079274 | 0.72 | 1.0105 | 0.0080108 | 1.0185 | 0.91696 | 0.0072691
257 | 0.063147 | 0.0018276 | 1.0945 | 0.069117 | 0.063147 | 0.0091123 | 0.032 | 0.97822 | 0.0089138 | 0.725 | 1.0175 | 0.0090702 | 1.0277 | 0.92527 | 0.0082477
258 | 0.068137 | 0.001785 | 1.069 | 0.072839 | 0.068137 | 0.0098325 | 0.033 | 0.97755 | 0.0096117 | 0.73 | 1.0246 | 0.0098478 | 1.037 | 0.9336 | 0.0089735
259 | 0.070538 | 0.0017649 | 1.057 | 0.074556 | 0.070538 | 0.010179 | 0.032 | 0.97822 | 0.0099571 | 0.735 | 1.0316 | 0.010272 | 1.0462 | 0.94192 | 0.0093788
260 | 0.070529 | 0.0017448 | 1.0449 | 0.073699 | 0.070529 | 0.010178 | 0.032 | 0.97822 | 0.0099559 | 0.725 | 1.0175 | 0.010131 | 1.0555 | 0.95025 | 0.0094606
261 | 0.068497 | 0.0016698 | 1 | 0.068497 | 0.068497 | 0.0098844 | 0.033 | 0.97755 | 0.0096624 | 0.7125 | 1 | 0.0096624 | 1.1107 | 1 | 0.0096624
262 | 0.063981 | 0.0015947 | 0.95505 | 0.061105 | 0.063981 | 0.0092327 | 0.03 | 0.97956 | 0.009044 | 0.7 | 0.98246 | 0.0088853 | 1.166 | 1.0497 | 0.0094938
263 | 0.057938 | 0.001563 | 0.93607 | 0.054235 | 0.057938 | 0.0083607 | 0.033 | 0.97755 | 0.008173 | 0.6925 | 0.97193 | 0.0079436 | 1.2212 | 1.0995 | 0.0089862
264 | 0.05108 | 0.0015314 | 0.9171 | 0.046845 | 0.05108 | 0.007371 | 0.031 | 0.97889 | 0.0072154 | 0.685 | 0.9614 | 0.0069369 | 1.2765 | 1.1492 | 0.0082923
265 | 0.044114 | 0.0014956 | 0.89567 | 0.039512 | 0.044114 | 0.0063658 | 0.03 | 0.97956 | 0.0062357 | 0.675 | 0.94737 | 0.0059075 | 1.3318 | 1.199 | 0.0074766
266 | 0.036884 | 0.0014598 | 0.87425 | 0.032246 | 0.036884 | 0.0053225 | 0.031 | 0.97889 | 0.0052101 | 0.662 | 0.92912 | 0.0048409 | 1.387 | 1.2487 | 0.0065062
267 | 0.030808 | 0.0014427 | 0.86402 | 0.026619 | 0.030808 | 0.0044457 | 0.031 | 0.97889 | 0.0043518 | 0.649 | 0.91088 | 0.003964 | 1.4423 | 1.2985 | 0.0056509
268 | 0.025452 | 0.0014257 | 0.85379 | 0.021731 | 0.025452 | 0.0036728 | 0.03 | 0.97956 | 0.0035977 | 0.635 | 0.89123 | 0.0032064 | 1.4975 | 1.3482 | 0.0048506
269 | 0.020534 | 0.0013689 | 0.81979 | 0.016834 | 0.020534 | 0.0029632 | 0.03 | 0.97956 | 0.0029026 | 0.622 | 0.87298 | 0.0025339 | 1.5528 | 1.398 | 0.0040578
270 | 0.016811 | 0.0013121 | 0.78579 | 0.01321 | 0.016811 | 0.002426 | 0.029 | 0.98023 | 0.002378 | 0.6 | 0.84211 | 0.0020025 | 1.608 | 1.4477 | 0.0034427
271 | 0.013566 | 0.0012774 | 0.76498 | 0.010378 | 0.013566 | 0.0019576 | 0.028 | 0.98091 | 0.0019202 | 0.58 | 0.81404 | 0.0015631 | 1.5746 | 1.4176 | 0.0027221
272 | 0.011099 | 0.0012426 | 0.74417 | 0.0082595 | 0.011099 | 0.0016016 | 0.029 | 0.98023 | 0.00157 | 0.56 | 0.78596 | 0.0012339 | 1.5411 | 1.3875 | 0.0021783
273 | 0.0089486 | 0.0012081 | 0.72349 | 0.0064742 | 0.0089486 | 0.0012913 | 0.029 | 0.98023 | 0.0012658 | 0.54 | 0.75789 | 0.00095933 | 1.5076 | 1.3573 | 0.0017181
274 | 0.0071493 | 0.0011736 | 0.70282 | 0.0050246 | 0.0071493 | 0.0010317 | 0.027 | 0.98158 | 0.0010127 | 0.52 | 0.72982 | 0.00073906 | 1.4742 | 1.3272 | 0.001344
275 | 0.0057916 | 0.0011338 | 0.679 | 0.0039325 | 0.0057916 | 0.00083575 | 0.027 | 0.98158 | 0.00082035 | 0.5 | 0.70175 | 0.00057569 | 1.4407 | 1.2971 | 0.001064
276 | 0.0047258 | 0.001094 | 0.65517 | 0.0030962 | 0.0047258 | 0.00068195 | 0.026 | 0.98225 | 0.00066984 | 0.48 | 0.67368 | 0.00045126 | 1.4072 | 1.2669 | 0.00084864
277 | 0.0037985 | 0.0010564 | 0.63267 | 0.0024032 | 0.0037985 | 0.00054814 | 0.026 | 0.98225 | 0.00053842 | 0.46 | 0.64561 | 0.00034761 | 1.3737 | 1.2368 | 0.0006659
278 | 0.0031776 | 0.0010188 | 0.61017 | 0.0019389 | 0.0031776 | 0.00045854 | 0.025 | 0.98293 | 0.00045071 | 0.44 | 0.61754 | 0.00027834 | 1.3402 | 1.2066 | 0.00054385
279 | 0.0025834 | 0.00097406 | 0.58334 | 0.001507 | 0.0025834 | 0.0003728 | 0.025 | 0.98293 | 0.00036643 | 0.42 | 0.58947 | 0.000216 | 1.3068 | 1.1765 | 0.00043111
280 | 0.0021306 | 0.00092926 | 0.55652 | 0.0011857 | 0.0021306 | 0.00030745 | 0.024 | 0.9836 | 0.00030241 | 0.4 | 0.5614 | 0.00016977 | 1.2733 | 1.1464 | 0.00034667
281 | 0.0017412 | 0.00087526 | 0.52417 | 0.00091267 | 0.0017412 | 0.00025126 | 0.023 | 0.98428 | 0.00024731 | 0.376 | 0.52772 | 0.00013051 | 1.2071 | 1.0868 | 0.00026877
282 | 0.0014787 | 0.00082126 | 0.49183 | 0.00072725 | 0.0014787 | 0.00021338 | 0.023 | 0.98428 | 0.00021002 | 0.352 | 0.49404 | 0.00010376 | 1.1409 | 1.0272 | 0.00021573
283 | 0.0012185 | 0.00078303 | 0.46894 | 0.00057141 | 0.0012185 | 0.00017584 | 0.022 | 0.98496 | 0.00017319 | 0.328 | 0.46035 | 7.9728e-05 | 1.0747 | 0.96759 | 0.00016758
284 | 0.0010461 | 0.00074479 | 0.44604 | 0.00046661 | 0.0010461 | 0.00015096 | 0.022 | 0.98496 | 0.00014869 | 0.305 | 0.42807 | 6.3648e-05 | 1.0085 | 0.90799 | 0.00013501
285 | 0.00086324 | 0.0007104 | 0.42544 | 0.00036726 | 0.00086324 | 0.00012457 | 0.02 | 0.98631 | 0.00012286 | 0.28 | 0.39298 | 4.8283e-05 | 0.94234 | 0.8484 | 0.00010424
286 | 0.00074218 | 0.000676 | 0.40484 | 0.00030046 | 0.00074218 | 0.0001071 | 0.019 | 0.98699 | 0.00010571 | 0.255 | 0.35789 | 3.7831e-05 | 0.87615 | 0.7888 | 8.3381e-05
287 | 0.00062816 | 0.00065196 | 0.39044 | 0.00024526 | 0.00062816 | 9.0646e-05 | 0.017 | 0.98835 | 8.959e-05 | 0.23 | 0.32281 | 2.892e-05 | 0.80996 | 0.72921 | 6.533e-05
288 | 0.0005812 | 0.00062791 | 0.37604 | 0.00021855 | 0.0005812 | 8.3869e-05 | 0.017 | 0.98835 | 8.2891e-05 | 0.21 | 0.29474 | 2.4431e-05 | 0.74376 | 0.66962 | 5.5505e-05
289 | 0.0004691 | 0.00059241 | 0.35478 | 0.00016643 | 0.0004691 | 6.7692e-05 | 0.016 | 0.98903 | 6.6949e-05 | 0.19 | 0.26667 | 1.7853e-05 | 0.67757 | 0.61002 | 4.0841e-05
290 | 0.00043638 | 0.0005569 | 0.33352 | 0.00014554 | 0.00043638 | 6.2972e-05 | 0.016 | 0.98903 | 6.2281e-05 | 0.17 | 0.2386 | 1.486e-05 | 0.61138 | 0.55043 | 3.4281e-05
291 | 0.0003639 | 0.00051359 | 0.30758 | 0.00011193 | 0.0003639 | 5.2512e-05 | 0.014 | 0.99039 | 5.2008e-05 | 0.15 | 0.21053 | 1.0949e-05 | 0.54519 | 0.49084 | 2.5527e-05
292 | 0.00032014 | 0.00047027 | 0.28164 | 9.0162e-05 | 0.00032014 | 4.6197e-05 | 0.014 | 0.99039 | 4.5753e-05 | 0.14 | 0.19649 | 8.9901e-06 | 0.47899 | 0.43124 | 1.9731e-05
293 | 0.00028506 | 0.0004505 | 0.26979 | 7.6908e-05 | 0.00028506 | 4.1135e-05 | 0.013 | 0.99107 | 4.0768e-05 | 0.125 | 0.17544 | 7.1523e-06 | 0.4128 | 0.37165 | 1.5151e-05
294 | 0.00026085 | 0.00043072 | 0.25795 | 6.7286e-05 | 0.00026085 | 3.7641e-05 | 0.012 | 0.99176 | 3.7331e-05 | 0.11 | 0.15439 | 5.7634e-06 | 0.34661 | 0.31206 | 1.1649e-05
295 | 0.00022155 | 0.00040247 | 0.24103 | 5.3399e-05 | 0.00022155 | 3.197e-05 | 0.011 | 0.99244 | 3.1728e-05 | 0.095 | 0.13333 | 4.2304e-06 | 0.28042 | 0.25246 | 8.0101e-06
296 | 0.00019812 | 0.00037422 | 0.22411 | 4.44e-05 | 0.00019812 | 2.8589e-05 | 0.011 | 0.99244 | 2.8373e-05 | 0.08 | 0.11228 | 3.1857e-06 | 0.21422 | 0.19287 | 5.4722e-06
297 | 0.00017012 | 0.00034845 | 0.20868 | 3.5501e-05 | 0.00017012 | 2.4549e-05 | 0.011 | 0.99244 | 2.4364e-05 | 0.071 | 0.099649 | 2.4278e-06 | 0.14803 | 0.13327 | 3.247e-06
298 | 0.00016255 | 0.00032269 | 0.19325 | 3.1412e-05 | 0.00016255 | 2.3456e-05 | 0.011 | 0.99244 | 2.3279e-05 | 0.062 | 0.087018 | 2.0257e-06 | 0.081839 | 0.07368 | 1.7152e-06
299 | 0.00014765 | 0.00030325 | 0.18161 | 2.6814e-05 | 0.00014765 | 2.1306e-05 | 0.01 | 0.99312 | 2.1159e-05 | 0.053 | 0.074386 | 1.574e-06 | 0.015646 | 0.014087 | 2.9806e-07
300 | 0.00012059 | 0.00028382 | 0.16998 | 2.0497e-05 | 0.00012059 | 1.7401e-05 | 0.01 | 0.99312 | 1.7282e-05 | 0.045 | 0.063158 | 1.0915e-06 | 0 | 0 | 0
TRLE: | 1 |  |  | 0.99966 |  | 0.1443 |  | Average Irradiance | 0.14115 |  | Germicidal | 0.13443 |  |  | 0.14888
|  |  | Lamp Correction Factor, fp | 1.0003
|  |  | Incident IR= | 0.1443

## Emission spectra
|  |  | Interpolation over germicidal range (200 to 300) |  |  |  | Calculating rated lamp energy (RLE)
Spectrometer Measurement |  |  | x | x1, x2 | y1, y2 | y |  | interpolated | RLE
200.11 | -9.7292e-05 | -0.019469 | 200.11 | 200.11 | -9.7292e-05 | -9.7292e-05 | 200.11 | -9.7292e-05 | -3.714e-05 | -0.019469
200.23 | 9.6387e-05 | 0.0193 |  |  |  |  | 201 | 0.0001444 | 5.5124e-05 | 0.029025
200.35 | -7.639e-05 | -0.015305 | 201 | 200.94 | -3.6476e-05 | 0.0001444 | 202 | -1.6798e-05 | -6.4124e-06 | -0.0033932
200.46 | 0 | 0 |  | 201.06 | 0.00032528 |  | 203 | -3.441e-05 | -1.3136e-05 | -0.0069853
200.58 | 0.00016871 | 0.03384 | 202 | 202 | -1.6798e-05 | -1.6798e-05 | 204 | -2.7681e-05 | -1.0567e-05 | -0.0056469
200.7 | 0.00018574 | 0.037278 |  | 202.12 | -1.6644e-05 |  | 205 | 0.00014949 | 5.7065e-05 | 0.030645
200.82 | -0.00012884 | -0.025874 | 203 | 202.95 | -0.0001404 | -3.441e-05 | 206 | 8.4342e-05 | 3.2196e-05 | 0.017374
200.94 | -3.6476e-05 | -0.0073295 |  | 203.06 | 9.2777e-05 |  | 207 | 0.00025354 | 9.6786e-05 | 0.052483
201.06 | 0.00032528 | 0.065401 | 204 | 203.89 | -1.4559e-05 | -2.7681e-05 | 208 | 0.00010407 | 3.9728e-05 | 0.021647
201.17 | 0.00010744 | 0.021614 |  | 204.01 | -2.8874e-05 |  | 209 | 4.8337e-05 | 1.8452e-05 | 0.010102
201.29 | 0.00014196 | 0.028575 | 205 | 204.95 | 9.4221e-05 | 0.00014949 | 210 | 5.0651e-05 | 1.9336e-05 | 0.010637
201.41 | 0.00029892 | 0.060205 |  | 205.07 | 0.00022686 |  | 211 | -3.2675e-07 | -1.2473e-07 | -6.8944e-05
201.53 | -5.2269e-05 | -0.010534 | 206 | 205.89 | 5.0338e-05 | 8.4342e-05 | 212 | 7.1169e-05 | 2.7168e-05 | 0.015088
201.65 | 0.0001726 | 0.034805 |  | 206.01 | 8.7433e-05 |  | 213 | -2.1495e-06 | -8.2053e-07 | -0.00045783
201.77 | 0.00037636 | 0.075938 | 207 | 206.95 | 0.00025955 | 0.00025354 | 214 | 4.0932e-05 | 1.5625e-05 | 0.0087594
201.88 | -3.3909e-05 | -0.0068455 |  | 207.07 | 4.6874e-05 |  | 215 | 7.43e-05 | 2.8363e-05 | 0.015974
202 | -1.6798e-05 | -0.0033932 | 208 | 207.9 | 0.00014537 | 0.00010407 | 216 | 6.1276e-05 | 2.3391e-05 | 0.013236
202.12 | -1.6644e-05 | -0.0033641 |  | 208.01 | 9.9942e-05 |  | 217 | 2.3986e-05 | 9.1565e-06 | 0.005205
202.24 | 4.9475e-05 | 0.010006 | 209 | 208.95 | 0.00010528 | 4.8337e-05 | 218 | 4.9464e-05 | 1.8882e-05 | 0.010783
202.36 | 3.2674e-05 | 0.0066119 |  | 209.07 | -3.1384e-05 |  | 219 | 4.9495e-05 | 1.8894e-05 | 0.010839
202.47 | 0.00045315 | 0.091749 | 210 | 209.89 | 6.0117e-05 | 5.0651e-05 | 220 | 1.9591e-05 | 7.4785e-06 | 0.0043099
202.59 | 0.00017636 | 0.035729 |  | 210.01 | 4.9791e-05 |  | 221 | 9.7486e-05 | 3.7214e-05 | 0.021544
202.71 | 0.00028592 | 0.057959 | 211 | 210.95 | -9.4903e-05 | -3.2675e-07 | 222 | 6.8367e-05 | 2.6098e-05 | 0.015177
202.83 | 6.2954e-05 | 0.012769 |  | 211.07 | 0.00013208 |  | 223 | 7.2672e-05 | 2.7742e-05 | 0.016206
202.95 | -0.0001404 | -0.028494 | 212 | 211.89 | 7.2437e-05 | 7.1169e-05 | 224 | 3.5088e-05 | 1.3394e-05 | 0.0078598
203.06 | 9.2777e-05 | 0.018839 |  | 212.01 | 0.00013506 |  | 225 | -4.8574e-07 | -1.8542e-07 | -0.00010929
203.18 | 0.00015327 | 0.031141 | 213 | 212.95 | 6.0221e-05 | -2.1495e-06 | 226 | 8.4134e-05 | 3.2117e-05 | 0.019014
203.3 | -7.5976e-05 | -0.015446 |  | 213.06 | -7.6994e-05 |  | 227 | 5.7617e-05 | 2.1994e-05 | 0.013079
203.42 | 6.0259e-05 | 0.012258 | 214 | 214 | 4.0932e-05 | 4.0932e-05 | 228 | 0.00010776 | 4.1136e-05 | 0.024569
203.54 | 0.00010454 | 0.021278 |  | 214.12 | 7.3285e-05 |  | 229 | 0.00011322 | 4.3219e-05 | 0.025926
203.65 | 0.00013327 | 0.02714 | 215 | 214.94 | 7.86e-06 | 7.43e-05 | 230 | 6.2158e-05 | 2.3728e-05 | 0.014296
203.77 | 8.8088e-05 | 0.01795 |  | 215.06 | 0.00014074 |  | 231 | 7.4578e-05 | 2.8469e-05 | 0.017227
203.89 | -1.4559e-05 | -0.0029684 | 216 | 215.99 | 6.7525e-05 | 6.1276e-05 | 232 | 0.00011896 | 4.5413e-05 | 0.0276
204.01 | -2.8874e-05 | -0.0058906 |  | 216.11 | -7.4633e-06 |  | 233 | 0.00011704 | 4.4679e-05 | 0.02727
204.13 | 4.2958e-05 | 0.008769 | 217 | 216.93 | -7.2039e-05 | 2.3986e-05 | 234 | 0.0001551 | 5.9208e-05 | 0.036293
204.24 | 7.0993e-05 | 0.0145 |  | 217.04 | 7.8858e-05 |  | 235 | 0.00011482 | 4.383e-05 | 0.026982
204.36 | -0.00011262 | -0.023015 | 218 | 217.98 | 5.5234e-05 | 4.9464e-05 | 236 | 0.00016574 | 6.3268e-05 | 0.039114
204.48 | 0.00015348 | 0.031384 |  | 218.1 | 2.0615e-05 |  | 237 | 0.00013455 | 5.1361e-05 | 0.031887
204.6 | 0.00023509 | 0.048099 | 219 | 218.91 | -5.9777e-05 | 4.9495e-05 | 238 | 0.00016779 | 6.4054e-05 | 0.039935
204.72 | 0.00012333 | 0.025248 |  | 219.03 | 8.5919e-05 |  | 239 | 0.00020847 | 7.958e-05 | 0.049824
204.83 | 0.00013581 | 0.027818 | 220 | 219.97 | 1.2708e-05 | 1.9591e-05 | 240 | 0.00031154 | 0.00011893 | 0.07477
204.95 | 9.4221e-05 | 0.019311 |  | 220.08 | 3.7944e-05 |  | 241 | 0.00043909 | 0.00016762 | 0.10582
205.07 | 0.00022686 | 0.046522 | 221 | 220.9 | 9.7863e-05 | 9.7486e-05 | 242 | 0.00051369 | 0.0001961 | 0.12431
205.19 | 7.9374e-05 | 0.016287 |  | 221.02 | 9.741e-05 |  | 243 | 0.00072333 | 0.00027612 | 0.17577
205.3 | 6.5573e-05 | 0.013462 | 222 | 221.95 | 2.3414e-05 | 6.8367e-05 | 244 | 0.0012005 | 0.00045828 | 0.29293
205.42 | 0.00015603 | 0.032052 |  | 222.06 | 0.00012231 |  | 245 | 0.0019619 | 0.00074895 | 0.48068
205.54 | 9.026e-05 | 0.018552 | 223 | 223 | 7.2672e-05 | 7.2672e-05 | 246 | 0.0033565 | 0.0012813 | 0.8257
205.66 | 3.8358e-05 | 0.0078887 |  | 223.11 | 0.00017243 |  | 247 | 0.0059487 | 0.0022709 | 1.4693
205.78 | 3.805e-05 | 0.0078299 | 224 | 223.93 | 5.9081e-05 | 3.5088e-05 | 248 | 0.010214 | 0.0038989 | 2.533
205.89 | 5.0338e-05 | 0.010364 |  | 224.04 | 2.1378e-05 |  | 249 | 0.016807 | 0.006416 | 4.185
206.01 | 8.7433e-05 | 0.018012 | 225 | 224.98 | -5.1339e-06 | -4.8574e-07 | 250 | 0.026192 | 0.0099985 | 6.548
206.13 | -7.4393e-05 | -0.015335 |  | 225.09 | 2.0431e-05 |  | 251 | 0.0399 | 0.015231 | 10.015
206.25 | 0.00016002 | 0.033004 | 226 | 225.91 | 0 | 8.4134e-05 | 252 | 0.057278 | 0.021865 | 14.4342
206.37 | -2.4441e-05 | -0.0050439 |  | 226.02 | 0.00010283 |  | 253 | 0.07839 | 0.029925 | 19.8327
206.48 | 0.00014558 | 0.030059 | 227 | 226.95 | 0.00012214 | 5.7617e-05 | 254 | 0.10211 | 0.038977 | 25.9347
206.6 | -0.00019276 | -0.039824 |  | 227.07 | -3.2716e-05 |  | 255 | 0.12551 | 0.047911 | 32.0046
206.72 | -7.1776e-05 | -0.014838 | 228 | 228 | 0.00010776 | 0.00010776 | 256 | 0.14732 | 0.056236 | 37.7128
206.84 | 0.00014255 | 0.029485 |  | 228.11 | 9.3817e-05 |  | 257 | 0.16542 | 0.063147 | 42.5127
206.95 | 0.00025955 | 0.053714 | 229 | 228.92 | 0.0001423 | 0.00011322 | 258 | 0.17849 | 0.068137 | 46.0511
207.07 | 4.6874e-05 | 0.0097062 |  | 229.04 | 9.8673e-05 |  | 259 | 0.18478 | 0.070538 | 47.858
207.19 | 0.0001979 | 0.041003 | 230 | 229.97 | 5.7755e-05 | 6.2158e-05 | 260 | 0.18476 | 0.070529 | 48.0371
207.31 | 0.00013877 | 0.028768 |  | 230.08 | 7.3898e-05 |  | 261 | 0.17943 | 0.068497 | 46.8324
207.42 | -2.2971e-05 | -0.0047646 | 231 | 230.89 | 6.7602e-05 | 7.4578e-05 | 262 | 0.1676 | 0.063981 | 43.9124
207.54 | -7.986e-05 | -0.016574 |  | 231.01 | 7.5212e-05 |  | 263 | 0.15178 | 0.057938 | 39.9169
207.66 | 0.00027199 | 0.056481 | 232 | 231.94 | 0.0001338 | 0.00011896 | 264 | 0.13381 | 0.05108 | 35.3256
207.78 | -2.2517e-05 | -0.0046786 |  | 232.05 | 0.0001066 |  | 265 | 0.11556 | 0.044114 | 30.6238
207.9 | 0.00014537 | 0.030222 | 233 | 232.98 | 0.00011444 | 0.00011704 | 266 | 0.096622 | 0.036884 | 25.7014
208.01 | 9.9942e-05 | 0.020789 |  | 233.09 | 0.00012874 |  | 267 | 0.080704 | 0.030808 | 21.5481
208.13 | 7.7203e-05 | 0.016068 | 234 | 233.9 | 0.0001077 | 0.0001551 | 268 | 0.066674 | 0.025452 | 17.8686
208.25 | 0.00015337 | 0.031939 |  | 234.02 | 0.00016458 |  | 269 | 0.053791 | 0.020534 | 14.4699
208.37 | 9.7932e-05 | 0.020406 | 235 | 234.94 | 9.7593e-05 | 0.00011482 | 270 | 0.044039 | 0.016811 | 11.8906
208.48 | 5.4039e-05 | 0.011266 |  | 235.06 | 0.00013204 |  | 271 | 0.035537 | 0.013566 | 9.6306
208.6 | -8.5879e-05 | -0.017914 | 236 | 235.98 | 0.00017315 | 0.00016574 | 272 | 0.029075 | 0.011099 | 7.9084
208.72 | 0.00011731 | 0.024485 |  | 236.1 | 0.00012867 |  | 273 | 0.023442 | 0.0089486 | 6.3996
208.84 | 7.4167e-05 | 0.015489 | 237 | 236.91 | 0.00011306 | 0.00013455 | 274 | 0.018728 | 0.0071493 | 5.1315
208.95 | 0.00010528 | 0.021998 |  | 237.02 | 0.00013932 |  | 275 | 0.015172 | 0.0057916 | 4.1722
209.07 | -3.1384e-05 | -0.0065615 | 238 | 237.94 | 0.00014351 | 0.00016779 | 276 | 0.01238 | 0.0047258 | 3.4168
209.19 | 0.00015595 | 0.032623 |  | 238.06 | 0.00019208 |  | 277 | 0.0099507 | 0.0037985 | 2.7563
209.31 | 0.00017567 | 0.036769 | 239 | 238.98 | 0.00019251 | 0.00020847 | 278 | 0.0083241 | 0.0031776 | 2.3141
209.42 | -0.00012325 | -0.025811 |  | 239.1 | 0.00028826 |  | 279 | 0.0067675 | 0.0025834 | 1.8881
209.54 | 0 | 0 | 240 | 239.9 | 0.0002882 | 0.00031154 | 280 | 0.0055812 | 0.0021306 | 1.5627
209.66 | 2.0288e-05 | 0.0042536 |  | 240.02 | 0.00031621 |  | 281 | 0.0045611 | 0.0017412 | 1.2817
209.78 | 0.00015121 | 0.031721 | 241 | 240.94 | 0.00042571 | 0.00043909 | 282 | 0.0038735 | 0.0014787 | 1.0923
209.89 | 6.0117e-05 | 0.012618 |  | 241.05 | 0.00045024 |  | 283 | 0.003192 | 0.0012185 | 0.90334
210.01 | 4.9791e-05 | 0.010457 | 242 | 241.97 | 0.00050222 | 0.00051369 | 284 | 0.0027404 | 0.0010461 | 0.77827
210.13 | 0.00020787 | 0.04368 |  | 242.09 | 0.00054812 |  | 285 | 0.0022613 | 0.00086324 | 0.64448
210.25 | -2.9514e-05 | -0.0062053 | 243 | 242.89 | 0.0007294 | 0.00072333 | 286 | 0.0019442 | 0.00074218 | 0.55604
210.36 | 0.00013689 | 0.028796 |  | 243.01 | 0.00072278 |  | 287 | 0.0016455 | 0.00062816 | 0.47227
210.48 | 0.00017495 | 0.036823 | 244 | 243.93 | 0.0011659 | 0.0012005 | 288 | 0.0015225 | 0.0005812 | 0.43848
210.6 | 0.00013525 | 0.028484 |  | 244.04 | 0.0012203 |  | 289 | 0.0012288 | 0.0004691 | 0.35514
210.72 | 9.603e-05 | 0.020235 | 245 | 244.96 | 0.0019258 | 0.0019619 | 290 | 0.0011432 | 0.00043638 | 0.33151
210.83 | -9.5455e-06 | -0.0020125 |  | 245.07 | 0.0020252 |  | 291 | 0.00095328 | 0.0003639 | 0.2774
210.95 | -9.4903e-05 | -0.02002 | 246 | 245.99 | 0.0033341 | 0.0033565 | 292 | 0.00083863 | 0.00032014 | 0.24488
211.07 | 0.00013208 | 0.027878 |  | 246.11 | 0.0036028 |  | 293 | 0.00074675 | 0.00028506 | 0.2188
211.19 | 1.8758e-05 | 0.0039615 | 247 | 246.91 | 0.0057734 | 0.0059487 | 294 | 0.00068332 | 0.00026085 | 0.2009
211.3 | 0.00013987 | 0.029555 |  | 247.02 | 0.0059877 |  | 295 | 0.00058036 | 0.00022155 | 0.17121
211.42 | 8.342e-05 | 0.017637 | 248 | 247.94 | 0.0098616 | 0.010214 | 296 | 0.00051899 | 0.00019812 | 0.15362
211.54 | 0.0001843 | 0.038987 |  | 248.05 | 0.010507 |  | 297 | 0.00044565 | 0.00017012 | 0.13236
211.66 | 0.00027486 | 0.058177 | 249 | 248.97 | 0.016594 | 0.016807 | 298 | 0.00042581 | 0.00016255 | 0.12689
211.77 | -0.00017305 | -0.036647 |  | 249.08 | 0.017376 |  | 299 | 0.00038677 | 0.00014765 | 0.11565
211.89 | 7.2437e-05 | 0.015349 | 250 | 250 | 0.026192 | 0.026192 | 300 | 0.00031589 | 0.00012059 | 0.094768
212.01 | 0.00013506 | 0.028634 |  | 250.11 | 0.027944 |  |  | 2.6196 | 1 | 683.7421
212.13 | 2.6862e-05 | 0.0056982 | 251 | 250.91 | 0.038503 | 0.0399 |  |  | Weighted Avg: | 261.0103
212.24 | 0.00015135 | 0.032123 |  | 251.03 | 0.040366
212.36 | -1.7703e-05 | -0.0037594 | 252 | 251.94 | 0.055935 | 0.057278
212.48 | 2.6402e-05 | 0.0056099 |  | 252.05 | 0.058398
212.59 | -8.7513e-06 | -0.0018604 | 253 | 252.97 | 0.077577 | 0.07839
212.71 | 8.7022e-06 | 0.001851 |  | 253.08 | 0.080559
212.83 | 8.6522e-05 | 0.018414 | 254 | 253.99 | 0.10191 | 0.10211
212.95 | 6.0221e-05 | 0.012824 |  | 254.11 | 0.10425
213.06 | -7.6994e-05 | -0.016404 | 255 | 254.9 | 0.1233 | 0.12551
213.18 | -0.00011059 | -0.023576 |  | 255.02 | 0.12595
213.3 | 0.00017763 | 0.037888 | 256 | 255.93 | 0.14568 | 0.14732
213.42 | 0.00015139 | 0.03231 |  | 256.04 | 0.14825
213.53 | -8.3635e-06 | -0.0017859 | 257 | 256.95 | 0.16439 | 0.16542
213.65 | 8.3182e-05 | 0.017772 |  | 257.07 | 0.16686
213.77 | 5.7922e-05 | 0.012382 | 258 | 257.98 | 0.17832 | 0.17849
213.88 | 4.938e-05 | 0.010561 |  | 258.09 | 0.17927
214 | 4.0932e-05 | 0.0087594 | 259 | 259 | 0.18478 | 0.18478
214.12 | 7.3285e-05 | 0.015692 |  | 259.11 | 0.18626
214.24 | -3.2403e-05 | -0.006942 | 260 | 259.91 | 0.18511 | 0.18476
214.35 | 0.00015316 | 0.03283 |  | 260.02 | 0.18468
214.47 | 7.2177e-05 | 0.01548 | 261 | 260.93 | 0.17946 | 0.17943
214.59 | -0.00010373 | -0.022259 |  | 261.04 | 0.17942
214.7 | 7.1453e-05 | 0.015341 | 262 | 261.95 | 0.1681 | 0.1676
214.82 | 1.58e-05 | 0.0033942 |  | 262.06 | 0.16701
214.94 | 7.86e-06 | 0.0016894 | 263 | 262.97 | 0.15234 | 0.15178
215.06 | 0.00014074 | 0.030268 |  | 263.08 | 0.15027
215.17 | 7.0006e-05 | 0.015063 | 264 | 263.99 | 0.13401 | 0.13381
215.29 | 8.5126e-05 | 0.018327 |  | 264.1 | 0.1318
215.41 | 6.1602e-05 | 0.01327 | 265 | 264.89 | 0.11701 | 0.11556
215.52 | 6.894e-05 | 0.014858 |  | 265.01 | 0.11543
215.64 | -7.6194e-05 | -0.01643 | 266 | 265.91 | 0.098744 | 0.096622
215.76 | 0.00013646 | 0.029443 |  | 266.02 | 0.09615
215.87 | 0.0001282 | 0.027675 | 267 | 266.93 | 0.081736 | 0.080704
215.99 | 6.7525e-05 | 0.014585 |  | 267.04 | 0.080115
216.11 | -7.4633e-06 | -0.0016129 | 268 | 267.94 | 0.067709 | 0.066674
216.23 | 1.4851e-05 | 0.0032112 |  | 268.06 | 0.065639
216.34 | 0.00014775 | 0.031964 | 269 | 268.96 | 0.054028 | 0.053791
216.46 | 0.00014699 | 0.031817 |  | 269.07 | 0.053377
216.58 | 0.00026324 | 0.057013 | 270 | 269.97 | 0.044353 | 0.044039
216.69 | 0.00013823 | 0.029953 |  | 270.08 | 0.043203
216.81 | 0.00010859 | 0.023543 | 271 | 270.99 | 0.035584 | 0.035537
216.93 | -7.2039e-05 | -0.015627 |  | 271.1 | 0.035069
217.04 | 7.8858e-05 | 0.017115 | 272 | 272 | 0.029075 | 0.029075
217.16 | -2.8538e-05 | -0.0061973 |  | 272.11 | 0.028475
217.28 | 9.9407e-05 | 0.021599 | 273 | 272.9 | 0.023958 | 0.023442
217.4 | 7.773e-05 | 0.016899 |  | 273.01 | 0.02339
217.51 | 0.00019694 | 0.042836 | 274 | 273.91 | 0.019413 | 0.018728
217.63 | 0.00012601 | 0.027424 |  | 274.02 | 0.018576
217.75 | 0.00014636 | 0.03187 | 275 | 274.92 | 0.01548 | 0.015172
217.86 | 7.6303e-05 | 0.016623 |  | 275.03 | 0.015056
217.98 | 5.5234e-05 | 0.01204 | 276 | 275.93 | 0.012498 | 0.01238
218.1 | 2.0615e-05 | 0.0044961 |  | 276.04 | 0.012312
218.21 | 0.0001231 | 0.026862 | 277 | 276.94 | 0.010113 | 0.0099507
218.33 | 4.764e-05 | 0.010401 |  | 277.05 | 0.0098154
218.45 | -4.0637e-05 | -0.0088772 | 278 | 277.95 | 0.008438 | 0.0083241
218.56 | 0.00012803 | 0.027982 |  | 278.06 | 0.0081874
218.68 | 7.378e-05 | 0.016134 | 279 | 278.96 | 0.0068266 | 0.0067675
218.8 | 3.3371e-05 | 0.0073016 |  | 279.07 | 0.0066642
218.91 | -5.9777e-05 | -0.013086 | 280 | 279.96 | 0.0056048 | 0.0055812
219.03 | 8.5919e-05 | 0.018819 |  | 280.08 | 0.005534
219.15 | 6.5772e-06 | 0.0014414 | 281 | 280.97 | 0.00456 | 0.0045611
219.26 | 0.00018979 | 0.041613 |  | 281.08 | 0.0045642
219.38 | 5.2099e-05 | 0.011429 | 282 | 281.98 | 0.0038929 | 0.0038735
219.5 | 7.1284e-05 | 0.015647 |  | 282.09 | 0.0037861
219.62 | 2.5794e-05 | 0.0056649 | 283 | 282.98 | 0.0031998 | 0.003192
219.73 | 0.00011548 | 0.025374 |  | 283.09 | 0.003157
219.85 | 6.3853e-05 | 0.014038 | 284 | 283.99 | 0.0027435 | 0.0027404
219.97 | 1.2708e-05 | 0.0027954 |  | 284.1 | 0.0027092
220.08 | 3.7944e-05 | 0.0083507 | 285 | 284.99 | 0.002267 | 0.0022613
220.2 | 5.0356e-05 | 0.011088 |  | 285.1 | 0.0022047
220.32 | 0.00025058 | 0.055208 | 286 | 285.99 | 0.0019552 | 0.0019442
220.43 | 0.00013715 | 0.030232 |  | 286.1 | 0.0018343
220.55 | -8.0656e-05 | -0.017789 | 287 | 286.99 | 0.0016486 | 0.0016455
220.67 | -1.235e-05 | -0.0027253 |  | 287.1 | 0.0016149
220.78 | 0.00011677 | 0.02578 | 288 | 287.99 | 0.0015345 | 0.0015225
220.9 | 9.7863e-05 | 0.021618 |  | 288.11 | 0.0013905
221.02 | 9.741e-05 | 0.02153 | 289 | 288.99 | 0.0012279 | 0.0012288
221.13 | 6.0583e-06 | 0.0013397 |  | 289.11 | 0.0012392
221.25 | 0.0001266 | 0.02801 | 290 | 289.99 | 0.0011521 | 0.0011432
221.37 | 0 | 0 |  | 290.11 | 0.0010447
221.48 | 9.5539e-05 | 0.02116 | 291 | 290.99 | 0.00095565 | 0.00095328
221.6 | 5.9418e-06 | 0.0013167 |  | 291.1 | 0.00092956
221.71 | 0.00010644 | 0.023599 | 292 | 291.99 | 0.0008431 | 0.00083863
221.83 | 8.8259e-05 | 0.019578 |  | 292.1 | 0.00079392
221.95 | 2.3414e-05 | 0.0051967 | 293 | 292.99 | 0.00074708 | 0.00074675
222.06 | 0.00012231 | 0.02716 |  | 293.1 | 0.00074343
222.18 | 9.2727e-05 | 0.020602 | 294 | 293.99 | 0.0006857 | 0.00068332
222.3 | 4.6131e-05 | 0.010255 |  | 294.1 | 0.0006595
222.41 | 9.1801e-05 | 0.020417 | 295 | 294.98 | 0.00057236 | 0.00058036
222.53 | 4.5664e-05 | 0.010162 |  | 295.09 | 0.00061636
222.65 | 4.5427e-05 | 0.010114 | 296 | 295.98 | 0.00052649 | 0.00051899
222.76 | 9.6018e-05 | 0.021389 |  | 296.09 | 0.00048522
222.88 | -3.9335e-05 | -0.008767 | 297 | 296.97 | 0.00045072 | 0.00044565
223 | 7.2672e-05 | 0.016206 |  | 297.08 | 0.00043214
223.11 | 0.00017243 | 0.038471 | 298 | 297.96 | 0.00044239 | 0.00042581
223.23 | 0.00011624 | 0.025948 |  | 298.07 | 0.00039679
223.35 | 0.0001432 | 0.031984 | 299 | 298.96 | 0.00039612 | 0.00038677
223.46 | 3.8361e-05 | 0.0085721 |  | 299.07 | 0.00037042
223.58 | 8.7246e-05 | 0.019506 | 300 | 299.95 | 0.00029707 | 0.00031589
223.7 | 4.8828e-05 | 0.010923 |  | 300.06 | 0.00033848
223.81 | 2.1594e-05 | 0.004833
223.93 | 5.9081e-05 | 0.01323
224.04 | 2.1378e-05 | 0.0047895
224.16 | 6.9129e-05 | 0.015496
224.28 | 0.00011113 | 0.024924
224.39 | 4.7394e-05 | 0.010635
224.51 | 6.2875e-05 | 0.014116
224.63 | 5.7339e-05 | 0.01288
224.74 | 8.2978e-05 | 0.018648
224.86 | 9.2876e-05 | 0.020884
224.98 | -5.1339e-06 | -0.001155
225.09 | 2.0431e-05 | 0.0045988
225.21 | 0.00013213 | 0.029757
225.32 | 7.5836e-05 | 0.017087
225.44 | 8.0479e-05 | 0.018143
225.56 | 7.5043e-05 | 0.016927
225.67 | 2.4878e-05 | 0.0056142
225.79 | 6.4338e-05 | 0.014527
225.91 | 0 | 0
226.02 | 0.00010283 | 0.023242
226.14 | 4.8709e-05 | 0.011015
226.25 | 8.2375e-05 | 0.018637
226.37 | 7.2307e-05 | 0.016368
226.49 | 1.918e-05 | 0.0043441
226.6 | 8.5862e-05 | 0.019456
226.72 | 9.4912e-06 | 0.0021518
226.83 | 0.00013692 | 0.031058
226.95 | 0.00012214 | 0.02772
227.07 | -3.2716e-05 | -0.0074288
227.18 | 5.1158e-05 | 0.011622
227.3 | 9.2561e-05 | 0.021039
227.42 | 0.00010591 | 0.024086
227.53 | 6.8721e-05 | 0.015636
227.65 | 1.3673e-05 | 0.0031127
227.76 | 3.6278e-05 | 0.0082627
227.88 | 9.9269e-05 | 0.022621
228 | 0.00010776 | 0.024569
228.11 | 9.3817e-05 | 0.021401
228.23 | 7.5569e-05 | 0.017247
228.34 | 6.1919e-05 | 0.014139
228.46 | -1.7599e-05 | -0.0040207
228.58 | 4.3774e-05 | 0.010006
228.69 | 3.92e-05 | 0.0089646
228.81 | 8.6671e-05 | 0.019831
228.92 | 0.0001423 | 0.032575
229.04 | 9.8673e-05 | 0.0226
229.16 | 0.00014513 | 0.033258
229.27 | 6.3704e-05 | 0.014605
229.39 | 4.2251e-05 | 0.009692
229.5 | 4.6247e-05 | 0.010614
229.62 | 0.0001088 | 0.024983
229.74 | 9.5794e-05 | 0.022008
229.85 | 7.8756e-05 | 0.018102
229.97 | 5.7755e-05 | 0.013282
230.08 | 7.3898e-05 | 0.017002
230.2 | 5.7208e-05 | 0.013169
230.32 | 9.7628e-05 | 0.022486
230.43 | 0.00011337 | 0.026124
230.55 | 8.4641e-05 | 0.019514
230.66 | 0.00012442 | 0.028699
230.78 | 7.5909e-05 | 0.017518
230.89 | 6.7602e-05 | 0.015609
231.01 | 7.5212e-05 | 0.017375
231.13 | 3.9405e-05 | 0.0091077
231.24 | 0.00015692 | 0.036286
231.36 | 0.00011719 | 0.027113
231.47 | 0.0001089 | 0.025207
231.59 | 0.00011618 | 0.026906
231.71 | 1.928e-05 | 0.0044674
231.82 | 6.5265e-05 | 0.01513
231.94 | 0.0001338 | 0.031034
232.05 | 0.0001066 | 0.024737
232.17 | 9.8577e-05 | 0.022887
232.28 | 5.287e-05 | 0.012281
232.4 | 0.00011285 | 0.026226
232.52 | 7.4949e-05 | 0.017427
232.63 | 9.333e-05 | 0.021711
232.75 | 0.00013016 | 0.030295
232.86 | 0.00010745 | 0.025021
232.98 | 0.00011444 | 0.026662
233.09 | 0.00012874 | 0.030008
233.21 | 0.00011361 | 0.026495
233.33 | 8.0345e-05 | 0.018747
233.44 | 9.4632e-05 | 0.022091
233.56 | 9.0675e-05 | 0.021178
233.67 | 3.615e-05 | 0.0084472
233.79 | 0.00015491 | 0.036216
233.9 | 0.0001077 | 0.025191
234.02 | 0.00016458 | 0.038515
234.13 | 7.8452e-05 | 0.018368
234.25 | 0.00011375 | 0.026646
234.37 | 0.00010628 | 0.024909
234.48 | 7.7684e-05 | 0.018215
234.6 | 0.00011966 | 0.028072
234.71 | 0.00011577 | 0.027172
234.83 | 0.00011889 | 0.027919
234.94 | 9.7593e-05 | 0.022928
235.06 | 0.00013204 | 0.031037
235.17 | 0.00014549 | 0.034215
235.29 | 0.00013127 | 0.030887
235.41 | 8.9548e-05 | 0.02108
235.52 | 0.00012362 | 0.029115
235.64 | 0.00017463 | 0.04115
235.75 | 5.1216e-05 | 0.012074
235.87 | 0.00010895 | 0.025698
235.98 | 0.00017315 | 0.04086
236.1 | 0.00012867 | 0.030379
236.21 | 0.00014522 | 0.034302
236.33 | 0.00013136 | 0.031044
236.44 | 0.00015452 | 0.036535
236.56 | 9.3801e-05 | 0.02219
236.68 | 0.00014035 | 0.033218
236.79 | 0.00013999 | 0.033148
236.91 | 0.00011306 | 0.026785
237.02 | 0.00013932 | 0.033022
237.14 | 0.00014562 | 0.034532
237.25 | 0.00019483 | 0.046223
237.37 | 0.00013839 | 0.03285
237.48 | 0.00020052 | 0.047619
237.6 | 9.8415e-05 | 0.023383
237.71 | 0.00018992 | 0.045146
237.83 | 0.0001634 | 0.038861
237.94 | 0.00014351 | 0.034147
238.06 | 0.00019208 | 0.045727
238.17 | 0.00021773 | 0.051857
238.29 | 0.00025625 | 0.061062
238.4 | 0.00020074 | 0.047856
238.52 | 0.00018096 | 0.043163
238.64 | 0.00016128 | 0.038488
238.75 | 0.00027049 | 0.064579
238.87 | 0.00016714 | 0.039925
238.98 | 0.00019251 | 0.046006
239.1 | 0.00028826 | 0.068923
239.21 | 0.00026539 | 0.063484
239.33 | 0.00028411 | 0.067996
239.44 | 0.00028678 | 0.068667
239.56 | 0.00025771 | 0.061737
239.67 | 0.00023825 | 0.057101
239.79 | 0.00032669 | 0.078337
239.9 | 0.0002882 | 0.069139
240.02 | 0.00031621 | 0.075897
240.13 | 0.00036312 | 0.087196
240.25 | 0.00039097 | 0.093931
240.36 | 0.00031172 | 0.074925
240.48 | 0.00038366 | 0.092263
240.59 | 0.00035491 | 0.085388
240.71 | 0.00035454 | 0.085341
240.82 | 0.00041054 | 0.098866
240.94 | 0.00042571 | 0.10257
241.05 | 0.00045024 | 0.10853
241.17 | 0.00040288 | 0.097163
241.28 | 0.00042746 | 0.10314
241.4 | 0.00048001 | 0.11587
241.51 | 0.00047952 | 0.11581
241.63 | 0.00046656 | 0.11273
241.74 | 0.00047535 | 0.11491
241.86 | 0.00045932 | 0.11109
241.97 | 0.00050222 | 0.12152
242.09 | 0.00054812 | 0.13269
242.2 | 0.00063111 | 0.15285
242.32 | 0.00067071 | 0.16253
242.43 | 0.00062682 | 0.15196
242.55 | 0.00063855 | 0.15488
242.66 | 0.00068115 | 0.16529
242.78 | 0.00070834 | 0.17197
242.89 | 0.0007294 | 0.17716
243.01 | 0.00072278 | 0.17564
243.12 | 0.00074379 | 0.18083
243.24 | 0.00081088 | 0.19724
243.35 | 0.00081949 | 0.19942
243.47 | 0.00088952 | 0.21657
243.58 | 0.00099927 | 0.2434
243.7 | 0.0010201 | 0.2486
243.81 | 0.0010442 | 0.25459
243.93 | 0.0011659 | 0.2844
244.04 | 0.0012203 | 0.2978
244.16 | 0.0012806 | 0.31267
244.27 | 0.0013472 | 0.32908
244.38 | 0.0013222 | 0.32312
244.5 | 0.0014773 | 0.3612
244.61 | 0.0015654 | 0.38291
244.73 | 0.0017354 | 0.4247
244.84 | 0.0018658 | 0.45682
244.96 | 0.0019258 | 0.47174
245.07 | 0.0020252 | 0.49632
245.19 | 0.0021155 | 0.5187
245.3 | 0.0022243 | 0.54562
245.42 | 0.0025337 | 0.62182
245.53 | 0.0025629 | 0.62927
245.65 | 0.0027136 | 0.6666
245.76 | 0.0029486 | 0.72465
245.88 | 0.0032413 | 0.79697
245.99 | 0.0033341 | 0.82016
246.11 | 0.0036028 | 0.88669
246.22 | 0.0039322 | 0.96819
246.33 | 0.0040371 | 0.99446
246.45 | 0.0043087 | 1.0619
246.56 | 0.0045371 | 1.1187
246.68 | 0.005014 | 1.2369
246.79 | 0.0050967 | 1.2578
246.91 | 0.0057734 | 1.4255
247.02 | 0.0059877 | 1.4791
247.14 | 0.0062781 | 1.5516
247.25 | 0.0067503 | 1.669
247.37 | 0.0072015 | 1.7814
247.48 | 0.0075444 | 1.8671
247.59 | 0.0080606 | 1.9957
247.71 | 0.0087123 | 2.1581
247.82 | 0.0092721 | 2.2978
247.94 | 0.0098616 | 2.4451
248.05 | 0.010507 | 2.6063
248.17 | 0.011188 | 2.7765
248.28 | 0.011586 | 2.8766
248.4 | 0.012446 | 3.0916
248.51 | 0.013061 | 3.2458
248.62 | 0.01384 | 3.4409
248.74 | 0.01488 | 3.7013
248.85 | 0.015691 | 3.9047
248.97 | 0.016594 | 4.1314
249.08 | 0.017376 | 4.328
249.2 | 0.018113 | 4.5138
249.31 | 0.019229 | 4.794
249.43 | 0.020539 | 5.123
249.54 | 0.021448 | 5.3521
249.65 | 0.022537 | 5.6264
249.77 | 0.023966 | 5.986
249.88 | 0.024803 | 6.1978
250 | 0.026192 | 6.548
250.11 | 0.027944 | 6.9891
250.23 | 0.029016 | 7.2607
250.34 | 0.030628 | 7.6674
250.45 | 0.031939 | 7.9991
250.57 | 0.033786 | 8.4658
250.68 | 0.035624 | 8.9302
250.8 | 0.037025 | 9.2859
250.91 | 0.038503 | 9.6608
251.03 | 0.040366 | 10.1331
251.14 | 0.042365 | 10.6395
251.25 | 0.043874 | 11.0233
251.37 | 0.046128 | 11.5952
251.48 | 0.048312 | 12.1495
251.6 | 0.049839 | 12.5395
251.71 | 0.052416 | 13.1936
251.82 | 0.05412 | 13.6285
251.94 | 0.055935 | 14.0923
252.05 | 0.058398 | 14.7192
252.17 | 0.060754 | 15.3203
252.28 | 0.063041 | 15.904
252.4 | 0.06571 | 16.5852
252.51 | 0.067953 | 17.1588
252.62 | 0.070457 | 17.7988
252.74 | 0.072625 | 18.3552
252.85 | 0.075499 | 19.0899
252.97 | 0.077577 | 19.6247
253.08 | 0.080559 | 20.3879
253.19 | 0.083081 | 21.0353
253.31 | 0.085609 | 21.6856
253.42 | 0.088201 | 22.3519
253.54 | 0.090689 | 22.9933
253.65 | 0.09347 | 23.7087
253.76 | 0.095753 | 24.2983
253.88 | 0.098468 | 24.9991
253.99 | 0.10191 | 25.8841
254.11 | 0.10425 | 26.491
254.22 | 0.10712 | 27.232
254.33 | 0.10998 | 27.9712
254.45 | 0.11212 | 28.5289
254.56 | 0.11457 | 29.1649
254.68 | 0.11757 | 29.9427
254.79 | 0.11936 | 30.4117
254.9 | 0.1233 | 31.4292
255.02 | 0.12595 | 32.1198
255.13 | 0.12801 | 32.6592
255.25 | 0.13054 | 33.3203
255.36 | 0.13366 | 34.1314
255.47 | 0.13546 | 34.606
255.59 | 0.13867 | 35.4427
255.7 | 0.14073 | 35.9847
255.81 | 0.14359 | 36.7318
255.93 | 0.14568 | 37.2839
256.04 | 0.14825 | 37.9579
256.16 | 0.15026 | 38.4906
256.27 | 0.15191 | 38.93
256.38 | 0.15511 | 39.7671
256.5 | 0.15743 | 40.3808
256.61 | 0.159 | 40.801
256.72 | 0.16112 | 41.3627
256.84 | 0.16305 | 41.8778
256.95 | 0.16439 | 42.24
257.07 | 0.16686 | 42.8947
257.18 | 0.16762 | 43.1085
257.29 | 0.16975 | 43.675
257.41 | 0.17105 | 44.03
257.52 | 0.17376 | 44.7467
257.63 | 0.1747 | 45.008
257.75 | 0.17626 | 45.431
257.86 | 0.17739 | 45.7418
257.98 | 0.17832 | 46.003
258.09 | 0.17927 | 46.2678
258.2 | 0.18106 | 46.7497
258.32 | 0.18128 | 46.8282
258.43 | 0.18262 | 47.1945
258.54 | 0.18271 | 47.2378
258.66 | 0.18312 | 47.3658
258.77 | 0.18362 | 47.5153
258.88 | 0.18474 | 47.8255
259 | 0.18478 | 47.858
259.11 | 0.18626 | 48.2618
259.23 | 0.18428 | 47.7709
259.34 | 0.18592 | 48.2165
259.45 | 0.18567 | 48.1721
259.57 | 0.18628 | 48.3527
259.68 | 0.1856 | 48.1966
259.79 | 0.18594 | 48.3054
259.91 | 0.18511 | 48.1119
260.02 | 0.18468 | 48.0205
260.13 | 0.18467 | 48.0382
260.25 | 0.18483 | 48.102
260.36 | 0.18299 | 47.6433
260.47 | 0.1831 | 47.6921
260.59 | 0.18176 | 47.3648
260.7 | 0.18246 | 47.5673
260.81 | 0.18025 | 47.011
260.93 | 0.17946 | 46.8265
261.04 | 0.17942 | 46.8358
261.15 | 0.17761 | 46.3829
261.27 | 0.17711 | 46.2735
261.38 | 0.17572 | 45.9297
261.49 | 0.1749 | 45.7346
261.61 | 0.17353 | 45.3972
261.72 | 0.17183 | 44.9713
261.83 | 0.17072 | 44.6996
261.95 | 0.1681 | 44.0338
262.06 | 0.16701 | 43.7666
262.17 | 0.1646 | 43.1532
262.29 | 0.16477 | 43.2175
262.4 | 0.16193 | 42.4904
262.51 | 0.15998 | 41.9963
262.63 | 0.15838 | 41.5953
262.74 | 0.15654 | 41.1293
262.85 | 0.15432 | 40.563
262.97 | 0.15234 | 40.0608
263.08 | 0.15027 | 39.533
263.19 | 0.14837 | 39.0495
263.31 | 0.14678 | 38.6486
263.42 | 0.14454 | 38.0747
263.53 | 0.14233 | 37.5082
263.65 | 0.14043 | 37.0244
263.76 | 0.13761 | 36.296
263.87 | 0.13539 | 35.7254
263.99 | 0.13401 | 35.3773
264.1 | 0.1318 | 34.8084
264.21 | 0.12966 | 34.2575
264.33 | 0.12749 | 33.6994
264.44 | 0.12565 | 33.2269
264.55 | 0.12336 | 32.6349
264.67 | 0.12186 | 32.2527
264.78 | 0.11902 | 31.5141
264.89 | 0.11701 | 30.9948
265.01 | 0.11543 | 30.5901
265.12 | 0.11271 | 29.8817
265.23 | 0.1103 | 29.2549
265.34 | 0.10851 | 28.792
265.46 | 0.10669 | 28.3219
265.57 | 0.10428 | 27.6936
265.68 | 0.10287 | 27.3305
265.8 | 0.10059 | 26.7368
265.91 | 0.098744 | 26.257
266.02 | 0.09615 | 25.5778
266.14 | 0.094909 | 25.2591
266.25 | 0.092631 | 24.663
266.36 | 0.090241 | 24.0366
266.47 | 0.089031 | 23.7241
266.59 | 0.086862 | 23.1565
266.7 | 0.085261 | 22.7391
266.81 | 0.083911 | 22.3883
266.93 | 0.081736 | 21.8178
267.04 | 0.080115 | 21.3939
267.15 | 0.078413 | 20.948
267.26 | 0.076626 | 20.4791
267.38 | 0.075008 | 20.0556
267.49 | 0.0738 | 19.7408
267.6 | 0.07119 | 19.0504
267.72 | 0.070116 | 18.7715
267.83 | 0.068639 | 18.3836
267.94 | 0.067709 | 18.1419
268.06 | 0.065639 | 17.5952
268.17 | 0.064095 | 17.1884
268.28 | 0.062701 | 16.8214
268.39 | 0.061385 | 16.4751
268.51 | 0.059977 | 16.1044
268.62 | 0.05841 | 15.6901
268.73 | 0.057077 | 15.3383
268.84 | 0.055737 | 14.9843
268.96 | 0.054028 | 14.5314
269.07 | 0.053377 | 14.3621
269.18 | 0.052022 | 14.0033
269.3 | 0.051151 | 13.775
269.41 | 0.049798 | 13.4161
269.52 | 0.048667 | 13.1167
269.63 | 0.047358 | 12.7691
269.75 | 0.046573 | 12.5631
269.86 | 0.045019 | 12.1488
269.97 | 0.044353 | 11.974
270.08 | 0.043203 | 11.6683
270.2 | 0.041923 | 11.3276
270.31 | 0.041372 | 11.1833
270.42 | 0.040416 | 10.9293
270.54 | 0.039253 | 10.6195
270.65 | 0.038421 | 10.3986
270.76 | 0.037266 | 10.0901
270.87 | 0.036829 | 9.9759
270.99 | 0.035584 | 9.6429
271.1 | 0.035069 | 9.5072
271.21 | 0.034429 | 9.3375
271.32 | 0.033188 | 9.0046
271.44 | 0.032332 | 8.7762
271.55 | 0.031588 | 8.5777
271.66 | 0.031063 | 8.4386
271.77 | 0.030218 | 8.2123
271.89 | 0.029477 | 8.0145
272 | 0.029075 | 7.9084
272.11 | 0.028475 | 7.7483
272.22 | 0.027489 | 7.4831
272.34 | 0.02688 | 7.3205
272.45 | 0.026263 | 7.1554
272.56 | 0.02546 | 6.9394
272.67 | 0.025087 | 6.8405
272.79 | 0.024175 | 6.5947
272.9 | 0.023958 | 6.5381
273.01 | 0.02339 | 6.3857
273.12 | 0.022774 | 6.22
273.24 | 0.022115 | 6.0427
273.35 | 0.021804 | 5.9601
273.46 | 0.021289 | 5.8217
273.57 | 0.020914 | 5.7214
273.69 | 0.020106 | 5.5028
273.8 | 0.019721 | 5.3996
273.91 | 0.019413 | 5.3174
274.02 | 0.018576 | 5.0902
274.13 | 0.018467 | 5.0624
274.25 | 0.017701 | 4.8545
274.36 | 0.017351 | 4.7604
274.47 | 0.016961 | 4.6553
274.58 | 0.016714 | 4.5893
274.7 | 0.01607 | 4.4144
274.81 | 0.016046 | 4.4096
274.92 | 0.01548 | 4.2558
275.03 | 0.015056 | 4.1409
275.15 | 0.014748 | 4.0579
275.26 | 0.014331 | 3.9448
275.37 | 0.014122 | 3.8888
275.48 | 0.013485 | 3.7148
275.59 | 0.013356 | 3.6808
275.71 | 0.013334 | 3.6763
275.82 | 0.012715 | 3.5071
275.93 | 0.012498 | 3.4486
276.04 | 0.012312 | 3.3986
276.16 | 0.011938 | 3.2968
276.27 | 0.011818 | 3.265
276.38 | 0.011535 | 3.188
276.49 | 0.011202 | 3.0972
276.6 | 0.010895 | 3.0136
276.72 | 0.010727 | 2.9684
276.83 | 0.01064 | 2.9455
276.94 | 0.010113 | 2.8007
277.05 | 0.0098154 | 2.7194
277.16 | 0.0096069 | 2.6626
277.28 | 0.0093905 | 2.6038
277.39 | 0.0093076 | 2.5818
277.5 | 0.0091436 | 2.5373
277.61 | 0.0087873 | 2.4394
277.73 | 0.0087257 | 2.4234
277.84 | 0.0085798 | 2.3838
277.95 | 0.008438 | 2.3453
278.06 | 0.0081874 | 2.2766
278.17 | 0.0079347 | 2.2072
278.29 | 0.0078586 | 2.187
278.4 | 0.0076503 | 2.1298
278.51 | 0.0074622 | 2.0783
278.62 | 0.00718 | 2.0005
278.73 | 0.0070909 | 1.9764
278.85 | 0.0070297 | 1.9602
278.96 | 0.0068266 | 1.9043
279.07 | 0.0066642 | 1.8598
279.18 | 0.0065477 | 1.828
279.29 | 0.0064055 | 1.789
279.4 | 0.0061985 | 1.7319
279.52 | 0.0061045 | 1.7063
279.63 | 0.005899 | 1.6495
279.74 | 0.0057754 | 1.6156
279.85 | 0.0056064 | 1.569
279.96 | 0.0056048 | 1.5691
280.08 | 0.005534 | 1.55
280.19 | 0.0052862 | 1.4811
280.3 | 0.0052074 | 1.4596
280.41 | 0.0051862 | 1.4543
280.52 | 0.0049795 | 1.3968
280.64 | 0.0049493 | 1.389
280.75 | 0.0047415 | 1.3312
280.86 | 0.0046775 | 1.3137
280.97 | 0.00456 | 1.2812
281.08 | 0.0045642 | 1.2829
281.19 | 0.0042699 | 1.2007
281.31 | 0.0044269 | 1.2453
281.42 | 0.0043366 | 1.2204
281.53 | 0.0039913 | 1.1237
281.64 | 0.0041228 | 1.1611
281.75 | 0.0038818 | 1.0937
281.86 | 0.0039105 | 1.1022
281.98 | 0.0038929 | 1.0977
282.09 | 0.0037861 | 1.068
282.2 | 0.0037282 | 1.0521
282.31 | 0.0035364 | 0.99836
282.42 | 0.00352 | 0.99412
282.53 | 0.0033823 | 0.9556
282.65 | 0.0034365 | 0.97133
282.76 | 0.0032488 | 0.91863
282.87 | 0.0032939 | 0.93175
282.98 | 0.0031998 | 0.90548
283.09 | 0.003157 | 0.89372
283.2 | 0.0030714 | 0.86982
283.32 | 0.0029885 | 0.8467
283.43 | 0.0029704 | 0.8419
283.54 | 0.0029521 | 0.83704
283.65 | 0.0029042 | 0.82378
283.76 | 0.0027898 | 0.79163
283.87 | 0.0028273 | 0.80259
283.99 | 0.0027435 | 0.77913
284.1 | 0.0027092 | 0.76968
284.21 | 0.0026676 | 0.75816
284.32 | 0.0025544 | 0.72627
284.43 | 0.0025743 | 0.73221
284.54 | 0.0024101 | 0.68577
284.65 | 0.0024306 | 0.69187
284.77 | 0.0024511 | 0.698
284.88 | 0.0023616 | 0.67277
284.99 | 0.002267 | 0.64607
285.1 | 0.0022047 | 0.62856
285.21 | 0.0021932 | 0.62552
285.32 | 0.0021407 | 0.61078
285.43 | 0.0021175 | 0.6044
285.55 | 0.0020606 | 0.5884
285.66 | 0.0020283 | 0.5794
285.77 | 0.0020297 | 0.58003
285.88 | 0.0019082 | 0.54552
285.99 | 0.0019552 | 0.55917
286.1 | 0.0018343 | 0.52479
286.21 | 0.0018835 | 0.53908
286.33 | 0.0018632 | 0.53349
286.44 | 0.0018241 | 0.5225
286.55 | 0.001792 | 0.5135
286.66 | 0.0018147 | 0.5202
286.77 | 0.0017638 | 0.5058
286.88 | 0.0017249 | 0.49484
286.99 | 0.0016486 | 0.47313
287.1 | 0.0016149 | 0.46364
287.22 | 0.0016495 | 0.47377
287.33 | 0.0016536 | 0.47513
287.44 | 0.0014823 | 0.42607
287.55 | 0.0015358 | 0.44162
287.66 | 0.0014817 | 0.42623
287.77 | 0.0015118 | 0.43505
287.88 | 0.0015045 | 0.43312
287.99 | 0.0015345 | 0.44192
288.11 | 0.0013905 | 0.40062
288.22 | 0.0013949 | 0.40204
288.33 | 0.001397 | 0.4028
288.44 | 0.0014406 | 0.41553
288.55 | 0.0013804 | 0.39831
288.66 | 0.0012927 | 0.37315
288.77 | 0.0013062 | 0.37719
288.88 | 0.0012899 | 0.37263
288.99 | 0.0012279 | 0.35485
289.11 | 0.0012392 | 0.35827
289.22 | 0.0012205 | 0.35299
289.33 | 0.0011634 | 0.33661
289.44 | 0.0012086 | 0.34982
289.55 | 0.0011611 | 0.3362
289.66 | 0.0011181 | 0.32387
289.77 | 0.001125 | 0.32599
289.88 | 0.0010465 | 0.30336
289.99 | 0.0011521 | 0.3341
290.11 | 0.0010447 | 0.30308
290.22 | 0.0010806 | 0.31361
290.33 | 0.0010294 | 0.29887
290.44 | 0.0010208 | 0.29648
290.55 | 0.0009634 | 0.27992
290.66 | 0.0010257 | 0.29813
290.77 | 0.001017 | 0.29571
290.88 | 0.00097746 | 0.28432
290.99 | 0.00095565 | 0.27808
291.1 | 0.00092956 | 0.27059
291.21 | 0.00097589 | 0.28419
291.33 | 0.00090817 | 0.26458
291.44 | 0.00088011 | 0.2565
291.55 | 0.00087399 | 0.25481
291.66 | 0.00090057 | 0.26266
291.77 | 0.00090296 | 0.26346
291.88 | 0.00083401 | 0.24343
291.99 | 0.0008431 | 0.24618
292.1 | 0.00079392 | 0.2319
292.21 | 0.00079876 | 0.23341
292.32 | 0.0008465 | 0.24745
292.43 | 0.00083184 | 0.24325
292.55 | 0.00074033 | 0.21658
292.66 | 0.00077515 | 0.22686
292.77 | 0.00074159 | 0.21712
292.88 | 0.00075497 | 0.22112
292.99 | 0.00074708 | 0.21889
293.1 | 0.00074343 | 0.2179
293.21 | 0.00071232 | 0.20886
293.32 | 0.00071927 | 0.21098
293.43 | 0.00073459 | 0.21555
293.54 | 0.00068479 | 0.20101
293.65 | 0.00067922 | 0.19945
293.76 | 0.0006695 | 0.19667
293.87 | 0.00067866 | 0.19944
293.99 | 0.0006857 | 0.20159
294.1 | 0.0006595 | 0.19396
294.21 | 0.00057744 | 0.16989
294.32 | 0.00059712 | 0.17574
294.43 | 0.00063327 | 0.18645
294.54 | 0.00058256 | 0.17159
294.65 | 0.00059807 | 0.17622
294.76 | 0.00056406 | 0.16626
294.87 | 0.00059396 | 0.17514
294.98 | 0.00057236 | 0.16883
295.09 | 0.00061636 | 0.18188
295.2 | 0.0005662 | 0.16714
295.31 | 0.0005651 | 0.16688
295.42 | 0.00055585 | 0.16421
295.53 | 0.00055887 | 0.16516
295.64 | 0.00049497 | 0.14633
295.76 | 0.00053454 | 0.1581
295.87 | 0.00057194 | 0.16922
295.98 | 0.00052649 | 0.15583
296.09 | 0.00048522 | 0.14367
296.2 | 0.00046026 | 0.13633
296.31 | 0.00050358 | 0.14922
296.42 | 0.00050462 | 0.14958
296.53 | 0.00046376 | 0.13752
296.64 | 0.00048685 | 0.14442
296.75 | 0.00042628 | 0.1265
296.86 | 0.00046932 | 0.13932
296.97 | 0.00045072 | 0.13385
297.08 | 0.00043214 | 0.12838
297.19 | 0.00048092 | 0.14292
297.3 | 0.0004268 | 0.12689
297.41 | 0.00043002 | 0.12789
297.52 | 0.00044114 | 0.13125
297.63 | 0.00040506 | 0.12056
297.74 | 0.00040637 | 0.12099
297.85 | 0.00044304 | 0.13196
297.96 | 0.00044239 | 0.13181
298.07 | 0.00039679 | 0.11827
298.19 | 0.00042552 | 0.12689
298.3 | 0.00041318 | 0.12325
298.41 | 0.00039507 | 0.11789
298.52 | 0.00039258 | 0.11719
298.63 | 0.00039591 | 0.11823
298.74 | 0.00039146 | 0.11694
298.85 | 0.00040445 | 0.12087
298.96 | 0.00039612 | 0.11842
299.07 | 0.00037042 | 0.11078
299.18 | 0.00034669 | 0.10372
299.29 | 0.0003615 | 0.10819
299.4 | 0.00042037 | 0.12586
299.51 | 0.00040249 | 0.12055
299.62 | 0.0003272 | 0.098036
299.73 | 0.00038013 | 0.11394
299.84 | 0.00032803 | 0.098357
299.95 | 0.00029707 | 0.089106
300.06 | 0.00033848 | 0.10156
| 23.1014 | 6030.1598
261.0298