## Input Output
1) Paste emission spectrum from Ocean Optics spectrometer in column I | [5 empty] | wavelength | Spectrum |  | Absorbance | Wavelength | Absorbance
2) Modify column D in the Germicidal Calculator tab to calculate radiometer | [5 empty] | 200.11 | -7.7834e-05 | 400 |  | 200 | 1.164
sensitivity relative to the weighted average wavelength |  |  |  |  | Peak | 200.23 | 0.00036627 | 399 |  | 201 | 0.979
3) Paste water absorbance spectrum (200 - 300 nm) in column L |  |  |  |  | 277.8818 | 200.35 | -0.00024827 | 398 |  | 202 | 0.813
4) Enter radiometer factor for weighted average wavelength into radiometer |  |  |  |  | Rad. Factor: | 200.46 | 1.8916e-05 | 397 |  | 203 | 0.684
5) Enter irradiance reading for center of petri dish (mW/cm2) |  |  |  |  | 0.0010188 | 200.58 | 1.8746e-05 | 396 |  | 204 | 0.574
6) Enter petri factor | [5 empty] | 200.7 | 0.00033433 | 395 |  | 205 | 0.484
7) Enter reflection factor for weighted average wavelength | [5 empty] | 200.82 | 3.6812e-05 | 394 |  | 206 | 0.407
8) Enter sample depth and distance from sample to beam below. | [5 empty] | 200.94 | -7.2952e-05 | 393 |  | 207 | 0.341
9) Enter desired UV doses | [5 empty] | 201.06 | 0.00036142 | 392 |  | 208 | 0.297
10) Outputs are in blue | [5 empty] | 201.17 | 0.00012534 | 391 |  | 209 | 0.251
[6 empty] | 201.29 | -0.00021294 | 390 |  | 210 | 0.222
MP UV Dose Calculations | [5 empty] | 201.41 | 0.00024617 | 389 |  | 211 | 0.202
Input: | Irradiance Reading at 278 |  | 0.314 | mW/cm2 |  | 201.53 | -0.00020908 | 388 |  | 212 | 0.178
| Petri factor = |  | 0.92509 |  |  | 201.65 | 0.00032795 | 387 |  | 213 | 0.165
| Sample depth in the petri dish = |  | 0.6 | cm |  | 201.77 | 0.00054744 | 386 |  | 214 | 0.152
|  | Distance from sample to beam = | 4 | cm |  | 201.88 | 0.00010173 | 385 |  | 215 | 0.139
| Reflection Factor = |  | 0.97745 | at 278 nm |  | 202 | -0.00015118 | 384 |  | 216 | 0.131
Output: | Divergence Factor = |  | 0.86957 |  |  | 202.12 | 8.322e-05 | 383 |  | 217 | 0.125
| UVT (254) = |  | 92.2571 |  |  | 202.24 | 4.9475e-05 | 382 |  | 218 | 0.119
| Incident irradiance = |  | 0.24854 | mW/cm2 |  | 202.36 | 0.00027773 | 381 |  | 219 | 0.113
| Average irradiance = |  | 0.24428 | mW/cm2 |  | 202.47 | 0.00045315 | 380 |  | 220 | 0.108
|  | Germicidal irradiance = | 0.24215 | mW/cm2 |  | 202.59 | 0.00032065 | 379 |  | 221 | 0.104
|  | Action Spectrum Irradiance = | 0.23731 | mW/cm2 |  | 202.71 | 0.00044476 | 378 |  | 222 | 0.1
| UV Dose Required |  |  |  |  | 202.83 | -3.1477e-05 | 377 |  | 223 | 0.097
|  | Exposure Time |  |  |  | 202.95 | -0.0001248 | 376 |  | 224 | 0.092
| (mJ/cm2) | (sec) | min | sec |  | 203.06 | 0.00018555 | 375 |  | 225 | 0.089
| 5 | 20.4681 | 0 | 21 |  | 203.18 | 0.00024523 | 374 |  | 226 | 0.085
| 10 | 40.9362 | 0 | 41 |  | 203.3 | -3.039e-05 | 373 |  | 227 | 0.081
| 15 | 61.4043 | 1 | 2 |  | 203.42 | -1.5065e-05 | 372 |  | 228 | 0.077
| 20 | 81.8724 | 1 | 22 |  | 203.54 | 0 | 371 |  | 229 | 0.074
| 30 | 122.8087 | 2 | 3 |  | 203.65 | -0.00013327 | 370 |  | 230 | 0.07
[6 empty] | 203.77 | 0.00030831 | 369 |  | 231 | 0.067
[6 empty] | 203.89 | 0.00016014 | 368 |  | 232 | 0.064
[6 empty] | 204.01 | 8.6623e-05 | 367 |  | 233 | 0.06
[6 empty] | 204.13 | 0 | 366 |  | 234 | 0.057
[6 empty] | 204.24 | 0.00031237 | 365 |  | 235 | 0.055
[6 empty] | 204.36 | 0.00014077 | 364 |  | 236 | 0.052
[6 empty] | 204.48 | 0.00016743 | 363 |  | 237 | 0.05
[6 empty] | 204.6 | 0.00031807 | 362 |  | 238 | 0.049
[6 empty] | 204.72 | 0.00023296 | 361 |  | 239 | 0.046
[6 empty] | 204.83 | 8.1488e-05 | 360 |  | 240 | 0.046
[6 empty] | 204.95 | 0.00018844 | 359 |  | 241 | 0.044
[6 empty] | 205.07 | 0.00013345 | 358 |  | 242 | 0.042
[6 empty] | 205.19 | -2.6458e-05 | 357 |  | 243 | 0.042
[6 empty] | 205.3 | 0.00019672 | 356 |  | 244 | 0.041
[6 empty] | 205.42 | 0.00024704 | 355 |  | 245 | 0.04
[6 empty] | 205.54 | 0.00018052 | 354 |  | 246 | 0.04
[6 empty] | 205.66 | 0.00024293 | 353 |  | 247 | 0.039
[6 empty] | 205.78 | 0.00027903 | 352 |  | 248 | 0.039
[6 empty] | 205.89 | 0 | 351 |  | 249 | 0.038
[6 empty] | 206.01 | 0.0001249 | 350 |  | 250 | 0.037
[6 empty] | 206.13 | -0.00026038 | 349 |  | 251 | 0.036
[6 empty] | 206.25 | 7.3858e-05 | 348 |  | 252 | 0.036
[6 empty] | 206.37 | 0.00017108 | 347 |  | 253 | 0.035
[6 empty] | 206.48 | 3.6394e-05 | 346 |  | 254 | 0.035
[6 empty] | 206.6 | -0.00018071 | 345 |  | 255 | 0.034
[6 empty] | 206.72 | 2.3925e-05 | 344 |  | 256 | 0.034
[6 empty] | 206.84 | 0.00019007 | 343 |  | 257 | 0.032
[6 empty] | 206.95 | 0.00024775 | 342 |  | 258 | 0.033
[6 empty] | 207.07 | 0.00017578 | 341 |  | 259 | 0.032
[6 empty] | 207.19 | 9.313e-05 | 340 |  | 260 | 0.032
[6 empty] | 207.31 | 4.6255e-05 | 339 |  | 261 | 0.033
[6 empty] | 207.42 | 0.00010337 | 338 |  | 262 | 0.03
[6 empty] | 207.54 | 4.5634e-05 | 337 |  | 263 | 0.033
[6 empty] | 207.66 | 0.00022666 | 336 |  | 264 | 0.031
[6 empty] | 207.78 | -6.755e-05 | 335 |  | 265 | 0.03
[6 empty] | 207.9 | 0.00023482 | 334 |  | 266 | 0.031
[6 empty] | 208.01 | 0.00013326 | 333 |  | 267 | 0.031
[6 empty] | 208.13 | 0 | 332 |  | 268 | 0.03
[6 empty] | 208.25 | 6.5732e-05 | 331 |  | 269 | 0.03
[6 empty] | 208.37 | 0.00020675 | 330 |  | 270 | 0.029
[6 empty] | 208.48 | 0.00010808 | 329 |  | 271 | 0.028
[6 empty] | 208.6 | -0.00016102 | 328 |  | 272 | 0.029
[6 empty] | 208.72 | 0.0001493 | 327 |  | 273 | 0.029
[6 empty] | 208.84 | 5.2977e-05 | 326 |  | 274 | 0.027
[6 empty] | 208.95 | 0.00014739 | 325 |  | 275 | 0.027
[6 empty] | 209.07 | 4.1845e-05 | 324 |  | 276 | 0.026
[6 empty] | 209.19 | 0.00013516 | 323 |  | 277 | 0.026
[6 empty] | 209.31 | 0.00013434 | 322 |  | 278 | 0.025
[6 empty] | 209.42 | -0.0001746 | 321 |  | 279 | 0.025
[6 empty] | 209.54 | 3.0623e-05 | 320 |  | 280 | 0.024
[6 empty] | 209.66 | 0.00018259 | 319 |  | 281 | 0.023
[6 empty] | 209.78 | 0.00029235 | 318 |  | 282 | 0.023
[6 empty] | 209.89 | 6.0117e-05 | 317 |  | 283 | 0.022
[6 empty] | 210.01 | 2.9875e-05 | 316 |  | 284 | 0.022
[6 empty] | 210.13 | 0.00024746 | 315 |  | 285 | 0.02
[6 empty] | 210.25 | -0.00013773 | 314 |  | 286 | 0.019
[6 empty] | 210.36 | -9.7781e-06 | 313 |  | 287 | 0.017
[6 empty] | 210.48 | 7.7755e-05 | 312 |  | 288 | 0.017
[6 empty] | 210.6 | 7.7288e-05 | 311 |  | 289 | 0.016
[6 empty] | 210.72 | 0.00011524 | 310 |  | 290 | 0.016
[6 empty] | 210.83 | 0.00011455 | 309 |  | 291 | 0.014
[6 empty] | 210.95 | 0.00012337 | 308 |  | 292 | 0.014
[6 empty] | 211.07 | 0.00013208 | 307 |  | 293 | 0.013
[6 empty] | 211.19 | 0.00011255 | 306 |  | 294 | 0.012
[6 empty] | 211.3 | 0.00013055 | 305 |  | 295 | 0.011
[6 empty] | 211.42 | 0.00010196 | 304 |  | 296 | 0.011
[6 empty] | 211.54 | 7.3721e-05 | 303 |  | 297 | 0.011
[6 empty] | 211.66 | 0.00021073 | 302 |  | 298 | 0.011
[6 empty] | 211.77 | -0.00013662 | 301 |  | 299 | 0.01
[6 empty] | 211.89 | -4.5273e-05 | 300 |  | 300 | 0.01
[6 empty] | 212.01 | 0.00010805 | 299 |  | 301
[6 empty] | 212.13 | 1.7908e-05 | 298 |  | 302
[6 empty] | 212.24 | 0.00015135 | 297 |  | 303
[6 empty] | 212.36 | 8.8517e-06 | 296 |  | 304
[6 empty] | 212.48 | 2.6402e-05 | 295 |  | 305
[6 empty] | 212.59 | 7.001e-05 | 294 |  | 306
[6 empty] | 212.71 | 2.6106e-05 | 293 |  | 307
[6 empty] | 212.83 | 0.00016439 | 292 |  | 308
[6 empty] | 212.95 | 3.4412e-05 | 291 |  | 309
[6 empty] | 213.06 | -5.133e-05 | 290 |  | 310
[6 empty] | 213.18 | 2.5521e-05 | 289 |  | 311
[6 empty] | 213.3 | 0.00010151 | 288 |  | 312
[6 empty] | 213.42 | 0.00010093 | 287 |  | 313
[6 empty] | 213.53 | 4.1817e-05 | 286 |  | 314
[6 empty] | 213.65 | 4.1591e-05 | 285 |  | 315
[6 empty] | 213.77 | 0.00015722 | 284 |  | 316
[6 empty] | 213.88 | 0.00010699 | 283 |  | 317
[6 empty] | 214 | 9.0051e-05 | 282 |  | 318
[6 empty] | 214.12 | 2.4428e-05 | 281 |  | 319
[6 empty] | 214.24 | -8.1008e-06 | 280 |  | 320
[6 empty] | 214.35 | 0.00020959 | 279 |  | 321
[6 empty] | 214.47 | 2.4059e-05 | 278 |  | 322
[6 empty] | 214.59 | 7.9792e-06 | 277 |  | 323
[6 empty] | 214.7 | 5.5575e-05 | 276 |  | 324
[6 empty] | 214.82 | 3.16e-05 | 275 |  | 325
[6 empty] | 214.94 | 7.86e-06 | 274 |  | 326
[6 empty] | 215.06 | 4.6914e-05 | 273 |  | 327
[6 empty] | 215.17 | 0.0001089 | 272 |  | 328
[6 empty] | 215.29 | -6.191e-05 | 271 |  | 329
[6 empty] | 215.41 | 7.7002e-05 | 270 |  | 330
[6 empty] | 215.52 | -7.66e-06 | 269 |  | 331
[6 empty] | 215.64 | 7.6194e-05 | 268 |  | 332
[6 empty] | 215.76 | 0.00014404 | 267 |  | 333
[6 empty] | 215.87 | 8.2954e-05 | 266 |  | 334
[6 empty] | 215.99 | 6.7525e-05 | 265 |  | 335
[6 empty] | 216.11 | 8.956e-05 | 264 |  | 336
[6 empty] | 216.23 | -1.4851e-05 | 263 |  | 337
[6 empty] | 216.34 | 0.00012559 | 262 |  | 338
[6 empty] | 216.46 | 8.0843e-05 | 261 |  | 339
[6 empty] | 216.58 | 0.00019743 | 260 |  | 340
[6 empty] | 216.69 | 6.5479e-05 | 259 |  | 341
[6 empty] | 216.81 | 9.4114e-05 | 258 |  | 342
[6 empty] | 216.93 | 0.00010085 | 257 |  | 343
[6 empty] | 217.04 | 0.00010036 | 256 |  | 344
[6 empty] | 217.16 | -0.00011415 | 255 |  | 345
[6 empty] | 217.28 | -1.4201e-05 | 254 |  | 346
[6 empty] | 217.4 | 4.2398e-05 | 253 |  | 347
[6 empty] | 217.51 | 4.9234e-05 | 252 |  | 348
[6 empty] | 217.63 | 0.00017502 | 251 |  | 349
[6 empty] | 217.75 | 0.00011151 | 250 |  | 350
[6 empty] | 217.86 | 0.00010405 | 249 |  | 351
[6 empty] | 217.98 | 1.3808e-05 | 248 |  | 352
[6 empty] | 218.1 | 6.8718e-06 | 247 |  | 353
[6 empty] | 218.21 | 0.00011626 | 246 |  | 354
[6 empty] | 218.33 | 0.00010889 | 245 |  | 355
[6 empty] | 218.45 | -6.7728e-06 | 244 |  | 356
[6 empty] | 218.56 | 0.00022238 | 243 |  | 357
[6 empty] | 218.68 | 4.6951e-05 | 242 |  | 358
[6 empty] | 218.8 | 5.3394e-05 | 241 |  | 359
[6 empty] | 218.91 | -1.3284e-05 | 240 |  | 360
[6 empty] | 219.03 | 1.3218e-05 | 239 |  | 361
[6 empty] | 219.15 | 7.2349e-05 | 238 |  | 362
[6 empty] | 219.26 | 1.9633e-05 | 237 |  | 363
[6 empty] | 219.38 | 0.00013676 | 236 |  | 364
[6 empty] | 219.5 | -1.9441e-05 | 235 |  | 365
[6 empty] | 219.62 | 1.9345e-05 | 234 |  | 366
[6 empty] | 219.73 | 0.00010265 | 233 |  | 367
[6 empty] | 219.85 | 8.9395e-05 | 232 |  | 368
[6 empty] | 219.97 | -6.3542e-06 | 231 |  | 369
[6 empty] | 220.08 | -1.8972e-05 | 230 |  | 370
[6 empty] | 220.2 | 0.00010071 | 229 |  | 371
[6 empty] | 220.32 | 9.3968e-05 | 228 |  | 372
[6 empty] | 220.43 | -6.234e-06 | 227 |  | 373
[6 empty] | 220.55 | -4.343e-05 | 226 |  | 374
[6 empty] | 220.67 | 0.00014203 | 225 |  | 375
[6 empty] | 220.78 | 9.8334e-05 | 224 |  | 376
[6 empty] | 220.9 | 2.4466e-05 | 223 |  | 377
[6 empty] | 221.02 | 6.0881e-06 | 222 |  | 378
[6 empty] | 221.13 | 0.00011511 | 221 |  | 379
[6 empty] | 221.25 | 8.4399e-05 | 220 |  | 380
[6 empty] | 221.37 | 6.0001e-05 | 219 |  | 381
[6 empty] | 221.48 | 0.00014928 | 218 |  | 382
[6 empty] | 221.6 | 5.9418e-05 | 217 |  | 383
[6 empty] | 221.71 | 1.774e-05 | 216 |  | 384
[6 empty] | 221.83 | 5.8839e-05 | 215 |  | 385
[6 empty] | 221.95 | 9.3658e-05 | 214 |  | 386
[6 empty] | 222.06 | 7.5714e-05 | 213 |  | 387
[6 empty] | 222.18 | 4.6363e-05 | 212 |  | 388
[6 empty] | 222.3 | 5.1898e-05 | 211 |  | 389
[6 empty] | 222.41 | 8.0325e-05 | 210 |  | 390
[6 empty] | 222.53 | 0 | 209 |  | 391
[6 empty] | 222.65 | -3.9749e-05 | 208 |  | 392
[6 empty] | 222.76 | 5.6481e-06 | 207 |  | 393
[6 empty] | 222.88 | 5.0573e-05 | 206 |  | 394
[6 empty] | 223 | 0.00010062 | 205 |  | 395
[6 empty] | 223.11 | 5.006e-05 | 204 |  | 396
[6 empty] | 223.23 | 0.0001107 | 203 |  | 397
[6 empty] | 223.35 | 9.3628e-05 | 202 |  | 398
[6 empty] | 223.46 | 8.2203e-05 | 201 |  | 399
[6 empty] | 223.58 | 8.1793e-05 | 200 |  | 400
[6 empty] | 223.7 | 0.00013021
[6 empty] | 223.81 | 3.7789e-05
[6 empty] | 223.93 | 4.2968e-05
[6 empty] | 224.04 | 6.4135e-05
[6 empty] | 224.16 | 3.1906e-05
[6 empty] | 224.28 | -2.6459e-05
[6 empty] | 224.39 | 1.0532e-05
[6 empty] | 224.51 | -5.2396e-05
[6 empty] | 224.63 | 4.6914e-05
[6 empty] | 224.74 | 2.0744e-05
[6 empty] | 224.86 | 0.0001032
[6 empty] | 224.98 | -2.5669e-05
[6 empty] | 225.09 | 0.0001277
[6 empty] | 225.21 | 0.00010672
[6 empty] | 225.32 | 2.5279e-05
[6 empty] | 225.44 | 8.0479e-05
[6 empty] | 225.56 | -3.502e-05
[6 empty] | 225.67 | 9.9511e-06
[6 empty] | 225.79 | -5.444e-05
[6 empty] | 225.91 | 6.4001e-05
[6 empty] | 226.02 | 8.8143e-05
[6 empty] | 226.14 | -5.845e-05
[6 empty] | 226.25 | 8.2375e-05
[6 empty] | 226.37 | 2.8923e-05
[6 empty] | 226.49 | -4.795e-05
[6 empty] | 226.6 | 4.7701e-06
[6 empty] | 226.72 | 3.7965e-05
[6 empty] | 226.83 | -1.8886e-05
[6 empty] | 226.95 | 9.3951e-06
[6 empty] | 227.07 | 3.739e-05
[6 empty] | 227.18 | -1.3952e-05
[6 empty] | 227.3 | 6.4792e-05
[6 empty] | 227.42 | 5.9861e-05
[6 empty] | 227.53 | 0
[6 empty] | 227.65 | 3.1904e-05
[6 empty] | 227.76 | -4.5347e-06
[6 empty] | 227.88 | 0.00013988
[6 empty] | 228 | 0.00012572
[6 empty] | 228.11 | 3.1272e-05
[6 empty] | 228.23 | 0
[6 empty] | 228.34 | 8.4033e-05
[6 empty] | 228.46 | 3.0799e-05
[6 empty] | 228.58 | 1.751e-05
[6 empty] | 228.69 | -6.9689e-05
[6 empty] | 228.81 | -3.4668e-05
[6 empty] | 228.92 | 2.1561e-05
[6 empty] | 229.04 | 8.5802e-06
[6 empty] | 229.16 | 5.5492e-05
[6 empty] | 229.27 | 2.1235e-05
[6 empty] | 229.39 | 3.3801e-05
[6 empty] | 229.5 | 8.4086e-06
[6 empty] | 229.62 | -1.2553e-05
[6 empty] | 229.74 | 2.499e-05
[6 empty] | 229.85 | -4.145e-06
[6 empty] | 229.97 | 6.1881e-05
[6 empty] | 230.08 | -8.2109e-06
[6 empty] | 230.2 | 4.0863e-05
[6 empty] | 230.32 | 6.5086e-05
[6 empty] | 230.43 | 8.9073e-05
[6 empty] | 230.55 | 4.0305e-06
[6 empty] | 230.66 | 9.2309e-05
[6 empty] | 230.78 | 1.5981e-05
[6 empty] | 230.89 | 3.9766e-06
[6 empty] | 231.01 | 3.5627e-05
[6 empty] | 231.13 | 5.1227e-05
[6 empty] | 231.24 | 7.0613e-05
[6 empty] | 231.36 | 5.8593e-05
[6 empty] | 231.47 | 5.8342e-05
[6 empty] | 231.59 | 9.6813e-05
[6 empty] | 231.71 | -1.1568e-05
[6 empty] | 231.82 | -4.9908e-05
[6 empty] | 231.94 | 5.7341e-05
[6 empty] | 232.05 | 5.7106e-05
[6 empty] | 232.17 | -2.654e-05
[6 empty] | 232.28 | -2.6435e-05
[6 empty] | 232.4 | 4.8903e-05
[6 empty] | 232.52 | 4.4969e-05
[6 empty] | 232.63 | 2.6132e-05
[6 empty] | 232.75 | 2.2314e-05
[6 empty] | 232.86 | 0.00010004
[6 empty] | 232.98 | 2.5841e-05
[6 empty] | 233.09 | 7.3566e-05
[6 empty] | 233.21 | 4.7642e-05
[6 empty] | 233.33 | -1.826e-05
[6 empty] | 233.44 | 1.4559e-05
[6 empty] | 233.56 | 5.4405e-05
[6 empty] | 233.67 | 2.892e-05
[6 empty] | 233.79 | 1.0807e-05
[6 empty] | 233.9 | 3.9492e-05
[6 empty] | 234.02 | 3.2199e-05
[6 empty] | 234.13 | 3.566e-06
[6 empty] | 234.25 | 1.0664e-05
[6 empty] | 234.37 | 3.1884e-05
[6 empty] | 234.48 | -2.8249e-05
[6 empty] | 234.6 | 1.4077e-05
[6 empty] | 234.71 | 7.0161e-06
[6 empty] | 234.83 | -6.9936e-06
[6 empty] | 234.94 | 5.5767e-05
[6 empty] | 235.06 | 4.517e-05
[6 empty] | 235.17 | 2.0785e-05
[6 empty] | 235.29 | 3.7999e-05
[6 empty] | 235.41 | 5.1662e-05
[6 empty] | 235.52 | 3.0905e-05
[6 empty] | 235.64 | 6.1634e-05
[6 empty] | 235.75 | -5.8045e-05
[6 empty] | 235.87 | -1.7024e-05
[6 empty] | 235.98 | 0.00010525
[6 empty] | 236.1 | 5.4176e-05
[6 empty] | 236.21 | 4.3903e-05
[6 empty] | 236.33 | 6.7363e-05
[6 empty] | 236.44 | 2.6872e-05
[6 empty] | 236.56 | 6.7001e-05
[6 empty] | 236.68 | 2.005e-05
[6 empty] | 236.79 | 5.9996e-05
[6 empty] | 236.91 | -4.3228e-05
[6 empty] | 237.02 | 6.6344e-06
[6 empty] | 237.14 | 9.9288e-06
[6 empty] | 237.25 | 1.9814e-05
[6 empty] | 237.37 | 6.5898e-06
[6 empty] | 237.48 | 5.2595e-05
[6 empty] | 237.6 | 3.2805e-05
[6 empty] | 237.71 | -9.8234e-06
[6 empty] | 237.83 | 2.6145e-05
[6 empty] | 237.94 | 5.5446e-05
[6 empty] | 238.06 | 6.5111e-06
[6 empty] | 238.17 | 1.9498e-05
[6 empty] | 238.29 | 4.8656e-05
[6 empty] | 238.4 | 5.8279e-05
[6 empty] | 238.52 | 4.2009e-05
[6 empty] | 238.64 | 2.2579e-05
[6 empty] | 238.75 | 5.7962e-05
[6 empty] | 238.87 | 3.8572e-05
[6 empty] | 238.98 | 2.8877e-05
[6 empty] | 239.1 | 3.2029e-06
[6 empty] | 239.21 | 4.7961e-05
[6 empty] | 239.33 | 7.6613e-05
[6 empty] | 239.44 | 6.0543e-05
[6 empty] | 239.56 | 1.5908e-05
[6 empty] | 239.67 | -1.5884e-05
[6 empty] | 239.79 | 4.4404e-05
[6 empty] | 239.9 | 4.4339e-05
[6 empty] | 240.02 | 5.6917e-05
[6 empty] | 240.13 | 2.5261e-05
[6 empty] | 240.25 | 7.5671e-05
[6 empty] | 240.36 | 1.2595e-05
[6 empty] | 240.48 | 4.7172e-05
[6 empty] | 240.59 | 1.5704e-05
[6 empty] | 240.71 | 7.2164e-05
[6 empty] | 240.82 | 6.5811e-05
[6 empty] | 240.94 | 6.5734e-05
[6 empty] | 241.05 | 3.752e-05
[6 empty] | 241.17 | 4.06e-05
[6 empty] | 241.28 | 5.3042e-05
[6 empty] | 241.4 | 7.169e-05
[6 empty] | 241.51 | 7.1616e-05
[6 empty] | 241.63 | 2.7994e-05
[6 empty] | 241.74 | 3.7283e-05
[6 empty] | 241.86 | 3.4139e-05
[6 empty] | 241.97 | 5.2702e-05
[6 empty] | 242.09 | -6.1935e-06
[6 empty] | 242.2 | 6.1873e-05
[6 empty] | 242.32 | 5.5635e-05
[6 empty] | 242.43 | 7.1019e-05
[6 empty] | 242.55 | 3.0848e-06
[6 empty] | 242.66 | 1.8493e-05
[6 empty] | 242.78 | 4.0037e-05
[6 empty] | 242.89 | -9.2329e-06
[6 empty] | 243.01 | 3.3832e-05
[6 empty] | 243.12 | 1.5368e-05
[6 empty] | 243.24 | 2.4572e-05
[6 empty] | 243.35 | 3.0693e-05
[6 empty] | 243.47 | 3.0673e-06
[6 empty] | 243.58 | 2.4522e-05
[6 empty] | 243.7 | 7.046e-05
[6 empty] | 243.81 | 3.062e-06
[6 empty] | 243.93 | 1.8361e-05
[6 empty] | 244.04 | 3.0583e-05
[6 empty] | 244.16 | 7.6408e-05
[6 empty] | 244.27 | 1.5274e-05
[6 empty] | 244.38 | 6.7178e-05
[6 empty] | 244.5 | 5.7994e-05
[6 empty] | 244.61 | 1.8308e-05
[6 empty] | 244.73 | -3.0499e-06
[6 empty] | 244.84 | 3.0488e-05
[6 empty] | 244.96 | 8.2275e-05
[6 empty] | 245.07 | 0
[6 empty] | 245.19 | 7.9142e-05
[6 empty] | 245.3 | 3.0428e-05
[6 empty] | 245.42 | 6.3875e-05
[6 empty] | 245.53 | -3.0403e-06
[6 empty] | 245.65 | 6.0774e-05
[6 empty] | 245.76 | 3.9477e-05
[6 empty] | 245.88 | 2.428e-05
[6 empty] | 245.99 | 7.281e-05
[6 empty] | 246.11 | 9.0979e-05
[6 empty] | 246.22 | 6.973e-05
[6 empty] | 246.33 | 5.4555e-05
[6 empty] | 246.45 | 1.212e-05
[6 empty] | 246.56 | 3.0288e-05
[6 empty] | 246.68 | 0
[6 empty] | 246.79 | 3.3292e-05
[6 empty] | 246.91 | 4.5388e-05
[6 empty] | 247.02 | 3.9333e-05
[6 empty] | 247.14 | 3.3281e-05
[6 empty] | 247.25 | 9.9848e-05
[6 empty] | 247.37 | 4.2362e-05
[6 empty] | 247.48 | -1.8157e-05
[6 empty] | 247.59 | 3.935e-05
[6 empty] | 247.71 | 2.4218e-05
[6 empty] | 247.82 | 6.9624e-05
[6 empty] | 247.94 | 6.0557e-06
[6 empty] | 248.05 | 5.4521e-05
[6 empty] | 248.17 | 3.6353e-05
[6 empty] | 248.28 | 4.8478e-05
[6 empty] | 248.4 | 6.3637e-05
[6 empty] | 248.51 | 3.9396e-05
[6 empty] | 248.62 | 9.0914e-06
[6 empty] | 248.74 | 7.2734e-05
[6 empty] | 248.85 | 3.6363e-05
[6 empty] | 248.97 | 7.2726e-05
[6 empty] | 249.08 | 0.00015763
[6 empty] | 249.2 | 7.8829e-05
[6 empty] | 249.31 | 7.5802e-05
[6 empty] | 249.43 | 2.7296e-05
[6 empty] | 249.54 | 4.854e-05
[6 empty] | 249.65 | 3.9443e-05
[6 empty] | 249.77 | 8.1942e-05
[6 empty] | 249.88 | 7.5887e-05
[6 empty] | 250 | 8.8045e-05
[6 empty] | 250.11 | 4.2514e-05
[6 empty] | 250.23 | 7.2889e-05
[6 empty] | 250.34 | 7.5924e-05
[6 empty] | 250.45 | 0.00010326
[6 empty] | 250.57 | 7.8967e-05
[6 empty] | 250.68 | 2.4298e-05
[6 empty] | 250.8 | 4.8601e-05
[6 empty] | 250.91 | 4.2534e-05
[6 empty] | 251.03 | 0.00010027
[6 empty] | 251.14 | 6.9905e-05
[6 empty] | 251.25 | 8.5127e-05
[6 empty] | 251.37 | 2.4331e-05
[6 empty] | 251.48 | 4.26e-05
[6 empty] | 251.6 | 7.6104e-05
[6 empty] | 251.71 | 0.00010048
[6 empty] | 251.82 | 0.00011576
[6 empty] | 251.94 | 9.1427e-06
[6 empty] | 252.05 | 6.7062e-05
[6 empty] | 252.17 | 7.9278e-05
[6 empty] | 252.28 | 5.4898e-05
[6 empty] | 252.4 | 7.3208e-05
[6 empty] | 252.51 | 7.0159e-05
[6 empty] | 252.62 | 8.541e-05
[6 empty] | 252.74 | 8.8474e-05
[6 empty] | 252.85 | 3.3563e-05
[6 empty] | 252.97 | 7.6287e-05
[6 empty] | 253.08 | 7.9332e-05
[6 empty] | 253.19 | 7.628e-05
[6 empty] | 253.31 | 6.7133e-05
[6 empty] | 253.42 | 6.7143e-05
[6 empty] | 253.54 | 5.7993e-05
[6 empty] | 253.65 | 3.9686e-05
[6 empty] | 253.76 | 4.8855e-05
[6 empty] | 253.88 | 0.0001069
[6 empty] | 253.99 | 0.00011913
[6 empty] | 254.11 | 7.6369e-05
[6 empty] | 254.22 | 9.7778e-05
[6 empty] | 254.33 | 8.5571e-05
[6 empty] | 254.45 | 0.00011309
[6 empty] | 254.56 | 6.4205e-05
[6 empty] | 254.68 | 0.00011619
[6 empty] | 254.79 | 9.1731e-05
[6 empty] | 254.9 | 9.4795e-05
[6 empty] | 255.02 | 8.5635e-05
[6 empty] | 255.13 | 0.00015902
[6 empty] | 255.25 | 7.0337e-05
[6 empty] | 255.36 | 0.00015596
[6 empty] | 255.47 | 0.00012538
[6 empty] | 255.59 | 7.6467e-05
[6 empty] | 255.7 | 0.00013767
[6 empty] | 255.81 | 7.6487e-05
[6 empty] | 255.93 | 0.00011932
[6 empty] | 256.04 | 0.00015298
[6 empty] | 256.16 | 0.00015297
[6 empty] | 256.27 | 5.2002e-05
[6 empty] | 256.38 | 9.7889e-05
[6 empty] | 256.5 | 0.00011626
[6 empty] | 256.61 | 8.8719e-05
[6 empty] | 256.72 | 0.0001499
[6 empty] | 256.84 | 0.00010094
[6 empty] | 256.95 | 0.0001468
[6 empty] | 257.07 | 0.00014371
[6 empty] | 257.18 | 9.784e-05
[6 empty] | 257.29 | 0.00015592
[6 empty] | 257.41 | 0.00012841
[6 empty] | 257.52 | 0.00012842
[6 empty] | 257.63 | 0.0001223
[6 empty] | 257.75 | 0.00013145
[6 empty] | 257.86 | 0.00010698
[6 empty] | 257.98 | 0.0001864
[6 empty] | 258.09 | 0.0001375
[6 empty] | 258.2 | 0.0001283
[6 empty] | 258.32 | 0.00011606
[6 empty] | 258.43 | 0.0001527
[6 empty] | 258.54 | 0.00013741
[6 empty] | 258.66 | 0.00012822
[6 empty] | 258.77 | 0.00016786
[6 empty] | 258.88 | 0.00018918
[6 empty] | 259 | 0.00018611
[6 empty] | 259.11 | 0.0001708
[6 empty] | 259.23 | 0.00021345
[6 empty] | 259.34 | 0.00014328
[6 empty] | 259.45 | 0.00021334
[6 empty] | 259.57 | 0.00019196
[6 empty] | 259.68 | 0.00019494
[6 empty] | 259.79 | 0.00018269
[6 empty] | 259.91 | 0.00011264
[6 empty] | 260.02 | 0.00020087
[6 empty] | 260.13 | 0.00023732
[6 empty] | 260.25 | 0.00017942
[6 empty] | 260.36 | 0.00019762
[6 empty] | 260.47 | 0.00024013
[6 empty] | 260.59 | 0.00024914
[6 empty] | 260.7 | 0.000243
[6 empty] | 260.81 | 0.00020953
[6 empty] | 260.93 | 0.00022464
[6 empty] | 261.04 | 0.00027316
[6 empty] | 261.15 | 0.00030339
[6 empty] | 261.27 | 0.00030028
[6 empty] | 261.38 | 0.00033962
[6 empty] | 261.49 | 0.00037904
[6 empty] | 261.61 | 0.000385
[6 empty] | 261.72 | 0.00039709
[6 empty] | 261.83 | 0.00043034
[6 empty] | 261.95 | 0.00046963
[6 empty] | 262.06 | 0.0004392
[6 empty] | 262.17 | 0.00056324
[6 empty] | 262.29 | 0.00058735
[6 empty] | 262.4 | 0.00062658
[6 empty] | 262.51 | 0.00079876
[6 empty] | 262.63 | 0.00071989
[6 empty] | 262.74 | 0.00076187
[6 empty] | 262.85 | 0.0008524
[6 empty] | 262.97 | 0.00093052
[6 empty] | 263.08 | 0.00097255
[6 empty] | 263.19 | 0.0010688
[6 empty] | 263.31 | 0.0012104
[6 empty] | 263.42 | 0.0013429
[6 empty] | 263.53 | 0.0014633
[6 empty] | 263.65 | 0.0015624
[6 empty] | 263.76 | 0.0016495
[6 empty] | 263.87 | 0.0017697
[6 empty] | 263.99 | 0.0019233
[6 empty] | 264.1 | 0.0020492
[6 empty] | 264.21 | 0.0021664
[6 empty] | 264.33 | 0.0023886
[6 empty] | 264.44 | 0.0025809
[6 empty] | 264.55 | 0.0027761
[6 empty] | 264.67 | 0.0028835
[6 empty] | 264.78 | 0.0032645
[6 empty] | 264.89 | 0.0033359
[6 empty] | 265.01 | 0.0035154
[6 empty] | 265.12 | 0.0037818
[6 empty] | 265.23 | 0.004069
[6 empty] | 265.34 | 0.0044225
[6 empty] | 265.46 | 0.00465
[6 empty] | 265.57 | 0.0048082
[6 empty] | 265.68 | 0.0052575
[6 empty] | 265.8 | 0.0055894
[6 empty] | 265.91 | 0.0057415
[6 empty] | 266.02 | 0.0061094
[6 empty] | 266.14 | 0.0066324
[6 empty] | 266.25 | 0.0070143
[6 empty] | 266.36 | 0.0072315
[6 empty] | 266.47 | 0.0076338
[6 empty] | 266.59 | 0.0082164
[6 empty] | 266.7 | 0.0084623
[6 empty] | 266.81 | 0.0088292
[6 empty] | 266.93 | 0.0094594
[6 empty] | 267.04 | 0.0098763
[6 empty] | 267.15 | 0.01026
[6 empty] | 267.26 | 0.010846
[6 empty] | 267.38 | 0.011035
[6 empty] | 267.49 | 0.011648
[6 empty] | 267.6 | 0.012295
[6 empty] | 267.72 | 0.012822
[6 empty] | 267.83 | 0.01344
[6 empty] | 267.94 | 0.014027
[6 empty] | 268.06 | 0.01496
[6 empty] | 268.17 | 0.015364
[6 empty] | 268.28 | 0.016076
[6 empty] | 268.39 | 0.016615
[6 empty] | 268.51 | 0.017414
[6 empty] | 268.62 | 0.018619
[6 empty] | 268.73 | 0.01935
[6 empty] | 268.84 | 0.020151
[6 empty] | 268.96 | 0.021046
[6 empty] | 269.07 | 0.022178
[6 empty] | 269.18 | 0.022997
[6 empty] | 269.3 | 0.023656
[6 empty] | 269.41 | 0.024988
[6 empty] | 269.52 | 0.026145
[6 empty] | 269.63 | 0.027264
[6 empty] | 269.75 | 0.028475
[6 empty] | 269.86 | 0.029522
[6 empty] | 269.97 | 0.031109
[6 empty] | 270.08 | 0.032479
[6 empty] | 270.2 | 0.033403
[6 empty] | 270.31 | 0.03524
[6 empty] | 270.42 | 0.036504
[6 empty] | 270.54 | 0.038683
[6 empty] | 270.65 | 0.039872
[6 empty] | 270.76 | 0.041704
[6 empty] | 270.87 | 0.043431
[6 empty] | 270.99 | 0.044993
[6 empty] | 271.1 | 0.047179
[6 empty] | 271.21 | 0.049024
[6 empty] | 271.32 | 0.051175
[6 empty] | 271.44 | 0.052579
[6 empty] | 271.55 | 0.054768
[6 empty] | 271.66 | 0.057099
[6 empty] | 271.77 | 0.059166
[6 empty] | 271.89 | 0.06155
[6 empty] | 272 | 0.063507
[6 empty] | 272.11 | 0.0661
[6 empty] | 272.22 | 0.068172
[6 empty] | 272.34 | 0.070005
[6 empty] | 272.45 | 0.072728
[6 empty] | 272.56 | 0.074899
[6 empty] | 272.67 | 0.077201
[6 empty] | 272.79 | 0.079744
[6 empty] | 272.9 | 0.082101
[6 empty] | 273.01 | 0.085249
[6 empty] | 273.12 | 0.086879
[6 empty] | 273.24 | 0.089383
[6 empty] | 273.35 | 0.092318
[6 empty] | 273.46 | 0.094695
[6 empty] | 273.57 | 0.096263
[6 empty] | 273.69 | 0.099324
[6 empty] | 273.8 | 0.10197
[6 empty] | 273.91 | 0.10377
[6 empty] | 274.02 | 0.10572
[6 empty] | 274.13 | 0.10818
[6 empty] | 274.25 | 0.10977
[6 empty] | 274.36 | 0.11211
[6 empty] | 274.47 | 0.11453
[6 empty] | 274.58 | 0.11567
[6 empty] | 274.7 | 0.11778
[6 empty] | 274.81 | 0.11945
[6 empty] | 274.92 | 0.12133
[6 empty] | 275.03 | 0.123
[6 empty] | 275.15 | 0.12391
[6 empty] | 275.26 | 0.12592
[6 empty] | 275.37 | 0.12784
[6 empty] | 275.48 | 0.12913
[6 empty] | 275.59 | 0.13032
[6 empty] | 275.71 | 0.13106
[6 empty] | 275.82 | 0.13175
[6 empty] | 275.93 | 0.13296
[6 empty] | 276.04 | 0.13366
[6 empty] | 276.16 | 0.13458
[6 empty] | 276.27 | 0.13479
[6 empty] | 276.38 | 0.13542
[6 empty] | 276.49 | 0.13514
[6 empty] | 276.6 | 0.13592
[6 empty] | 276.72 | 0.13514
[6 empty] | 276.83 | 0.13492
[6 empty] | 276.94 | 0.13507
[6 empty] | 277.05 | 0.13461
[6 empty] | 277.16 | 0.13486
[6 empty] | 277.28 | 0.13444
[6 empty] | 277.39 | 0.13367
[6 empty] | 277.5 | 0.13207
[6 empty] | 277.61 | 0.13203
[6 empty] | 277.73 | 0.13089
[6 empty] | 277.84 | 0.12973
[6 empty] | 277.95 | 0.12853
[6 empty] | 278.06 | 0.12776
[6 empty] | 278.17 | 0.12639
[6 empty] | 278.29 | 0.12562
[6 empty] | 278.4 | 0.1236
[6 empty] | 278.51 | 0.12239
[6 empty] | 278.62 | 0.12085
[6 empty] | 278.73 | 0.1192
[6 empty] | 278.85 | 0.1178
[6 empty] | 278.96 | 0.11569
[6 empty] | 279.07 | 0.11416
[6 empty] | 279.18 | 0.11238
[6 empty] | 279.29 | 0.11037
[6 empty] | 279.4 | 0.10903
[6 empty] | 279.52 | 0.10723
[6 empty] | 279.63 | 0.10582
[6 empty] | 279.74 | 0.10377
[6 empty] | 279.85 | 0.10139
[6 empty] | 279.96 | 0.099917
[6 empty] | 280.08 | 0.097639
[6 empty] | 280.19 | 0.095841
[6 empty] | 280.3 | 0.094178
[6 empty] | 280.41 | 0.091876
[6 empty] | 280.52 | 0.089917
[6 empty] | 280.64 | 0.088224
[6 empty] | 280.75 | 0.086742
[6 empty] | 280.86 | 0.084764
[6 empty] | 280.97 | 0.082634
[6 empty] | 281.08 | 0.0812
[6 empty] | 281.19 | 0.079557
[6 empty] | 281.31 | 0.07681
[6 empty] | 281.42 | 0.075487
[6 empty] | 281.53 | 0.073705
[6 empty] | 281.64 | 0.072067
[6 empty] | 281.75 | 0.070591
[6 empty] | 281.86 | 0.068989
[6 empty] | 281.98 | 0.067541
[6 empty] | 282.09 | 0.065458
[6 empty] | 282.2 | 0.063986
[6 empty] | 282.31 | 0.062572
[6 empty] | 282.42 | 0.060517
[6 empty] | 282.53 | 0.059364
[6 empty] | 282.65 | 0.057907
[6 empty] | 282.76 | 0.056489
[6 empty] | 282.87 | 0.055074
[6 empty] | 282.98 | 0.053491
[6 empty] | 283.09 | 0.052285
[6 empty] | 283.2 | 0.051008
[6 empty] | 283.32 | 0.04953
[6 empty] | 283.43 | 0.048251
[6 empty] | 283.54 | 0.046577
[6 empty] | 283.65 | 0.04584
[6 empty] | 283.76 | 0.044677
[6 empty] | 283.87 | 0.043424
[6 empty] | 283.99 | 0.042715
[6 empty] | 284.1 | 0.04132
[6 empty] | 284.21 | 0.040417
[6 empty] | 284.32 | 0.039292
[6 empty] | 284.43 | 0.03813
[6 empty] | 284.54 | 0.037448
[6 empty] | 284.65 | 0.036375
[6 empty] | 284.77 | 0.035724
[6 empty] | 284.88 | 0.034296
[6 empty] | 284.99 | 0.033574
[6 empty] | 285.1 | 0.032635
[6 empty] | 285.21 | 0.031815
[6 empty] | 285.32 | 0.030598
[6 empty] | 285.43 | 0.030192
[6 empty] | 285.55 | 0.029133
[6 empty] | 285.66 | 0.028774
[6 empty] | 285.77 | 0.027561
[6 empty] | 285.88 | 0.027527
[6 empty] | 285.99 | 0.026574
[6 empty] | 286.1 | 0.025747
[6 empty] | 286.21 | 0.025313
[6 empty] | 286.33 | 0.024611
[6 empty] | 286.44 | 0.023831
[6 empty] | 286.55 | 0.023347
[6 empty] | 286.66 | 0.02307
[6 empty] | 286.77 | 0.022369
[6 empty] | 286.88 | 0.021779
[6 empty] | 286.99 | 0.021127
[6 empty] | 287.1 | 0.020549
[6 empty] | 287.22 | 0.020062
[6 empty] | 287.33 | 0.019566
[6 empty] | 287.44 | 0.019143
[6 empty] | 287.55 | 0.018871
[6 empty] | 287.66 | 0.018074
[6 empty] | 287.77 | 0.01796
[6 empty] | 287.88 | 0.017461
[6 empty] | 287.99 | 0.016975
[6 empty] | 288.11 | 0.016616
[6 empty] | 288.22 | 0.0162
[6 empty] | 288.33 | 0.015998
[6 empty] | 288.44 | 0.015434
[6 empty] | 288.55 | 0.015244
[6 empty] | 288.66 | 0.014723
[6 empty] | 288.77 | 0.014368
[6 empty] | 288.88 | 0.014237
[6 empty] | 288.99 | 0.013859
[6 empty] | 289.11 | 0.013431
[6 empty] | 289.22 | 0.013228
[6 empty] | 289.33 | 0.012875
[6 empty] | 289.44 | 0.012595
[6 empty] | 289.55 | 0.012338
[6 empty] | 289.66 | 0.012148
[6 empty] | 289.77 | 0.011797
[6 empty] | 289.88 | 0.011614
[6 empty] | 289.99 | 0.011351
[6 empty] | 290.11 | 0.011029
[6 empty] | 290.22 | 0.01086
[6 empty] | 290.33 | 0.010564
[6 empty] | 290.44 | 0.010365
[6 empty] | 290.55 | 0.010238
[6 empty] | 290.66 | 0.0099442
[6 empty] | 290.77 | 0.0097586
[6 empty] | 290.88 | 0.0093598
[6 empty] | 290.99 | 0.0092438
[6 empty] | 291.1 | 0.0090561
[6 empty] | 291.21 | 0.0090089
[6 empty] | 291.33 | 0.008655
[6 empty] | 291.44 | 0.0084429
[6 empty] | 291.55 | 0.0083759
[6 empty] | 291.66 | 0.0082421
[6 empty] | 291.77 | 0.0081266
[6 empty] | 291.88 | 0.0078635
[6 empty] | 291.99 | 0.0077976
[6 empty] | 292.1 | 0.0075574
[6 empty] | 292.21 | 0.0075742
[6 empty] | 292.32 | 0.0072425
[6 empty] | 292.43 | 0.0071221
[6 empty] | 292.55 | 0.0069432
[6 empty] | 292.66 | 0.0068824
[6 empty] | 292.77 | 0.0065337
[6 empty] | 292.88 | 0.0065544
[6 empty] | 292.99 | 0.0063671
[6 empty] | 293.1 | 0.006263
[6 empty] | 293.21 | 0.0061044
[6 empty] | 293.32 | 0.0059503
[6 empty] | 293.43 | 0.0060199
[6 empty] | 293.54 | 0.0057976
[6 empty] | 293.65 | 0.0057021
[6 empty] | 293.76 | 0.0055945
[6 empty] | 293.87 | 0.0055191
[6 empty] | 293.99 | 0.005223
[6 empty] | 294.1 | 0.0051928
[6 empty] | 294.21 | 0.0050785
[6 empty] | 294.32 | 0.0049532
[6 empty] | 294.43 | 0.0048592
[6 empty] | 294.54 | 0.0046522
[6 empty] | 294.65 | 0.0046711
[6 empty] | 294.76 | 0.004601
[6 empty] | 294.87 | 0.0045667
[6 empty] | 294.98 | 0.0044045
[6 empty] | 295.09 | 0.0043432
[6 empty] | 295.2 | 0.0041739
[6 empty] | 295.31 | 0.004123
[6 empty] | 295.42 | 0.0040294
[6 empty] | 295.53 | 0.0039304
[6 empty] | 295.64 | 0.0037792
[6 empty] | 295.76 | 0.0037195
[6 empty] | 295.87 | 0.0036196
[6 empty] | 295.98 | 0.0035926
[6 empty] | 296.09 | 0.0034469
[6 empty] | 296.2 | 0.0033062
[6 empty] | 296.31 | 0.0034769
[6 empty] | 296.42 | 0.0033161
[6 empty] | 296.53 | 0.0032003
[6 empty] | 296.64 | 0.0031466
[6 empty] | 296.75 | 0.0030417
[6 empty] | 296.86 | 0.0030685
[6 empty] | 296.97 | 0.0029982
[6 empty] | 297.08 | 0.00291
[6 empty] | 297.19 | 0.0028024
[6 empty] | 297.3 | 0.002788
[6 empty] | 297.41 | 0.0026906
[6 empty] | 297.52 | 0.0026882
[6 empty] | 297.63 | 0.0025601
[6 empty] | 297.74 | 0.0025894
[6 empty] | 297.85 | 0.0025093
[6 empty] | 297.96 | 0.0024723
[6 empty] | 298.07 | 0.0022967
[6 empty] | 298.19 | 0.0022603
[6 empty] | 298.3 | 0.0021926
[6 empty] | 298.41 | 0.0022478
[6 empty] | 298.52 | 0.002198
[6 empty] | 298.63 | 0.0021077
[6 empty] | 298.74 | 0.0020406
[6 empty] | 298.85 | 0.0020126
[6 empty] | 298.96 | 0.0019168
[6 empty] | 299.07 | 0.0019505
[6 empty] | 299.18 | 0.0018606
[6 empty] | 299.29 | 0.0018267
[6 empty] | 299.4 | 0.0017583
[6 empty] | 299.51 | 0.0017786
[6 empty] | 299.62 | 0.0017278
[6 empty] | 299.73 | 0.0016294
[6 empty] | 299.84 | 0.0015848
[6 empty] | 299.95 | 0.0016224
[6 empty] | 300.06 | 0.0015669

## Petri
[8 empty] | UV Dose Calculations for a Low Pressure UV Lamp | [14 empty] | Calculations for carrying out a Dose-Reponse Curve
[8 empty] | Jim Bolton - Bolton Photosciences Inc., 92 Main St., Ayr, ON, Canada N0B 1E0 | [14 empty] | for a low pressure UV lamp
[8 empty] | Tel: 519-741-6283; Fax: 519-632-8941; Email: jbolton@boltonuv.com | [14 empty] | Programmed by Jim Bolton (519-741-6283)
Steps:
1. Draw a 0.5 cm x 0.5 cm grid and place the center of the grid at the center of the collimated beam
2. Measure the UV irradiance (with a radiometer) every 0.5 cm in the x and y directions and place into columns C and H below.
3. Place the Petri dish "inner" radius into Cell Q27.
4. The Petri factor is calculated in Cell Q20
UV Distribution across the Petri Dish for a low pressure UV Lamp
|  | Meter |  |  |  |  | Meter
x | y | Reading |  |  | x | y | Reading
|  | mW/cm2 | Ratio |  |  |  | mW/cm2 | Ratio
|  | y |  |  |  |  | x
0 | -6 |  | 0 |  | -6 | 0 |  | 0
0 | -5.5 |  | 0 |  | -5.5 | 0 |  | 0
0 | -5 |  | 0 |  | -5 | 0 |  | 0
0 | -4.5 |  | 0 |  | -4.5 | 0 |  | 0 | [6 empty] | Petri factor = | 0.92509
0 | -4 |  | 0 |  | -4 | 0 |  | 0
0 | -3.5 |  | 0 |  | -3.5 | 0 |  | 0
0 | -3 |  | 0 |  | -3 | 0 |  | 0
0 | -2.5 | 1.62 | 0.52941 |  | -2.5 | 0 | 2.2 | 0.7074
0 | -2 | 1.94 | 0.63399 |  | -2 | 0 | 2.53 | 0.8135
0 | -1.5 | 2.28 | 0.7451 |  | -1.5 | 0 | 2.78 | 0.89389
0 | -1 | 2.54 | 0.83007 |  | -1 | 0 | 2.98 | 0.9582 | [6 empty] | Petri Dish Radius = | 1.75 | cm
0 | -0.5 | 2.81 | 0.9183 |  | -0.5 | 0 | 3.11 | 1
0 | 0 | 3.06 | 1 |  | 0 | 0 | 3.11 | 1
0 | 0.5 | 3.18 | 1.0392 |  | 0.5 | 0 | 3.04 | 0.97749
0 | 1 | 3.21 | 1.049 |  | 1 | 0 | 2.96 | 0.95177
0 | 1.5 | 3.14 | 1.0261 |  | 1.5 | 0 | 2.81 | 0.90354
0 | 2 | 2.98 | 0.97386 |  | 2 | 0 | 2.58 | 0.82958
0 | 2.5 | 2.69 | 0.87908 |  | 2.5 | 0 | 2.24 | 0.72026
0 | 3 |  | 0 |  | 3 | 0 |  | 0
0 | 3.5 |  | 0 |  | 3.5 | 0 |  | 0
0 | 4 |  | 0 |  | 4 | 0 |  | 0
0 | 4.5 |  | 0 |  | 4.5 | 0 |  | 0
0 | 5 |  | 0 |  | 5 | 0 |  | 0
0 | 5.5 |  | 0 |  | 5.5 | 0 |  | 0
0 | 6 |  | 0 |  | 6 | 0 |  | 0
[21 empty] | 3.5 | 4 | 4.5 | 5
x y |  | -6 | -5.5 | 5 | -4.5 | -4 | -3.5 | -3 | -2.5 | -2 | -1.5 | -1 | -0.5 | 0 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 |  |  |  |  | 5.5 | 6
-6 | [27 empty] | 0 | 0
-5.5 | [27 empty] | 0 | 0
-5 | [27 empty] | 0 | 0
-4.5 | [27 empty] | 0 | 0
-4 | [27 empty] | 0 | 0
-3.5 | [27 empty] | 0 | 0
-3 | [27 empty] | 0 | 0
-2.5 | [27 empty] | 0 | 0
-2 | [27 empty] | 0 | 0
-1.5 | [12 empty] | 0.82086 | 0.89389 | 0.92895 | [12 empty] | 2.6437 | 3
-1 | [11 empty] | 0.79537 | 0.87992 | 0.9582 | 0.99578 | 1.0052 | [11 empty] | 4.6344 | 5
-0.5 | [10 empty] | 0.7451 | 0.83007 | 0.9183 | 1 | 1.0392 | 1.049 | 1.0261 | [10 empty] | 6.6078 | 7
0 | [10 empty] | 0.7451 | 0.83007 | 0.9183 | 1 | 1.0392 | 1.049 | 1.0261 | [10 empty] | 6.6078 | 7
0.5 | [10 empty] | 0.72833 | 0.81138 | 0.89763 | 0.97749 | 1.0158 | 1.0254 | 1.003 | [10 empty] | 6.4591 | 7
1 | [11 empty] | 0.79003 | 0.87401 | 0.95177 | 0.98909 | 0.99842 | [11 empty] | 4.6033 | 5
1.5 | [12 empty] | 0.82972 | 0.90354 | 0.93897 | [12 empty] | 2.6722 | 3
2 | [27 empty] | 0 | 0
2.5 | [27 empty] | 0 | 0
3 | [27 empty] | 0 | 0
3.5 | [27 empty] | 0 | 0
4 | [27 empty] | 0 | 0
4.5 | [27 empty] | 0 | 0
5 | [27 empty] | 0 | 0
5.5 | [27 empty] | 0 | 0
6 | [27 empty] | 0 | 0
[28 empty] | 34.2285 | 37
[27 empty] | Average Petri factor = | 0.92509

## Germicidal Calculator
Petri Factor = | 0.92509 |  | Incident Irradiance = | 0.314
Sample Depth = | 0.6 | cm | Average Irradiance = | 0.24428
Sample to beam = | 4 | cm
Divergence Factor = | 0.86957 |  | TRLE | 1
Reflection Factor = | 0.97745
Sensor Factor = | 1.0067
Wavelength (nm) | Rated Lamp Energy (RLE) over 200-300 nm | Radiometer Calibration Scan Data (Linden) | Radiometer Sensitivity (fr) (relative to 278 nm) | RLE x radioeter sensitivity | L(l) RLE/TRLE (Same as column B) | Ii | Water Absorbance | Water Factor | Average Irradiance Calc | Absorbance DNA | Relative Absorbance | Germicidal Irradiance | Action Spectrum |  | Action Spectrum Absorbance
200 | -5.053e-05 | 0.0015617 | 1.5328 | -7.7451e-05 | -5.053e-05 | -1.2559e-05 | 1.164 | 0.49731 | -6.2455e-06 | 0.83 | 1.8864 | -1.1781e-05 | 23.5339 | 17.5593 | -0.00010967
201 | 9.3637e-05 | 0.0016241 | 1.5941 | 0.00014926 | 9.3637e-05 | 2.3272e-05 | 0.979 | 0.54817 | 1.2757e-05 | 0.82 | 1.8636 | 2.3774e-05 | 22.9553 | 17.1276 | 0.0002185
202 | -9.8146e-05 | 0.0016865 | 1.6553 | -0.00016247 | -9.8146e-05 | -2.4393e-05 | 0.813 | 0.60075 | -1.4654e-05 | 0.81 | 1.8409 | -2.6977e-05 | 22.3766 | 16.6958 | -0.00024466
203 | 1.0561e-05 | 0.0017625 | 1.7299 | 1.827e-05 | 1.0561e-05 | 2.6249e-06 | 0.684 | 0.64691 | 1.698e-06 | 0.8 | 1.8182 | 3.0874e-06 | 21.7979 | 16.2641 | 2.7617e-05
204 | 6.0213e-05 | 0.0018384 | 1.8044 | 0.00010865 | 6.0213e-05 | 1.4965e-05 | 0.574 | 0.69043 | 1.0332e-05 | 0.79 | 1.7955 | 1.8551e-05 | 21.2192 | 15.8323 | 0.00016359
205 | 0.00010746 | 0.0018895 | 1.8546 | 0.00019929 | 0.00010746 | 2.6708e-05 | 0.484 | 0.72922 | 1.9476e-05 | 0.776 | 1.7636 | 3.4349e-05 | 20.6406 | 15.4006 | 0.00029994
206 | 7.4328e-05 | 0.0019407 | 1.9048 | 0.00014158 | 7.4328e-05 | 1.8473e-05 | 0.407 | 0.7649 | 1.413e-05 | 0.762 | 1.7318 | 2.4471e-05 | 20.0619 | 14.9688 | 0.00021151
207 | 0.0001604 | 0.0019779 | 1.9413 | 0.00031138 | 0.0001604 | 3.9865e-05 | 0.341 | 0.79746 | 3.1791e-05 | 0.748 | 1.7 | 5.4044e-05 | 19.4832 | 14.537 | 0.00046214
208 | 9.2506e-05 | 0.0020151 | 1.9779 | 0.00018296 | 9.2506e-05 | 2.2991e-05 | 0.297 | 0.82024 | 1.8858e-05 | 0.735 | 1.6705 | 3.1502e-05 | 18.9045 | 14.1052 | 0.000266
209 | 6.7136e-05 | 0.0020608 | 2.0226 | 0.00013579 | 6.7136e-05 | 1.6686e-05 | 0.251 | 0.84503 | 1.41e-05 | 0.71 | 1.6136 | 2.2752e-05 | 18.3258 | 13.6734 | 0.00019279
210 | 2.1031e-05 | 0.0021064 | 2.0674 | 4.348e-05 | 2.1031e-05 | 5.227e-06 | 0.222 | 0.86119 | 4.5014e-06 | 0.685 | 1.5568 | 7.0079e-06 | 17.7472 | 13.2417 | 5.9607e-05
211 | 8.2448e-05 | 0.0021453 | 2.1056 | 0.0001736 | 8.2448e-05 | 2.0491e-05 | 0.202 | 0.87259 | 1.788e-05 | 0.66 | 1.5 | 2.6821e-05 | 17.1685 | 12.8099 | 0.00022905
212 | -2.4023e-05 | 0.0021842 | 2.1438 | -5.1499e-05 | -2.4023e-05 | -5.9706e-06 | 0.178 | 0.88653 | -5.2931e-06 | 0.645 | 1.4659 | -7.7592e-06 | 16.5898 | 12.3781 | -6.5519e-05
213 | -2.9614e-06 | 0.0022142 | 2.1732 | -6.4358e-06 | -2.9614e-06 | -7.3602e-07 | 0.165 | 0.89421 | -6.5816e-07 | 0.624 | 1.4182 | -9.3339e-07 | 16.0111 | 11.9464 | -7.8626e-06
214 | 5.8461e-05 | 0.0022442 | 2.2027 | 0.00012877 | 5.8461e-05 | 1.453e-05 | 0.152 | 0.90198 | 1.3106e-05 | 0.603 | 1.3705 | 1.7961e-05 | 15.4325 | 11.5146 | 0.00015091
215 | 1.778e-05 | 0.0022661 | 2.2242 | 3.9545e-05 | 1.778e-05 | 4.4189e-06 | 0.139 | 0.90984 | 4.0205e-06 | 0.582 | 1.3227 | 5.318e-06 | 14.8538 | 11.0829 | 4.4559e-05
216 | 4.5029e-05 | 0.0022879 | 2.2456 | 0.00010112 | 4.5029e-05 | 1.1191e-05 | 0.131 | 0.91473 | 1.0237e-05 | 0.56 | 1.2727 | 1.3029e-05 | 14.2751 | 10.6511 | 0.00010904
217 | 6.5269e-05 | 0.0023127 | 2.2699 | 0.00014816 | 6.5269e-05 | 1.6222e-05 | 0.125 | 0.91842 | 1.4898e-05 | 0.54 | 1.2273 | 1.8284e-05 | 13.6964 | 10.2193 | 0.00015225
218 | 8.2136e-06 | 0.0023375 | 2.2943 | 1.8844e-05 | 8.2136e-06 | 2.0414e-06 | 0.119 | 0.92212 | 1.8824e-06 | 0.52 | 1.1818 | 2.2247e-06 | 13.1177 | 9.7875 | 1.8424e-05
219 | 4.2798e-06 | 0.002371 | 2.3271 | 9.9598e-06 | 4.2798e-06 | 1.0637e-06 | 0.113 | 0.92585 | 9.8483e-07 | 0.5 | 1.1364 | 1.1191e-06 | 12.5391 | 9.3558 | 9.2139e-06
220 | -6.3592e-06 | 0.0024045 | 2.36 | -1.5008e-05 | -6.3592e-06 | -1.5805e-06 | 0.108 | 0.92897 | -1.4682e-06 | 0.48 | 1.0909 | -1.6017e-06 | 11.9604 | 8.924 | -1.3103e-05
221 | 5.9409e-06 | 0.0023973 | 2.3529 | 1.3978e-05 | 5.9409e-06 | 1.4765e-06 | 0.104 | 0.93148 | 1.3754e-06 | 0.453 | 1.0295 | 1.416e-06 | 11.3552 | 8.4724 | 1.1653e-05
222 | 5.5508e-05 | 0.00239 | 2.3458 | 0.00013021 | 5.5508e-05 | 1.3796e-05 | 0.1 | 0.934 | 1.2885e-05 | 0.426 | 0.96818 | 1.2475e-05 | 10.7501 | 8.021 | 0.00010335
223 | 6.5322e-05 | 0.0023994 | 2.355 | 0.00015384 | 6.5322e-05 | 1.6235e-05 | 0.097 | 0.93589 | 1.5194e-05 | 0.399 | 0.90682 | 1.3778e-05 | 10.1449 | 7.5694 | 0.00011501
224 | 3.6639e-05 | 0.0024088 | 2.3643 | 8.6625e-05 | 3.6639e-05 | 9.1063e-06 | 0.092 | 0.93906 | 8.5513e-06 | 0.37 | 0.84091 | 7.1909e-06 | 9.5398 | 7.1179 | 6.0867e-05
225 | 1.4388e-06 | 0.0024061 | 2.3616 | 3.3979e-06 | 1.4388e-06 | 3.576e-07 | 0.089 | 0.94097 | 3.3649e-07 | 0.35 | 0.79545 | 2.6766e-07 | 8.9346 | 6.6664 | 2.2431e-06
226 | 5.4373e-05 | 0.0024035 | 2.359 | 0.00012827 | 5.4373e-05 | 1.3514e-05 | 0.085 | 0.94352 | 1.275e-05 | 0.337 | 0.76591 | 9.7656e-06 | 8.3294 | 6.2148 | 7.9241e-05
227 | 1.3672e-05 | 0.0024126 | 2.368 | 3.2375e-05 | 1.3672e-05 | 3.398e-06 | 0.081 | 0.94608 | 3.2147e-06 | 0.324 | 0.73636 | 2.3672e-06 | 7.7243 | 5.7633 | 1.8528e-05
228 | 8.1617e-05 | 0.0024217 | 2.3769 | 0.000194 | 8.1617e-05 | 2.0285e-05 | 0.077 | 0.94865 | 1.9243e-05 | 0.31 | 0.70455 | 1.3558e-05 | 7.1191 | 5.3118 | 0.00010222
229 | 8.3793e-06 | 0.0024115 | 2.3669 | 1.9833e-05 | 8.3793e-06 | 2.0826e-06 | 0.074 | 0.95058 | 1.9797e-06 | 0.305 | 0.69318 | 1.3723e-06 | 6.514 | 4.8603 | 9.6216e-06
230 | 2.7763e-05 | 0.0024013 | 2.3569 | 6.5433e-05 | 2.7763e-05 | 6.9001e-06 | 0.07 | 0.95317 | 6.577e-06 | 0.3 | 0.68182 | 4.4843e-06 | 5.9088 | 4.4087 | 2.8996e-05
231 | 2.1417e-05 | 0.0024014 | 2.357 | 5.0479e-05 | 2.1417e-05 | 5.3229e-06 | 0.067 | 0.95511 | 5.0839e-06 | 0.3 | 0.68182 | 3.4663e-06 | 5.4152 | 4.0404 | 2.0541e-05
232 | 3.7143e-05 | 0.0024016 | 2.3571 | 8.755e-05 | 3.7143e-05 | 9.2313e-06 | 0.064 | 0.95707 | 8.835e-06 | 0.3 | 0.68182 | 6.0238e-06 | 4.9216 | 3.6721 | 3.2443e-05
233 | 2.2409e-05 | 0.0023984 | 2.3541 | 5.2753e-05 | 2.2409e-05 | 5.5695e-06 | 0.06 | 0.95968 | 5.3449e-06 | 0.305 | 0.69318 | 3.705e-06 | 4.428 | 3.3038 | 1.7659e-05
234 | 2.1693e-05 | 0.0023953 | 2.351 | 5.0999e-05 | 2.1693e-05 | 5.3914e-06 | 0.057 | 0.96164 | 5.1846e-06 | 0.31 | 0.70455 | 3.6528e-06 | 3.9343 | 2.9355 | 1.522e-05
235 | 3.2764e-05 | 0.0023887 | 2.3445 | 7.6816e-05 | 3.2764e-05 | 8.1431e-06 | 0.055 | 0.96295 | 7.8414e-06 | 0.32 | 0.72727 | 5.7028e-06 | 3.4407 | 2.5672 | 2.0131e-05
236 | 6.2802e-05 | 0.0023821 | 2.3381 | 0.00014684 | 6.2802e-05 | 1.5609e-05 | 0.052 | 0.96492 | 1.5061e-05 | 0.33 | 0.75 | 1.1296e-05 | 2.9471 | 2.1989 | 3.3118e-05
237 | -1.5785e-06 | 0.0023659 | 2.3221 | -3.6655e-06 | -1.5785e-06 | -3.9232e-07 | 0.05 | 0.96624 | -3.7908e-07 | 0.35 | 0.79545 | -3.0154e-07 | 2.4535 | 1.8306 | -6.9394e-07
238 | 2.0111e-05 | 0.0023496 | 2.3061 | 4.6379e-05 | 2.0111e-05 | 4.9984e-06 | 0.049 | 0.9669 | 4.833e-06 | 0.37 | 0.84091 | 4.0641e-06 | 1.9599 | 1.4623 | 7.0673e-06
239 | 1.5969e-05 | 0.0023429 | 2.2995 | 3.6721e-05 | 1.5969e-05 | 3.9689e-06 | 0.046 | 0.96889 | 3.8454e-06 | 0.39 | 0.88636 | 3.4084e-06 | 1.4662 | 1.094 | 4.2069e-06
240 | 3.559e-05 | 0.0023361 | 2.2929 | 8.1603e-05 | 3.559e-05 | 8.8453e-06 | 0.046 | 0.96889 | 8.5701e-06 | 0.41 | 0.93182 | 7.9858e-06 | 0.9726 | 0.72569 | 6.2192e-06
241 | 3.2684e-05 | 0.0023177 | 2.2748 | 7.4348e-05 | 3.2684e-05 | 8.1231e-06 | 0.044 | 0.97021 | 7.8811e-06 | 0.433 | 0.98409 | 7.7558e-06 | 0.97456 | 0.72715 | 5.7308e-06
242 | 2.4655e-05 | 0.0022992 | 2.2567 | 5.5639e-05 | 2.4655e-05 | 6.1278e-06 | 0.042 | 0.97154 | 5.9534e-06 | 0.456 | 1.0364 | 6.1699e-06 | 0.97652 | 0.72861 | 4.3377e-06
243 | 1.9634e-05 | 0.0022817 | 2.2395 | 4.3969e-05 | 1.9634e-05 | 4.8798e-06 | 0.042 | 0.97154 | 4.7409e-06 | 0.479 | 1.0886 | 5.1611e-06 | 0.97847 | 0.73007 | 3.4612e-06
244 | 1.6969e-05 | 0.0022641 | 2.2222 | 3.7709e-05 | 1.6969e-05 | 4.2175e-06 | 0.041 | 0.97221 | 4.1003e-06 | 0.5 | 1.1364 | 4.6594e-06 | 0.98043 | 0.73153 | 2.9995e-06
245 | 3.399e-05 | 0.0022321 | 2.1908 | 7.4466e-05 | 3.399e-05 | 8.4478e-06 | 0.04 | 0.97287 | 8.2186e-06 | 0.52 | 1.1818 | 9.7129e-06 | 0.98239 | 0.73299 | 6.0241e-06
246 | 4.8251e-05 | 0.0022001 | 2.1594 | 0.00010419 | 4.8251e-05 | 1.1992e-05 | 0.04 | 0.97287 | 1.1667e-05 | 0.54 | 1.2273 | 1.4318e-05 | 0.98435 | 0.73445 | 8.5687e-06
247 | 2.625e-05 | 0.0021777 | 2.1374 | 5.6106e-05 | 2.625e-05 | 6.524e-06 | 0.039 | 0.97354 | 6.3514e-06 | 0.56 | 1.2727 | 8.0836e-06 | 0.9863 | 0.73591 | 4.674e-06
248 | 2.1093e-05 | 0.0021553 | 2.1154 | 4.462e-05 | 2.1093e-05 | 5.2425e-06 | 0.039 | 0.97354 | 5.1038e-06 | 0.59 | 1.3409 | 6.8437e-06 | 0.98826 | 0.73737 | 3.7633e-06
249 | 6.2246e-05 | 0.0021371 | 2.0976 | 0.00013057 | 6.2246e-05 | 1.5471e-05 | 0.038 | 0.9742 | 1.5071e-05 | 0.61 | 1.3864 | 2.0895e-05 | 0.99022 | 0.73883 | 1.1135e-05
250 | 5.7159e-05 | 0.002119 | 2.0798 | 0.00011888 | 5.7159e-05 | 1.4206e-05 | 0.037 | 0.97487 | 1.3849e-05 | 0.63 | 1.4318 | 1.9829e-05 | 0.99217 | 0.74029 | 1.0252e-05
251 | 5.5725e-05 | 0.0020712 | 2.0329 | 0.00011328 | 5.5725e-05 | 1.385e-05 | 0.036 | 0.97554 | 1.3511e-05 | 0.65 | 1.4773 | 1.9959e-05 | 0.99413 | 0.74175 | 1.0022e-05
252 | 2.6445e-05 | 0.0020234 | 1.986 | 5.252e-05 | 2.6445e-05 | 6.5726e-06 | 0.036 | 0.97554 | 6.4118e-06 | 0.665 | 1.5114 | 9.6906e-06 | 0.99609 | 0.74321 | 4.7653e-06
253 | 5.0065e-05 | 0.0019932 | 1.9563 | 9.7943e-05 | 5.0065e-05 | 1.2443e-05 | 0.035 | 0.97621 | 1.2147e-05 | 0.68 | 1.5455 | 1.8772e-05 | 0.99804 | 0.74467 | 9.0454e-06
254 | 7.5026e-05 | 0.001963 | 1.9267 | 0.00014455 | 7.5026e-05 | 1.8647e-05 | 0.035 | 0.97621 | 1.8203e-05 | 0.695 | 1.5795 | 2.8753e-05 | 1 | 0.74613 | 1.3582e-05
255 | 5.6585e-05 | 0.0019167 | 1.8812 | 0.00010645 | 5.6585e-05 | 1.4064e-05 | 0.034 | 0.97688 | 1.3738e-05 | 0.71 | 1.6136 | 2.2169e-05 | 1.0092 | 0.75302 | 1.0345e-05
256 | 9.1368e-05 | 0.0018703 | 1.8357 | 0.00016772 | 9.1368e-05 | 2.2708e-05 | 0.034 | 0.97688 | 2.2183e-05 | 0.72 | 1.6364 | 3.63e-05 | 1.0185 | 0.75993 | 1.6858e-05
257 | 9.4467e-05 | 0.0018276 | 1.7938 | 0.00016946 | 9.4467e-05 | 2.3478e-05 | 0.032 | 0.97822 | 2.2967e-05 | 0.725 | 1.6477 | 3.7843e-05 | 1.0277 | 0.76682 | 1.7612e-05
258 | 0.00011524 | 0.001785 | 1.752 | 0.0002019 | 0.00011524 | 2.8641e-05 | 0.033 | 0.97755 | 2.7998e-05 | 0.73 | 1.6591 | 4.6451e-05 | 1.037 | 0.77372 | 2.1663e-05
259 | 0.00012082 | 0.0017649 | 1.7323 | 0.0002093 | 0.00012082 | 3.0029e-05 | 0.032 | 0.97822 | 2.9375e-05 | 0.735 | 1.6705 | 4.9069e-05 | 1.0462 | 0.78062 | 2.293e-05
260 | 0.00011999 | 0.0017448 | 1.7126 | 0.00020549 | 0.00011999 | 2.9822e-05 | 0.032 | 0.97822 | 2.9172e-05 | 0.725 | 1.6477 | 4.8068e-05 | 1.0555 | 0.78752 | 2.2974e-05
261 | 0.00016588 | 0.0016698 | 1.6389 | 0.00027186 | 0.00016588 | 4.1228e-05 | 0.033 | 0.97755 | 4.0302e-05 | 0.7125 | 1.6193 | 6.5262e-05 | 1.1107 | 0.82875 | 3.34e-05
262 | 0.0002959 | 0.0015947 | 1.5652 | 0.00046316 | 0.0002959 | 7.3543e-05 | 0.03 | 0.97956 | 7.204e-05 | 0.7 | 1.5909 | 0.00011461 | 1.166 | 0.86997 | 6.2673e-05
263 | 0.00061153 | 0.001563 | 1.5341 | 0.00093817 | 0.00061153 | 0.00015199 | 0.033 | 0.97755 | 0.00014858 | 0.6925 | 1.5739 | 0.00023384 | 1.2212 | 0.9112 | 0.00013538
264 | 0.001256 | 0.0015314 | 1.503 | 0.0018879 | 0.001256 | 0.00031217 | 0.031 | 0.97889 | 0.00030558 | 0.685 | 1.5568 | 0.00047573 | 1.2765 | 0.95243 | 0.00029105
265 | 0.0022725 | 0.0014956 | 1.4679 | 0.0033358 | 0.0022725 | 0.0005648 | 0.03 | 0.97956 | 0.00055325 | 0.675 | 1.5341 | 0.00084874 | 1.3318 | 0.99367 | 0.00054975
266 | 0.0039228 | 0.0014598 | 1.4328 | 0.0056206 | 0.0039228 | 0.00097496 | 0.031 | 0.97889 | 0.00095438 | 0.662 | 1.5045 | 0.0014359 | 1.387 | 1.0349 | 0.00098768
267 | 0.0063133 | 0.0014427 | 1.416 | 0.0089399 | 0.0063133 | 0.0015691 | 0.031 | 0.97889 | 0.001536 | 0.649 | 1.475 | 0.0022655 | 1.4423 | 1.0761 | 0.0016529
268 | 0.0094092 | 0.0014257 | 1.3993 | 0.013166 | 0.0094092 | 0.0023385 | 0.03 | 0.97956 | 0.0022907 | 0.635 | 1.4432 | 0.0033059 | 1.4975 | 1.1174 | 0.0025596
269 | 0.01393 | 0.0013689 | 1.3436 | 0.018716 | 0.01393 | 0.0034622 | 0.03 | 0.97956 | 0.0033914 | 0.622 | 1.4136 | 0.0047942 | 1.5528 | 1.1586 | 0.0039292
270 | 0.020438 | 0.0013121 | 1.2878 | 0.026321 | 0.020438 | 0.0050797 | 0.029 | 0.98023 | 0.0049793 | 0.6 | 1.3636 | 0.00679 | 1.608 | 1.1998 | 0.0059743
271 | 0.029338 | 0.0012774 | 1.2537 | 0.036782 | 0.029338 | 0.0072917 | 0.028 | 0.98091 | 0.0071525 | 0.58 | 1.3182 | 0.0094282 | 1.5746 | 1.1748 | 0.008403
272 | 0.041229 | 0.0012426 | 1.2196 | 0.050283 | 0.041229 | 0.010247 | 0.029 | 0.98023 | 0.010044 | 0.56 | 1.2727 | 0.012784 | 1.5411 | 1.1499 | 0.01155
273 | 0.055158 | 0.0012081 | 1.1857 | 0.065402 | 0.055158 | 0.013709 | 0.029 | 0.98023 | 0.013438 | 0.54 | 1.2273 | 0.016492 | 1.5076 | 1.1249 | 0.015116
274 | 0.068403 | 0.0011736 | 1.1518 | 0.07879 | 0.068403 | 0.017001 | 0.027 | 0.98158 | 0.016688 | 0.52 | 1.1818 | 0.019722 | 1.4742 | 1.0999 | 0.018355
275 | 0.079556 | 0.0011338 | 1.1128 | 0.08853 | 0.079556 | 0.019773 | 0.027 | 0.98158 | 0.019408 | 0.5 | 1.1364 | 0.022055 | 1.4407 | 1.0749 | 0.020863
276 | 0.086607 | 0.001094 | 1.0738 | 0.092995 | 0.086607 | 0.021525 | 0.026 | 0.98225 | 0.021143 | 0.48 | 1.0909 | 0.023065 | 1.4072 | 1.05 | 0.022199
277 | 0.087524 | 0.0010564 | 1.0369 | 0.090752 | 0.087524 | 0.021753 | 0.026 | 0.98225 | 0.021367 | 0.46 | 1.0455 | 0.022338 | 1.3737 | 1.025 | 0.021901
278 | 0.083214 | 0.0010188 | 1 | 0.083214 | 0.083214 | 0.020682 | 0.025 | 0.98293 | 0.020329 | 0.44 | 1 | 0.020329 | 1.3402 | 1 | 0.020329
279 | 0.074745 | 0.00097406 | 0.95604 | 0.071459 | 0.074745 | 0.018577 | 0.025 | 0.98293 | 0.01826 | 0.42 | 0.95455 | 0.01743 | 1.3068 | 0.97503 | 0.017804
280 | 0.064373 | 0.00092926 | 0.91207 | 0.058713 | 0.064373 | 0.015999 | 0.024 | 0.9836 | 0.015737 | 0.4 | 0.90909 | 0.014306 | 1.2733 | 0.95005 | 0.014951
281 | 0.053392 | 0.00087526 | 0.85907 | 0.045867 | 0.053392 | 0.01327 | 0.023 | 0.98428 | 0.013061 | 0.376 | 0.85455 | 0.011161 | 1.2071 | 0.90066 | 0.011764
282 | 0.043602 | 0.00082126 | 0.80606 | 0.035146 | 0.043602 | 0.010837 | 0.023 | 0.98428 | 0.010666 | 0.352 | 0.8 | 0.008533 | 1.1409 | 0.85127 | 0.0090799
283 | 0.034584 | 0.00078303 | 0.76854 | 0.026579 | 0.034584 | 0.0085954 | 0.022 | 0.98496 | 0.0084661 | 0.328 | 0.74545 | 0.0063111 | 1.0747 | 0.80189 | 0.0067889
284 | 0.027648 | 0.00074479 | 0.73101 | 0.020211 | 0.027648 | 0.0068716 | 0.022 | 0.98496 | 0.0067682 | 0.305 | 0.69318 | 0.0046916 | 1.0085 | 0.75249 | 0.0050931
285 | 0.021741 | 0.0007104 | 0.69725 | 0.015159 | 0.021741 | 0.0054034 | 0.02 | 0.98631 | 0.0053294 | 0.28 | 0.63636 | 0.0033915 | 0.94234 | 0.70311 | 0.0037472
286 | 0.017203 | 0.000676 | 0.66349 | 0.011414 | 0.017203 | 0.0042756 | 0.019 | 0.98699 | 0.00422 | 0.255 | 0.57955 | 0.0024457 | 0.87615 | 0.65372 | 0.0027587
287 | 0.013682 | 0.00065196 | 0.63989 | 0.0087547 | 0.013682 | 0.0034004 | 0.017 | 0.98835 | 0.0033607 | 0.23 | 0.52273 | 0.0017568 | 0.80996 | 0.60433 | 0.002031
288 | 0.011001 | 0.00062791 | 0.6163 | 0.0067797 | 0.011001 | 0.0027341 | 0.017 | 0.98835 | 0.0027022 | 0.21 | 0.47727 | 0.0012897 | 0.74376 | 0.55494 | 0.0014996
289 | 0.0089741 | 0.00059241 | 0.58145 | 0.005218 | 0.0089741 | 0.0022304 | 0.016 | 0.98903 | 0.0022059 | 0.19 | 0.43182 | 0.00095256 | 0.67757 | 0.50556 | 0.0011152
290 | 0.0073516 | 0.0005569 | 0.5466 | 0.0040184 | 0.0073516 | 0.0018272 | 0.016 | 0.98903 | 0.0018071 | 0.17 | 0.38636 | 0.0006982 | 0.61138 | 0.45617 | 0.00082434
291 | 0.00599 | 0.00051359 | 0.50408 | 0.0030195 | 0.00599 | 0.0014887 | 0.014 | 0.99039 | 0.0014744 | 0.15 | 0.34091 | 0.00050265 | 0.54519 | 0.40678 | 0.00059977
292 | 0.005048 | 0.00047027 | 0.46157 | 0.00233 | 0.005048 | 0.0012546 | 0.014 | 0.99039 | 0.0012426 | 0.14 | 0.31818 | 0.00039536 | 0.47899 | 0.35739 | 0.00044408
293 | 0.0041274 | 0.0004505 | 0.44216 | 0.001825 | 0.0041274 | 0.0010258 | 0.013 | 0.99107 | 0.0010166 | 0.125 | 0.28409 | 0.00028882 | 0.4128 | 0.308 | 0.00031313
294 | 0.003389 | 0.00043072 | 0.42276 | 0.0014327 | 0.003389 | 0.00084229 | 0.012 | 0.99176 | 0.00083535 | 0.11 | 0.25 | 0.00020884 | 0.34661 | 0.25862 | 0.00021603
295 | 0.0028522 | 0.00040247 | 0.39502 | 0.0011267 | 0.0028522 | 0.00070887 | 0.011 | 0.99244 | 0.00070351 | 0.095 | 0.21591 | 0.00015189 | 0.28042 | 0.20923 | 0.00014719
296 | 0.0023151 | 0.00037422 | 0.36729 | 0.00085033 | 0.0023151 | 0.00057539 | 0.011 | 0.99244 | 0.00057104 | 0.08 | 0.18182 | 0.00010383 | 0.21422 | 0.15984 | 9.1275e-05
297 | 0.0019308 | 0.00034845 | 0.34201 | 0.00066035 | 0.0019308 | 0.00047988 | 0.011 | 0.99244 | 0.00047625 | 0.071 | 0.16136 | 7.685e-05 | 0.14803 | 0.11045 | 5.2602e-05
298 | 0.0015636 | 0.00032269 | 0.31672 | 0.00049521 | 0.0015636 | 0.0003886 | 0.011 | 0.99244 | 0.00038567 | 0.062 | 0.14091 | 5.4344e-05 | 0.081839 | 0.061062 | 2.355e-05
299 | 0.0012523 | 0.00030325 | 0.29764 | 0.00037275 | 0.0012523 | 0.00031125 | 0.01 | 0.99312 | 0.00030911 | 0.053 | 0.12045 | 3.7234e-05 | 0.015646 | 0.011674 | 3.6087e-06
300 | 0.0010369 | 0.00028382 | 0.27857 | 0.00028885 | 0.0010369 | 0.0002577 | 0.01 | 0.99312 | 0.00025593 | 0.045 | 0.10227 | 2.6175e-05 | 0 | 0 | 0
TRLE: | 1 |  |  | 0.99339 |  | 0.24854 |  | Average Irradiance | 0.24428 |  | Germicidal | 0.24215 |  |  | 0.23731
|  |  | Lamp Correction Factor, fp | 1.0067
|  |  | Incident IR= | 0.24854

## Emission spectra
|  |  | Interpolation over germicidal range (200 to 300) |  |  |  | Calculating rated lamp energy (RLE)
Spectrometer Measurement |  |  | x | x1, x2 | y1, y2 | y |  | interpolated | RLE
200.11 | -7.7834e-05 | -0.015575 | 200.11 | 200.11 | -7.7834e-05 | -7.7834e-05 | 200.11 | -7.7834e-05 | -5.053e-05 | -0.015575
200.23 | 0.00036627 | 0.073338 |  |  |  |  | 201 | 0.00014423 | 9.3637e-05 | 0.028991
200.35 | -0.00024827 | -0.049741 | 201 | 200.94 | -7.2952e-05 | 0.00014423 | 202 | -0.00015118 | -9.8146e-05 | -0.030538
200.46 | 1.8916e-05 | 0.0037919 |  | 201.06 | 0.00036142 |  | 203 | 1.6268e-05 | 1.0561e-05 | 0.0033024
200.58 | 1.8746e-05 | 0.0037601 | 202 | 202 | -0.00015118 | -0.00015118 | 204 | 9.2749e-05 | 6.0213e-05 | 0.018921
200.7 | 0.00033433 | 0.0671 |  | 202.12 | 8.322e-05 |  | 205 | 0.00016553 | 0.00010746 | 0.033933
200.82 | 3.6812e-05 | 0.0073926 | 203 | 202.95 | -0.0001248 | 1.6268e-05 | 206 | 0.00011449 | 7.4328e-05 | 0.023585
200.94 | -7.2952e-05 | -0.014659 |  | 203.06 | 0.00018555 |  | 207 | 0.00024707 | 0.0001604 | 0.051143
201.06 | 0.00036142 | 0.072667 | 204 | 203.89 | 0.00016014 | 9.2749e-05 | 208 | 0.00014249 | 9.2506e-05 | 0.029638
201.17 | 0.00012534 | 0.025215 |  | 204.01 | 8.6623e-05 |  | 209 | 0.00010341 | 6.7136e-05 | 0.021613
201.29 | -0.00021294 | -0.042863 | 205 | 204.95 | 0.00018844 | 0.00016553 | 210 | 3.2395e-05 | 2.1031e-05 | 0.006803
201.41 | 0.00024617 | 0.049581 |  | 205.07 | 0.00013345 |  | 211 | 0.000127 | 8.2448e-05 | 0.026797
201.53 | -0.00020908 | -0.042136 | 206 | 205.89 | 0 | 0.00011449 | 212 | -3.7004e-05 | -2.4023e-05 | -0.0078448
201.65 | 0.00032795 | 0.066131 |  | 206.01 | 0.0001249 |  | 213 | -4.5616e-06 | -2.9614e-06 | -0.00097163
201.77 | 0.00054744 | 0.11046 | 207 | 206.95 | 0.00024775 | 0.00024707 | 214 | 9.0051e-05 | 5.8461e-05 | 0.019271
201.88 | 0.00010173 | 0.020537 |  | 207.07 | 0.00017578 |  | 215 | 2.7387e-05 | 1.778e-05 | 0.0058882
202 | -0.00015118 | -0.030538 | 208 | 207.9 | 0.00023482 | 0.00014249 | 216 | 6.9361e-05 | 4.5029e-05 | 0.014982
202.12 | 8.322e-05 | 0.01682 |  | 208.01 | 0.00013326 |  | 217 | 0.00010054 | 6.5269e-05 | 0.021817
202.24 | 4.9475e-05 | 0.010006 | 209 | 208.95 | 0.00014739 | 0.00010341 | 218 | 1.2652e-05 | 8.2136e-06 | 0.0027581
202.36 | 0.00027773 | 0.056201 |  | 209.07 | 4.1845e-05 |  | 219 | 6.5925e-06 | 4.2798e-06 | 0.0014438
202.47 | 0.00045315 | 0.091749 | 210 | 209.89 | 6.0117e-05 | 3.2395e-05 | 220 | -9.7954e-06 | -6.3592e-06 | -0.002155
202.59 | 0.00032065 | 0.06496 |  | 210.01 | 2.9875e-05 |  | 221 | 9.1511e-06 | 5.9409e-06 | 0.0020224
202.71 | 0.00044476 | 0.090157 | 211 | 210.95 | 0.00012337 | 0.000127 | 222 | 8.5502e-05 | 5.5508e-05 | 0.018981
202.83 | -3.1477e-05 | -0.0063845 |  | 211.07 | 0.00013208 |  | 223 | 0.00010062 | 6.5322e-05 | 0.022438
202.95 | -0.0001248 | -0.025328 | 212 | 211.89 | -4.5273e-05 | -3.7004e-05 | 224 | 5.6438e-05 | 3.6639e-05 | 0.012642
203.06 | 0.00018555 | 0.037678 |  | 212.01 | 0.00010805 |  | 225 | 2.2163e-06 | 1.4388e-06 | 0.00049866
203.18 | 0.00024523 | 0.049826 | 213 | 212.95 | 3.4412e-05 | -4.5616e-06 | 226 | 8.3754e-05 | 5.4373e-05 | 0.018928
203.3 | -3.039e-05 | -0.0061783 |  | 213.06 | -5.133e-05 |  | 227 | 2.106e-05 | 1.3672e-05 | 0.0047805
203.42 | -1.5065e-05 | -0.0030645 | 214 | 214 | 9.0051e-05 | 9.0051e-05 | 228 | 0.00012572 | 8.1617e-05 | 0.028664
203.54 | 0 | 0 |  | 214.12 | 2.4428e-05 |  | 229 | 1.2907e-05 | 8.3793e-06 | 0.0029557
203.65 | -0.00013327 | -0.02714 | 215 | 214.94 | 7.86e-06 | 2.7387e-05 | 230 | 4.2765e-05 | 2.7763e-05 | 0.009836
203.77 | 0.00030831 | 0.062824 |  | 215.06 | 4.6914e-05 |  | 231 | 3.2989e-05 | 2.1417e-05 | 0.0076206
203.89 | 0.00016014 | 0.032651 | 216 | 215.99 | 6.7525e-05 | 6.9361e-05 | 232 | 5.7213e-05 | 3.7143e-05 | 0.013273
204.01 | 8.6623e-05 | 0.017672 |  | 216.11 | 8.956e-05 |  | 233 | 3.4518e-05 | 2.2409e-05 | 0.0080428
204.13 | 0 | 0 | 217 | 216.93 | 0.00010085 | 0.00010054 | 234 | 3.3415e-05 | 2.1693e-05 | 0.007819
204.24 | 0.00031237 | 0.063798 |  | 217.04 | 0.00010036 |  | 235 | 5.0468e-05 | 3.2764e-05 | 0.01186
204.36 | 0.00014077 | 0.028768 | 218 | 217.98 | 1.3808e-05 | 1.2652e-05 | 236 | 9.6738e-05 | 6.2802e-05 | 0.02283
204.48 | 0.00016743 | 0.034236 |  | 218.1 | 6.8718e-06 |  | 237 | -2.4315e-06 | -1.5785e-06 | -0.00057626
204.6 | 0.00031807 | 0.065077 | 219 | 218.91 | -1.3284e-05 | 6.5925e-06 | 238 | 3.0979e-05 | 2.0111e-05 | 0.0073729
204.72 | 0.00023296 | 0.047692 |  | 219.03 | 1.3218e-05 |  | 239 | 2.4598e-05 | 1.5969e-05 | 0.0058789
204.83 | 8.1488e-05 | 0.016691 | 220 | 219.97 | -6.3542e-06 | -9.7954e-06 | 240 | 5.4821e-05 | 3.559e-05 | 0.013157
204.95 | 0.00018844 | 0.038621 |  | 220.08 | -1.8972e-05 |  | 241 | 5.0345e-05 | 3.2684e-05 | 0.012133
205.07 | 0.00013345 | 0.027367 | 221 | 220.9 | 2.4466e-05 | 9.1511e-06 | 242 | 3.7978e-05 | 2.4655e-05 | 0.0091907
205.19 | -2.6458e-05 | -0.0054289 |  | 221.02 | 6.0881e-06 |  | 243 | 3.0243e-05 | 1.9634e-05 | 0.0073491
205.3 | 0.00019672 | 0.040387 | 222 | 221.95 | 9.3658e-05 | 8.5502e-05 | 244 | 2.6139e-05 | 1.6969e-05 | 0.0063778
205.42 | 0.00024704 | 0.050747 |  | 222.06 | 7.5714e-05 |  | 245 | 5.2357e-05 | 3.399e-05 | 0.012827
205.54 | 0.00018052 | 0.037104 | 223 | 223 | 0.00010062 | 0.00010062 | 246 | 7.4324e-05 | 4.8251e-05 | 0.018284
205.66 | 0.00024293 | 0.049961 |  | 223.11 | 5.006e-05 |  | 247 | 4.0434e-05 | 2.625e-05 | 0.0099872
205.78 | 0.00027903 | 0.057419 | 224 | 223.93 | 4.2968e-05 | 5.6438e-05 | 248 | 3.2491e-05 | 2.1093e-05 | 0.0080578
205.89 | 0 | 0 |  | 224.04 | 6.4135e-05 |  | 249 | 9.5882e-05 | 6.2246e-05 | 0.023875
206.01 | 0.0001249 | 0.025731 | 225 | 224.98 | -2.5669e-05 | 2.2163e-06 | 250 | 8.8045e-05 | 5.7159e-05 | 0.022011
206.13 | -0.00026038 | -0.053672 |  | 225.09 | 0.0001277 |  | 251 | 8.5836e-05 | 5.5725e-05 | 0.021545
206.25 | 7.3858e-05 | 0.015233 | 226 | 225.91 | 6.4001e-05 | 8.3754e-05 | 252 | 4.0735e-05 | 2.6445e-05 | 0.010265
206.37 | 0.00017108 | 0.035306 |  | 226.02 | 8.8143e-05 |  | 253 | 7.7117e-05 | 5.0065e-05 | 0.019511
206.48 | 3.6394e-05 | 0.0075146 | 227 | 226.95 | 9.3951e-06 | 2.106e-05 | 254 | 0.00011557 | 7.5026e-05 | 0.029354
206.6 | -0.00018071 | -0.037335 |  | 227.07 | 3.739e-05 |  | 255 | 8.7162e-05 | 5.6585e-05 | 0.022226
206.72 | 2.3925e-05 | 0.0049458 | 228 | 228 | 0.00012572 | 0.00012572 | 256 | 0.00014074 | 9.1368e-05 | 0.036029
206.84 | 0.00019007 | 0.039314 |  | 228.11 | 3.1272e-05 |  | 257 | 0.00014551 | 9.4467e-05 | 0.037397
206.95 | 0.00024775 | 0.051272 | 229 | 228.92 | 2.1561e-05 | 1.2907e-05 | 258 | 0.00017751 | 0.00011524 | 0.045797
207.07 | 0.00017578 | 0.036399 |  | 229.04 | 8.5802e-06 |  | 259 | 0.00018611 | 0.00012082 | 0.048202
207.19 | 9.313e-05 | 0.019296 | 230 | 229.97 | 6.1881e-05 | 4.2765e-05 | 260 | 0.00018483 | 0.00011999 | 0.048055
207.31 | 4.6255e-05 | 0.0095891 |  | 230.08 | -8.2109e-06 |  | 261 | 0.00025552 | 0.00016588 | 0.06669
207.42 | 0.00010337 | 0.021441 | 231 | 230.89 | 3.9766e-06 | 3.2989e-05 | 262 | 0.0004558 | 0.0002959 | 0.11942
207.54 | 4.5634e-05 | 0.0094709 |  | 231.01 | 3.5627e-05 |  | 263 | 0.00094198 | 0.00061153 | 0.24774
207.66 | 0.00022666 | 0.047068 | 232 | 231.94 | 5.7341e-05 | 5.7213e-05 | 264 | 0.0019347 | 0.001256 | 0.51077
207.78 | -6.755e-05 | -0.014036 |  | 232.05 | 5.7106e-05 |  | 265 | 0.0035004 | 0.0022725 | 0.92762
207.9 | 0.00023482 | 0.048819 | 233 | 232.98 | 2.5841e-05 | 3.4518e-05 | 266 | 0.0060425 | 0.0039228 | 1.6073
208.01 | 0.00013326 | 0.027719 |  | 233.09 | 7.3566e-05 |  | 267 | 0.0097247 | 0.0063133 | 2.5965
208.13 | 0 | 0 | 234 | 233.9 | 3.9492e-05 | 3.3415e-05 | 268 | 0.014493 | 0.0094092 | 3.8843
208.25 | 6.5732e-05 | 0.013689 |  | 234.02 | 3.2199e-05 |  | 269 | 0.021458 | 0.01393 | 5.7721
208.37 | 0.00020675 | 0.04308 | 235 | 234.94 | 5.5767e-05 | 5.0468e-05 | 270 | 0.031483 | 0.020438 | 8.5003
208.48 | 0.00010808 | 0.022533 |  | 235.06 | 4.517e-05 |  | 271 | 0.045192 | 0.029338 | 12.247
208.6 | -0.00016102 | -0.033589 | 236 | 235.98 | 0.00010525 | 9.6738e-05 | 272 | 0.063507 | 0.041229 | 17.2739
208.72 | 0.0001493 | 0.031162 |  | 236.1 | 5.4176e-05 |  | 273 | 0.084963 | 0.055158 | 23.1948
208.84 | 5.2977e-05 | 0.011064 | 237 | 236.91 | -4.3228e-05 | -2.4315e-06 | 274 | 0.10537 | 0.068403 | 28.8701
208.95 | 0.00014739 | 0.030797 |  | 237.02 | 6.6344e-06 |  | 275 | 0.12254 | 0.079556 | 33.6998
209.07 | 4.1845e-05 | 0.0087485 | 238 | 237.94 | 5.5446e-05 | 3.0979e-05 | 276 | 0.13341 | 0.086607 | 36.8199
209.19 | 0.00013516 | 0.028274 |  | 238.06 | 6.5111e-06 |  | 277 | 0.13482 | 0.087524 | 37.3449
209.31 | 0.00013434 | 0.028119 | 239 | 238.98 | 2.8877e-05 | 2.4598e-05 | 278 | 0.12818 | 0.083214 | 35.634
209.42 | -0.0001746 | -0.036565 |  | 239.1 | 3.2029e-06 |  | 279 | 0.11513 | 0.074745 | 32.1223
209.54 | 3.0623e-05 | 0.0064167 | 240 | 239.9 | 4.4339e-05 | 5.4821e-05 | 280 | 0.099158 | 0.064373 | 27.7641
209.66 | 0.00018259 | 0.038282 |  | 240.02 | 5.6917e-05 |  | 281 | 0.082243 | 0.053392 | 23.1103
209.78 | 0.00029235 | 0.061329 | 241 | 240.94 | 6.5734e-05 | 5.0345e-05 | 282 | 0.067162 | 0.043602 | 18.9398
209.89 | 6.0117e-05 | 0.012618 |  | 241.05 | 3.752e-05 |  | 283 | 0.053272 | 0.034584 | 15.0759
210.01 | 2.9875e-05 | 0.006274 | 242 | 241.97 | 5.2702e-05 | 3.7978e-05 | 284 | 0.042588 | 0.027648 | 12.095
210.13 | 0.00024746 | 0.051999 |  | 242.09 | -6.1935e-06 |  | 285 | 0.033489 | 0.021741 | 9.5443
210.25 | -0.00013773 | -0.028958 | 243 | 242.89 | -9.2329e-06 | 3.0243e-05 | 286 | 0.026499 | 0.017203 | 7.5787
210.36 | -9.7781e-06 | -0.0020569 |  | 243.01 | 3.3832e-05 |  | 287 | 0.021074 | 0.013682 | 6.0484
210.48 | 7.7755e-05 | 0.016366 | 244 | 243.93 | 1.8361e-05 | 2.6139e-05 | 288 | 0.016945 | 0.011001 | 4.8802
210.6 | 7.7288e-05 | 0.016277 |  | 244.04 | 3.0583e-05 |  | 289 | 0.013823 | 0.0089741 | 3.9949
210.72 | 0.00011524 | 0.024283 | 245 | 244.96 | 8.2275e-05 | 5.2357e-05 | 290 | 0.011324 | 0.0073516 | 3.284
210.83 | 0.00011455 | 0.024151 |  | 245.07 | 0 |  | 291 | 0.0092267 | 0.00599 | 2.685
210.95 | 0.00012337 | 0.026025 | 246 | 245.99 | 7.281e-05 | 7.4324e-05 | 292 | 0.0077758 | 0.005048 | 2.2705
211.07 | 0.00013208 | 0.027878 |  | 246.11 | 9.0979e-05 |  | 293 | 0.0063576 | 0.0041274 | 1.8628
211.19 | 0.00011255 | 0.023769 | 247 | 246.91 | 4.5388e-05 | 4.0434e-05 | 294 | 0.0052203 | 0.003389 | 1.5348
211.3 | 0.00013055 | 0.027585 |  | 247.02 | 3.9333e-05 |  | 295 | 0.0043934 | 0.0028522 | 1.296
211.42 | 0.00010196 | 0.021556 | 248 | 247.94 | 6.0557e-06 | 3.2491e-05 | 296 | 0.0035661 | 0.0023151 | 1.0556
211.54 | 7.3721e-05 | 0.015595 |  | 248.05 | 5.4521e-05 |  | 297 | 0.0029741 | 0.0019308 | 0.88332
211.66 | 0.00021073 | 0.044603 | 249 | 248.97 | 7.2726e-05 | 9.5882e-05 | 298 | 0.0024084 | 0.0015636 | 0.71772
211.77 | -0.00013662 | -0.028932 |  | 249.08 | 0.00015763 |  | 299 | 0.0019291 | 0.0012523 | 0.57679
211.89 | -4.5273e-05 | -0.0095929 | 250 | 250 | 8.8045e-05 | 8.8045e-05 | 300 | 0.0015972 | 0.0010369 | 0.47915
212.01 | 0.00010805 | 0.022908 |  | 250.11 | 4.2514e-05 |  |  | 1.5404 | 1 | 428.0379
212.13 | 1.7908e-05 | 0.0037988 | 251 | 250.91 | 4.2534e-05 | 8.5836e-05 |  |  | Weighted Avg: | 277.8818
212.24 | 0.00015135 | 0.032123 |  | 251.03 | 0.00010027
212.36 | 8.8517e-06 | 0.0018797 | 252 | 251.94 | 9.1427e-06 | 4.0735e-05
212.48 | 2.6402e-05 | 0.0056099 |  | 252.05 | 6.7062e-05
212.59 | 7.001e-05 | 0.014883 | 253 | 252.97 | 7.6287e-05 | 7.7117e-05
212.71 | 2.6106e-05 | 0.005553 |  | 253.08 | 7.9332e-05
212.83 | 0.00016439 | 0.034987 | 254 | 253.99 | 0.00011913 | 0.00011557
212.95 | 3.4412e-05 | 0.007328 |  | 254.11 | 7.6369e-05
213.06 | -5.133e-05 | -0.010936 | 255 | 254.9 | 9.4795e-05 | 8.7162e-05
213.18 | 2.5521e-05 | 0.0054406 |  | 255.02 | 8.5635e-05
213.3 | 0.00010151 | 0.021652 | 256 | 255.93 | 0.00011932 | 0.00014074
213.42 | 0.00010093 | 0.02154 |  | 256.04 | 0.00015298
213.53 | 4.1817e-05 | 0.0089292 | 257 | 256.95 | 0.0001468 | 0.00014551
213.65 | 4.1591e-05 | 0.0088859 |  | 257.07 | 0.00014371
213.77 | 0.00015722 | 0.033609 | 258 | 257.98 | 0.0001864 | 0.00017751
213.88 | 0.00010699 | 0.022883 |  | 258.09 | 0.0001375
214 | 9.0051e-05 | 0.019271 | 259 | 259 | 0.00018611 | 0.00018611
214.12 | 2.4428e-05 | 0.0052305 |  | 259.11 | 0.0001708
214.24 | -8.1008e-06 | -0.0017355 | 260 | 259.91 | 0.00011264 | 0.00018483
214.35 | 0.00020959 | 0.044926 |  | 260.02 | 0.00020087
214.47 | 2.4059e-05 | 0.0051599 | 261 | 260.93 | 0.00022464 | 0.00025552
214.59 | 7.9792e-06 | 0.0017123 |  | 261.04 | 0.00027316
214.7 | 5.5575e-05 | 0.011932 | 262 | 261.95 | 0.00046963 | 0.0004558
214.82 | 3.16e-05 | 0.0067883 |  | 262.06 | 0.0004392
214.94 | 7.86e-06 | 0.0016894 | 263 | 262.97 | 0.00093052 | 0.00094198
215.06 | 4.6914e-05 | 0.010089 |  | 263.08 | 0.00097255
215.17 | 0.0001089 | 0.023432 | 264 | 263.99 | 0.0019233 | 0.0019347
215.29 | -6.191e-05 | -0.013329 |  | 264.1 | 0.0020492
215.41 | 7.7002e-05 | 0.016587 | 265 | 264.89 | 0.0033359 | 0.0035004
215.52 | -7.66e-06 | -0.0016509 |  | 265.01 | 0.0035154
215.64 | 7.6194e-05 | 0.01643 | 266 | 265.91 | 0.0057415 | 0.0060425
215.76 | 0.00014404 | 0.031078 |  | 266.02 | 0.0061094
215.87 | 8.2954e-05 | 0.017907 | 267 | 266.93 | 0.0094594 | 0.0097247
215.99 | 6.7525e-05 | 0.014585 |  | 267.04 | 0.0098763
216.11 | 8.956e-05 | 0.019355 | 268 | 267.94 | 0.014027 | 0.014493
216.23 | -1.4851e-05 | -0.0032112 |  | 268.06 | 0.01496
216.34 | 0.00012559 | 0.02717 | 269 | 268.96 | 0.021046 | 0.021458
216.46 | 8.0843e-05 | 0.017499 |  | 269.07 | 0.022178
216.58 | 0.00019743 | 0.042759 | 270 | 269.97 | 0.031109 | 0.031483
216.69 | 6.5479e-05 | 0.014189 |  | 270.08 | 0.032479
216.81 | 9.4114e-05 | 0.020405 | 271 | 270.99 | 0.044993 | 0.045192
216.93 | 0.00010085 | 0.021877 |  | 271.1 | 0.047179
217.04 | 0.00010036 | 0.021782 | 272 | 272 | 0.063507 | 0.063507
217.16 | -0.00011415 | -0.024789 |  | 272.11 | 0.0661
217.28 | -1.4201e-05 | -0.0030856 | 273 | 272.9 | 0.082101 | 0.084963
217.4 | 4.2398e-05 | 0.0092173 |  | 273.01 | 0.085249
217.51 | 4.9234e-05 | 0.010709 | 274 | 273.91 | 0.10377 | 0.10537
217.63 | 0.00017502 | 0.03809 |  | 274.02 | 0.10572
217.75 | 0.00011151 | 0.024281 | 275 | 274.92 | 0.12133 | 0.12254
217.86 | 0.00010405 | 0.022668 |  | 275.03 | 0.123
217.98 | 1.3808e-05 | 0.0030099 | 276 | 275.93 | 0.13296 | 0.13341
218.1 | 6.8718e-06 | 0.0014987 |  | 276.04 | 0.13366
218.21 | 0.00011626 | 0.025369 | 277 | 276.94 | 0.13507 | 0.13482
218.33 | 0.00010889 | 0.023774 |  | 277.05 | 0.13461
218.45 | -6.7728e-06 | -0.0014795 | 278 | 277.95 | 0.12853 | 0.12818
218.56 | 0.00022238 | 0.048603 |  | 278.06 | 0.12776
218.68 | 4.6951e-05 | 0.010267 | 279 | 278.96 | 0.11569 | 0.11513
218.8 | 5.3394e-05 | 0.011683 |  | 279.07 | 0.11416
218.91 | -1.3284e-05 | -0.002908 | 280 | 279.96 | 0.099917 | 0.099158
219.03 | 1.3218e-05 | 0.0028951 |  | 280.08 | 0.097639
219.15 | 7.2349e-05 | 0.015855 | 281 | 280.97 | 0.082634 | 0.082243
219.26 | 1.9633e-05 | 0.0043047 |  | 281.08 | 0.0812
219.38 | 0.00013676 | 0.030002 | 282 | 281.98 | 0.067541 | 0.067162
219.5 | -1.9441e-05 | -0.0042673 |  | 282.09 | 0.065458
219.62 | 1.9345e-05 | 0.0042485 | 283 | 282.98 | 0.053491 | 0.053272
219.73 | 0.00010265 | 0.022555 |  | 283.09 | 0.052285
219.85 | 8.9395e-05 | 0.019653 | 284 | 283.99 | 0.042715 | 0.042588
219.97 | -6.3542e-06 | -0.0013977 |  | 284.1 | 0.04132
220.08 | -1.8972e-05 | -0.0041754 | 285 | 284.99 | 0.033574 | 0.033489
220.2 | 0.00010071 | 0.022176 |  | 285.1 | 0.032635
220.32 | 9.3968e-05 | 0.020703 | 286 | 285.99 | 0.026574 | 0.026499
220.43 | -6.234e-06 | -0.0013742 |  | 286.1 | 0.025747
220.55 | -4.343e-05 | -0.0095785 | 287 | 286.99 | 0.021127 | 0.021074
220.67 | 0.00014203 | 0.031342 |  | 287.1 | 0.020549
220.78 | 9.8334e-05 | 0.02171 | 288 | 287.99 | 0.016975 | 0.016945
220.9 | 2.4466e-05 | 0.0054045 |  | 288.11 | 0.016616
221.02 | 6.0881e-06 | 0.0013456 | 289 | 288.99 | 0.013859 | 0.013823
221.13 | 0.00011511 | 0.025454 |  | 289.11 | 0.013431
221.25 | 8.4399e-05 | 0.018673 | 290 | 289.99 | 0.011351 | 0.011324
221.37 | 6.0001e-05 | 0.013282 |  | 290.11 | 0.011029
221.48 | 0.00014928 | 0.033063 | 291 | 290.99 | 0.0092438 | 0.0092267
221.6 | 5.9418e-05 | 0.013167 |  | 291.1 | 0.0090561
221.71 | 1.774e-05 | 0.0039331 | 292 | 291.99 | 0.0077976 | 0.0077758
221.83 | 5.8839e-05 | 0.013052 |  | 292.1 | 0.0075574
221.95 | 9.3658e-05 | 0.020787 | 293 | 292.99 | 0.0063671 | 0.0063576
222.06 | 7.5714e-05 | 0.016813 |  | 293.1 | 0.006263
222.18 | 4.6363e-05 | 0.010301 | 294 | 293.99 | 0.005223 | 0.0052203
222.3 | 5.1898e-05 | 0.011537 |  | 294.1 | 0.0051928
222.41 | 8.0325e-05 | 0.017865 | 295 | 294.98 | 0.0044045 | 0.0043934
222.53 | 0 | 0 |  | 295.09 | 0.0043432
222.65 | -3.9749e-05 | -0.0088501 | 296 | 295.98 | 0.0035926 | 0.0035661
222.76 | 5.6481e-06 | 0.0012582 |  | 296.09 | 0.0034469
222.88 | 5.0573e-05 | 0.011272 | 297 | 296.97 | 0.0029982 | 0.0029741
223 | 0.00010062 | 0.022438 |  | 297.08 | 0.00291
223.11 | 5.006e-05 | 0.011169 | 298 | 297.96 | 0.0024723 | 0.0024084
223.23 | 0.0001107 | 0.024712 |  | 298.07 | 0.0022967
223.35 | 9.3628e-05 | 0.020912 | 299 | 298.96 | 0.0019168 | 0.0019291
223.46 | 8.2203e-05 | 0.018369 |  | 299.07 | 0.0019505
223.58 | 8.1793e-05 | 0.018287 | 300 | 299.95 | 0.0016224 | 0.0015972
223.7 | 0.00013021 | 0.029128 |  | 300.06 | 0.0015669
223.81 | 3.7789e-05 | 0.0084576
223.93 | 4.2968e-05 | 0.0096218
224.04 | 6.4135e-05 | 0.014369
224.16 | 3.1906e-05 | 0.007152
224.28 | -2.6459e-05 | -0.0059342
224.39 | 1.0532e-05 | 0.0023633
224.51 | -5.2396e-05 | -0.011763
224.63 | 4.6914e-05 | 0.010538
224.74 | 2.0744e-05 | 0.004662
224.86 | 0.0001032 | 0.023206
224.98 | -2.5669e-05 | -0.005775
225.09 | 0.0001277 | 0.028744
225.21 | 0.00010672 | 0.024034
225.32 | 2.5279e-05 | 0.0056959
225.44 | 8.0479e-05 | 0.018143
225.56 | -3.502e-05 | -0.0078991
225.67 | 9.9511e-06 | 0.0022457
225.79 | -5.444e-05 | -0.012292
225.91 | 6.4001e-05 | 0.014458
226.02 | 8.8143e-05 | 0.019922
226.14 | -5.845e-05 | -0.013218
226.25 | 8.2375e-05 | 0.018637
226.37 | 2.8923e-05 | 0.0065473
226.49 | -4.795e-05 | -0.01086
226.6 | 4.7701e-06 | 0.0010809
226.72 | 3.7965e-05 | 0.0086074
226.83 | -1.8886e-05 | -0.0042839
226.95 | 9.3951e-06 | 0.0021322
227.07 | 3.739e-05 | 0.0084901
227.18 | -1.3952e-05 | -0.0031696
227.3 | 6.4792e-05 | 0.014727
227.42 | 5.9861e-05 | 0.013614
227.53 | 0 | 0
227.65 | 3.1904e-05 | 0.0072629
227.76 | -4.5347e-06 | -0.0010328
227.88 | 0.00013988 | 0.031876
228 | 0.00012572 | 0.028664
228.11 | 3.1272e-05 | 0.0071335
228.23 | 0 | 0
228.34 | 8.4033e-05 | 0.019188
228.46 | 3.0799e-05 | 0.0070363
228.58 | 1.751e-05 | 0.0040024
228.69 | -6.9689e-05 | -0.015937
228.81 | -3.4668e-05 | -0.0079324
228.92 | 2.1561e-05 | 0.0049357
229.04 | 8.5802e-06 | 0.0019652
229.16 | 5.5492e-05 | 0.012717
229.27 | 2.1235e-05 | 0.0048685
229.39 | 3.3801e-05 | 0.0077536
229.5 | 8.4086e-06 | 0.0019298
229.62 | -1.2553e-05 | -0.0028824
229.74 | 2.499e-05 | 0.0057412
229.85 | -4.145e-06 | -0.00095273
229.97 | 6.1881e-05 | 0.014231
230.08 | -8.2109e-06 | -0.0018892
230.2 | 4.0863e-05 | 0.0094067
230.32 | 6.5086e-05 | 0.014991
230.43 | 8.9073e-05 | 0.020525
230.55 | 4.0305e-06 | 0.00092923
230.66 | 9.2309e-05 | 0.021292
230.78 | 1.5981e-05 | 0.0036881
230.89 | 3.9766e-06 | 0.00091816
231.01 | 3.5627e-05 | 0.0082302
231.13 | 5.1227e-05 | 0.01184
231.24 | 7.0613e-05 | 0.016329
231.36 | 5.8593e-05 | 0.013556
231.47 | 5.8342e-05 | 0.013504
231.59 | 9.6813e-05 | 0.022421
231.71 | -1.1568e-05 | -0.0026804
231.82 | -4.9908e-05 | -0.01157
231.94 | 5.7341e-05 | 0.0133
232.05 | 5.7106e-05 | 0.013251
232.17 | -2.654e-05 | -0.0061618
232.28 | -2.6435e-05 | -0.0061403
232.4 | 4.8903e-05 | 0.011365
232.52 | 4.4969e-05 | 0.010456
232.63 | 2.6132e-05 | 0.0060791
232.75 | 2.2314e-05 | 0.0051936
232.86 | 0.00010004 | 0.023295
232.98 | 2.5841e-05 | 0.0060204
233.09 | 7.3566e-05 | 0.017147
233.21 | 4.7642e-05 | 0.011111
233.33 | -1.826e-05 | -0.0042606
233.44 | 1.4559e-05 | 0.0033987
233.56 | 5.4405e-05 | 0.012707
233.67 | 2.892e-05 | 0.0067577
233.79 | 1.0807e-05 | 0.0025266
233.9 | 3.9492e-05 | 0.0092372
234.02 | 3.2199e-05 | 0.0075352
234.13 | 3.566e-06 | 0.00083491
234.25 | 1.0664e-05 | 0.002498
234.37 | 3.1884e-05 | 0.0074727
234.48 | -2.8249e-05 | -0.0066238
234.6 | 1.4077e-05 | 0.0033025
234.71 | 7.0161e-06 | 0.0016467
234.83 | -6.9936e-06 | -0.0016423
234.94 | 5.5767e-05 | 0.013102
235.06 | 4.517e-05 | 0.010618
235.17 | 2.0785e-05 | 0.004888
235.29 | 3.7999e-05 | 0.0089408
235.41 | 5.1662e-05 | 0.012162
235.52 | 3.0905e-05 | 0.0072787
235.64 | 6.1634e-05 | 0.014523
235.75 | -5.8045e-05 | -0.013684
235.87 | -1.7024e-05 | -0.0040155
235.98 | 0.00010525 | 0.024837
236.1 | 5.4176e-05 | 0.012791
236.21 | 4.3903e-05 | 0.01037
236.33 | 6.7363e-05 | 0.01592
236.44 | 2.6872e-05 | 0.0063536
236.56 | 6.7001e-05 | 0.01585
236.68 | 2.005e-05 | 0.0047454
236.79 | 5.9996e-05 | 0.014206
236.91 | -4.3228e-05 | -0.010241
237.02 | 6.6344e-06 | 0.0015725
237.14 | 9.9288e-06 | 0.0023545
237.25 | 1.9814e-05 | 0.0047009
237.37 | 6.5898e-06 | 0.0015642
237.48 | 5.2595e-05 | 0.01249
237.6 | 3.2805e-05 | 0.0077945
237.71 | -9.8234e-06 | -0.0023351
237.83 | 2.6145e-05 | 0.0062181
237.94 | 5.5446e-05 | 0.013193
238.06 | 6.5111e-06 | 0.00155
238.17 | 1.9498e-05 | 0.0046438
238.29 | 4.8656e-05 | 0.011594
238.4 | 5.8279e-05 | 0.013894
238.52 | 4.2009e-05 | 0.01002
238.64 | 2.2579e-05 | 0.0053883
238.75 | 5.7962e-05 | 0.013838
238.87 | 3.8572e-05 | 0.0092137
238.98 | 2.8877e-05 | 0.006901
239.1 | 3.2029e-06 | 0.00076581
239.21 | 4.7961e-05 | 0.011473
239.33 | 7.6613e-05 | 0.018336
239.44 | 6.0543e-05 | 0.014496
239.56 | 1.5908e-05 | 0.0038109
239.67 | -1.5884e-05 | -0.0038069
239.79 | 4.4404e-05 | 0.010648
239.9 | 4.4339e-05 | 0.010637
240.02 | 5.6917e-05 | 0.013661
240.13 | 2.5261e-05 | 0.0060659
240.25 | 7.5671e-05 | 0.01818
240.36 | 1.2595e-05 | 0.0030273
240.48 | 4.7172e-05 | 0.011344
240.59 | 1.5704e-05 | 0.0037782
240.71 | 7.2164e-05 | 0.017371
240.82 | 6.5811e-05 | 0.015849
240.94 | 6.5734e-05 | 0.015838
241.05 | 3.752e-05 | 0.0090442
241.17 | 4.06e-05 | 0.0097915
241.28 | 5.3042e-05 | 0.012798
241.4 | 7.169e-05 | 0.017306
241.51 | 7.1616e-05 | 0.017296
241.63 | 2.7994e-05 | 0.0067642
241.74 | 3.7283e-05 | 0.0090128
241.86 | 3.4139e-05 | 0.0082569
241.97 | 5.2702e-05 | 0.012752
242.09 | -6.1935e-06 | -0.0014994
242.2 | 6.1873e-05 | 0.014986
242.32 | 5.5635e-05 | 0.013481
242.43 | 7.1019e-05 | 0.017217
242.55 | 3.0848e-06 | 0.00074822
242.66 | 1.8493e-05 | 0.0044875
242.78 | 4.0037e-05 | 0.0097202
242.89 | -9.2329e-06 | -0.0022426
243.01 | 3.3832e-05 | 0.0082215
243.12 | 1.5368e-05 | 0.0037363
243.24 | 2.4572e-05 | 0.0059769
243.35 | 3.0693e-05 | 0.0074691
243.47 | 3.0673e-06 | 0.0007468
243.58 | 2.4522e-05 | 0.0059731
243.7 | 7.046e-05 | 0.017171
243.81 | 3.062e-06 | 0.00074655
243.93 | 1.8361e-05 | 0.0044788
244.04 | 3.0583e-05 | 0.0074635
244.16 | 7.6408e-05 | 0.018656
244.27 | 1.5274e-05 | 0.003731
244.38 | 6.7178e-05 | 0.016417
244.5 | 5.7994e-05 | 0.01418
244.61 | 1.8308e-05 | 0.0044783
244.73 | -3.0499e-06 | -0.0007464
244.84 | 3.0488e-05 | 0.0074647
244.96 | 8.2275e-05 | 0.020154
245.07 | 0 | 0
245.19 | 7.9142e-05 | 0.019405
245.3 | 3.0428e-05 | 0.007464
245.42 | 6.3875e-05 | 0.015676
245.53 | -3.0403e-06 | -0.00074648
245.65 | 6.0774e-05 | 0.014929
245.76 | 3.9477e-05 | 0.0097019
245.88 | 2.428e-05 | 0.00597
245.99 | 7.281e-05 | 0.017911
246.11 | 9.0979e-05 | 0.022391
246.22 | 6.973e-05 | 0.017169
246.33 | 5.4555e-05 | 0.013439
246.45 | 1.212e-05 | 0.002987
246.56 | 3.0288e-05 | 0.0074678
246.68 | 0 | 0
246.79 | 3.3292e-05 | 0.0082161
246.91 | 4.5388e-05 | 0.011207
247.02 | 3.9333e-05 | 0.009716
247.14 | 3.3281e-05 | 0.0082251
247.25 | 9.9848e-05 | 0.024687
247.37 | 4.2362e-05 | 0.010479
247.48 | -1.8157e-05 | -0.0044935
247.59 | 3.935e-05 | 0.0097427
247.71 | 2.4218e-05 | 0.005999
247.82 | 6.9624e-05 | 0.017254
247.94 | 6.0557e-06 | 0.0015015
248.05 | 5.4521e-05 | 0.013524
248.17 | 3.6353e-05 | 0.0090217
248.28 | 4.8478e-05 | 0.012036
248.4 | 6.3637e-05 | 0.015807
248.51 | 3.9396e-05 | 0.0097903
248.62 | 9.0914e-06 | 0.0022603
248.74 | 7.2734e-05 | 0.018092
248.85 | 3.6363e-05 | 0.0090489
248.97 | 7.2726e-05 | 0.018107
249.08 | 0.00015763 | 0.039262
249.2 | 7.8829e-05 | 0.019644
249.31 | 7.5802e-05 | 0.018898
249.43 | 2.7296e-05 | 0.0068084
249.54 | 4.854e-05 | 0.012113
249.65 | 3.9443e-05 | 0.0098469
249.77 | 8.1942e-05 | 0.020467
249.88 | 7.5887e-05 | 0.018963
250 | 8.8045e-05 | 0.022011
250.11 | 4.2514e-05 | 0.010633
250.23 | 7.2889e-05 | 0.018239
250.34 | 7.5924e-05 | 0.019007
250.45 | 0.00010326 | 0.025861
250.57 | 7.8967e-05 | 0.019787
250.68 | 2.4298e-05 | 0.006091
250.8 | 4.8601e-05 | 0.012189
250.91 | 4.2534e-05 | 0.010672
251.03 | 0.00010027 | 0.025171
251.14 | 6.9905e-05 | 0.017556
251.25 | 8.5127e-05 | 0.021388
251.37 | 2.4331e-05 | 0.0061161
251.48 | 4.26e-05 | 0.010713
251.6 | 7.6104e-05 | 0.019148
251.71 | 0.00010048 | 0.025292
251.82 | 0.00011576 | 0.029151
251.94 | 9.1427e-06 | 0.0023034
252.05 | 6.7062e-05 | 0.016903
252.17 | 7.9278e-05 | 0.019992
252.28 | 5.4898e-05 | 0.01385
252.4 | 7.3208e-05 | 0.018478
252.51 | 7.0159e-05 | 0.017716
252.62 | 8.541e-05 | 0.021576
252.74 | 8.8474e-05 | 0.022361
252.85 | 3.3563e-05 | 0.0084864
252.97 | 7.6287e-05 | 0.019298
253.08 | 7.9332e-05 | 0.020077
253.19 | 7.628e-05 | 0.019313
253.31 | 6.7133e-05 | 0.017005
253.42 | 6.7143e-05 | 0.017015
253.54 | 5.7993e-05 | 0.014704
253.65 | 3.9686e-05 | 0.010066
253.76 | 4.8855e-05 | 0.012397
253.88 | 0.0001069 | 0.02714
253.99 | 0.00011913 | 0.030258
254.11 | 7.6369e-05 | 0.019406
254.22 | 9.7778e-05 | 0.024857
254.33 | 8.5571e-05 | 0.021763
254.45 | 0.00011309 | 0.028776
254.56 | 6.4205e-05 | 0.016344
254.68 | 0.00011619 | 0.029591
254.79 | 9.1731e-05 | 0.023372
254.9 | 9.4795e-05 | 0.024163
255.02 | 8.5635e-05 | 0.021839
255.13 | 0.00015902 | 0.040571
255.25 | 7.0337e-05 | 0.017954
255.36 | 0.00015596 | 0.039826
255.47 | 0.00012538 | 0.032031
255.59 | 7.6467e-05 | 0.019544
255.7 | 0.00013767 | 0.035202
255.81 | 7.6487e-05 | 0.019566
255.93 | 0.00011932 | 0.030538
256.04 | 0.00015298 | 0.039169
256.16 | 0.00015297 | 0.039185
256.27 | 5.2002e-05 | 0.013327
256.38 | 9.7889e-05 | 0.025097
256.5 | 0.00011626 | 0.029821
256.61 | 8.8719e-05 | 0.022766
256.72 | 0.0001499 | 0.038482
256.84 | 0.00010094 | 0.025925
256.95 | 0.0001468 | 0.03772
257.07 | 0.00014371 | 0.036944
257.18 | 9.784e-05 | 0.025162
257.29 | 0.00015592 | 0.040117
257.41 | 0.00012841 | 0.033054
257.52 | 0.00012842 | 0.033071
257.63 | 0.0001223 | 0.031508
257.75 | 0.00013145 | 0.033881
257.86 | 0.00010698 | 0.027586
257.98 | 0.0001864 | 0.048087
258.09 | 0.0001375 | 0.035487
258.2 | 0.0001283 | 0.033127
258.32 | 0.00011606 | 0.029981
258.43 | 0.0001527 | 0.039462
258.54 | 0.00013741 | 0.035526
258.66 | 0.00012822 | 0.033165
258.77 | 0.00016786 | 0.043437
258.88 | 0.00018918 | 0.048975
259 | 0.00018611 | 0.048202
259.11 | 0.0001708 | 0.044256
259.23 | 0.00021345 | 0.055333
259.34 | 0.00014328 | 0.037158
259.45 | 0.00021334 | 0.055351
259.57 | 0.00019196 | 0.049827
259.68 | 0.00019494 | 0.050622
259.79 | 0.00018269 | 0.047461
259.91 | 0.00011264 | 0.029276
260.02 | 0.00020087 | 0.05223
260.13 | 0.00023732 | 0.061734
260.25 | 0.00017942 | 0.046694
260.36 | 0.00019762 | 0.051452
260.47 | 0.00024013 | 0.062547
260.59 | 0.00024914 | 0.064923
260.7 | 0.000243 | 0.06335
260.81 | 0.00020953 | 0.054648
260.93 | 0.00022464 | 0.058615
261.04 | 0.00027316 | 0.071306
261.15 | 0.00030339 | 0.07923
261.27 | 0.00030028 | 0.078454
261.38 | 0.00033962 | 0.08877
261.49 | 0.00037904 | 0.099115
261.61 | 0.000385 | 0.10072
261.72 | 0.00039709 | 0.10393
261.83 | 0.00043034 | 0.11268
261.95 | 0.00046963 | 0.12302
262.06 | 0.0004392 | 0.1151
262.17 | 0.00056324 | 0.14766
262.29 | 0.00058735 | 0.15406
262.4 | 0.00062658 | 0.16441
262.51 | 0.00079876 | 0.20968
262.63 | 0.00071989 | 0.18906
262.74 | 0.00076187 | 0.20017
262.85 | 0.0008524 | 0.22405
262.97 | 0.00093052 | 0.2447
263.08 | 0.00097255 | 0.25586
263.19 | 0.0010688 | 0.2813
263.31 | 0.0012104 | 0.31871
263.42 | 0.0013429 | 0.35375
263.53 | 0.0014633 | 0.38562
263.65 | 0.0015624 | 0.41193
263.76 | 0.0016495 | 0.43507
263.87 | 0.0017697 | 0.46697
263.99 | 0.0019233 | 0.50773
264.1 | 0.0020492 | 0.54119
264.21 | 0.0021664 | 0.57238
264.33 | 0.0023886 | 0.63138
264.44 | 0.0025809 | 0.68249
264.55 | 0.0027761 | 0.73442
264.67 | 0.0028835 | 0.76318
264.78 | 0.0032645 | 0.86437
264.89 | 0.0033359 | 0.88365
265.01 | 0.0035154 | 0.93162
265.12 | 0.0037818 | 1.0026
265.23 | 0.004069 | 1.0792
265.34 | 0.0044225 | 1.1735
265.46 | 0.00465 | 1.2344
265.57 | 0.0048082 | 1.2769
265.68 | 0.0052575 | 1.3968
265.8 | 0.0055894 | 1.4857
265.91 | 0.0057415 | 1.5267
266.02 | 0.0061094 | 1.6252
266.14 | 0.0066324 | 1.7651
266.25 | 0.0070143 | 1.8676
266.36 | 0.0072315 | 1.9262
266.47 | 0.0076338 | 2.0342
266.59 | 0.0082164 | 2.1904
266.7 | 0.0084623 | 2.2569
266.81 | 0.0088292 | 2.3557
266.93 | 0.0094594 | 2.525
267.04 | 0.0098763 | 2.6374
267.15 | 0.01026 | 2.741
267.26 | 0.010846 | 2.8987
267.38 | 0.011035 | 2.9505
267.49 | 0.011648 | 3.1157
267.6 | 0.012295 | 3.2901
267.72 | 0.012822 | 3.4327
267.83 | 0.01344 | 3.5996
267.94 | 0.014027 | 3.7584
268.06 | 0.01496 | 4.0102
268.17 | 0.015364 | 4.1202
268.28 | 0.016076 | 4.3129
268.39 | 0.016615 | 4.4593
268.51 | 0.017414 | 4.6758
268.62 | 0.018619 | 5.0014
268.73 | 0.01935 | 5.1999
268.84 | 0.020151 | 5.4174
268.96 | 0.021046 | 5.6605
269.07 | 0.022178 | 5.9674
269.18 | 0.022997 | 6.1903
269.3 | 0.023656 | 6.3706
269.41 | 0.024988 | 6.732
269.52 | 0.026145 | 7.0466
269.63 | 0.027264 | 7.3512
269.75 | 0.028475 | 7.6811
269.86 | 0.029522 | 7.9668
269.97 | 0.031109 | 8.3985
270.08 | 0.032479 | 8.7719
270.2 | 0.033403 | 9.0255
270.31 | 0.03524 | 9.5257
270.42 | 0.036504 | 9.8714
270.54 | 0.038683 | 10.4653
270.65 | 0.039872 | 10.7914
270.76 | 0.041704 | 11.2918
270.87 | 0.043431 | 11.7642
270.99 | 0.044993 | 12.1927
271.1 | 0.047179 | 12.7902
271.21 | 0.049024 | 13.2958
271.32 | 0.051175 | 13.8848
271.44 | 0.052579 | 14.272
271.55 | 0.054768 | 14.8723
271.66 | 0.057099 | 15.5115
271.77 | 0.059166 | 16.0795
271.89 | 0.06155 | 16.7348
272 | 0.063507 | 17.2739
272.11 | 0.0661 | 17.9865
272.22 | 0.068172 | 18.5578
272.34 | 0.070005 | 19.0652
272.45 | 0.072728 | 19.8147
272.56 | 0.074899 | 20.4145
272.67 | 0.077201 | 21.0504
272.79 | 0.079744 | 21.7534
272.9 | 0.082101 | 22.4054
273.01 | 0.085249 | 23.2738
273.12 | 0.086879 | 23.7284
273.24 | 0.089383 | 24.423
273.35 | 0.092318 | 25.2351
273.46 | 0.094695 | 25.8953
273.57 | 0.096263 | 26.3347
273.69 | 0.099324 | 27.184
273.8 | 0.10197 | 27.9194
273.91 | 0.10377 | 28.4236
274.02 | 0.10572 | 28.9694
274.13 | 0.10818 | 29.6554
274.25 | 0.10977 | 30.1044
274.36 | 0.11211 | 30.7585
274.47 | 0.11453 | 31.435
274.58 | 0.11567 | 31.7607
274.7 | 0.11778 | 32.3542
274.81 | 0.11945 | 32.8261
274.92 | 0.12133 | 33.356
275.03 | 0.123 | 33.8287
275.15 | 0.12391 | 34.0938
275.26 | 0.12592 | 34.6607
275.37 | 0.12784 | 35.2033
275.48 | 0.12913 | 35.5727
275.59 | 0.13032 | 35.9149
275.71 | 0.13106 | 36.1346
275.82 | 0.13175 | 36.3393
275.93 | 0.13296 | 36.6877
276.04 | 0.13366 | 36.8955
276.16 | 0.13458 | 37.1656
276.27 | 0.13479 | 37.2384
276.38 | 0.13542 | 37.4274
276.49 | 0.13514 | 37.3649
276.6 | 0.13592 | 37.5955
276.72 | 0.13514 | 37.3959
276.83 | 0.13492 | 37.3499
276.94 | 0.13507 | 37.4063
277.05 | 0.13461 | 37.2937
277.16 | 0.13486 | 37.3778
277.28 | 0.13444 | 37.2775
277.39 | 0.13367 | 37.0787
277.5 | 0.13207 | 36.6494
277.61 | 0.13203 | 36.6528
277.73 | 0.13089 | 36.3521
277.84 | 0.12973 | 36.0442
277.95 | 0.12853 | 35.7249
278.06 | 0.12776 | 35.5249
278.17 | 0.12639 | 35.1579
278.29 | 0.12562 | 34.9588
278.4 | 0.1236 | 34.4102
278.51 | 0.12239 | 34.0868
278.62 | 0.12085 | 33.6712
278.73 | 0.1192 | 33.2246
278.85 | 0.1178 | 32.8485
278.96 | 0.11569 | 32.2729
279.07 | 0.11416 | 31.8586
279.18 | 0.11238 | 31.3742
279.29 | 0.11037 | 30.8252
279.4 | 0.10903 | 30.463
279.52 | 0.10723 | 29.9729
279.63 | 0.10582 | 29.5904
279.74 | 0.10377 | 29.0286
279.85 | 0.10139 | 28.374
279.96 | 0.099917 | 27.9728
280.08 | 0.097639 | 27.3467
280.19 | 0.095841 | 26.8537
280.3 | 0.094178 | 26.3981
280.41 | 0.091876 | 25.7629
280.52 | 0.089917 | 25.2235
280.64 | 0.088224 | 24.7592
280.75 | 0.086742 | 24.3528
280.86 | 0.084764 | 23.8068
280.97 | 0.082634 | 23.2177
281.08 | 0.0812 | 22.8237
281.19 | 0.079557 | 22.3706
281.31 | 0.07681 | 21.6074
281.42 | 0.075487 | 21.2436
281.53 | 0.073705 | 20.7502
281.64 | 0.072067 | 20.2969
281.75 | 0.070591 | 19.889
281.86 | 0.068989 | 19.4452
281.98 | 0.067541 | 19.0452
282.09 | 0.065458 | 18.465
282.2 | 0.063986 | 18.0568
282.31 | 0.062572 | 17.6647
282.42 | 0.060517 | 17.0912
282.53 | 0.059364 | 16.7721
282.65 | 0.057907 | 16.3674
282.76 | 0.056489 | 15.9728
282.87 | 0.055074 | 15.5788
282.98 | 0.053491 | 15.1369
283.09 | 0.052285 | 14.8014
283.2 | 0.051008 | 14.4455
283.32 | 0.04953 | 14.0328
283.43 | 0.048251 | 13.6758
283.54 | 0.046577 | 13.2064
283.65 | 0.04584 | 13.0025
283.76 | 0.044677 | 12.6775
283.87 | 0.043424 | 12.3268
283.99 | 0.042715 | 12.1306
284.1 | 0.04132 | 11.739
284.21 | 0.040417 | 11.4869
284.32 | 0.039292 | 11.1715
284.43 | 0.03813 | 10.8453
284.54 | 0.037448 | 10.6555
284.65 | 0.036375 | 10.3541
284.77 | 0.035724 | 10.1731
284.88 | 0.034296 | 9.7702
284.99 | 0.033574 | 9.5683
285.1 | 0.032635 | 9.3042
285.21 | 0.031815 | 9.074
285.32 | 0.030598 | 8.7302
285.43 | 0.030192 | 8.6177
285.55 | 0.029133 | 8.3189
285.66 | 0.028774 | 8.2196
285.77 | 0.027561 | 7.8761
285.88 | 0.027527 | 7.8694
285.99 | 0.026574 | 7.5999
286.1 | 0.025747 | 7.3662
286.21 | 0.025313 | 7.2448
286.33 | 0.024611 | 7.0469
286.44 | 0.023831 | 6.8262
286.55 | 0.023347 | 6.6901
286.66 | 0.02307 | 6.6132
286.77 | 0.022369 | 6.4148
286.88 | 0.021779 | 6.248
286.99 | 0.021127 | 6.0632
287.1 | 0.020549 | 5.8996
287.22 | 0.020062 | 5.7622
287.33 | 0.019566 | 5.6219
287.44 | 0.019143 | 5.5025
287.55 | 0.018871 | 5.4264
287.66 | 0.018074 | 5.1992
287.77 | 0.01796 | 5.1683
287.88 | 0.017461 | 5.0267
287.99 | 0.016975 | 4.8886
288.11 | 0.016616 | 4.7872
288.22 | 0.0162 | 4.6692
288.33 | 0.015998 | 4.6127
288.44 | 0.015434 | 4.4518
288.55 | 0.015244 | 4.3987
288.66 | 0.014723 | 4.2499
288.77 | 0.014368 | 4.149
288.88 | 0.014237 | 4.1128
288.99 | 0.013859 | 4.0051
289.11 | 0.013431 | 3.883
289.22 | 0.013228 | 3.8258
289.33 | 0.012875 | 3.7251
289.44 | 0.012595 | 3.6455
289.55 | 0.012338 | 3.5725
289.66 | 0.012148 | 3.5188
289.77 | 0.011797 | 3.4184
289.88 | 0.011614 | 3.3667
289.99 | 0.011351 | 3.2917
290.11 | 0.011029 | 3.1996
290.22 | 0.01086 | 3.1518
290.33 | 0.010564 | 3.067
290.44 | 0.010365 | 3.0104
290.55 | 0.010238 | 2.9747
290.66 | 0.0099442 | 2.8904
290.77 | 0.0097586 | 2.8375
290.88 | 0.0093598 | 2.7226
290.99 | 0.0092438 | 2.6899
291.1 | 0.0090561 | 2.6362
291.21 | 0.0090089 | 2.6235
291.33 | 0.008655 | 2.5215
291.44 | 0.0084429 | 2.4606
291.55 | 0.0083759 | 2.442
291.66 | 0.0082421 | 2.4039
291.77 | 0.0081266 | 2.3711
291.88 | 0.0078635 | 2.2952
291.99 | 0.0077976 | 2.2768
292.1 | 0.0075574 | 2.2075
292.21 | 0.0075742 | 2.2133
292.32 | 0.0072425 | 2.1171
292.43 | 0.0071221 | 2.0827
292.55 | 0.0069432 | 2.0312
292.66 | 0.0068824 | 2.0142
292.77 | 0.0065337 | 1.9129
292.88 | 0.0065544 | 1.9197
292.99 | 0.0063671 | 1.8655
293.1 | 0.006263 | 1.8357
293.21 | 0.0061044 | 1.7899
293.32 | 0.0059503 | 1.7453
293.43 | 0.0060199 | 1.7664
293.54 | 0.0057976 | 1.7018
293.65 | 0.0057021 | 1.6744
293.76 | 0.0055945 | 1.6434
293.87 | 0.0055191 | 1.6219
293.99 | 0.005223 | 1.5355
294.1 | 0.0051928 | 1.5272
294.21 | 0.0050785 | 1.4941
294.32 | 0.0049532 | 1.4578
294.43 | 0.0048592 | 1.4307
294.54 | 0.0046522 | 1.3703
294.65 | 0.0046711 | 1.3763
294.76 | 0.004601 | 1.3562
294.87 | 0.0045667 | 1.3466
294.98 | 0.0044045 | 1.2992
295.09 | 0.0043432 | 1.2816
295.2 | 0.0041739 | 1.2321
295.31 | 0.004123 | 1.2176
295.42 | 0.0040294 | 1.1904
295.53 | 0.0039304 | 1.1616
295.64 | 0.0037792 | 1.1173
295.76 | 0.0037195 | 1.1001
295.87 | 0.0036196 | 1.0709
295.98 | 0.0035926 | 1.0633
296.09 | 0.0034469 | 1.0206
296.2 | 0.0033062 | 0.9793
296.31 | 0.0034769 | 1.0302
296.42 | 0.0033161 | 0.98296
296.53 | 0.0032003 | 0.94898
296.64 | 0.0031466 | 0.93341
296.75 | 0.0030417 | 0.90262
296.86 | 0.0030685 | 0.91091
296.97 | 0.0029982 | 0.89038
297.08 | 0.00291 | 0.8645
297.19 | 0.0028024 | 0.83285
297.3 | 0.002788 | 0.82887
297.41 | 0.0026906 | 0.80021
297.52 | 0.0026882 | 0.79979
297.63 | 0.0025601 | 0.76196
297.74 | 0.0025894 | 0.77097
297.85 | 0.0025093 | 0.7474
297.96 | 0.0024723 | 0.73665
298.07 | 0.0022967 | 0.68458
298.19 | 0.0022603 | 0.674
298.3 | 0.0021926 | 0.65405
298.41 | 0.0022478 | 0.67077
298.52 | 0.002198 | 0.65615
298.63 | 0.0021077 | 0.62942
298.74 | 0.0020406 | 0.60961
298.85 | 0.0020126 | 0.60147
298.96 | 0.0019168 | 0.57305
299.07 | 0.0019505 | 0.58334
299.18 | 0.0018606 | 0.55665
299.29 | 0.0018267 | 0.54671
299.4 | 0.0017583 | 0.52644
299.51 | 0.0017786 | 0.53271
299.62 | 0.0017278 | 0.51768
299.73 | 0.0016294 | 0.48838
299.84 | 0.0015848 | 0.47519
299.95 | 0.0016224 | 0.48664
300.06 | 0.0015669 | 0.47016
| 13.7417 | 3818.6258
277.8856