## Data Tracking
Data found in L:\Priv\Cin\NERL\PattyCarol\method 200.8\All work Prior to 11-8-17\Figure 1\7-19-16 abundance sensitivity data 2.5 mL He.xlsx
Raw data collected on 7700 and is saved at: D:\Agilent\ICPMH\1\DATA\7-19-16 narrow peak 2.b
Raw data collected on 7700 and is saved at: D:\Agilent\ICPMH\1\DATA\7-19-16 standard res 2.b
The Acquisition method is at L:\Lab\Lablan\iAs Group\7700_551\Method 200.8\7-19-16 narrow peak 2.pdf
The Acquisition method is at L:\Lab\Lablan\iAs Group\7700_551\Method 200.8\7-19-16 standard res 2.pdf

## Sheet2
| 3BW34A | 3BW17 (C-J) | [7 empty] | 3BW34C | 3BW34B
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1000 ppb | Solution 2 | Solution 1
| 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg
| CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS
| 0 | 1 | 5 | 10 | 20 | 40 | 60 | 80 | 100 | 500 | 1000 | Solution 2 | Solution 1
71.5 (0.4 AMU) | 0 | 0 | 0 | 0 | 0 | 6.6667 | 6.67 | 2.22 | 6.6667 | 45.5667 | 44124.13 | 1771.5233 | 2.22 | Narrow mode | Narrow resolution, 20 pts, 10 replicates, 2.5 mL/min He
72 (0.4 AMU) | 4.4433 | 1671.4867 | 7977.3867 | 15873.82 | 31056.31 | 62847.71 | 94782.6933 | 126766.06 | 157794.2233 | 805966.2467 | 1703688.2167 | 4067.93 | 353.3967 | Narrow mode
143 (0.4 AMU) | 2.2233 | 4.4467 | 2.22 | 0 | 2.2233 | 23.3733 | 8.8933 | 3.3333 | 13.3467 | 4.4467 | 6.67 | 187115.3633 | 40.0033 | Narrow mode
71.5 (0.8 AMU) | 0 | 83.35 | 336.73 | 700.1733 | 1401.5867 | 2701.2667 | 3863.3867 | 4790.0367 | 6295.5033 | 32489.12 | 66285.3833 | 1982.7033 | 4.4467 | Standard mode | Standard resolution, 20 pts, 10 replicates, 2.5 mL/min He
72 (0.8 AMU) | 12.2233 | 1918.2233 | 9642.4567 | 19455.8467 | 38345.2467 | 76382.98 | 114229.62 | 151214.03 | 189446.9667 | 962612.1033 | 2039370.42 | 4561.5033 | 510.0967 | Standard mode
143 (0.8 AMU) | 4.4433 | 12.2267 | 6.6667 | 3.3333 | 3.3333 | 0 | 4.4433 | 7.78 | 11.1133 | 2.22 | 11.1133 | 226735.79 | 66.6767 | Standard mode
71.5 (0.4amu)/72 (0.4amu) | 0 | 0 | 0 | 0 | 0 | 0.00010608 | 7.0371e-05 | 1.7513e-05 | 4.2249e-05 | 5.6537e-05 | 0.025899 | 0.43549
71.5 (0.8amu)/72 (0.8amu) | 0 | 0.043452 | 0.034922 | 0.035988 | 0.036552 | 0.035365 | 0.033821 | 0.031677 | 0.033231 | 0.033751 | 0.032503 | 0.43466
71.5/143 (Solution 2) (0.4 AMU) | 0 | 0 | 0 | 0 | 0 | 3.5629e-05 | 3.5646e-05 | 1.1864e-05 | 3.5629e-05 | 0.00024352 | 0.23581 | 0.0094675 |  | Narrow mode
71.5/143 (Solution 2) (0.8 AMU) | 0 | 0.00036761 | 0.0014851 | 0.0030881 | 0.0061816 | 0.011914 | 0.017039 | 0.021126 | 0.027766 | 0.14329 | 0.29235 | 0.0087446 |  | Standard mode
cps 72 (.8AMU)/cps 72(.4 AMU) | 2.7509 | 1.1476 | 1.2087 | 1.2257 | 1.2347 | 1.2154 | 1.2052 | 1.1929 | 1.2006 | 1.1944 | 1.197
Figure S2B
Percent Error of 71.5/143 as a function of Ge Concentration in both Narrow and Standard Resolution
Theoretical 71.5/143 (0.4 AMU) | 0.0094675
Experimental 71.5/143 (0.4 AMU) | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0095032 | 0.0095032 | 0.0094794 | 0.0095032 | 0.0097111
Theoretical 71.5/143 (0.8 AMU) | 0.0087446
Experimental 71.5/143 (0.8 AMU) | 0.0087446 | 0.0091122 | 0.01023 | 0.011833 | 0.014926 | 0.020658 | 0.025784 | 0.029871 | 0.03651 | 0.15204
Percent error of 71.5/143 ratio
ppb Ge | 0 | 1 | 5 | 10 | 20 | 40 | 60 | 80 | 100 | 500
cps at 72 (0.4 AMU) | 4.4433 | 1671.4867 | 7977.3867 | 15873.82 | 31056.31 | 62847.71 | 94782.6933 | 126766.06 | 157794.2233 | 805966.2467
cps x 104 at 72 (0.4 AMU) | 0.00044433 | 0.16715 | 0.79774 | 1.5874 | 3.1056 | 6.2848 | 9.4783 | 12.6766 | 15.7794 | 80.5966
Narrow Resolution (0.4 AMU) | 0 | 0 | 0 | 0 | 0 | 0.37632 | 0.37651 | 0.12532 | 0.37632 | 2.5722
Standard Resolution (0.8 AMU) | 0 | 4.2039 | 16.9834 | 35.3141 | 70.6907 | 136.2416 | 194.8545 | 241.5912 | 317.5212 | 1638.6274
(71.5 std + 71.5 Ge)]/71.5 std
| 0 | 1 | 5 | 10 | 20 | 40 | 60 | 80 | 100 | 500
Narrow Resolution (0.4 AMU) | 1 | 1 | 1 | 1 | 1 | 1.0038 | 1.0038 | 1.0013 | 1.0038 | 1.0257
Standard Resolution (0.8 AMU) | 1 | 1.042 | 1.1698 | 1.3531 | 1.7069 | 2.3624 | 2.9485 | 3.4159 | 4.1752 | 17.3863
Percent error of 71.5 signal | 0 | 1 | 5 | 10 | 20 | 40 | 60 | 80 | 100 | 500
Narrow Resolution (0.4 AMU) | 0 | 0 | 0 | 0 | 0 | 0.37632 | 0.37651 | 0.12532 | 0.37632 | 2.5722
Standard Resolution (0.8 AMU) | 0 | 4.2039 | 16.9834 | 35.3141 | 70.6907 | 136.2416 | 194.8545 | 241.5912 | 317.5212 | 1638.6274
| 3BW34A | 3BW17 (C-J) | [7 empty] | 3BW34C | 3BW34B
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1000 ppb | Solution 2 | Solution 1
| 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg
| CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS
| 0 | 1 | 5 | 10 | 20 | 40 | 60 | 80 | 100 | 500 | 1000 | Solution 2 | Solution 1
71.5 (0.8 AMU) | 1.11 | 2.2233 | 4.4467 | 11.11 | 12.2267 | 31.1167 | 36.6733 | 54.4567 | 75.5667 | 34129.2133 | 39709.7467 | 4.4467 | 2077.17 | Standard mode
72 (0.8 AMU) | 48.9 | 1251.3633 | 5892.1467 | 11977.8767 | 23622.5667 | 47431.12 | 70619.9733 | 92494.0033 | 116712.1567 | 577966.3233 | 1238239.7233 | 1313.62 | 1372.53 | Standard mode
143 (0.8 AMU) | 26.6733 | 3.3333 | 2.2233 | 7.78 | 4.4467 | 3.3333 | 3.3333 | 2.2233 | 1.11 | 0 | 14.45 | 246.72 | 613481.7167 | Standard mode
ppb Ge | 0 | 1 | 5 | 10 | 20 | 40 | 60 | 80 | 100 | 500
cps x 104 at 72 (0.8 AMU) | 0.0012223 | 0.19182 | 0.96425 | 1.9456 | 3.8345 | 7.6383 | 11.423 | 15.1214 | 18.9447 | 96.2612
cps x 104 at 72 (0.4 AMU) | 0.00044433 | 0.16715 | 0.79774 | 1.5874 | 3.1056 | 6.2848 | 9.4783 | 12.6766 | 15.7794 | 80.5966
Theoretical 71.5/143 (0.4 AMU) | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675
Experimental 71.5/143 (0.4 AMU) | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0094675 | 0.0095032 | 0.0095032 | 0.0094794 | 0.0095032 | 0.0097111
Theoretical 71.5/143 (0.8 AMU) | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446
Experimental 71.5/143 (0.8 AMU) | 0.0087446 | 0.0091122 | 0.01023 | 0.011833 | 0.014926 | 0.020658 | 0.025784 | 0.029871 | 0.03651 | 0.15204
71.5/143 (0.8 AMU) without Ge | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446 | 0.0087446
| 3BW34A | 3BW17 (C-J) | [7 empty] | 3BW34C | 3BW34B
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1000 ppb | Solution 2 | Solution 1
| 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg | 3 point Avg
| CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS
| 0 | 1 | 5 | 10 | 20 | 40 | 60 | 80 | 100 | 500 | 1000 | Solution 2 | Solution 1
71.5 (0.4 AMU) | 0 | 0 | 0 | 0 | 0 | 6.6667 | 6.67 | 2.22 | 6.6667 | 45.5667 | 44124.13 | 1771.5233 | 2.22 | Narrow mode | Narrow resolution, 20 pts, 10 replicates, 2.5 mL/min He
72 (0.4 AMU) | 4.4433 | 1671.4867 | 7977.3867 | 15873.82 | 31056.31 | 62847.71 | 94782.6933 | 126766.06 | 157794.2233 | 805966.2467 | 1703688.2167 | 4067.93 | 353.3967 | Narrow mode
143 (0.4 AMU) | 2.2233 | 4.4467 | 2.22 | 0 | 2.2233 | 23.3733 | 8.8933 | 3.3333 | 13.3467 | 4.4467 | 6.67 | 187115.3633 | 40.0033 | Narrow mode
71.5 (0.8 AMU) | 0 | 83.35 | 336.73 | 700.1733 | 1401.5867 | 2701.2667 | 3863.3867 | 4790.0367 | 6295.5033 | 32489.12 | 66285.3833 | 1982.7033 | 4.4467 | Standard mode | Standard resolution, 20 pts, 10 replicates, 2.5 mL/min He
72 (0.8 AMU) | 12.2233 | 1918.2233 | 9642.4567 | 19455.8467 | 38345.2467 | 76382.98 | 114229.62 | 151214.03 | 189446.9667 | 962612.1033 | 2039370.42 | 4561.5033 | 510.0967 | Standard mode
143 (0.8 AMU) | 4.4433 | 12.2267 | 6.6667 | 3.3333 | 3.3333 | 0 | 4.4433 | 7.78 | 11.1133 | 2.22 | 11.1133 | 226735.79 | 66.6767 | Standard mode

## Narrow Peak
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Narrow resolution, 20 pts, 10 replicates, 2.5 mL/min He
Mass Hunter file Agilent 7700 ICP-MS: 7-19-16 narrow peak 2
| 3BW34A | 3BW17 (C-J) | [7 empty] | 3BW34C | 3BW34B
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppb | Solution 2 | Solution 1
Mass | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS
64.55 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0
64.6 | 3.33 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 0
64.65 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 0 | 6.67 | 0 | 3.33 | 0 | 0 | 0
64.7 | 3.33 | 0 | 0 | 0 | 16.67 | 3.33 | 0 | 3.33 | 0 | 6.67 | 0 | 3.33 | 10
64.75 | 23.34 | 66.68 | 76.68 | 20 | 123.35 | 10 | 56.68 | 56.68 | 23.34 | 40.01 | 96.69 | 73.34 | 110.02
64.8 | 126.69 | 253.37 | 203.37 | 180.03 | 630.13 | 150.03 | 313.39 | 200.03 | 190.04 | 390.07 | 500.09 | 520.09 | 580.12
64.85 | 290.05 | 463.42 | 326.73 | 266.71 | 920.18 | 313.39 | 326.72 | 340.06 | 310.05 | 556.77 | 573.44 | 926.85 | 910.17
64.9 | 386.73 | 486.76 | 316.72 | 316.72 | 956.85 | 356.73 | 483.42 | 316.72 | 390.07 | 636.78 | 860.16 | 1100.22 | 1070.22
64.95 | 333.39 | 536.77 | 376.74 | 283.38 | 1050.21 | 376.73 | 456.75 | 453.42 | 400.06 | 780.14 | 943.51 | 1146.91 | 1210.25
65 | 403.41 | 453.41 | 443.41 | 270.04 | 1260.25 | 253.38 | 510.09 | 423.4 | 436.74 | 776.8 | 990.19 | 1260.28 | 1296.94
65.05 | 393.41 | 500.09 | 316.72 | 320.06 | 1043.54 | 353.4 | 540.1 | 390.07 | 433.41 | 590.11 | 963.53 | 1036.88 | 1270.27
65.1 | 353.39 | 486.76 | 333.39 | 233.37 | 956.88 | 310.06 | 430.08 | 370.07 | 393.41 | 636.78 | 760.15 | 1020.2 | 1006.86
65.15 | 6.67 | 30 | 20 | 13.34 | 20 | 6.67 | 20 | 3.33 | 3.33 | 20 | 10 | 50.01 | 66.68
65.2 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
65.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 0
65.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 6.67 | 13.34 | 0 | 0
65.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 30.01 | 0 | 0
65.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 13.33 | 0 | 0
65.45 | 0 | 0 | 0 | 0 | 0 | 13.34 | 6.67 | 0 | 13.34 | 6.67 | 63937.92 | 6.67 | 0
65.5 | 0 | 6.67 | 0 | 6.67 | 0 | 0 | 10 | 10 | 0 | 56.67 | 44377.33 | 0 | 0
65.55 | 0 | 0 | 6.67 | 0 | 0 | 6.67 | 23.34 | 36.67 | 16.67 | 90.01 | 44015.47 | 0 | 0
65.6 | 0 | 0 | 6.67 | 6.67 | 20 | 56.67 | 46.67 | 73.34 | 43.34 | 423.41 | 45054.22 | 0 | 0
65.65 | 0 | 0 | 23.34 | 20 | 46.67 | 86.68 | 173.36 | 250.04 | 323.39 | 1567.03 | 43554.28 | 0 | 0
65.7 | 0 | 13.34 | 70.01 | 116.69 | 176.7 | 573.44 | 613.45 | 930.2 | 1186.93 | 5772.7 | 46357.84 | 3.33 | 10
65.75 | 0 | 120.02 | 723.48 | 1443.66 | 2777.61 | 5922.99 | 8949 | 12042.6 | 14643.5 | 68293.87 | 79609.86 | 83.35 | 43.34
65.8 | 23.34 | 1466.99 | 6875.98 | 13935.44 | 27924.07 | 55245.55 | 82097.23 | 107563.14 | 134693.92 | 622881.51 | 1035416.79 | 686.78 | 410.07
65.85 | 60.01 | 2070.48 | 9487.86 | 18708.03 | 36738.11 | 73167.69 | 110090.01 | 146702.33 | 182526.92 | 829653.14 | 1732793.14 | 1176.89 | 590.1
65.9 | 76.68 | 2230.52 | 10108.34 | 20303.79 | 40156.46 | 82561.37 | 123875.39 | 161472.06 | 201467.34 | 923760.83 | 1967711.5 | 1293.6 | 646.79
65.95 | 40 | 2500.61 | 11289.35 | 22677.42 | 44496.23 | 86608.88 | 130742.14 | 174539.89 | 216429.81 | 990443.9 | 2166091.82 | 1400.28 | 773.48
66 | 26.67 | 2677.35 | 11019.05 | 22560.77 | 45687.36 | 89324.17 | 132742.33 | 177197.88 | 221546.32 | 1002740.66 | 2249373.72 | 1416.97 | 780.15
66.05 | 36.67 | 2380.58 | 10538.81 | 21642.53 | 42945.59 | 85584.41 | 127914.21 | 167143.67 | 211468.33 | 962231.38 | 2134928.54 | 1310.28 | 690.13
66.1 | 6.67 | 2203.84 | 10088.49 | 18260.54 | 37112.47 | 72093.29 | 106929.85 | 140644.78 | 176614.92 | 798755.68 | 2007738.35 | 1073.57 | 516.76
66.15 | 0 | 63.35 | 190.03 | 323.39 | 620.12 | 1243.61 | 1587.08 | 1977.21 | 2694.15 | 11622.07 | 255605.56 | 10 | 43.34
66.2 | 0 | 0 | 6.67 | 6.67 | 13.33 | 20 | 50.01 | 50.01 | 66.68 | 340.06 | 44343.65 | 0 | 0
66.25 | 0 | 0 | 6.67 | 0 | 0 | 3.33 | 13.34 | 6.67 | 36.67 | 43.34 | 42544.21 | 0 | 0
66.3 | 0 | 0 | 0 | 3.33 | 0 | 0 | 6.67 | 3.33 | 0 | 6.67 | 43413.03 | 0 | 0
66.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 43473.93 | 0 | 0
66.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 45513.3 | 6.67 | 0
66.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 3.33 | 0
66.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 10 | 20 | 6.67 | 0
66.55 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 6.67 | 0 | 0 | 26.67 | 30 | 0 | 0
66.6 | 0 | 0 | 0 | 0 | 0 | 6.67 | 13.34 | 10 | 13.34 | 63.34 | 110.02 | 6.67 | 0
66.65 | 0 | 3.33 | 10 | 0 | 20 | 30 | 50.01 | 60.01 | 93.35 | 376.74 | 850.16 | 0 | 0
66.7 | 0 | 6.67 | 10 | 30 | 66.68 | 166.69 | 313.39 | 356.73 | 483.44 | 2394.09 | 4939.29 | 3.33 | 6.67
66.75 | 3.33 | 50.01 | 216.71 | 490.09 | 1176.92 | 2420.69 | 3324.31 | 4638.3 | 5388.7 | 25506.74 | 53256.35 | 56.68 | 20
66.8 | 0 | 253.39 | 1020.21 | 2090.48 | 3914.5 | 8336.98 | 12293.83 | 16652.25 | 20214.06 | 94113.71 | 197714.72 | 313.39 | 56.68
66.85 | 3.33 | 273.38 | 1240.24 | 2810.76 | 5585.43 | 10852.44 | 15533.9 | 21355.81 | 26801.73 | 118568.46 | 248614.61 | 363.39 | 66.68
66.9 | 20 | 296.72 | 1607.01 | 3070.79 | 6058.9 | 12136.93 | 17876.89 | 24664.4 | 30820.77 | 139608.56 | 293453.42 | 536.76 | 60.01
66.95 | 3.33 | 440.09 | 1717.03 | 3490.99 | 6765.92 | 13987.23 | 20584 | 28050.61 | 33649.14 | 156495.95 | 331414.65 | 610.11 | 106.68
67 | 6.67 | 310.06 | 1777.06 | 3337.55 | 6909.44 | 13528.29 | 20053.44 | 26875.04 | 33195.71 | 153143.57 | 323290.12 | 723.48 | 93.35 | Zn (4.1)
67.05 | 10 | 430.08 | 1623.67 | 3174.2 | 6265.73 | 13258.15 | 19559.13 | 26467.88 | 32644.06 | 149332.53 | 318184.62 | 673.46 | 103.35
67.1 | 6.67 | 313.39 | 1390.27 | 3057.44 | 6048.88 | 12350.39 | 16875.77 | 22774.55 | 28628.48 | 132012.03 | 276017.23 | 526.77 | 140.03
67.15 | 0 | 6.67 | 63.34 | 83.35 | 156.69 | 336.72 | 556.77 | 643.45 | 876.85 | 3154.32 | 6409.61 | 233.38 | 16.67
67.2 | 0 | 0 | 6.67 | 0 | 6.67 | 0 | 6.67 | 3.33 | 30 | 66.68 | 116.69 | 6.67 | 0
67.25 | 0 | 3.33 | 3.33 | 0 | 0 | 3.33 | 10 | 0 | 0 | 16.67 | 10 | 103.35 | 0
67.3 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 30.02 | 10 | 390.07 | 0
67.35 | 0 | 10 | 3.33 | 0 | 6.67 | 0 | 3.33 | 0 | 0 | 10 | 0 | 700.13 | 0
67.4 | 0 | 0 | 43.39 | 0 | 6.67 | 0 | 0 | 0 | 241.84 | 3.33 | 13.33 | 1016.88 | 0
67.45 | 0 | 0 | 0 | 16.67 | 0 | 0 | 6.67 | 6.67 | 3.33 | 3.33 | 53507.02 | 973.53 | 6.67
67.5 | 6.67 | 0 | 0 | 10 | 0 | 6.67 | 0 | 3.33 | 33.36 | 60.01 | 38737.31 | 1063.55 | 6.67 | Ba-135 2+
67.55 | 0 | 0 | 3.33 | 10 | 6.67 | 10 | 3.33 | 16.67 | 3.33 | 143.39 | 38771.74 | 1230.25 | 0
67.6 | 33.37 | 3.33 | 6.67 | 23.34 | 20 | 16.67 | 33.34 | 20 | 26.67 | 200.04 | 39654.62 | 1196.92 | 0
67.65 | 0 | 13.34 | 6.67 | 13.34 | 20 | 36.67 | 76.68 | 73.35 | 103.35 | 523.43 | 39197.07 | 893.51 | 0
67.7 | 0 | 13.33 | 46.67 | 96.68 | 246.71 | 646.79 | 940.21 | 1276.97 | 1413.66 | 7183.91 | 44474.19 | 13.34 | 0
67.75 | 0 | 130.02 | 486.75 | 1053.56 | 2237.32 | 4661.93 | 6533.03 | 9496.18 | 11104.98 | 52838.82 | 84721.93 | 70.01 | 33.34
67.8 | 30.01 | 893.51 | 3624.38 | 6992.79 | 14359.66 | 29137.39 | 44061.35 | 58718.02 | 71314.79 | 331755.77 | 535447.76 | 763.49 | 276.72
67.85 | 13.34 | 1393.64 | 6228.91 | 12467.32 | 24440.5 | 48292.94 | 71791.99 | 96363.69 | 119369.43 | 546308.42 | 1118380.64 | 1670.38 | 283.38
67.9 | 50.04 | 1663.71 | 6715.91 | 13942.14 | 28231.41 | 56512.03 | 83231.38 | 108614.53 | 135952.54 | 628287.73 | 1324858.96 | 2033.82 | 476.76
67.95 | 80.03 | 1813.74 | 8050.1 | 15810.85 | 31230.82 | 61442.29 | 92311.18 | 122719.29 | 151743 | 699328.77 | 1505980.55 | 2297.2 | 516.76
68 | 13.33 | 1790.37 | 8467.04 | 16640.3 | 33857.08 | 66531.6 | 98920.54 | 130787.51 | 164781 | 757142.98 | 1626806.82 | 2507.29 | 476.75 | Zn (18.8)
68.05 | 43.34 | 1753.72 | 8440.48 | 16411.65 | 32240.23 | 66067.42 | 98175.72 | 130502.73 | 161776.2 | 737512.2 | 1616915.83 | 2570.69 | 490.09
68.1 | 13.34 | 1793.74 | 7676.54 | 15350.61 | 30258.99 | 59853.37 | 89813.92 | 119061.84 | 149202.18 | 675713.33 | 1544268 | 2640.71 | 586.79
68.15 | 16.67 | 436.75 | 2037.19 | 3977.89 | 7476.61 | 14787.37 | 21103.29 | 26517.05 | 34131.55 | 141931.49 | 840422.68 | 1366.96 | 86.68
68.2 | 6.67 | 3.33 | 10 | 20 | 26.67 | 76.68 | 110.02 | 133.36 | 196.7 | 773.47 | 42456.97 | 16.67 | 0
68.25 | 0 | 0 | 3.33 | 0 | 3.33 | 3.33 | 0 | 20 | 26.67 | 63.34 | 43872.66 | 20 | 0
68.3 | 0 | 6.67 | 0 | 0 | 10 | 6.67 | 3.33 | 0 | 0 | 13.34 | 42949.46 | 606.78 | 0
68.35 | 13.34 | 0 | 269 | 0 | 0 | 3.33 | 16.67 | 13.34 | 6.67 | 40.01 | 43390.99 | 1343.63 | 0
68.4 | 0 | 3.33 | 96.97 | 10 | 46.73 | 3.33 | 0 | 16.67 | 0 | 10 | 43861.15 | 1853.76 | 0
68.45 | 3.33 | 6.67 | 13.34 | 20.01 | 16.67 | 0 | 16.67 | 6.67 | 10 | 63377.36 | 44396.06 | 2030.45 | 3.33
68.5 | 0 | 3.33 | 3.33 | 0 | 0 | 6.67 | 10 | 6.67 | 13.34 | 41253.7 | 41594.83 | 2213.85 | 0 | Ba-137 2+
68.55 | 0 | 6.67 | 0 | 0 | 20 | 6.67 | 40.01 | 46.67 | 36.67 | 44841.35 | 42339.48 | 2073.85 | 0
68.6 | 3.33 | 3.33 | 23.34 | 20 | 36.67 | 76.75 | 100.02 | 83.35 | 110.02 | 47217.47 | 44615.66 | 1967.13 | 0
68.65 | 0 | 3.33 | 26.67 | 40.01 | 70.01 | 180.04 | 200.03 | 296.72 | 477.16 | 45688.29 | 45819.2 | 820.16 | 0
68.7 | 0 | 16.67 | 86.68 | 236.71 | 350.06 | 796.83 | 1453.69 | 1997.24 | 2287.28 | 48383.01 | 48372.69 | 40.01 | 0
68.75 | 0 | 653.45 | 2990.86 | 6015.73 | 12314.83 | 25860.73 | 39260.93 | 52837.2 | 63726.14 | 133529.77 | 242548.98 | 636.79 | 46.68
68.8 | 33.36 | 2485.16 | 11676.53 | 22791.22 | 46075.71 | 92283.05 | 138319.29 | 184631.8 | 229228.22 | 1006987.46 | 2012595.44 | 5902.19 | 76.68
68.85 | 6.67 | 3124.26 | 14940.05 | 30278.76 | 60579.78 | 120065.4 | 180787.88 | 243567.16 | 304626.58 | 1493971.21 | 3009464.18 | 9868.18 | 153.36
68.9 | 6.67 | 3287.77 | 17035.82 | 34051.07 | 67222.15 | 135733.85 | 205476.12 | 273324.36 | 341450.94 | 1739500.05 | 3450356.33 | 11993.62 | 150.03
68.95 | 66.73 | 3831.16 | 18448.05 | 37162.81 | 72601.65 | 147646.49 | 218839.58 | 289521.17 | 369493.36 | 1906373.08 | 3824213.99 | 13608.47 | 216.7
69 | 0 | 3691.11 | 18774.9 | 37145.34 | 74923.1 | 149688.16 | 224166.64 | 300072.36 | 375463.7 | 2000195.49 | 3977346 | 14519.54 | 240.04 | Ga (60.11)
69.05 | 6.67 | 3892.33 | 17933.49 | 35678.2 | 72177.82 | 142763.45 | 217902.44 | 286286.48 | 358018.54 | 1912932.57 | 3852048.84 | 14472.53 | 180.03
69.1 | 93.5 | 3458.34 | 15644.25 | 30463.03 | 61822.12 | 121816.27 | 182615.36 | 240604.99 | 306199.65 | 1775313.32 | 3556137.92 | 13154.73 | 133.36
69.15 | 6.67 | 100.05 | 336.73 | 783.5 | 1286.94 | 2287.32 | 3087.65 | 3854.77 | 5475.86 | 271906.91 | 495590.76 | 5191.85 | 6.67
69.2 | 0 | 0 | 13.34 | 0 | 50.01 | 36.67 | 60.01 | 86.68 | 150.03 | 47239.34 | 44764.6 | 13.34 | 0
69.25 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 16.67 | 6.67 | 10 | 45671.49 | 41675.74 | 3.33 | 0
69.3 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 48041 | 43608.06 | 0 | 0
69.35 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 47500.7 | 43664.49 | 6.67 | 0
69.4 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 6.67 | 45980.42 | 44118.74 | 6.67 | 0
69.45 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 3.33 | 13.34 | 50.01 | 20 | 3.33
69.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 13.33 | 30 | 6.67 | 0
69.55 | 0 | 0 | 0 | 3.33 | 0 | 0 | 13.34 | 6.67 | 16.67 | 53.34 | 123.35 | 0 | 3.33
69.6 | 0 | 0 | 3.33 | 13.34 | 6.67 | 0 | 26.67 | 13.34 | 36.67 | 186.7 | 373.4 | 6.67 | 0
69.65 | 6.67 | 0 | 16.67 | 26.67 | 56.68 | 80.01 | 166.7 | 196.7 | 276.71 | 1410.32 | 2594.07 | 16.67 | 0
69.7 | 0 | 10 | 50.01 | 133.36 | 346.74 | 656.8 | 893.52 | 1156.94 | 1380.32 | 8188.47 | 15491.5 | 20 | 0
69.75 | 0 | 160.03 | 983.54 | 1693.74 | 3997.96 | 7670.24 | 11493.82 | 16663.68 | 19859.33 | 104367.38 | 213590.08 | 233.38 | 10
69.8 | 3.33 | 636.78 | 3264.25 | 6669.25 | 13204.76 | 25712.99 | 39258.65 | 52067.15 | 64632.18 | 330329.4 | 667344.16 | 1823.76 | 30
69.85 | 0 | 723.46 | 4001.18 | 8463.88 | 17255.98 | 33015.29 | 49821.84 | 66074.46 | 83939.41 | 427042.87 | 858514.8 | 2774.04 | 26.67
69.9 | 3.33 | 916.83 | 5171.84 | 9968.3 | 19619.68 | 40759.04 | 60981.72 | 79416.35 | 100325.75 | 506517.85 | 1031346.84 | 3507.64 | 23.34
69.95 | 16.67 | 1186.93 | 5451.87 | 11209.29 | 23201.71 | 45850.9 | 68152.89 | 89437.05 | 113926.05 | 571401.33 | 1180696.71 | 4194.67 | 40.01
70 | 0 | 1306.92 | 5705.39 | 11506.36 | 23098.31 | 44891.87 | 67710.93 | 89703.43 | 112943.32 | 572256.05 | 1179020.51 | 4581.5 | 46.67
70.05 | 0 | 1166.9 | 5498.52 | 11199.35 | 22490.79 | 43704.88 | 65541.89 | 86615.26 | 109640.05 | 554806.3 | 1138623.32 | 4648.26 | 43.34
70.1 | 13.34 | 1046.88 | 4818.29 | 9374.5 | 19269 | 38522.26 | 57552.03 | 76210.2 | 96550.8 | 480604.73 | 993414.33 | 4554.88 | 46.67
70.15 | 6.67 | 26.67 | 193.36 | 323.39 | 610.13 | 1053.56 | 1617.08 | 1817.14 | 2484.08 | 11154.28 | 50511.31 | 1867.12 | 0
70.2 | 0 | 0 | 0 | 3.33 | 20 | 16.67 | 10 | 30.01 | 63.34 | 220.04 | 9025.17 | 10 | 0
70.25 | 0 | 0 | 0 | 0 | 0 | 10 | 10 | 10 | 3.33 | 23.34 | 8805.59 | 3.33 | 0
70.3 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 7808.69 | 0 | 0
70.35 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 8717 | 6.67 | 0
70.4 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 6.67 | 0 | 20 | 8199.19 | 3.33 | 0
70.45 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 10 | 6.67 | 5356.32 | 64247.15 | 13.34 | 0
70.5 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 3.33 | 4605.86 | 40960.13 | 23.34 | 0
70.55 | 0 | 3.33 | 3.33 | 3.33 | 0 | 6.67 | 10 | 13.34 | 13.34 | 4636.84 | 44139.56 | 16.67 | 0
70.6 | 0 | 0 | 0 | 0 | 0 | 30 | 40.01 | 53.34 | 40.01 | 4902.62 | 42863.54 | 0 | 0
70.65 | 0 | 3.33 | 10 | 26.67 | 76.68 | 120.02 | 166.7 | 233.38 | 270.05 | 5869.4 | 45341.93 | 13.33 | 0
70.7 | 0 | 20 | 60.01 | 230.05 | 426.75 | 770.15 | 1243.63 | 1817.19 | 2137.36 | 14161.81 | 50538.07 | 16.67 | 0
70.75 | 0 | 313.39 | 1410.3 | 2670.74 | 5852.6 | 11564.49 | 18077.84 | 24103.72 | 29166.71 | 153155.1 | 135270.95 | 116.69 | 23.34
70.8 | 0 | 1163.58 | 5448.66 | 11289.62 | 22367.61 | 46499.46 | 67440.82 | 91510.14 | 113717.59 | 577119.53 | 884322.52 | 803.49 | 26.67
70.85 | 0 | 1943.8 | 9814.89 | 19532.62 | 39204.75 | 79288.72 | 118451.29 | 157293.17 | 196981.26 | 1016192.57 | 1910532.96 | 2760.75 | 70.01
70.9 | 0 | 2440.62 | 11616.49 | 23893.35 | 48085.93 | 95835.14 | 145662.51 | 194853.9 | 243271.84 | 1249208.61 | 2460931.07 | 3721.1 | 113.35
70.95 | 6.67 | 2767.35 | 13838.62 | 27643.43 | 54632.46 | 110712.69 | 165066.31 | 219251.35 | 274956.5 | 1449714.75 | 2884120.06 | 4659.38 | 100.02
71 | 0 | 3184.22 | 14345.98 | 28815.92 | 56666.45 | 115811.39 | 172096.38 | 225017 | 285707.88 | 1527172.6 | 3036044.52 | 4961.67 | 156.69
71.05 | 0 | 2770.74 | 13932.12 | 27262.62 | 56127.43 | 111469.21 | 164122.95 | 219656.12 | 279823.75 | 1472273.37 | 2937448.76 | 4601.49 | 136.69
71.1 | 0 | 2567.29 | 13264.85 | 25869.85 | 50794.99 | 101323.17 | 150672.13 | 202890.97 | 258415.18 | 1389365.41 | 2787909.93 | 4251.39 | 136.69
71.15 | 0 | 630.11 | 3294.26 | 5995.71 | 11273.39 | 21528.1 | 31205.63 | 39299.96 | 50448.28 | 740963.69 | 1418337.95 | 3694.47 | 20
71.2 | 0 | 10 | 16.67 | 10 | 50.01 | 116.68 | 133.36 | 170.03 | 260.05 | 47755.88 | 43388.56 | 76.68 | 0
71.25 | 0 | 0 | 0 | 0 | 0 | 13.34 | 6.67 | 0 | 13.33 | 46612.23 | 45806.67 | 26.67 | 0
71.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 46386.83 | 42739.15 | 326.72 | 0
71.35 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 45081.35 | 43257.75 | 1086.87 | 0
71.4 | 0 | 0 | 0 | 6.67 | 0 | 10 | 0 | 0 | 0 | 42803.86 | 43551.06 | 1380.3 | 0
71.45 | 0 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 3.33 | 0 | 13.34 | 45589.09 | 1577.01 | 3.33
71.5 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 10 | 33.34 | 43049.46 | 1763.74 | 3.33
71.55 | 0 | 0 | 0 | 0 | 0 | 10 | 13.34 | 3.33 | 10 | 90.02 | 43733.84 | 1973.82 | 0
71.6 | 0 | 0 | 0 | 3.33 | 13.34 | 10 | 26.67 | 33.34 | 43.34 | 246.71 | 42682 | 1647.04 | 0
71.65 | 0 | 0 | 20 | 3.33 | 10 | 73.35 | 106.69 | 100.02 | 140.02 | 786.82 | 44740.81 | 1406.97 | 0
71.7 | 0 | 0 | 76.68 | 60.01 | 220.04 | 346.73 | 620.13 | 870.19 | 800.15 | 5025.34 | 45502.18 | 23.34 | 0
71.75 | 0 | 60.01 | 300.06 | 590.13 | 1387 | 2560.82 | 3938.21 | 5559.28 | 6246.58 | 36153.64 | 65501.34 | 36.67 | 10
71.8 | 10 | 640.12 | 2904.08 | 6082.26 | 12160.39 | 24564.92 | 36586.6 | 50162.57 | 62703.78 | 322730.24 | 411615.69 | 563.44 | 60.01
71.85 | 6.67 | 1393.63 | 6415.86 | 12697.51 | 25345.6 | 50550.56 | 74525.62 | 100836.33 | 125539.67 | 640052.81 | 1240228.07 | 2910.79 | 213.37
71.9 | 3.33 | 1627.03 | 7029.5 | 14339.27 | 29269.59 | 57336.4 | 86746.64 | 114860.99 | 144527.06 | 735411.38 | 1501662.57 | 3244.28 | 310.05
71.95 | 0 | 1550.35 | 7819.99 | 15632.52 | 30933.92 | 61788.63 | 93042.8 | 126759.3 | 155758.07 | 797287.2 | 1669978.5 | 4138 | 366.73
72 | 3.33 | 1817.07 | 8006.85 | 16234.71 | 31832.47 | 65241.36 | 97790.75 | 130063.71 | 162927.7 | 829617.1 | 1760195.89 | 3991.19 | 333.4
72.05 | 10 | 1647.04 | 8105.32 | 15754.23 | 30402.54 | 61513.14 | 93514.53 | 123475.17 | 154696.9 | 790994.44 | 1680890.26 | 4074.6 | 360.06
72.1 | 0 | 1530.33 | 7516.48 | 14746.37 | 29400.69 | 59270.14 | 87768.88 | 116879.31 | 146436 | 746800.43 | 1620773.86 | 4284.71 | 363.4
72.15 | 6.67 | 236.71 | 830.16 | 1637.07 | 3167.64 | 6105.86 | 8581.21 | 9672.08 | 13322.38 | 61059.73 | 672673.21 | 2607.39 | 280.05
72.2 | 0 | 0 | 10 | 6.67 | 10 | 40.01 | 66.68 | 60.01 | 123.35 | 410.07 | 43303.53 | 20 | 0
72.25 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 10 | 3.33 | 6.67 | 43.34 | 44237.05 | 6.67 | 0
72.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 41579.15 | 273.39 | 0
72.35 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 43179.75 | 736.8 | 0
72.4 | 0 | 0 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 0 | 0 | 42065.28 | 966.86 | 0
72.45 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 1286.96 | 0
72.5 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 10 | 0 | 26.67 | 30 | 1450.35 | 0
72.55 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 13.34 | 0 | 20 | 46.67 | 1286.94 | 0
72.6 | 0 | 0 | 0 | 10 | 0 | 3.33 | 13.34 | 6.67 | 20.01 | 53.34 | 116.68 | 1316.93 | 0
72.65 | 0 | 0 | 6.67 | 0 | 6.67 | 3.33 | 13.33 | 33.34 | 40.01 | 233.37 | 446.75 | 1036.89 | 0
72.7 | 0 | 0 | 3.33 | 13.34 | 46.67 | 73.35 | 100.02 | 110.02 | 136.69 | 763.48 | 1443.65 | 43.34 | 0
72.75 | 0 | 40.01 | 250.04 | 503.42 | 1066.88 | 2177.25 | 3104.3 | 4641.76 | 5335.45 | 28636.38 | 61104.66 | 66.68 | 0
72.8 | 0 | 236.7 | 1226.92 | 2353.91 | 4931.61 | 9741.5 | 14629.7 | 19803.45 | 24000.44 | 127084.64 | 256226.67 | 763.49 | 0
72.85 | 0 | 363.39 | 1763.74 | 3407.61 | 6859.42 | 13888.84 | 20757.94 | 27613.38 | 34291.9 | 178143.05 | 357819.77 | 1697.07 | 6.67
72.9 | 0 | 393.4 | 2147.21 | 4294.6 | 8734.03 | 17253.02 | 24928.24 | 33362.57 | 41213.44 | 213293.58 | 428464.21 | 2170.54 | 13.33
72.95 | 0 | 446.75 | 2343.9 | 4451.43 | 9521.22 | 18124.01 | 27205.67 | 36657.46 | 45115.07 | 233488.08 | 473412.42 | 2820.74 | 13.33
73 | 3.33 | 496.76 | 2430.58 | 4331.33 | 9264.36 | 19425.71 | 28091 | 37005.33 | 47510.23 | 239744.84 | 483763.47 | 2797.4 | 13.34
73.05 | 6.67 | 490.09 | 2403.9 | 4601.42 | 9264.28 | 18674.93 | 28404.95 | 37794.24 | 46456.37 | 239497.54 | 481050.29 | 2937.49 | 10
73.1 | 0 | 450.08 | 2090.49 | 4591.42 | 9014.2 | 16895.78 | 25886.78 | 35183.64 | 43704.18 | 224304.31 | 454647.27 | 2727.37 | 20
73.15 | 0 | 80.01 | 516.76 | 976.87 | 1887.11 | 3454.38 | 4761.71 | 5869.06 | 8120.51 | 37272.23 | 76073.55 | 2343.9 | 0
73.2 | 0 | 10 | 0 | 6.67 | 3.33 | 16.67 | 16.67 | 20 | 23.34 | 196.7 | 293.38 | 20 | 6.67
73.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 46.67 | 3.33 | 20
73.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 26.67 | 3.33 | 236.71
73.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 6.67 | 3.33 | 1043.56
73.4 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 20 | 10 | 0 | 1270.26
73.45 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 6.67 | 33.34 | 63088.95 | 0 | 1567.01
73.5 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 16.67 | 3.33 | 30.01 | 42467.32 | 0 | 1663.71
73.55 | 3.33 | 3.33 | 0 | 0 | 0 | 3.33 | 23.34 | 10 | 3.33 | 73.34 | 43002.59 | 0 | 1763.77
73.6 | 0 | 0 | 0 | 0 | 10 | 23.34 | 40.01 | 23.34 | 63.34 | 260.05 | 44166.66 | 0 | 1700.4
73.65 | 0 | 0 | 20 | 33.34 | 26.67 | 73.34 | 136.69 | 180.03 | 260.05 | 1303.63 | 43681.94 | 0 | 1223.61
73.7 | 0 | 16.67 | 40.01 | 116.69 | 266.72 | 540.1 | 826.84 | 1040.23 | 1427.03 | 7357.67 | 47228.49 | 6.67 | 20
73.75 | 0 | 170.03 | 916.86 | 1780.43 | 3604.64 | 7100.38 | 11238.17 | 15561.01 | 18790.48 | 96408.56 | 94347.37 | 56.68 | 3.33
73.8 | 10 | 813.5 | 3824.56 | 8257.08 | 16338.83 | 33478.63 | 51167.72 | 68305.06 | 82315.27 | 434246.85 | 607948.23 | 293.38 | 156.7
73.85 | 13.34 | 1613.69 | 7893.37 | 16381.84 | 32196.65 | 64193.34 | 95965.11 | 126752.99 | 159713.68 | 819938.44 | 1527188.99 | 830.16 | 530.1
73.9 | 6.67 | 2163.86 | 9814.76 | 20370.64 | 40394.68 | 80145.26 | 120082.38 | 160803.56 | 202106.66 | 1027470.37 | 2043236.13 | 1310.27 | 1070.22
73.95 | 6.67 | 2477.27 | 11936.67 | 23582.48 | 47934.59 | 95836.12 | 143280.47 | 190855.31 | 238900.74 | 1243585.06 | 2465251.02 | 1787.1 | 1446.99
74 | 3.33 | 2337.24 | 11753.3 | 23896.67 | 47506.42 | 95837.35 | 142433.45 | 190412.55 | 237789.29 | 1270443.33 | 2600225.55 | 1683.72 | 1520.36
74.05 | 10 | 2403.92 | 11653.16 | 22951.45 | 46235.86 | 93327.35 | 138704.06 | 186065.76 | 234975.28 | 1245429.05 | 2518032.43 | 1910.46 | 1370.29
74.1 | 6.67 | 2260.54 | 11032.46 | 21863.15 | 44146.04 | 86776.63 | 131612.66 | 176042.57 | 217386.16 | 1183715.35 | 2409908.49 | 1750.41 | 1233.59
74.15 | 0 | 483.42 | 2380.62 | 4141.34 | 7666.85 | 15054.54 | 21771.82 | 26792.35 | 36049.51 | 495551.19 | 1170966.82 | 993.54 | 1163.58
74.2 | 0 | 0 | 13.34 | 3.33 | 56.68 | 83.35 | 140.02 | 126.69 | 213.37 | 40508.63 | 43339.43 | 20 | 10
74.25 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 6.67 | 16.67 | 13.34 | 38725.22 | 45257.03 | 0 | 26.67
74.3 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 39957.52 | 44096.06 | 0 | 270.04
74.35 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 40779.24 | 43346.41 | 0 | 886.85
74.4 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 41376.12 | 43705.68 | 3.33 | 1173.57
74.45 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 3.33 | 16.67 | 0 | 1580.36
74.5 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 3.33 | 43.34 | 0 | 1613.69
74.55 | 0 | 0 | 3.33 | 0 | 13.34 | 3.33 | 0 | 0 | 0 | 23.34 | 96.68 | 0 | 1750.39
74.6 | 0 | 0 | 0 | 0 | 10 | 13.34 | 10 | 30 | 46.68 | 106.68 | 236.7 | 0 | 1833.77
74.65 | 0 | 0 | 3.33 | 6.67 | 10 | 16.67 | 56.68 | 83.35 | 86.68 | 386.73 | 890.19 | 0 | 1450.3
74.7 | 0 | 6.67 | 10 | 53.34 | 80.01 | 223.37 | 276.71 | 383.4 | 373.4 | 2433.97 | 5328.82 | 3.33 | 40.01
74.75 | 0 | 33.34 | 143.35 | 276.71 | 656.81 | 1096.9 | 1973.91 | 2757.6 | 3434.53 | 17128.35 | 36802.36 | 3.33 | 20
74.8 | 0 | 336.73 | 1637.03 | 3137.53 | 6309.15 | 13044.88 | 19193 | 26185.19 | 32208.13 | 164930.63 | 344163.25 | 100.02 | 126.69
74.85 | 3.33 | 866.83 | 3667.72 | 7706.62 | 15454.13 | 30543.05 | 44069.34 | 60211.27 | 74826.47 | 375250.76 | 781849.59 | 670.13 | 680.13
74.9 | 0 | 976.85 | 4594.82 | 9010.8 | 17650.28 | 36082.94 | 53906.47 | 72173.52 | 91163.81 | 451205.52 | 943092.57 | 916.85 | 883.5
74.95 | 6.67 | 1016.87 | 4908.3 | 9421.18 | 18571.49 | 38432.3 | 57784.72 | 76619.98 | 95514.73 | 481756.12 | 1004608.74 | 1070.21 | 1023.53
75 | 0 | 1043.55 | 4708.18 | 9716.47 | 20120.43 | 40003.82 | 59679.84 | 80354.14 | 98789.87 | 497207.36 | 1047394.09 | 1143.57 | 1070.22
75.05 | 0 | 900.18 | 4774.89 | 9511.19 | 19859.97 | 39361.52 | 58984 | 78401.51 | 97510.94 | 488583.09 | 1021973.58 | 1153.57 | 1200.26
75.1 | 16.67 | 946.87 | 4848.37 | 9297.76 | 18290.88 | 36233 | 53976.84 | 72990.47 | 92545.06 | 456981.28 | 951710.53 | 1216.93 | 1016.86
75.15 | 0 | 66.68 | 513.43 | 936.85 | 1830.47 | 3451.04 | 4878.49 | 5865.76 | 8364.23 | 36806.74 | 72824.34 | 896.86 | 756.82
75.2 | 0 | 0 | 3.33 | 6.67 | 13.34 | 26.67 | 33.34 | 60.01 | 60.01 | 270.05 | 626.78 | 3.33 | 23.34
75.25 | 6.67 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 26.67 | 83.35 | 0 | 0
75.3 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 0
75.35 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 6.67 | 0 | 3.33 | 3.33
75.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 3.33 | 3.33
75.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 26.67 | 0 | 0
75.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 30 | 0 | 0
75.55 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 13.34 | 23.34 | 40.01 | 0 | 0
75.6 | 0 | 0 | 0 | 6.67 | 3.33 | 10 | 3.33 | 16.67 | 6.67 | 76.68 | 110.02 | 0 | 3.33
75.65 | 0 | 0 | 0 | 0 | 0 | 10 | 26.67 | 33.34 | 33.34 | 233.37 | 540.1 | 3.33 | 3.33
75.7 | 0 | 3.33 | 13.34 | 33.34 | 40.01 | 106.68 | 120.02 | 210.03 | 233.37 | 1123.57 | 2240.6 | 0 | 10
75.75 | 6.67 | 80.01 | 340.06 | 640.13 | 1263.6 | 2657.44 | 3941.45 | 5668.94 | 6513.02 | 34012.07 | 73206.05 | 30.01 | 46.67
75.8 | 20 | 313.39 | 1517 | 2800.75 | 5738.84 | 11403.17 | 17537.03 | 22774.79 | 28519.04 | 146610.87 | 300902.14 | 96.69 | 673.47
75.85 | 50.01 | 516.76 | 2473.96 | 4444.66 | 8960.85 | 18448.06 | 27022.07 | 35387.91 | 45065.82 | 230410.56 | 468346.1 | 223.37 | 2080.48
75.9 | 73.35 | 520.1 | 2780.72 | 5218.44 | 10598.85 | 20457.24 | 31418.5 | 41146.75 | 52297.76 | 263221 | 538026.43 | 313.4 | 2684.02
75.95 | 70.01 | 610.11 | 2937.42 | 6172.27 | 11946.66 | 22911.59 | 34903.31 | 45831.45 | 57831.37 | 291953.06 | 597220.68 | 283.39 | 3160.88
76 | 60.01 | 740.14 | 3217.53 | 6265.62 | 11976.85 | 24704.58 | 36701.88 | 47882.04 | 60559.41 | 308198.73 | 625160.57 | 333.4 | 3704.43
76.05 | 50.01 | 586.78 | 3157.51 | 5912.14 | 11676.59 | 23789.89 | 34636.04 | 47377.45 | 58601.19 | 298061.22 | 608318.35 | 243.38 | 3451.02
76.1 | 80.02 | 583.44 | 2867.43 | 5581.93 | 11186.25 | 21542.68 | 33078.92 | 44115.6 | 54201.01 | 276666.37 | 563427.65 | 293.38 | 3244.26
76.15 | 23.34 | 113.35 | 546.77 | 906.84 | 1660.37 | 3427.66 | 5128.58 | 5865.56 | 8310.81 | 36565.78 | 72497.87 | 56.68 | 3194.21
76.2 | 0 | 0 | 3.33 | 3.33 | 16.67 | 13.34 | 26.67 | 40.01 | 56.68 | 203.37 | 356.73 | 0 | 73.34
76.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 16.67 | 46.67 | 0 | 0
76.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0
76.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0
76.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0
76.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0
76.5 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 6.67
76.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 6.67
76.6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
76.65 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 3.33 | 16.67 | 46.67 | 0 | 6.67
76.7 | 3.33 | 0 | 0 | 0 | 0 | 13.34 | 10 | 3.33 | 20 | 110.02 | 213.37 | 0 | 6.67
76.75 | 0 | 3.33 | 16.67 | 20 | 23.34 | 110.02 | 193.37 | 233.37 | 300.05 | 1313.61 | 3201.06 | 0 | 36.67
76.8 | 0 | 20 | 96.68 | 193.37 | 320.05 | 630.12 | 826.82 | 1270.25 | 1500.34 | 7816.78 | 16825.72 | 36.67 | 506.77
76.85 | 0 | 33.34 | 126.69 | 286.72 | 423.41 | 943.52 | 1350.28 | 1827.07 | 2357.24 | 11476.22 | 25295.44 | 23.34 | 1386.99
76.9 | 0 | 26.67 | 166.69 | 346.73 | 630.12 | 1320.29 | 1977.13 | 2650.69 | 3157.54 | 16268.53 | 33753.53 | 76.68 | 2870.82
76.95 | 3.33 | 56.68 | 170.03 | 370.06 | 580.1 | 1333.61 | 2183.87 | 2930.78 | 3514.35 | 17306.53 | 37753.46 | 60.01 | 3340.96
77 | 0 | 30 | 160.03 | 386.73 | 783.48 | 1403.64 | 2047.15 | 2747.37 | 3711.08 | 17167.78 | 37646.71 | 36.67 | 3294.3
77.05 | 0 | 46.67 | 233.38 | 330.06 | 706.81 | 1390.28 | 2270.59 | 2610.68 | 3554.4 | 16822.29 | 36738.21 | 33.34 | 3601.08
77.1 | 0 | 20 | 186.7 | 366.73 | 706.81 | 1400.31 | 2133.87 | 2547.31 | 3410.93 | 15780.9 | 34124.76 | 30 | 3637.77
77.15 | 0 | 13.33 | 46.67 | 80.01 | 110.02 | 203.37 | 270.05 | 376.73 | 513.43 | 2327.31 | 4718.34 | 6.67 | 2884.11
77.2 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 3.33 | 0 | 20 | 13.33 | 3.33 | 43.34
77.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.67
77.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 133.36
77.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 710.14
77.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 996.88
77.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 1260.26
77.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1486.99
77.55 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 10 | 0 | 1487
77.6 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 0 | 6.67 | 33.34 | 3.33 | 1537.01
77.65 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 13.34 | 6.67 | 50.01 | 146.69 | 3.33 | 1220.26
77.7 | 0 | 0 | 0 | 6.67 | 0 | 13.34 | 10 | 36.67 | 33.34 | 183.36 | 483.43 | 0 | 63.34
77.75 | 0 | 3.33 | 20.01 | 40.01 | 46.67 | 203.37 | 363.42 | 386.74 | 473.42 | 2507.39 | 5472.61 | 3.33 | 16.67
77.8 | 6.67 | 40.01 | 326.73 | 563.44 | 1233.58 | 2423.95 | 3504.3 | 4531.44 | 5665.41 | 27222.96 | 60443.74 | 73.35 | 433.41
77.85 | 6.67 | 110.02 | 380.07 | 963.53 | 1953.79 | 3641.03 | 5245.14 | 6999.46 | 8643.94 | 42872.76 | 94077.3 | 86.68 | 1243.61
77.9 | 26.67 | 116.69 | 503.42 | 1116.89 | 2137.18 | 4341.32 | 6515.83 | 8537.12 | 10385.29 | 51744.98 | 111533.51 | 163.36 | 1870.41
77.95 | 13.34 | 146.69 | 550.1 | 1093.56 | 2217.19 | 4518.12 | 6732.67 | 8934.18 | 11413.02 | 55272.02 | 120483.73 | 153.36 | 2307.26
78 | 26.67 | 186.7 | 566.77 | 1306.97 | 2423.91 | 4728.23 | 7286.29 | 8974.24 | 11833.42 | 57115.04 | 124161.59 | 200.03 | 2420.62
78.05 | 36.67 | 126.69 | 653.47 | 1250.28 | 2330.56 | 4848.27 | 6942.88 | 9291.04 | 11729.91 | 55010.83 | 122322.55 | 180.03 | 2373.93
78.1 | 26.67 | 123.35 | 596.78 | 1066.87 | 2217.18 | 4227.99 | 6335.72 | 8480.51 | 10065.03 | 51314.1 | 107788.86 | 173.36 | 2310.59
78.15 | 0 | 0 | 76.68 | 86.68 | 113.35 | 213.37 | 303.38 | 500.09 | 583.46 | 2624.06 | 5635.74 | 6.67 | 1767.08
78.2 | 0 | 3.33 | 0 | 0 | 6.67 | 6.67 | 3.33 | 0 | 0 | 16.67 | 26.67 | 0 | 66.68
78.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 16.67
78.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 110.02
78.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 706.82
78.4 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1560.36
78.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 1760.42
78.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 1567.01
78.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1757.07
78.6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1550.36
78.65 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 1356.94
78.7 | 0 | 3.33 | 0 | 3.33 | 6.67 | 0 | 0 | 6.67 | 0 | 6.67 | 3.33 | 0 | 40.01
78.75 | 10 | 3.33 | 13.34 | 13.34 | 10 | 16.67 | 30.01 | 10 | 6.67 | 13.34 | 3.33 | 30.01 | 23.34
78.8 | 66.68 | 93.36 | 80.03 | 90.02 | 73.36 | 50.01 | 93.35 | 83.35 | 43.34 | 96.68 | 80.01 | 150.03 | 333.39
78.85 | 100.02 | 126.69 | 93.35 | 40.01 | 100.02 | 126.69 | 140.03 | 120.02 | 106.7 | 86.68 | 70.01 | 266.73 | 1410.31
78.9 | 116.69 | 160.05 | 93.35 | 113.35 | 150.03 | 166.7 | 203.38 | 170.04 | 163.37 | 120.02 | 116.69 | 280.06 | 2000.48
78.95 | 213.39 | 186.73 | 183.38 | 153.38 | 176.7 | 183.38 | 210.05 | 153.37 | 173.37 | 166.71 | 143.37 | 266.72 | 2387.26
79 | 180.04 | 193.39 | 153.36 | 110.02 | 160.04 | 180.04 | 190.04 | 166.71 | 223.41 | 136.7 | 150.03 | 436.77 | 2960.83
79.05 | 126.7 | 140.03 | 180.04 | 153.37 | 166.72 | 196.71 | 196.72 | 230.07 | 163.38 | 126.7 | 90.02 | 366.76 | 2900.75
79.1 | 140.03 | 140.03 | 163.37 | 116.7 | 190.05 | 193.38 | 136.7 | 140.04 | 143.37 | 140.03 | 126.69 | 376.76 | 2827.43
79.15 | 36.67 | 36.67 | 36.67 | 20 | 16.67 | 46.68 | 20 | 6.67 | 23.34 | 23.34 | 20 | 46.68 | 2387.28
79.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 143.36
79.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 10 | 0
79.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 10
79.35 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0
79.4 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 0
79.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 0 | 0
79.5 | 0 | 3.33 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 6.67
134.55 | 0 | 6056.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
134.6 | 0 | 4653.6 | 0 | 0 | 3.33 | 6.67 | 0 | 0 | 0 | 0 | 0 | 20 | 0
134.65 | 0 | 5071.15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 0
134.7 | 0 | 5510.31 | 3.33 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 70.02 | 0
134.75 | 0 | 5012.12 | 0 | 3.33 | 3.33 | 0 | 0 | 6.67 | 0 | 0 | 0 | 700.14 | 0
134.8 | 6.67 | 5228.44 | 53.43 | 0 | 0 | 3.33 | 16.67 | 10 | 3.33 | 154.11 | 535.94 | 6719.48 | 0
134.85 | 70.07 | 5792.27 | 230.66 | 10 | 40.03 | 26.68 | 3.33 | 76.78 | 184.29 | 16.68 | 80.08 | 12984.91 | 6.67
134.9 | 240.71 | 5417.13 | 20 | 0 | 220.91 | 16.67 | 23.34 | 26.68 | 40.01 | 16.67 | 46.69 | 18515.33 | 10
134.95 | 50.04 | 20 | 110.37 | 10 | 1158.84 | 0 | 50.02 | 10 | 6.67 | 33.37 | 20 | 22230.94 | 16.67
135 | 66.7 | 292.52 | 40.01 | 10 | 598.25 | 6.67 | 10 | 96.93 | 3.33 | 16.67 | 46.71 | 23089.42 | 16.67
135.05 | 36.69 | 50.06 | 13.34 | 26.68 | 46.69 | 20 | 0 | 23.34 | 13.34 | 10 | 100.09 | 23399.89 | 30.01
135.1 | 16.67 | 36.69 | 245.1 | 43.36 | 110.17 | 20 | 10 | 16.67 | 33.34 | 33.35 | 103.57 | 20304.59 | 13.34
135.15 | 3.33 | 228.43 | 106.99 | 16.67 | 23.34 | 0 | 0 | 0 | 6.67 | 50.06 | 0 | 2323.95 | 0
135.2 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 0 | 23.34 | 0
135.25 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 0 | 0 | 3.33 | 0
135.3 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0
135.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 3.33 | 0
135.4 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33
135.45 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0
135.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 0
136.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0
136.6 | 3.33 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0
136.65 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 40.01 | 0
136.7 | 0 | 0 | 0 | 0 | 6.67 | 0 | 6.67 | 0 | 6.67 | 0 | 6.67 | 243.38 | 0
136.75 | 26.69 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 53.36 | 53.4 | 1847.18 | 10
136.8 | 23.34 | 0 | 13.34 | 3.33 | 0 | 23.34 | 0 | 43.35 | 6171.99 | 26.67 | 636.9 | 11253.54 | 10
136.85 | 23.34 | 16.67 | 20 | 66.72 | 13.33 | 16.67 | 46.69 | 322.73 | 861.65 | 272.22 | 56.72 | 26973.5 | 6.67
136.9 | 16.67 | 40.03 | 23.34 | 23.34 | 20 | 33.34 | 33.34 | 30.01 | 106.93 | 20.01 | 43.34 | 34250.11 | 20
136.95 | 10 | 46.68 | 156.84 | 23.34 | 23.34 | 46.68 | 26.67 | 167.34 | 36.69 | 36.67 | 234.63 | 41073.06 | 33.34
137 | 110.21 | 86.74 | 33.35 | 13.34 | 36.67 | 36.67 | 20 | 5837.68 | 33.34 | 26.67 | 40.01 | 43500.32 | 40.01
137.05 | 33.34 | 133.75 | 110.18 | 30.01 | 16.67 | 83.47 | 43.37 | 5445.74 | 30 | 1199.84 | 60.02 | 42999.16 | 10
137.1 | 36.69 | 70.05 | 43.35 | 26.67 | 20 | 103.47 | 43.35 | 1331 | 23.34 | 184.29 | 46.68 | 38197.52 | 30.01
137.15 | 10 | 3.33 | 3.33 | 13.34 | 46.69 | 23.35 | 6.67 | 143.62 | 0 | 0 | 0 | 4761.72 | 0
137.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 10 | 0 | 26.67 | 0
137.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0
137.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
137.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33
137.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
137.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0
137.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0
142.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 0 | 3.33 | 0
142.6 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 46.67 | 0
142.65 | 0 | 3.33 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 203.37 | 0
142.7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2267.56 | 0
142.75 | 0 | 6.67 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 19163.22 | 6.67
142.8 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 10 | 62628.67 | 20
142.85 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 13.34 | 0 | 13.34 | 103487.58 | 36.67
142.9 | 0 | 0 | 0 | 6.67 | 0 | 3.33 | 0 | 3.33 | 3.33 | 3.33 | 6.67 | 158088.46 | 33.34
142.95 | 0 | 6.67 | 0 | 0 | 6.67 | 0 | 6.67 | 0 | 36.71 | 0 | 6.67 | 183954.73 | 20
143 | 0 | 6.67 | 3.33 | 0 | 0 | 13.34 | 6.67 | 6.67 | 0 | 6.67 | 6.67 | 187256.11 | 46.67
143.05 | 6.67 | 0 | 3.33 | 0 | 0 | 56.78 | 13.34 | 3.33 | 3.33 | 6.67 | 6.67 | 190135.25 | 53.34
143.1 | 0 | 0 | 0 | 6.67 | 3.33 | 3.33 | 13.34 | 6.67 | 50.08 | 3.33 | 6.67 | 135898.96 | 40.01
143.15 | 0 | 0 | 0 | 0 | 3.33 | 0 | 6.67 | 0 | 0 | 3.33 | 0 | 10443.26 | 3.33
143.2 | 0 | 0 | 0 | 6.67 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 0 | 53.34 | 0
143.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 13.34 | 0
143.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0
143.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
143.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
143.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 3.33
143.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 16.67 | 0
144.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 13.34 | 0
144.6 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 23.34 | 0
144.65 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 116.69 | 0
144.7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 2244.07 | 3.33
144.75 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 13573.92 | 6.67
144.8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 45684.51 | 16.67
144.85 | 3.33 | 10 | 10 | 0 | 3.33 | 6.67 | 13.34 | 0 | 6.67 | 0 | 3.33 | 79613.34 | 23.34
144.9 | 0 | 6.67 | 0 | 0 | 0 | 3.33 | 0 | 6.67 | 0 | 0 | 13.34 | 100292.55 | 20
144.95 | 0 | 6.67 | 3.33 | 0 | 0 | 0 | 0 | 0 | 16.67 | 0 | 6.67 | 131698.63 | 50.01
145 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 0 | 0 | 3.33 | 0 | 0 | 138184.65 | 43.34
145.05 | 0 | 0 | 3.33 | 0 | 6.67 | 0 | 110.42 | 0 | 6.67 | 6.67 | 6.67 | 135867.09 | 36.67
145.1 | 0 | 6.67 | 0 | 0 | 0 | 0 | 26.69 | 0 | 0 | 0 | 3.33 | 94587.92 | 33.34
145.15 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 8100.8 | 3.33
145.2 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 73.35 | 0
145.25 | 0 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 0
145.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
145.35 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0
145.4 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
145.45 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
145.5 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 16.67 | 0
149.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67
149.6 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 46.67 | 36.67
149.65 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 73.35 | 116.69
149.7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1273.74 | 1657.1
149.75 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 10099.78 | 14982.61
149.8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 35001.82 | 45082.3
149.85 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 30.01 | 3.33 | 0 | 3.33 | 55198.07 | 69414.97
149.9 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 6.67 | 0 | 13.34 | 10 | 77630.69 | 96188.34
149.95 | 0 | 13.34 | 0 | 0 | 0 | 0 | 0 | 10 | 6.67 | 0 | 3.33 | 101752.69 | 125877.59
150 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 13.34 | 10 | 6.67 | 0 | 106592.15 | 127897.74
150.05 | 6.67 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 107757.28 | 130962.84
150.1 | 6.67 | 3.33 | 0 | 0 | 0 | 0 | 0 | 30.03 | 3.33 | 3.33 | 13.34 | 86798.72 | 100362.58
150.15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 11056.88 | 8834.88
150.2 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 80.01 | 116.69
150.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 3.33
150.3 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
150.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
150.4 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
150.45 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0
150.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
154.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 26.67
154.6 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 96.68
154.65 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 3.33 | 0 | 0 | 790.18
154.7 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 10 | 7301.73
154.75 | 10 | 0 | 0 | 10 | 6.67 | 0 | 0 | 0 | 6.67 | 0 | 0 | 16.67 | 54674.26
154.8 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 50.01 | 139874.98
154.85 | 0 | 0 | 0 | 3.33 | 0 | 6.67 | 0 | 0 | 0 | 0 | 3.33 | 140.03 | 250254.22
154.9 | 0 | 13.34 | 0 | 6.67 | 13.34 | 0 | 0 | 0 | 13.34 | 3.33 | 0 | 223.37 | 341318.17
154.95 | 10 | 6.67 | 0 | 26.68 | 3.33 | 10 | 0 | 0 | 3.33 | 3.33 | 3.33 | 163.36 | 433914.52
155 | 6.67 | 6.67 | 6.67 | 6.67 | 0 | 0 | 36.71 | 0 | 0 | 0 | 6.67 | 246.71 | 459179.27
155.05 | 0 | 0 | 0 | 13.34 | 3.33 | 0 | 0 | 6.67 | 0 | 0 | 20.01 | 206.71 | 468721.33
155.1 | 3.33 | 20.01 | 3.33 | 0 | 0 | 73.5 | 6.67 | 0 | 0 | 13.34 | 0 | 113.35 | 358937.5
155.15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 10 | 50107.83
155.2 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 473.42
155.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 0 | 3.33 | 26.67
155.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33
155.35 | 3.33 | 6.67 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
155.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33
155.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
155.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 3.33
155.55 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 90.02
155.6 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 6.67 | 173.36
155.65 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 0 | 6.67 | 790.18
155.7 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 3.33 | 0 | 3.33 | 70.01 | 7119.42
155.75 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 313.4 | 41126.92
155.8 | 0 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 793.51 | 161485.83
155.85 | 6.67 | 3.33 | 6.67 | 10 | 0 | 3.33 | 0 | 43.4 | 3.33 | 0 | 10 | 1673.85 | 385460.84
155.9 | 20.01 | 6.67 | 0 | 20.01 | 6.67 | 16.68 | 23.35 | 46.72 | 53.43 | 6.67 | 3.33 | 2120.65 | 491364.26
155.95 | 0 | 0 | 60.12 | 6.67 | 6.67 | 33.35 | 3.33 | 13.34 | 3.33 | 10 | 20 | 2660.98 | 603754.35
156 | 3.33 | 282.74 | 6.67 | 13.34 | 3.33 | 6.67 | 3.33 | 0 | 3.33 | 3.33 | 0 | 2470.87 | 649288.1
156.05 | 0 | 6.67 | 6.67 | 6.67 | 0 | 43.37 | 3.33 | 0 | 6.67 | 0 | 20.01 | 2400.78 | 643951.31
156.1 | 0 | 0 | 0 | 6.67 | 6.67 | 0 | 0 | 0 | 10 | 0 | 33.37 | 1360.4 | 542270.56
156.15 | 0 | 0 | 0 | 0 | 126.95 | 0 | 0 | 0 | 3.33 | 0 | 6.67 | 33.34 | 49690.96
156.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 420.08
156.25 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 3.33 | 23.34
156.3 | 0 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
156.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
156.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33
156.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 10
156.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 0 | 0 | 6.67
156.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 40.01
156.6 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 103.35
156.65 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 500.1
156.7 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 2464.26
156.75 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 26.69 | 6.67 | 3.33 | 0 | 10 | 55978.5
156.8 | 0 | 0 | 6.67 | 0 | 30.03 | 0 | 33.37 | 0 | 3.33 | 3.33 | 0 | 76.68 | 174351.68
156.85 | 0 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 130.02 | 274108.44
156.9 | 6.67 | 0 | 174.38 | 13.34 | 0 | 6.67 | 6.67 | 0 | 20.01 | 0 | 0 | 253.38 | 373381.4
156.95 | 0 | 0 | 66.81 | 0 | 0 | 0 | 0 | 0 | 10 | 0 | 10 | 200.04 | 457643.76
157 | 0 | 13.34 | 23.35 | 30.03 | 0 | 0 | 0 | 63.46 | 0 | 10 | 0 | 250.05 | 505855.77
157.05 | 10 | 10 | 0 | 0 | 0 | 3.33 | 0 | 20.01 | 6.67 | 0 | 0 | 176.7 | 515246.1
157.1 | 16.67 | 6.67 | 1510.85 | 0 | 10 | 0 | 0 | 6.67 | 33.36 | 0 | 0 | 133.36 | 406348.23
157.15 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 13.33 | 50102.03
157.2 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 6.67 | 0 | 0 | 0 | 450.09
157.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 30.01
157.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67
157.35 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0
157.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
157.45 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 13.34
157.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10
Printed:7/20/2016 9:48:52 AM

## Standard Resolution
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Standard resolution, 20 pts, 10 replicates, 2.5 mL/min He
Mass Hunter file Agilent 7700 ICP-MS: 7-19-16 standard res 2
| 3BW34A | 3BW17 (C-J) | [7 empty] | 3BW34C | 3BW34B
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1
Mass | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS | CPS
64.55 | 60.01 | 73.35 | 40.01 | 60.01 | 133.35 | 26.67 | 46.67 | 53.34 | 60.01 | 86.68 | 146.69 | 160.03 | 146.69
64.6 | 236.71 | 313.39 | 186.69 | 150.03 | 503.42 | 146.69 | 210.04 | 236.71 | 186.7 | 393.41 | 326.72 | 626.78 | 666.79
64.65 | 263.38 | 373.4 | 293.39 | 233.38 | 930.18 | 266.71 | 313.39 | 310.06 | 273.38 | 460.09 | 573.44 | 876.83 | 933.51
64.7 | 296.71 | 506.75 | 350.07 | 253.38 | 1110.22 | 396.74 | 430.07 | 360.07 | 423.41 | 686.79 | 926.86 | 1123.54 | 1333.6
64.75 | 323.38 | 443.42 | 393.4 | 296.72 | 1140.23 | 333.39 | 480.09 | 413.41 | 503.43 | 736.8 | 1026.91 | 1290.26 | 1470.32
64.8 | 420.07 | 503.43 | 420.07 | 360.06 | 1266.9 | 463.42 | 540.1 | 513.43 | 466.75 | 813.48 | 1150.23 | 1280.25 | 1446.99
64.85 | 376.73 | 570.11 | 356.73 | 303.39 | 1406.97 | 303.39 | 500.09 | 426.74 | 433.41 | 870.16 | 1170.24 | 1376.96 | 1470.32
64.9 | 436.74 | 600.1 | 420.07 | 386.74 | 1430.31 | 430.08 | 720.15 | 333.39 | 536.76 | 883.51 | 1073.55 | 1373.61 | 1510.32
64.95 | 350.06 | 706.8 | 410.07 | 373.41 | 1400.27 | 420.08 | 580.11 | 590.1 | 536.77 | 953.53 | 1093.55 | 1356.94 | 1483.63
65 | 353.4 | 653.45 | 506.75 | 413.4 | 1323.59 | 336.72 | 643.46 | 450.08 | 510.09 | 866.83 | 1213.57 | 1456.96 | 1510.32
65.05 | 416.74 | 720.13 | 306.71 | 416.74 | 1283.6 | 416.74 | 500.09 | 460.08 | 546.77 | 960.2 | 1036.86 | 1386.98 | 1620.36
65.1 | 380.06 | 530.09 | 433.41 | 370.06 | 1136.9 | 303.38 | 593.43 | 393.4 | 476.76 | 720.14 | 1146.9 | 1266.91 | 1500.33
65.15 | 353.39 | 453.41 | 420.08 | 303.38 | 1203.58 | 380.07 | 483.42 | 390.07 | 423.41 | 800.15 | 1020.21 | 1080.22 | 1303.6
65.2 | 263.38 | 526.77 | 330.06 | 206.7 | 933.52 | 276.71 | 473.41 | 313.39 | 386.73 | 676.8 | 753.48 | 1113.56 | 1393.62
65.25 | 16.67 | 13.34 | 0 | 13.33 | 23.34 | 6.67 | 10 | 20 | 23.34 | 43.34 | 50.01 | 23.34 | 60.01
65.3 | 0 | 3.33 | 0 | 0 | 3.33 | 3.33 | 13.34 | 3.33 | 23.34 | 80.01 | 126.68 | 0 | 3.33
65.35 | 0 | 0 | 10 | 6.67 | 3.33 | 3.33 | 33.34 | 16.67 | 43.34 | 210.03 | 413.4 | 0 | 0
65.4 | 0 | 0 | 3.33 | 10 | 10 | 33.34 | 70.01 | 93.35 | 83.35 | 483.42 | 1043.56 | 0 | 0
65.45 | 0 | 0 | 26.67 | 40.01 | 63.34 | 156.69 | 206.7 | 243.37 | 370.07 | 43976.22 | 73185.84 | 0 | 0
65.5 | 0 | 20 | 160.02 | 270.05 | 556.78 | 983.55 | 1453.68 | 1823.78 | 2274.02 | 41170.47 | 55927.41 | 23.34 | 0
65.55 | 13.34 | 950.19 | 4544.77 | 8554.04 | 17647.14 | 33364.26 | 49693.92 | 62575.56 | 81587.46 | 271586.49 | 372239.71 | 510.09 | 260.04
65.6 | 43.34 | 1703.69 | 7603.21 | 15544.15 | 29988.6 | 57981.5 | 85746.25 | 112897.29 | 140151.79 | 618571.97 | 1266781.77 | 766.81 | 553.44
65.65 | 36.67 | 2087.14 | 9454.61 | 18881.79 | 37336.47 | 74662.45 | 107854.85 | 142125.29 | 179189.94 | 806974.65 | 1686519.73 | 1016.86 | 756.84
65.7 | 43.34 | 2370.58 | 11202.7 | 21545.55 | 43938.98 | 86355.75 | 127145.99 | 164867.89 | 210223.73 | 958076.59 | 2039673.31 | 1416.96 | 916.85
65.75 | 33.34 | 2483.98 | 11394.58 | 23298.66 | 46533.03 | 92509.94 | 138422.85 | 179270.21 | 227504.6 | 1047268.01 | 2230764.79 | 1370.3 | 796.8
65.8 | 60.01 | 2870.77 | 13224.74 | 25672.73 | 49647.19 | 99960.01 | 147263.03 | 194773.17 | 244601.75 | 1124032.9 | 2406307.87 | 1627.04 | 880.16
65.85 | 56.68 | 2844.08 | 13775.35 | 27012.06 | 53871.86 | 107625.45 | 160714.91 | 210576.17 | 261145.83 | 1220933.98 | 2582600.5 | 1697.05 | 1000.2
65.9 | 46.67 | 2970.78 | 14269.27 | 28104.17 | 56773.53 | 109910.6 | 164714.28 | 215836.42 | 269900.1 | 1258340.56 | 2684526.31 | 1633.67 | 983.52
65.95 | 56.68 | 2824.05 | 13875.39 | 28023.67 | 55954 | 110237.71 | 165174.57 | 215469 | 272226.34 | 1301361.89 | 2721501.96 | 1596.99 | 1010.21
66 | 70.01 | 3017.47 | 14102.31 | 27108.57 | 54983.36 | 109051.73 | 164681.71 | 214369.12 | 269846.71 | 1274537.15 | 2689950.06 | 1607.01 | 1076.89
66.05 | 46.67 | 2990.77 | 13421.76 | 26842.13 | 53717.57 | 104714.87 | 156542.71 | 205328.98 | 258624.5 | 1240970.28 | 2644605.37 | 1460.3 | 996.87
66.1 | 33.34 | 2830.74 | 13141.22 | 25322.24 | 51420.52 | 100536.67 | 151083.76 | 193276.82 | 247208.16 | 1192728.56 | 2513855.97 | 1613.69 | 880.17
66.15 | 50.01 | 2353.92 | 11326.27 | 23248.52 | 45285.88 | 90581.83 | 133466.74 | 176359.49 | 221928.31 | 1082958.92 | 2310988.62 | 1380.29 | 850.16
66.2 | 36.67 | 2143.84 | 9651.4 | 18671.51 | 37549.44 | 74964.19 | 111768.67 | 148829.55 | 187034.47 | 951546.42 | 2057344.19 | 1100.21 | 693.48
66.25 | 0 | 56.68 | 236.71 | 486.76 | 1106.91 | 2253.99 | 3180.98 | 4368.32 | 5355.44 | 193968 | 443293.26 | 43.34 | 50.01
66.3 | 0 | 0 | 6.67 | 0 | 0 | 36.67 | 46.68 | 93.35 | 120.02 | 49626.49 | 51587.5 | 0 | 0
66.35 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 3.33 | 16.67 | 20 | 50207.62 | 51179.62 | 0 | 0
66.4 | 0 | 0 | 0 | 0 | 6.67 | 10 | 13.34 | 16.67 | 13.34 | 50241.61 | 51675.3 | 0 | 0
66.45 | 0 | 3.33 | 0 | 16.67 | 26.67 | 60.01 | 66.68 | 60.01 | 120.02 | 516.77 | 1020.22 | 10 | 3.33
66.5 | 0 | 23.34 | 93.35 | 123.35 | 320.06 | 653.46 | 1006.9 | 1046.9 | 1490.38 | 6456.91 | 13669.68 | 16.67 | 6.67
66.55 | 3.33 | 116.69 | 476.75 | 1036.87 | 2173.86 | 3971.24 | 5588.75 | 7346.64 | 9248.04 | 39583.27 | 85313.33 | 103.35 | 33.34
66.6 | 10 | 240.04 | 913.51 | 2297.24 | 4554.76 | 8790.84 | 12697.44 | 17356.44 | 21111.88 | 93937.53 | 201261.18 | 390.07 | 40.01
66.65 | 0 | 333.39 | 1416.95 | 2830.69 | 5575.33 | 11019.34 | 16662.07 | 21812.64 | 27560.18 | 123507.64 | 259629.43 | 350.06 | 86.68
66.7 | 10 | 276.71 | 1607.02 | 3247.55 | 6672.6 | 13051.28 | 19112.15 | 25382.27 | 31895.79 | 139970.78 | 296071.17 | 403.4 | 156.7
66.75 | 10 | 423.41 | 2020.47 | 3834.46 | 7873.43 | 14269.14 | 21902.96 | 28611.68 | 35835.67 | 164667.01 | 348005.03 | 596.77 | 143.36
66.8 | 3.33 | 526.76 | 1987.13 | 3924.47 | 8016.95 | 15263.62 | 23559.04 | 30599.15 | 38636.38 | 174172.58 | 366594.12 | 600.11 | 110.02
66.85 | 13.34 | 460.08 | 2173.86 | 4021.22 | 8403.75 | 16141.47 | 24140.21 | 31615.28 | 39137 | 180225.36 | 377853.73 | 583.44 | 133.35
66.9 | 3.33 | 506.77 | 2233.87 | 4287.99 | 8286.97 | 17316.14 | 24954.58 | 32891.15 | 41249.51 | 184824.44 | 393605.11 | 630.12 | 223.38
66.95 | 0 | 453.41 | 2210.56 | 4067.91 | 8520.45 | 16518.59 | 24935.04 | 33262.49 | 40611.88 | 187248.6 | 396073.25 | 663.46 | 103.35
67 | 0 | 450.08 | 2060.47 | 4197.95 | 8585.48 | 16732.37 | 25172.17 | 32363.56 | 40812.26 | 184798.32 | 389953.13 | 663.46 | 166.7
67.05 | 0 | 346.73 | 1883.76 | 3934.49 | 8286.97 | 15874.31 | 24383.93 | 31127.39 | 39495.13 | 179792.93 | 379692.28 | 856.84 | 136.69
67.1 | 13.34 | 453.41 | 1897.11 | 4084.51 | 7863.3 | 16041.3 | 23435.37 | 30242.29 | 38783.17 | 171795.02 | 363162 | 1146.88 | 140.02
67.15 | 13.34 | 426.74 | 1807.06 | 3454.26 | 6839.36 | 13842.04 | 20718.66 | 27112.61 | 34361.4 | 153628.3 | 326935.53 | 1320.29 | 136.69
67.2 | 10 | 283.38 | 1630.37 | 3174.2 | 5922.1 | 12317.13 | 18341.06 | 24317.15 | 30539.44 | 135208.75 | 287649.58 | 1416.96 | 126.69
67.25 | 3.33 | 13.34 | 100.02 | 216.7 | 446.75 | 950.19 | 1233.6 | 1857.1 | 2193.91 | 9682.25 | 20835.78 | 1253.61 | 26.67
67.3 | 0 | 0 | 0 | 6.67 | 13.34 | 23.34 | 23.34 | 33.34 | 36.68 | 140.02 | 316.72 | 1146.91 | 0
67.35 | 0 | 0 | 3.33 | 6.67 | 26.67 | 16.67 | 26.67 | 33.34 | 30 | 150.02 | 373.4 | 1206.92 | 0
67.4 | 0 | 3.33 | 10 | 13.34 | 20 | 53.34 | 63.34 | 60.01 | 113.35 | 516.76 | 1006.86 | 1313.6 | 0
67.45 | 6.67 | 6.67 | 10 | 13.34 | 16.67 | 66.68 | 43.34 | 110.02 | 176.7 | 693.48 | 71083.9 | 1343.62 | 0
67.5 | 6.67 | 40.01 | 170.03 | 236.71 | 606.79 | 1003.57 | 1377.02 | 2040.65 | 2214.02 | 9509.77 | 54304.16 | 1246.92 | 6.67
67.55 | 6.67 | 463.42 | 2013.83 | 3661.09 | 7466.55 | 13886.33 | 19851.46 | 26331.06 | 32960.24 | 143442.41 | 157805.37 | 1496.98 | 163.36
67.6 | 23.34 | 760.15 | 4247.98 | 8287.03 | 16555.16 | 33186.37 | 48691.98 | 63919.82 | 79842.63 | 362177.51 | 635838.67 | 1933.75 | 300.06
67.65 | 26.67 | 1240.25 | 6532.52 | 12290.36 | 24887.81 | 48915.68 | 73000.69 | 96057.64 | 120110.92 | 541476.29 | 1107452.8 | 2794.05 | 516.76
67.7 | 16.67 | 1650.37 | 7139.64 | 14769.77 | 30088.53 | 59123.63 | 88453.32 | 113081.56 | 144978.04 | 646317.5 | 1358009.61 | 2990.8 | 570.11
67.75 | 30.01 | 2010.48 | 8497.19 | 16655.34 | 34264.8 | 67522.65 | 98763.63 | 129142.86 | 162442.13 | 735298.72 | 1576884.24 | 2970.76 | 663.45
67.8 | 50.01 | 2010.45 | 9077.55 | 19141.76 | 38331.97 | 72802.66 | 109664.54 | 143726.69 | 181408.71 | 817791.56 | 1750154.38 | 2357.28 | 676.81
67.85 | 43.34 | 2197.2 | 10125.15 | 19722.7 | 40143.77 | 76993.47 | 115218.89 | 152319.29 | 190210.45 | 862363.89 | 1879472.39 | 2553.99 | 650.12
67.9 | 36.67 | 2017.1 | 9544.66 | 19752.9 | 40438.03 | 80091.02 | 119281.55 | 158408.56 | 196485.43 | 888263.36 | 1932413.79 | 2784.06 | 783.48
67.95 | 53.34 | 2313.93 | 10398.68 | 19996.56 | 42142.66 | 79099.25 | 120692.29 | 159304.32 | 198170.83 | 892194.15 | 1976465.03 | 2653.99 | 673.47
68 | 33.34 | 2340.54 | 9678.03 | 19963.29 | 40758.93 | 80878.61 | 118419.32 | 156616.87 | 197161.18 | 886798.98 | 1962594.78 | 2784.06 | 723.48
68.05 | 16.67 | 2217.2 | 9874.79 | 20063.38 | 40187.15 | 78486.35 | 117082.28 | 155095.5 | 192613.39 | 873197.71 | 1950824.34 | 2634 | 683.47
68.1 | 46.69 | 2187.17 | 9658.16 | 19746.5 | 38566.13 | 74875.89 | 113155.53 | 148106.56 | 184441.75 | 840071.85 | 1878082.54 | 3094.14 | 640.12
68.15 | 20 | 1937.11 | 8760.68 | 17753.34 | 34763 | 68024.86 | 103059.17 | 135212.69 | 171250.86 | 775195.94 | 1740888.27 | 3691.09 | 620.11
68.2 | 20 | 1703.7 | 8003.44 | 16289.89 | 31331.11 | 62310.04 | 93241.28 | 122613.52 | 153680.6 | 694591.63 | 1582396.65 | 3724.41 | 696.78
68.25 | 16.67 | 390.06 | 1997.12 | 4304.7 | 8784.14 | 17664.29 | 26914.08 | 36382.86 | 45003.74 | 198467.58 | 976279.3 | 3237.6 | 190.03
68.3 | 0 | 6.67 | 20 | 16.67 | 33.34 | 83.35 | 120.02 | 170.03 | 256.71 | 823.49 | 49498.34 | 1890.43 | 0
68.35 | 0 | 0 | 3.33 | 0 | 3.33 | 26.67 | 20 | 26.67 | 46.68 | 313.39 | 51901.21 | 2170.53 | 0
68.4 | 0 | 0 | 6.67 | 20 | 13.34 | 60.01 | 123.36 | 123.36 | 170.03 | 720.13 | 52447.85 | 2273.91 | 0
68.45 | 0 | 3.33 | 36.67 | 110.02 | 76.68 | 276.72 | 370.07 | 306.72 | 500.09 | 71006.39 | 52423.92 | 2250.54 | 3.33
68.5 | 0 | 66.68 | 220.04 | 636.79 | 1153.61 | 2047.27 | 2897.65 | 3664.85 | 4405.2 | 55084.39 | 60288.5 | 2230.55 | 10
68.55 | 16.67 | 1273.59 | 7279.68 | 14276.1 | 28051.68 | 54943.75 | 80214.13 | 104677.62 | 130193.42 | 332589.69 | 668326.94 | 4214.65 | 80.01
68.6 | 46.67 | 2517.32 | 11751.52 | 23799.63 | 46457.52 | 92991.19 | 137239 | 181432.93 | 228382.44 | 1064724.24 | 2148734.8 | 8283.67 | 143.36
68.65 | 23.34 | 2867.4 | 15183.56 | 29868.09 | 61268.55 | 120981.91 | 179491.5 | 235519.52 | 300390.74 | 1475600.75 | 2940918.4 | 11105.93 | 203.37
68.7 | 76.7 | 3624.36 | 18087.24 | 35858.65 | 71824.95 | 141830.06 | 211647.19 | 276613.98 | 350755.85 | 1769039.31 | 3555692.58 | 13418.15 | 230.04
68.75 | 76.68 | 4081.19 | 19549.32 | 38710.32 | 76541.99 | 152514.19 | 227869.97 | 299765.48 | 376648.21 | 1946522.19 | 3878361.28 | 13965.63 | 263.38
68.8 | 30.01 | 3991.27 | 20039.89 | 41694.65 | 82107.3 | 163961.05 | 245169.17 | 320385.86 | 404620.81 | 2093577.6 | 4162562.25 | 14539.44 | 243.4
68.85 | 86.71 | 4404.78 | 21652.69 | 44116.18 | 87946.44 | 178352.41 | 264299.43 | 347306.67 | 435654.63 | 2267741.27 | 4521998.64 | 14803.2 | 363.42
68.9 | 100.04 | 4261.31 | 22674.47 | 45232.66 | 91846.39 | 180641.91 | 270646.68 | 357309.9 | 451805.7 | 2346143.18 | 4666851.88 | 15777.61 | 353.4
68.95 | 50.01 | 4281.33 | 22500.72 | 44784.17 | 90817.35 | 179925.45 | 267555.52 | 351694.9 | 446464.74 | 2364069.31 | 4691264.74 | 15707.61 | 310.05
69 | 23.34 | 4601.47 | 22183.7 | 44735.92 | 90820.66 | 179259.22 | 270388.01 | 354417.03 | 445559.81 | 2340859.35 | 4717877.08 | 16381.53 | 256.71
69.05 | 60.02 | 4384.72 | 21826.42 | 43790.94 | 86756.1 | 174259.48 | 259493.77 | 341636.99 | 429439.93 | 2288978.22 | 4530097.45 | 16251.69 | 276.71
69.1 | 30 | 4154.6 | 21168.8 | 41327.28 | 83163.91 | 165677.38 | 246336.69 | 327110.78 | 416578.52 | 2227274.11 | 4406122.7 | 15020.06 | 276.72
69.15 | 33.34 | 3557.73 | 18574.54 | 36998.14 | 74983.86 | 147951.57 | 222992.58 | 291767.59 | 366323.78 | 2006932.08 | 4015757.09 | 14015.42 | 333.41
69.2 | 73.38 | 3364.27 | 15951.24 | 32430.98 | 63896.24 | 128990.6 | 192903.07 | 252590.56 | 320202.02 | 1787336.26 | 3570334.3 | 12097 | 220.05
69.25 | 3.33 | 120.02 | 396.73 | 980.2 | 1777.1 | 3644.62 | 5939.37 | 8205 | 10176.76 | 444731.36 | 836560.01 | 6052.34 | 13.34
69.3 | 0 | 0 | 0 | 16.67 | 20 | 66.68 | 70.01 | 93.35 | 163.36 | 51945.4 | 51217.47 | 23.34 | 0
69.35 | 0 | 6.67 | 3.33 | 10 | 3.33 | 6.67 | 16.67 | 16.67 | 20 | 51173.74 | 51684.42 | 6.67 | 0
69.4 | 0 | 0 | 3.33 | 3.33 | 10 | 16.67 | 23.34 | 10 | 43.34 | 50365.4 | 52086.47 | 6.67 | 0
69.45 | 3.33 | 3.33 | 6.67 | 60.01 | 60.01 | 180.03 | 200.03 | 346.73 | 350.06 | 2013.86 | 53184.95 | 6.67 | 0
69.5 | 0 | 43.34 | 360.07 | 460.08 | 980.22 | 1897.22 | 2374.14 | 3077.7 | 3911.72 | 19555.42 | 57728.86 | 50.01 | 3.33
69.55 | 6.67 | 340.06 | 1783.72 | 3627.73 | 6709.32 | 13285.1 | 19082.9 | 24963.63 | 32369.9 | 159317.92 | 194687.79 | 436.75 | 23.34
69.6 | 6.67 | 636.78 | 3507.71 | 7032.8 | 13888.78 | 27326.22 | 41160.02 | 54686.8 | 69195.35 | 341497.91 | 627295.23 | 1867.09 | 43.34
69.65 | 6.67 | 900.18 | 4688.21 | 9311.16 | 18741.65 | 37753.93 | 54297.55 | 71850.16 | 90554.94 | 449959.52 | 898994.02 | 3074.16 | 36.67
69.7 | 10 | 1090.22 | 5268.51 | 10562.13 | 20657.58 | 41089.72 | 61355.75 | 80864.06 | 102124.8 | 515258.76 | 1056311.92 | 3324.27 | 63.34
69.75 | 16.67 | 1160.24 | 6152.22 | 12023.65 | 23899.79 | 47449.73 | 71666.15 | 94605.61 | 117912.87 | 599726.12 | 1233322.49 | 4051.21 | 100.02
69.8 | 33.34 | 1380.28 | 6402.5 | 12447.27 | 25906.67 | 51071.74 | 75234.82 | 100929.29 | 127435.27 | 639504.87 | 1337822.62 | 4378.08 | 73.34
69.85 | 36.67 | 1153.57 | 6879.4 | 13521.66 | 26848.45 | 52888.65 | 79431.76 | 104980.28 | 129684.11 | 665881.51 | 1406990.67 | 5141.81 | 80.01
69.9 | 16.67 | 1406.97 | 6846.01 | 13435.11 | 27125.35 | 53551.74 | 79904.22 | 107735.59 | 133720.43 | 679862.07 | 1434753.01 | 4808.26 | 73.34
69.95 | 6.67 | 1310.27 | 6969.42 | 13778.67 | 27582.97 | 54170.86 | 81565.46 | 107972.55 | 133350.86 | 676014.92 | 1457767.98 | 5161.76 | 63.34
70 | 13.34 | 1213.57 | 6986.15 | 13661.87 | 27246.02 | 54817.18 | 81297.51 | 107054.91 | 134320.37 | 682151.58 | 1458330.57 | 5392.06 | 93.35
70.05 | 10 | 1310.26 | 6682.53 | 12650.7 | 26263.92 | 52443.61 | 79102.32 | 103345.23 | 131440.2 | 650923.75 | 1396968.68 | 4784.91 | 80.01
70.1 | 16.67 | 1320.29 | 5955.42 | 12824.2 | 24433.86 | 49988.88 | 73639.12 | 97739.79 | 124159.57 | 617059.33 | 1347302.35 | 4628.25 | 80.01
70.15 | 33.34 | 960.19 | 5341.87 | 11476.12 | 22764.53 | 44774.7 | 66602.47 | 88864.38 | 112767.83 | 560612.54 | 1240515.61 | 4124.58 | 30
70.2 | 20 | 873.5 | 4921.63 | 10008.44 | 19556.1 | 38419.57 | 57882.05 | 76758.81 | 95838.63 | 481923.72 | 1112980.72 | 4021.26 | 33.34
70.25 | 0 | 50.01 | 210.04 | 393.4 | 923.53 | 1703.77 | 2847.5 | 3627.86 | 4358.32 | 23366.65 | 312595.11 | 2107.19 | 3.33
70.3 | 0 | 3.33 | 3.33 | 0 | 6.67 | 26.67 | 46.68 | 30 | 53.34 | 286.71 | 51479.39 | 30.01 | 0
70.35 | 0 | 0 | 0 | 10 | 3.33 | 10 | 56.68 | 43.34 | 36.67 | 230.04 | 51058.77 | 3.33 | 0
70.4 | 0 | 3.33 | 0 | 20 | 6.67 | 80.02 | 100.02 | 73.34 | 120.02 | 780.17 | 51218.8 | 23.34 | 0
70.45 | 0 | 0 | 33.34 | 73.35 | 73.34 | 226.71 | 336.72 | 446.75 | 510.1 | 67384.57 | 52813.95 | 10 | 0
70.5 | 0 | 33.34 | 150.02 | 343.4 | 683.48 | 1326.99 | 1870.54 | 2080.63 | 2770.97 | 54608.31 | 56915.34 | 33.34 | 0
70.55 | 3.33 | 610.11 | 2573.99 | 5265.33 | 9745.4 | 18710.34 | 27026.4 | 34951.28 | 43375.45 | 121376.34 | 201809.56 | 210.03 | 20
70.6 | 13.34 | 1433.65 | 7713.34 | 15512.48 | 29975.5 | 59265.79 | 88399.96 | 115053.58 | 144766.6 | 621777.12 | 1263039.68 | 1463.67 | 106.69
70.65 | 13.33 | 2010.47 | 10305.37 | 19949.77 | 40281.04 | 78901.17 | 118607.34 | 156324 | 196785.8 | 969997.19 | 1952806.18 | 2440.64 | 133.36 | Solution 2
70.7 | 33.34 | 2227.19 | 12614.03 | 24557.58 | 48570.82 | 97608.58 | 143864.99 | 188545.8 | 241407.73 | 1197787.42 | 2417121.82 | 3250.99 | 146.69 | 70.55/71
70.75 | 60.02 | 2937.41 | 13912.15 | 29537.1 | 58557.34 | 116373.3 | 173686.19 | 228190.83 | 284879.21 | 1460930.02 | 2917650.61 | 4394.83 | 190.04 | 0.039294
70.8 | 46.68 | 2950.77 | 15240.27 | 30796.83 | 61717.12 | 122732.01 | 184571.11 | 239484.91 | 302999.77 | 1592861.51 | 3169322.36 | 4464.78 | 190.03
70.85 | 30.01 | 3354.28 | 15957.86 | 32009.56 | 65553.95 | 130751.22 | 192140.68 | 256950.76 | 322297.78 | 1670488.44 | 3342489.89 | 4658.25 | 196.7
70.9 | 36.67 | 3317.6 | 17012.7 | 34024.87 | 67724.73 | 134476.14 | 201412.22 | 262003.59 | 334295.87 | 1745175.47 | 3488014.19 | 5071.82 | 240.04
70.95 | 40.01 | 3264.22 | 16902.46 | 34305.34 | 66200.58 | 134273.09 | 201782.05 | 266321.24 | 334783.18 | 1774559.68 | 3529408.28 | 5118.43 | 193.36
71 | 33.34 | 3397.58 | 16368.53 | 33887.16 | 67583.93 | 134715.86 | 201526.16 | 265917.6 | 338952.55 | 1767621.72 | 3564916.25 | 5345.13 | 216.71
71.05 | 16.67 | 3034.19 | 16585.18 | 33319.81 | 66478.83 | 131812.35 | 197542.7 | 258903.86 | 328928.52 | 1731763.75 | 3466303.59 | 5451.92 | 233.38
71.1 | 33.34 | 3230.9 | 15207.13 | 32323.52 | 63142.93 | 125734.27 | 188098.41 | 252114.92 | 314922.48 | 1687319.35 | 3383876.85 | 5138.43 | 300.07
71.15 | 13.34 | 3054.2 | 14409.13 | 28879.21 | 58803.23 | 116373.25 | 172254.01 | 228407.85 | 289017.12 | 1569033.62 | 3145203.56 | 5742.16 | 206.72
71.2 | 16.67 | 2580.66 | 13228.04 | 26047.25 | 51925.7 | 104167.17 | 156144.97 | 205376.34 | 258496.42 | 1416147.38 | 2840068.53 | 5718.8 | 226.71
71.25 | 3.33 | 793.48 | 4368.11 | 8178.71 | 17753.98 | 35968.96 | 54802.87 | 73406.69 | 92471.06 | 968131.84 | 1890382.03 | 5141.77 | 43.34
71.3 | 0 | 0 | 16.67 | 36.67 | 76.68 | 90.02 | 136.69 | 243.38 | 333.39 | 50508.7 | 53383.2 | 1807.06 | 0
71.35 | 0 | 0 | 0 | 6.67 | 16.67 | 13.34 | 10 | 20 | 50.01 | 50195.56 | 50817.13 | 1957.13 | 0
71.4 | 0 | 0 | 0 | 16.67 | 13.34 | 36.67 | 53.34 | 86.68 | 86.68 | 51154.36 | 52075.37 | 1827.09 | 6.67
71.45 | 0 | 0 | 3.33 | 16.67 | 56.68 | 103.35 | 120.02 | 160.03 | 143.35 | 983.52 | 52325.22 | 1973.8 | 0
71.5 | 0 | 6.67 | 70.01 | 126.69 | 316.73 | 573.45 | 760.15 | 990.22 | 1103.57 | 5972.83 | 54254.15 | 2033.83 | 0
71.55 | 0 | 243.38 | 936.85 | 1957.16 | 3831.35 | 7427 | 10709.99 | 13219.86 | 17639.59 | 90511.01 | 92276.78 | 1940.48 | 13.34
71.6 | 6.67 | 726.8 | 3324.28 | 6869.46 | 13111.3 | 26194.42 | 38370.32 | 49599.24 | 63343.17 | 310290.63 | 498306.52 | 2687.37 | 106.68
71.65 | 16.67 | 1463.65 | 6489.03 | 12573.84 | 25419.34 | 50246.7 | 73901.77 | 98301.37 | 123019.63 | 622654.06 | 1202471.71 | 4261.31 | 250.04
71.7 | 33.34 | 1507 | 7596.44 | 15313.82 | 30820.35 | 59939.66 | 89573.23 | 116416.43 | 147817.43 | 743196.98 | 1517027.59 | 4788.3 | 360.06
71.75 | 3.33 | 1530.34 | 8036.79 | 16638.53 | 32290.47 | 64198.92 | 96806.27 | 129109.41 | 161959.88 | 817142 | 1692551.93 | 4898.36 | 386.73
71.8 | 3.33 | 1660.37 | 8667.25 | 17060.96 | 34682.04 | 70371.26 | 104871.69 | 136597.71 | 174794.4 | 874867.67 | 1826535.21 | 3854.49 | 440.09
71.85 | 16.67 | 1860.44 | 8964.13 | 18574.65 | 36547.3 | 73341.76 | 109815.33 | 143441.94 | 180859.27 | 915666.04 | 1914030.39 | 3904.53 | 476.75
71.9 | 6.67 | 1973.79 | 9651.34 | 18831.66 | 37336.09 | 75664.21 | 112333.04 | 150397.38 | 187538.06 | 939192.05 | 1996664.78 | 4201.33 | 450.08 | 0 | 0.00034974 | 0.00085272 | 0.0014469 | 0.0013378 | 0.001043 | 0.0010451 | 0.00074749
71.95 | 20 | 2010.46 | 9938.21 | 19699.45 | 38121.44 | 76659.77 | 115572.12 | 152229.97 | 190807.7 | 972218.32 | 2045267.17 | 4651.56 | 580.11
72 | 10 | 1873.78 | 9521.27 | 19549.28 | 39174.46 | 77254.62 | 115068.48 | 153127.72 | 191774.69 | 973794.55 | 2068479.12 | 4608.18 | 513.43 | 0.12989 | 0.098395 | 0.10011 | 0.097802 | 0.096137 | 0.093075 | 0.086332 | 0.091981
72.05 | 6.67 | 1870.43 | 9467.89 | 19118.81 | 37739.84 | 75234.55 | 112048.26 | 148284.4 | 185758.51 | 941823.44 | 2004364.97 | 4424.77 | 436.75
72.1 | 16.67 | 1883.77 | 8877.38 | 18157.52 | 36259.95 | 71948.38 | 108303.14 | 142997.44 | 179173.39 | 911479.64 | 1952287.47 | 4474.81 | 493.43
72.15 | 16.67 | 1783.75 | 8123.57 | 15937.81 | 33563.09 | 66682.56 | 99514.49 | 133243.64 | 164818.56 | 839934.22 | 1814319.1 | 4701.56 | 490.09
72.2 | 0 | 1446.97 | 7393.14 | 15313.67 | 30145.49 | 60121.64 | 90347.82 | 118001.14 | 148180.57 | 759905.08 | 1643377.39 | 4694.88 | 366.73
72.25 | 0 | 233.37 | 970.19 | 2003.83 | 4394.85 | 8711.09 | 13152.11 | 18576.49 | 22262.72 | 116527.26 | 879638.29 | 3784.44 | 266.71
72.3 | 0 | 6.67 | 20.01 | 10 | 26.67 | 26.67 | 73.34 | 93.35 | 120.02 | 616.79 | 49427.18 | 1406.98 | 10
72.35 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 13.34 | 10 | 13.33 | 60.01 | 50517.08 | 1400.31 | 0
72.4 | 0 | 0 | 3.33 | 0 | 0 | 0 | 6.67 | 20 | 26.67 | 80.01 | 51222.89 | 1537.02 | 6.67
72.45 | 0 | 3.33 | 0 | 20 | 3.33 | 16.67 | 46.68 | 46.68 | 66.68 | 213.37 | 503.42 | 1390.3 | 0
72.5 | 0 | 6.67 | 33.34 | 30 | 106.68 | 130.02 | 186.7 | 240.05 | 273.38 | 1260.29 | 3197.76 | 1440.33 | 0
72.55 | 0 | 110.02 | 453.42 | 946.84 | 1767.06 | 3257.66 | 5121.88 | 6465.97 | 7977.03 | 42626.56 | 83203.32 | 1470.37 | 3.33
72.6 | 10 | 360.06 | 1166.91 | 2567.34 | 5238.38 | 9784.8 | 14729.79 | 19913.58 | 24724.87 | 125750.43 | 254189.49 | 2327.25 | 13.34
72.65 | 0 | 330.06 | 1853.8 | 3681.07 | 7149.51 | 14065.65 | 21161.85 | 28612.1 | 34859.66 | 177180.71 | 358775.13 | 3084.21 | 16.67
72.7 | 0 | 340.06 | 2187.18 | 4017.9 | 8000.14 | 16533.55 | 24621.21 | 32189.82 | 41986.21 | 208921.85 | 420221.01 | 3270.95 | 30.01
72.75 | 6.67 | 513.42 | 2487.27 | 4874.9 | 9741.48 | 19282.4 | 27667.15 | 36962.27 | 47281.93 | 238095.85 | 478658.87 | 3641.08 | 16.67
72.8 | 0 | 526.76 | 2827.4 | 5368.58 | 10201.93 | 20523.94 | 30649.88 | 39994.12 | 49527.13 | 255207.02 | 521273.6 | 2533.97 | 13.34
72.85 | 6.67 | 646.78 | 2904.11 | 5642.04 | 11199.27 | 22974.83 | 33218.92 | 44085.97 | 55612.09 | 279012.85 | 560738.51 | 3050.86 | 30.01
72.9 | 6.67 | 650.12 | 2924.1 | 5588.57 | 11122.67 | 23108.38 | 33723.43 | 45182.58 | 57101.88 | 285136.2 | 575556.46 | 3207.56 | 33.34
72.95 | 3.33 | 530.09 | 2947.49 | 6165.66 | 11889.88 | 23151.69 | 34646.1 | 45179.16 | 57040.75 | 293392.62 | 583966.07 | 3370.97 | 26.67
73 | 6.67 | 503.43 | 2880.76 | 5795.42 | 11422.87 | 22891.68 | 34809.7 | 45760.96 | 57232.17 | 291186.66 | 587551.31 | 3280.95 | 33.34
73.05 | 0 | 606.78 | 2700.66 | 5685.46 | 11583.02 | 22417.15 | 33539.74 | 44396.73 | 55228.24 | 279707.29 | 567567.43 | 3037.45 | 40.01
73.1 | 3.33 | 563.44 | 2664.06 | 5398.54 | 10995.82 | 21405.41 | 32460.77 | 42018.71 | 53182.78 | 274592.12 | 550562.74 | 3127.51 | 346.74
73.15 | 0 | 536.76 | 2457.31 | 4924.94 | 9833.06 | 20016.69 | 29146.63 | 38572.48 | 48619.77 | 248638.01 | 501874.57 | 3004.18 | 1120.23
73.2 | 6.67 | 440.08 | 2323.91 | 4458 | 9204.33 | 17423.05 | 26331.12 | 35211.1 | 43407.28 | 221099.65 | 445975.37 | 2490.64 | 1467.01
73.25 | 0 | 146.7 | 503.43 | 1053.56 | 2137.23 | 4288.02 | 6872.9 | 8610.76 | 10792.82 | 54284.9 | 109182.53 | 2093.86 | 1797.08
73.3 | 0 | 0 | 0 | 6.67 | 16.67 | 23.34 | 13.33 | 26.67 | 63.34 | 313.39 | 470.09 | 63.34 | 1913.78
73.35 | 0 | 0 | 3.33 | 0 | 0 | 10 | 23.34 | 6.67 | 23.34 | 140.03 | 333.4 | 0 | 1930.48
73.4 | 0 | 0 | 0 | 3.33 | 3.33 | 30.01 | 20 | 36.67 | 73.35 | 256.71 | 563.44 | 0 | 1997.13
73.45 | 0 | 6.67 | 26.67 | 36.67 | 130.02 | 233.37 | 320.06 | 406.74 | 496.76 | 69426.61 | 70684.7 | 0 | 2397.33
73.5 | 0 | 36.67 | 176.7 | 296.72 | 656.81 | 1260.32 | 1873.88 | 2264.09 | 2927.71 | 54586.6 | 56381.37 | 10 | 2250.59
73.55 | 13.34 | 446.76 | 2130.51 | 4294.81 | 7860.45 | 14817.63 | 21836.15 | 27921.85 | 35949.71 | 98806.17 | 165565.45 | 160.03 | 2307.25
73.6 | 3.33 | 1210.23 | 6175.68 | 13057.99 | 25763.43 | 49511.52 | 75141.78 | 97315.02 | 121629.25 | 512035.22 | 1066730.54 | 623.45 | 2644.06
73.65 | 13.34 | 1937.12 | 8507.15 | 17753.51 | 33580.32 | 67620.89 | 100811.82 | 132889.75 | 167236.09 | 821053.95 | 1646889.85 | 996.88 | 2907.48
73.7 | 16.67 | 2093.82 | 10051.72 | 21609.04 | 41865.57 | 83567.25 | 122989.24 | 162772.77 | 205574.84 | 1029334.27 | 2082007.64 | 1273.6 | 3034.21
73.75 | 16.67 | 2300.54 | 12560.68 | 25282.05 | 50510.3 | 97826.59 | 146779.78 | 192717.05 | 244869.45 | 1250473.23 | 2494628.65 | 1713.71 | 3167.58 | Solution 2
73.8 | 46.67 | 2627.36 | 12580.77 | 26070.38 | 52928.84 | 106696.7 | 157253.34 | 205531.83 | 261128.17 | 1369923.21 | 2736475.76 | 1710.39 | 1813.76 | 73.55/74
73.85 | 13.34 | 2934.12 | 13962.06 | 27222.6 | 54076.62 | 109264.89 | 163473.99 | 213289.38 | 267251.41 | 1415464.64 | 2863582.88 | 1847.09 | 1747.06 | 0.079334
73.9 | 30 | 2990.83 | 14192.38 | 28388.35 | 56228.01 | 112646.68 | 167482.23 | 221532.26 | 277695.49 | 1479285.87 | 2984588.56 | 1933.82 | 1813.77
73.95 | 23.34 | 2814.07 | 14476.07 | 28742.21 | 57059.27 | 113656.3 | 170117.54 | 224448.61 | 283463.17 | 1504140.01 | 3009796.53 | 1923.83 | 1753.74
74 | 30.01 | 2517.29 | 14349.28 | 28832.65 | 56653.24 | 115023.9 | 169755.1 | 224421.05 | 279889.51 | 1500044.98 | 3013541.06 | 2017.18 | 1767.07
74.05 | 23.34 | 2850.76 | 14202.35 | 28345.22 | 55204.62 | 111710.86 | 166178.51 | 218117.5 | 273296.46 | 1471543.67 | 2948898.3 | 1860.45 | 1910.46
74.1 | 26.67 | 2647.38 | 13201.22 | 27403.24 | 53775.65 | 106986.77 | 159902.05 | 211145.46 | 265597.2 | 1432817.06 | 2888419.49 | 1770.39 | 2377.25
74.15 | 33.34 | 2417.25 | 12387.28 | 25032.04 | 47800.85 | 97752.51 | 146171.06 | 192670.74 | 244434.21 | 1333264.91 | 2671343.97 | 1617.05 | 2533.99
74.2 | 36.67 | 2273.89 | 11192.69 | 22243.19 | 44313.06 | 88960.22 | 132590.26 | 174440.35 | 217604.46 | 1210480.5 | 2427526.8 | 1513.67 | 2860.79
74.25 | 13.33 | 520.09 | 2990.8 | 6425.79 | 12430.97 | 25354.04 | 38880.93 | 51999.93 | 66940.43 | 752480.08 | 1512605.52 | 1056.89 | 2944.16
74.3 | 0 | 0 | 13.33 | 26.67 | 63.34 | 90.01 | 110.02 | 240.04 | 233.38 | 47905.57 | 52779.2 | 56.68 | 1773.73
74.35 | 3.33 | 0 | 0 | 3.33 | 0 | 13.34 | 26.67 | 23.34 | 36.67 | 51160.16 | 51937.62 | 13.34 | 1890.44
74.4 | 0 | 0 | 0 | 10 | 13.34 | 10 | 23.34 | 40.01 | 46.67 | 50537.34 | 49815.99 | 0 | 1963.8
74.45 | 0 | 6.67 | 13.34 | 20 | 23.34 | 70.01 | 60.01 | 100.02 | 143.36 | 690.14 | 17375.63 | 0 | 2377.27
74.5 | 0 | 6.67 | 40.01 | 83.35 | 176.7 | 303.39 | 390.08 | 513.44 | 673.47 | 2904.32 | 20570.98 | 0 | 2167.17
74.55 | 0 | 106.68 | 650.13 | 1323.64 | 2507.31 | 4604.99 | 6713 | 8878.12 | 11444.18 | 53975.93 | 110224.23 | 33.34 | 2303.93
74.6 | 13.33 | 500.09 | 2420.62 | 4831.57 | 9404.44 | 18805.36 | 27253.54 | 35532.78 | 44806.34 | 224987.44 | 451440.35 | 283.38 | 2280.59
74.65 | 33.34 | 780.15 | 3807.76 | 7753.34 | 15340.44 | 31641.9 | 46035.5 | 60279.11 | 75380.61 | 376800.19 | 783836.76 | 846.83 | 2824.1
74.7 | 66.68 | 996.86 | 4714.91 | 9581.36 | 18524.42 | 36801.48 | 54803.55 | 72592.74 | 90576.43 | 448194.61 | 940879.75 | 896.86 | 2747.43
74.75 | 30 | 1070.21 | 4921.69 | 10715.46 | 20851.67 | 41150.06 | 61775.15 | 79860.92 | 101165.31 | 502514.18 | 1064517.76 | 1120.22 | 2704.06
74.8 | 53.34 | 1206.94 | 5755.42 | 10765.64 | 21792.64 | 43624.8 | 64559.57 | 83157.58 | 107184.04 | 532306.52 | 1118097.78 | 1136.91 | 1203.58
74.85 | 56.68 | 1183.56 | 5875.48 | 11503.17 | 22804.55 | 46824.3 | 68946.67 | 89413.12 | 112976.07 | 559737.51 | 1189779.65 | 1223.59 | 1340.3
74.9 | 56.68 | 1226.91 | 5805.39 | 12180.32 | 23485.58 | 46964.73 | 71054.81 | 93612.14 | 117672.02 | 580860.41 | 1234164.48 | 1206.93 | 1316.96
74.95 | 33.34 | 1343.62 | 5995.57 | 12250.14 | 24577.57 | 48504.33 | 71858.44 | 94456.75 | 119488.38 | 593301.05 | 1290775.04 | 1510.34 | 1460.33
75 | 33.34 | 1303.59 | 6409.09 | 12025.26 | 24230 | 49119.32 | 71737.78 | 96018.53 | 120669.81 | 589861.46 | 1298202.44 | 1293.62 | 1467
75.05 | 43.34 | 1223.58 | 6188.98 | 11669.74 | 23569.17 | 47106.06 | 70542.41 | 93147.52 | 115877.12 | 574077 | 1258949.96 | 1326.97 | 1426.97
75.1 | 73.35 | 1213.6 | 5572.04 | 11235.95 | 23275.87 | 45252.54 | 68110.18 | 88580.09 | 111427.95 | 555756.7 | 1215418.46 | 1376.99 | 1190.27
75.15 | 40.01 | 1143.57 | 5465.2 | 11182.66 | 21148.62 | 41618.55 | 62709.1 | 81465.3 | 102998.1 | 511489.46 | 1149595.96 | 1190.25 | 1420.31
75.2 | 26.67 | 980.21 | 4541.42 | 8810.72 | 18110.56 | 37045.14 | 55011.77 | 72585.5 | 90172.9 | 447517.42 | 1013302.54 | 1073.57 | 1210.25
75.25 | 0 | 140.02 | 680.13 | 1190.25 | 2163.87 | 5101.89 | 7633.68 | 10209.15 | 11927.42 | 59174.87 | 399810.2 | 896.85 | 976.87
75.3 | 0 | 3.33 | 0 | 6.67 | 13.34 | 16.67 | 36.67 | 70.01 | 60.01 | 356.73 | 46048.63 | 33.34 | 13.34
75.35 | 0 | 0 | 0 | 10 | 0 | 0 | 13.34 | 10 | 6.67 | 73.34 | 45827.71 | 0 | 0
75.4 | 0 | 0 | 6.67 | 0 | 6.67 | 10 | 26.67 | 16.67 | 6.67 | 80.01 | 45375.06 | 0 | 3.33
75.45 | 0 | 3.33 | 6.67 | 0 | 13.34 | 26.67 | 23.34 | 50.01 | 40.01 | 256.71 | 483.43 | 0 | 3.33
75.5 | 3.33 | 26.67 | 73.35 | 153.37 | 286.71 | 503.43 | 620.12 | 813.5 | 1106.91 | 5246.14 | 10810.01 | 3.33 | 6.67
75.55 | 50.01 | 156.69 | 756.82 | 1523.65 | 2930.8 | 5838.96 | 8384.02 | 10642.69 | 13045.13 | 69221.49 | 139989.78 | 96.68 | 163.36
75.6 | 73.35 | 346.73 | 1773.71 | 3334.31 | 6219.01 | 12404.09 | 18491.12 | 24520.78 | 30961.18 | 156215.75 | 316161.56 | 236.71 | 846.84
75.65 | 93.35 | 563.44 | 2587.33 | 4461.4 | 8724.01 | 17696.89 | 26134.08 | 35327.56 | 43594.88 | 220356.8 | 448648.5 | 243.38 | 2143.85
75.7 | 120.02 | 616.78 | 3064.14 | 5735.36 | 10832.42 | 22080.09 | 32179.9 | 41892.38 | 52721.51 | 265344.55 | 544854.11 | 256.71 | 3014.18
75.75 | 83.35 | 810.16 | 2894.09 | 5705.34 | 12337.21 | 23709.16 | 35408.09 | 47273.08 | 58316.76 | 297665.16 | 599957.67 | 266.71 | 3561.05
75.8 | 140.02 | 890.17 | 3440.93 | 6859.36 | 12992.87 | 26818.39 | 40194.56 | 52330.57 | 65945.76 | 333538.96 | 677381.04 | 393.41 | 4064.6
75.85 | 123.36 | 683.45 | 3531 | 7182.88 | 13675.31 | 28318.13 | 42451.07 | 56054.64 | 69429.81 | 353477.39 | 716672.89 | 356.73 | 4544.86
75.9 | 113.35 | 760.14 | 3731.07 | 7282.94 | 14973.29 | 28221.72 | 42303.03 | 57774.49 | 71711.92 | 358737.32 | 726187.71 | 396.74 | 4488.14
75.95 | 136.69 | 783.47 | 3651.05 | 7319.65 | 14603.05 | 29126.49 | 43310.23 | 57256.46 | 72378.78 | 361205.76 | 734791.7 | 426.75 | 4858.32
76 | 150.03 | 783.48 | 3701.04 | 7326.29 | 14696.35 | 29303.58 | 43012 | 58076.09 | 72077.02 | 363112.81 | 736924.28 | 406.74 | 4871.67
76.05 | 63.34 | 820.16 | 3374.23 | 7803.3 | 13955.53 | 27874.18 | 41805.55 | 55673.4 | 69760.88 | 352820.71 | 708360.82 | 443.42 | 4895.06
76.1 | 123.35 | 833.5 | 3607.69 | 6865.96 | 13334.75 | 27470.28 | 40224.06 | 53022.81 | 66934.23 | 339102.91 | 685663.37 | 380.07 | 4998.37
76.15 | 113.35 | 823.5 | 3274.22 | 6449.1 | 12430.67 | 24601.06 | 36717.99 | 48864.56 | 61517 | 313098.28 | 627417.18 | 343.4 | 4328.1
76.2 | 80.01 | 533.43 | 2713.98 | 5545.24 | 11332.8 | 21829.84 | 33299.36 | 44239.63 | 55516.18 | 274609.2 | 561489.2 | 303.38 | 3884.53
76.25 | 23.34 | 113.35 | 660.13 | 1126.92 | 2383.93 | 4774.97 | 7393.23 | 10549.16 | 12831.55 | 62506.48 | 126485.06 | 86.68 | 3487.72
76.3 | 0 | 0 | 0 | 3.33 | 23.34 | 20 | 56.68 | 30 | 36.67 | 310.05 | 606.78 | 0 | 286.72
76.35 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 3.33 | 0 | 40.01 | 83.35 | 0 | 10
76.4 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 13.34 | 0 | 6.67
76.45 | 0 | 0 | 3.33 | 6.67 | 0 | 0 | 0 | 0 | 6.67 | 36.67 | 50.01 | 0 | 0
76.5 | 0 | 0 | 10 | 3.33 | 30.01 | 26.67 | 30.01 | 40.01 | 30.01 | 206.7 | 363.41 | 0 | 10
76.55 | 0 | 10 | 20 | 60.01 | 100.02 | 213.37 | 250.04 | 196.7 | 253.37 | 1490.35 | 3741.22 | 6.67 | 33.34
76.6 | 30.01 | 23.34 | 63.34 | 150.03 | 370.06 | 636.8 | 1013.54 | 1136.91 | 1476.98 | 7286.32 | 15958.19 | 66.68 | 536.77
76.65 | 20 | 40.01 | 76.68 | 300.06 | 533.43 | 1093.56 | 1590.33 | 1857.06 | 2457.3 | 11693.29 | 25719.47 | 76.68 | 2037.14
76.7 | 26.67 | 66.68 | 206.7 | 330.06 | 676.8 | 1336.93 | 1873.77 | 2623.98 | 3344.26 | 14913.08 | 33479.11 | 140.02 | 3050.84 | Solution1
76.75 | 23.34 | 100.02 | 233.37 | 326.72 | 816.82 | 1577.01 | 2067.18 | 2920.76 | 3647.7 | 17423.51 | 37546.08 | 150.03 | 3594.43 | 76.55/77
76.8 | 20 | 93.35 | 213.37 | 370.06 | 906.83 | 1533.68 | 2317.25 | 2854.06 | 3884.45 | 18845.08 | 40614.52 | 196.7 | 4041.27 | 0.0074284
76.85 | 30 | 63.34 | 286.72 | 570.11 | 780.15 | 1553.68 | 2197.16 | 3157.49 | 3891.23 | 19165.32 | 42356.37 | 143.36 | 4274.74
76.9 | 46.68 | 73.34 | 196.7 | 496.76 | 733.48 | 1780.42 | 2540.65 | 2987.43 | 3994.52 | 19619.18 | 43439.9 | 176.7 | 4411.4
76.95 | 43.34 | 36.67 | 213.37 | 560.11 | 820.15 | 1777.08 | 2664.03 | 3360.94 | 4091.28 | 20781.15 | 44887.72 | 160.03 | 4804.96
77 | 20 | 80.01 | 226.71 | 400.07 | 753.47 | 1803.72 | 2320.57 | 3324.29 | 4014.55 | 20463.97 | 43918.11 | 183.37 | 4488.17
77.05 | 13.33 | 73.35 | 226.7 | 460.09 | 900.18 | 1620.36 | 2320.56 | 3404.27 | 4024.53 | 20357.28 | 43686.89 | 93.35 | 4588.19
77.1 | 43.34 | 63.34 | 253.37 | 423.41 | 856.84 | 1553.65 | 2537.32 | 3140.82 | 3921.19 | 19379.32 | 41504.07 | 140.02 | 4975.02
77.15 | 20 | 46.67 | 300.05 | 393.4 | 750.14 | 1553.68 | 2337.26 | 2987.49 | 3721.13 | 17379.69 | 38438.67 | 136.69 | 5258.61
77.2 | 16.67 | 73.35 | 153.36 | 316.72 | 673.46 | 1386.96 | 1920.47 | 2730.73 | 3450.99 | 15083.34 | 35100.05 | 96.68 | 4885
77.25 | 10 | 20 | 46.67 | 100.01 | 183.37 | 313.39 | 543.43 | 773.49 | 910.17 | 4421.37 | 9401.44 | 33.34 | 4738.33
77.3 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 13.34 | 23.34 | 0 | 1870.44
77.35 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 0 | 3.33 | 6.67 | 0 | 1897.16
77.4 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 3.33 | 23.34 | 20 | 6.67 | 1833.77
77.45 | 0 | 0 | 0 | 3.33 | 0 | 10 | 0 | 3.33 | 10 | 116.69 | 166.7 | 0 | 1973.8
77.5 | 0 | 0 | 3.33 | 26.67 | 30.01 | 53.34 | 56.68 | 66.68 | 80.01 | 390.08 | 813.54 | 6.67 | 1960.47
77.55 | 23.34 | 10 | 76.68 | 170.03 | 280.05 | 506.76 | 630.12 | 856.83 | 1230.27 | 5085.37 | 11567.94 | 70.01 | 2017.15
77.6 | 43.34 | 96.68 | 410.07 | 923.52 | 1527.01 | 2764.07 | 3864.5 | 5378.54 | 6876.24 | 32327.23 | 69597.91 | 190.03 | 2697.39
77.65 | 80.01 | 216.7 | 466.75 | 1106.9 | 1907.1 | 3601.05 | 5345.13 | 6819.33 | 8537.2 | 41150.16 | 91105.36 | 250.04 | 3207.61
77.7 | 40.01 | 230.04 | 713.47 | 1126.89 | 2097.16 | 4191.32 | 6465.86 | 8417.17 | 10785.77 | 50547.37 | 109201.87 | 310.05 | 3624.38
77.75 | 76.68 | 173.37 | 670.12 | 1310.31 | 2637.35 | 4965.07 | 6966.16 | 9654.8 | 12140.4 | 57512.39 | 124130.84 | 353.39 | 3831.2
77.8 | 103.35 | 176.7 | 706.8 | 1313.6 | 2570.64 | 5018.34 | 7913.44 | 9768.21 | 12143.72 | 60218.66 | 130650.99 | 400.07 | 2717.4
77.85 | 133.36 | 290.06 | 800.16 | 1523.66 | 2954.13 | 5248.48 | 7903.41 | 10372 | 13201.34 | 64368.05 | 139057.58 | 473.42 | 2887.44
77.9 | 123.35 | 223.37 | 866.82 | 1593.69 | 2944.14 | 5938.77 | 8777.41 | 11289.51 | 13815.59 | 66625.72 | 145437.56 | 446.75 | 2780.74
77.95 | 106.68 | 246.71 | 750.14 | 1567.03 | 2884.08 | 5802.18 | 8473.89 | 11115.91 | 13808.76 | 68890.35 | 148491.36 | 470.08 | 3010.8
78 | 106.68 | 196.7 | 826.83 | 1667.02 | 3124.16 | 5675.42 | 8714.06 | 10955.89 | 14202.32 | 67149.2 | 146726.07 | 410.08 | 2850.78
78.05 | 110.02 | 240.05 | 703.46 | 1426.98 | 2844.07 | 5785.37 | 8417.13 | 10678.95 | 13488.3 | 65856.31 | 142518.86 | 473.42 | 2837.41
78.1 | 96.68 | 223.37 | 796.82 | 1350.29 | 3117.49 | 5615.35 | 8233.75 | 10435.5 | 13651.9 | 63161.58 | 138150.37 | 423.41 | 3404.28
78.15 | 116.68 | 196.7 | 763.49 | 1430.3 | 2720.74 | 5041.67 | 7593.26 | 9881.65 | 11960.39 | 58634.36 | 126726.33 | 376.73 | 3787.79
78.2 | 113.35 | 203.37 | 573.44 | 1110.22 | 2163.87 | 4598.1 | 6199.01 | 8417.15 | 10361.96 | 50956.06 | 110378.04 | 373.41 | 3644.39
78.25 | 30.01 | 10 | 50.01 | 140.02 | 270.04 | 443.42 | 763.49 | 926.85 | 1280.31 | 5492.14 | 12110.93 | 13.33 | 3814.51
78.3 | 0 | 0 | 0 | 0 | 0 | 3.33 | 10 | 3.33 | 0 | 16.67 | 40.01 | 0 | 2107.16
78.35 | 0 | 0 | 0 | 0 | 6.67 | 10 | 0 | 0 | 6.67 | 3.33 | 6.67 | 0 | 1807.09
78.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2210.61
78.45 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 2020.47
78.5 | 6.67 | 0 | 6.67 | 3.33 | 6.67 | 0 | 3.33 | 0 | 0 | 13.34 | 3.33 | 16.67 | 2260.58
78.55 | 40.01 | 60.01 | 43.34 | 26.67 | 33.34 | 63.34 | 36.67 | 66.68 | 36.67 | 16.67 | 13.34 | 103.35 | 2380.59
78.6 | 103.35 | 73.35 | 96.69 | 73.35 | 56.68 | 116.69 | 70.01 | 76.68 | 123.36 | 63.35 | 40.01 | 140.03 | 2824.07
78.65 | 163.38 | 170.05 | 116.69 | 153.37 | 136.7 | 126.7 | 216.72 | 113.36 | 106.69 | 86.68 | 80.01 | 290.07 | 3627.72
78.7 | 186.71 | 146.7 | 173.38 | 176.71 | 143.36 | 126.7 | 210.04 | 170.05 | 130.03 | 140.03 | 76.68 | 266.73 | 4151.26
78.75 | 126.7 | 200.05 | 196.73 | 120.02 | 156.7 | 173.38 | 173.37 | 176.71 | 160.03 | 136.7 | 120.02 | 326.73 | 4154.62
78.8 | 230.07 | 213.39 | 173.38 | 170.04 | 160.03 | 190.04 | 203.38 | 203.38 | 193.39 | 140.02 | 176.71 | 383.43 | 3317.66
78.85 | 270.07 | 200.06 | 220.09 | 216.72 | 223.4 | 240.05 | 226.73 | 223.38 | 243.4 | 153.36 | 103.35 | 400.09 | 3694.45
78.9 | 213.4 | 176.72 | 213.39 | 193.38 | 213.4 | 186.72 | 223.38 | 170.04 | 163.37 | 136.7 | 106.69 | 383.41 | 3907.87
78.95 | 183.39 | 230.05 | 226.73 | 260.09 | 193.38 | 200.05 | 233.39 | 233.41 | 263.41 | 163.36 | 156.7 | 496.78 | 3777.82
79 | 230.06 | 243.4 | 193.38 | 166.71 | 220.06 | 183.37 | 256.74 | 283.43 | 186.72 | 196.71 | 193.37 | 326.74 | 3631.09
79.05 | 273.4 | 186.72 | 203.38 | 256.74 | 180.05 | 213.39 | 203.38 | 253.39 | 260.06 | 166.7 | 170.03 | 403.43 | 3731.17
79.1 | 180.04 | 230.05 | 190.05 | 210.06 | 216.72 | 240.07 | 216.71 | 160.04 | 176.7 | 226.73 | 153.36 | 333.4 | 3741.1
79.15 | 150.04 | 156.7 | 243.41 | 176.72 | 176.7 | 196.71 | 216.72 | 193.39 | 243.39 | 103.35 | 120.02 | 460.11 | 3554.37
79.2 | 133.36 | 163.38 | 150.03 | 176.72 | 130.03 | 186.72 | 166.7 | 153.36 | 150.03 | 223.38 | 126.69 | 366.75 | 2994.14
79.25 | 13.34 | 36.67 | 23.34 | 16.67 | 10 | 30 | 30.01 | 23.34 | 73.35 | 43.34 | 13.34 | 40.01 | 2577.32
79.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.67 | 6.67 | 266.71
79.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 10 | 6.67 | 0
79.4 | 3.33 | 6.67 | 0 | 0 | 3.33 | 6.67 | 6.67 | 13.34 | 0 | 30.01 | 53.34 | 6.67 | 23.34
79.45 | 6.67 | 13.33 | 13.34 | 6.67 | 26.67 | 23.34 | 56.68 | 36.67 | 53.34 | 203.37 | 436.75 | 13.34 | 13.34
79.5 | 103.35 | 103.35 | 156.69 | 173.36 | 163.37 | 183.37 | 310.06 | 336.72 | 396.74 | 1443.7 | 3287.92 | 53.34 | 76.68
134.55 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 3351.07 | 3.33
134.6 | 3.33 | 6.67 | 3.33 | 6.67 | 6.67 | 3.33 | 13.34 | 0 | 13.34 | 13.34 | 13.34 | 9291.32 | 13.34
134.65 | 10 | 10 | 6.67 | 10 | 16.67 | 3.33 | 16.67 | 0 | 10 | 26.68 | 0 | 14165.93 | 16.67
134.7 | 30.01 | 6.67 | 13.34 | 13.34 | 16.67 | 6.67 | 0 | 3.33 | 13.34 | 13.34 | 10 | 17981.15 | 16.67
134.75 | 0 | 3.33 | 13.34 | 13.33 | 10 | 6.67 | 0 | 6.67 | 26.67 | 10 | 3.33 | 21593.01 | 3.33
134.8 | 66.78 | 6.67 | 16.67 | 16.67 | 13.34 | 6.67 | 20 | 6.67 | 20 | 10 | 23.34 | 23773.72 | 6.67
134.85 | 13.34 | 16.67 | 16.67 | 10 | 20.01 | 20 | 6.67 | 3.33 | 26.67 | 6.67 | 30 | 24585.29 | 6.67
134.9 | 13.34 | 6.67 | 20 | 10 | 76.81 | 10 | 6.67 | 16.67 | 23.34 | 23.34 | 16.67 | 27230.11 | 20
134.95 | 6.67 | 30 | 26.67 | 33.34 | 6.67 | 33.34 | 13.34 | 20 | 6.67 | 10 | 10 | 27618.11 | 6.67
135 | 13.34 | 6.67 | 60.06 | 10 | 10 | 16.67 | 76.79 | 6.67 | 6.67 | 6.67 | 20 | 27684.82 | 23.34
135.05 | 26.68 | 6.67 | 46.68 | 6.67 | 26.68 | 16.67 | 218.14 | 20 | 137.16 | 36.67 | 6.67 | 27835.1 | 26.67
135.1 | 6.67 | 26.67 | 126.95 | 33.34 | 20 | 20 | 83.46 | 6.67 | 113.56 | 3.33 | 36.67 | 27086.95 | 6.67
135.15 | 6.67 | 20 | 6.67 | 30.01 | 20 | 10 | 10 | 6.67 | 10 | 10 | 23.34 | 24738.88 | 26.67
135.2 | 6.67 | 33.34 | 6.67 | 0 | 20 | 6.67 | 6.67 | 10 | 6.67 | 16.67 | 20 | 21042.31 | 16.67
135.25 | 6.67 | 6.67 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 20 | 3.33 | 4121.39 | 3.33
135.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 33.34 | 0
135.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 0 | 0
135.4 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 0
135.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 23.34 | 0
135.5 | 0 | 6.67 | 0 | 3.33 | 0 | 0 | 3.33 | 6.67 | 3.33 | 0 | 0 | 316.72 | 0
136.55 | 0 | 0 | 3.33 | 16.67 | 0 | 0 | 0 | 3.33 | 0 | 30.03 | 3.33 | 5292.09 | 0
136.6 | 6.67 | 16.67 | 13.34 | 6.67 | 6.67 | 3.33 | 10 | 0 | 33.36 | 3.33 | 3.33 | 10142.28 | 10
136.65 | 16.67 | 20 | 23.34 | 13.34 | 30.01 | 3.33 | 40.02 | 16.67 | 23.34 | 143.94 | 36.67 | 25277.03 | 0
136.7 | 10 | 23.34 | 23.34 | 20 | 20 | 20 | 137.28 | 26.67 | 10 | 170.74 | 6.67 | 32890.59 | 13.34
136.75 | 26.68 | 157.22 | 40.01 | 36.68 | 83.47 | 23.34 | 16.67 | 16.67 | 16.67 | 110.3 | 23.34 | 38602.03 | 46.67
136.8 | 66.7 | 489.96 | 23.34 | 60.03 | 53.35 | 80.03 | 13.34 | 33.34 | 16.67 | 56.69 | 30 | 43567.22 | 20
136.85 | 23.34 | 66.69 | 33.34 | 33.35 | 20 | 46.68 | 0 | 66.7 | 56.7 | 53.37 | 23.34 | 45870.98 | 20
136.9 | 26.67 | 33.34 | 46.68 | 76.72 | 160.4 | 66.69 | 10 | 40.01 | 16.67 | 53.35 | 16.67 | 49380.25 | 23.34
136.95 | 16.67 | 26.67 | 46.68 | 16.67 | 83.44 | 40.01 | 36.68 | 20 | 23.34 | 16.67 | 140.42 | 51424.07 | 10
137 | 26.67 | 23.34 | 46.68 | 16.67 | 10 | 43.34 | 23.34 | 20 | 50.01 | 26.67 | 113.44 | 51056.39 | 10
137.05 | 210.7 | 40.01 | 86.69 | 46.69 | 33.34 | 30 | 3.33 | 103.63 | 30.01 | 20 | 120.23 | 50912.86 | 40.01
137.1 | 13.34 | 43.35 | 26.67 | 23.34 | 60.03 | 26.67 | 23.34 | 3.33 | 30.01 | 93.49 | 36.67 | 49788.72 | 26.67
137.15 | 20 | 53.35 | 36.68 | 6.67 | 164.04 | 20 | 33.34 | 30 | 36.67 | 36.67 | 33.34 | 46683.86 | 20
137.2 | 6.67 | 86.81 | 10 | 6.67 | 13.34 | 20 | 10 | 6.67 | 13.34 | 13.34 | 40.01 | 41203.49 | 26.68
137.25 | 3.33 | 10 | 6.67 | 0 | 6.67 | 3.33 | 3.33 | 0 | 3.33 | 6.67 | 6.67 | 5775.67 | 6.67
137.3 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 40.01 | 0
137.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 0
137.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 50.01 | 0
137.45 | 3.33 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 206.7 | 0
137.5 | 0 | 0 | 0 | 3.33 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 1453.77 | 0
142.55 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 36883.32 | 16.67
142.6 | 3.33 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 72241.78 | 33.34
142.65 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 6.67 | 0 | 109982.39 | 56.68
142.7 | 3.33 | 0 | 6.67 | 0 | 0 | 26.69 | 3.33 | 3.33 | 6.67 | 0 | 6.67 | 155837.22 | 33.34
142.75 | 0 | 6.67 | 3.33 | 10 | 0 | 0 | 10 | 10 | 0 | 6.67 | 6.67 | 182812.27 | 56.68
142.8 | 3.33 | 3.33 | 0 | 6.67 | 0 | 0 | 0 | 6.67 | 6.67 | 10 | 16.67 | 201332.08 | 80.02
142.85 | 3.33 | 10 | 0 | 3.33 | 10 | 0 | 6.67 | 16.67 | 0 | 6.67 | 3.33 | 208429.96 | 116.69
142.9 | 6.67 | 6.67 | 6.67 | 13.34 | 3.33 | 3.33 | 3.33 | 0 | 3.33 | 16.67 | 3.33 | 219435.92 | 83.35
142.95 | 10 | 16.67 | 6.67 | 0 | 0 | 0 | 10 | 16.67 | 0 | 3.33 | 3.33 | 228269.64 | 80.01
143 | 0 | 6.67 | 10 | 3.33 | 10 | 0 | 0 | 6.67 | 23.34 | 0 | 6.67 | 226427.32 | 53.34
143.05 | 3.33 | 13.34 | 3.33 | 6.67 | 0 | 0 | 3.33 | 0 | 10 | 3.33 | 23.34 | 225510.41 | 66.68
143.1 | 10 | 6.67 | 0 | 6.67 | 3.33 | 3.33 | 0 | 0 | 0 | 6.67 | 6.67 | 218035.37 | 136.69
143.15 | 0 | 10 | 0 | 0 | 3.33 | 0 | 0 | 0 | 3.33 | 3.33 | 3.33 | 201703 | 60.01
143.2 | 0 | 0 | 10 | 0 | 0 | 6.67 | 0 | 6.67 | 6.67 | 3.33 | 6.67 | 148963.26 | 63.34
143.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 15819.93 | 0
143.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 86.68 | 0
143.35 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 26.67 | 6.67
143.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 90.01 | 10
143.45 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 853.56 | 113.35
143.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 6974.53 | 910.21
144.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 32195.53 | 3.33
144.6 | 0 | 3.33 | 6.67 | 3.33 | 3.33 | 3.33 | 0 | 0 | 6.67 | 0 | 0 | 49615.83 | 13.33
144.65 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 79552.44 | 23.34
144.7 | 6.67 | 6.67 | 6.67 | 3.33 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 107120.2 | 66.68
144.75 | 0 | 0 | 0 | 0 | 10 | 6.67 | 10 | 0 | 3.33 | 10 | 6.67 | 131013.63 | 50.01
144.8 | 0 | 3.33 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 6.67 | 13.34 | 6.67 | 144829.79 | 53.34
144.85 | 3.33 | 0 | 0 | 0 | 0 | 0 | 10 | 3.33 | 0 | 6.67 | 0 | 154587.38 | 93.35
144.9 | 13.34 | 0 | 0 | 0 | 0 | 0 | 10 | 6.67 | 13.34 | 3.33 | 3.33 | 160625.95 | 70.01
144.95 | 10 | 10 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 163732.94 | 70.01
145 | 0 | 0 | 6.67 | 3.33 | 3.33 | 3.33 | 0 | 10 | 3.33 | 3.33 | 0 | 165464.58 | 70.01
145.05 | 0 | 3.33 | 13.34 | 0 | 0 | 0 | 0 | 3.33 | 13.34 | 0 | 3.33 | 163783.58 | 56.68
145.1 | 0 | 10 | 0 | 0 | 0 | 3.33 | 0 | 20.01 | 0 | 16.67 | 6.67 | 159759.11 | 46.67
145.15 | 3.33 | 3.33 | 0 | 0 | 3.33 | 10 | 0 | 3.33 | 0 | 6.67 | 13.34 | 144109.86 | 56.68
145.2 | 6.67 | 3.33 | 0 | 0 | 0 | 10 | 6.67 | 0 | 3.33 | 0 | 3.33 | 109814.74 | 13.33
145.25 | 0 | 0 | 3.33 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 10489.98 | 0
145.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 90.02 | 10
145.35 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 46.67 | 0
145.4 | 0 | 3.33 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 126.69 | 0
145.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 456.76 | 0
145.5 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5132.62 | 3.33
149.55 | 0 | 0 | 3.33 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 6.67 | 19558.93 | 28998.16
149.6 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 3.33 | 37685.14 | 54571.82
149.65 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 6.67 | 0 | 0 | 0 | 61521.87 | 85942.99
149.7 | 3.33 | 13.34 | 0 | 3.33 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 76366.81 | 107532.6
149.75 | 3.33 | 3.33 | 13.34 | 0 | 13.34 | 0 | 3.33 | 0 | 0 | 3.33 | 13.34 | 100876.6 | 141150.99
149.8 | 26.67 | 0 | 0 | 3.33 | 0 | 6.67 | 0 | 0 | 3.33 | 6.67 | 6.67 | 112445.39 | 155734.96
149.85 | 13.34 | 0 | 0 | 6.67 | 6.67 | 0 | 0 | 3.33 | 6.67 | 3.33 | 0 | 117469.69 | 168356.65
149.9 | 16.67 | 0 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 6.67 | 3.33 | 3.33 | 126302.58 | 175574.48
149.95 | 3.33 | 0 | 6.67 | 13.34 | 0 | 0 | 6.67 | 3.33 | 6.67 | 0 | 10 | 128582 | 183924.89
150 | 16.67 | 6.67 | 3.33 | 0 | 3.33 | 0 | 3.33 | 3.33 | 0 | 0 | 10 | 128352.03 | 179376.58
150.05 | 26.67 | 3.33 | 0 | 6.67 | 0 | 3.33 | 6.67 | 6.67 | 0 | 10 | 10 | 127974.34 | 180895.38
150.1 | 10 | 3.33 | 13.34 | 6.67 | 6.67 | 0 | 0 | 3.33 | 0 | 10 | 0 | 125077.03 | 176407.34
150.15 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 3.33 | 6.67 | 117326.28 | 162686.66
150.2 | 6.67 | 6.67 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 3.33 | 10 | 92377.22 | 125756.19
150.25 | 0 | 6.67 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16353.3 | 17312.01
150.3 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 110.02 | 120.02
150.35 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33
150.4 | 0 | 3.33 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 0
150.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
150.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33
154.55 | 6.67 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 70.01 | 105112.04
154.6 | 16.67 | 10 | 3.33 | 0 | 0 | 3.33 | 0 | 10 | 3.33 | 0 | 0 | 63.34 | 192049.94
154.65 | 20 | 0 | 3.33 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 0 | 0 | 93.35 | 284596.65
154.7 | 26.67 | 0 | 6.67 | 3.33 | 0 | 3.33 | 6.67 | 20.01 | 6.67 | 0 | 10 | 100.02 | 382876.63
154.75 | 53.34 | 6.67 | 16.67 | 23.35 | 0 | 0 | 16.67 | 16.68 | 10 | 6.67 | 6.67 | 236.71 | 499622.57
154.8 | 20 | 3.33 | 0 | 3.33 | 6.67 | 6.67 | 3.33 | 23.34 | 0 | 3.33 | 3.33 | 236.71 | 543075.76
154.85 | 30.01 | 0 | 0 | 6.67 | 0 | 0 | 0 | 6.67 | 3.33 | 0 | 23.35 | 216.71 | 587651.79
154.9 | 30.01 | 10 | 6.67 | 10 | 6.67 | 10 | 3.33 | 0 | 0 | 0 | 0 | 250.06 | 617821.61
154.95 | 20 | 0 | 0 | 6.67 | 6.67 | 0 | 6.67 | 0 | 0 | 0 | 20 | 273.39 | 634971.75
155 | 33.34 | 0 | 0 | 13.34 | 3.33 | 3.33 | 0 | 6.67 | 3.33 | 6.67 | 0 | 260.04 | 642564.52
155.05 | 40.01 | 16.67 | 0 | 0 | 3.33 | 13.34 | 6.67 | 0 | 3.33 | 3.33 | 3.33 | 296.72 | 639995.66
155.1 | 10 | 3.33 | 3.33 | 20.01 | 0 | 6.67 | 3.33 | 3.33 | 0 | 0 | 0 | 190.04 | 623024.1
155.15 | 26.67 | 3.33 | 0 | 16.67 | 10 | 3.33 | 10 | 0 | 0 | 6.67 | 0 | 150.02 | 587255.62
155.2 | 36.67 | 0 | 10 | 13.34 | 6.67 | 0 | 3.33 | 23.34 | 0 | 3.33 | 3.33 | 143.36 | 476117.73
155.25 | 13.34 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 3.33 | 0 | 0 | 10 | 87301.24
155.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 743.5
155.35 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 3.33 | 0 | 83.35
155.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 336.73
155.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 1063.61
155.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 93.35 | 11341.31
155.55 | 16.67 | 0 | 6.67 | 3.33 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 360.07 | 81719.86
155.6 | 20 | 13.34 | 3.33 | 6.67 | 0 | 0 | 0 | 13.34 | 0 | 6.67 | 0 | 823.53 | 217303.22
155.65 | 10 | 6.67 | 0 | 0 | 3.33 | 3.33 | 6.67 | 30.02 | 10 | 3.33 | 0 | 1423.7 | 455453.05
155.7 | 30 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 3.33 | 3.33 | 6.67 | 3.33 | 1980.56 | 592759.68
155.75 | 6.67 | 13.34 | 0 | 0 | 3.33 | 20.01 | 10 | 0 | 20 | 3.33 | 0 | 2467.47 | 685679.4
155.8 | 30 | 6.67 | 10 | 0 | 6.67 | 0 | 3.33 | 3.33 | 13.34 | 10 | 13.34 | 2487.58 | 768275.5
155.85 | 26.67 | 16.67 | 33.34 | 3.33 | 0 | 3.33 | 13.34 | 6.67 | 56.74 | 10 | 23.34 | 2981.1 | 825757.52
155.9 | 50.01 | 6.67 | 46.7 | 3.33 | 6.67 | 6.67 | 26.69 | 3.33 | 10 | 0 | 3.33 | 2804.32 | 864038.24
155.95 | 50.01 | 20 | 6.67 | 3.33 | 0 | 50.02 | 26.68 | 3.33 | 6.67 | 6.67 | 16.67 | 2877.7 | 902857.16
156 | 50.01 | 6.67 | 0 | 0 | 10 | 6.67 | 23.34 | 0 | 13.34 | 6.67 | 0 | 2680.9 | 911461.29
156.05 | 46.68 | 6.67 | 3.33 | 3.33 | 6.67 | 6.67 | 20.01 | 3.33 | 6.67 | 0 | 6.67 | 2657.53 | 907222.76
156.1 | 56.68 | 13.34 | 0 | 6.67 | 0 | 3.33 | 0 | 16.67 | 26.69 | 0 | 0 | 2620.85 | 879648.5
156.15 | 36.67 | 3.33 | 0 | 0 | 20 | 6.67 | 0 | 0 | 26.69 | 6.67 | 0 | 2650.95 | 828028.27
156.2 | 43.34 | 3.33 | 0 | 10 | 0 | 0 | 6.67 | 6.67 | 6.67 | 0 | 0 | 1510.43 | 687579.4
156.25 | 10 | 6.67 | 0 | 0 | 0 | 3.33 | 6.67 | 3.33 | 3.33 | 6.67 | 3.33 | 46.67 | 86735.17
156.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 530.1
156.35 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 110.02
156.4 | 3.33 | 0 | 0 | 0 | 3.33 | 6.67 | 0 | 0 | 0 | 0 | 6.67 | 0 | 150.03
156.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 790.26
156.5 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 0 | 6648.72
156.55 | 13.34 | 0 | 0 | 0 | 6.67 | 0 | 0 | 6.67 | 0 | 0 | 0 | 26.67 | 78584.9
156.6 | 20 | 10 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 103.35 | 206896.88
156.65 | 26.67 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 110.02 | 331359.48
156.7 | 13.34 | 13.34 | 3.33 | 0 | 0 | 23.34 | 0 | 0 | 0 | 56.78 | 6.67 | 183.37 | 417970.26
156.75 | 23.34 | 6.67 | 0 | 6.67 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 153.36 | 514748.6
156.8 | 70.06 | 0 | 10 | 3.33 | 13.34 | 16.67 | 3.33 | 3.33 | 0 | 6.67 | 0 | 170.03 | 611653.37
156.85 | 36.68 | 0 | 0 | 3.33 | 10 | 0 | 30.02 | 0 | 3.33 | 6.67 | 3.33 | 226.71 | 648568.47
156.9 | 136.98 | 23.34 | 16.68 | 16.67 | 6.67 | 10 | 30.01 | 3.33 | 3.33 | 6.67 | 6.67 | 223.38 | 689175.13
156.95 | 56.68 | 20 | 13.34 | 13.34 | 0 | 3.33 | 3.33 | 3.33 | 0 | 6.67 | 10 | 243.38 | 716840.23
157 | 33.34 | 0 | 30.03 | 3.33 | 13.34 | 13.34 | 6.67 | 13.34 | 13.34 | 3.33 | 6.67 | 310.06 | 713493.63
157.05 | 53.35 | 10 | 23.34 | 6.67 | 3.33 | 3.33 | 0 | 0 | 10 | 3.33 | 16.67 | 200.03 | 704084.28
157.1 | 40.01 | 3.33 | 6.67 | 3.33 | 3.33 | 6.67 | 26.69 | 0 | 13.34 | 10 | 13.34 | 236.71 | 693134.89
157.15 | 26.67 | 3.33 | 10 | 6.67 | 0 | 6.67 | 3.33 | 3.33 | 10 | 13.34 | 10 | 226.71 | 652437.46
157.2 | 20 | 13.34 | 0 | 6.67 | 3.33 | 0 | 3.33 | 0 | 0 | 3.33 | 13.34 | 83.35 | 509523.38
157.25 | 10 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 89391.53
157.3 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 0 | 0 | 0 | 0 | 946.9
157.35 | 0 | 0 | 6.67 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 6.67 | 0 | 96.68
157.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 366.74
157.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 2398.04
157.5 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 106.69 | 16488.08
Printed:7/20/2016 10:02:22 AM

## 71.5
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Narrow
Narrow
71 | 0 | 3184.22 | 14345.98 | 28815.92 | 56666.45 | 115811.39 | 172096.38 | 225017 | 285707.88 | 1527172.6 | 3036044.52 | 4961.67 | 156.69 |  | 0.05 AMU | 0.15 AMU
71.05 | 0 | 2770.74 | 13932.12 | 27262.62 | 56127.43 | 111469.21 | 164122.95 | 219656.12 | 279823.75 | 1472273.37 | 2937448.76 | 4601.49 | 136.69 | 1 | 0 | 0
71.1 | 0 | 2567.29 | 13264.85 | 25869.85 | 50794.99 | 101323.17 | 150672.13 | 202890.97 | 258415.18 | 1389365.41 | 2787909.93 | 4251.39 | 136.69 | 5 | 0 | 0
71.15 | 0 | 630.11 | 3294.26 | 5995.71 | 11273.39 | 21528.1 | 31205.63 | 39299.96 | 50448.28 | 740963.69 | 1418337.95 | 3694.47 | 20 | 10 | 0 | 0
71.2 | 0 | 10 | 16.67 | 10 | 50.01 | 116.68 | 133.36 | 170.03 | 260.05 | 47755.88 | 43388.56 | 76.68 | 0 | 20 | 0 | 0
71.25 | 0 | 0 | 0 | 0 | 0 | 13.34 | 6.67 | 0 | 13.33 | 46612.23 | 45806.67 | 26.67 | 0 | 40 | 0.001888 | 0.0037632
71.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 46386.83 | 42739.15 | 326.72 | 0 | 60 | 0 | 0.0037651
71.35 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 45081.35 | 43257.75 | 1086.87 | 0 | 80 | 0 | 0.0012532
71.4 | 0 | 0 | 0 | 6.67 | 0 | 10 | 0 | 0 | 0 | 42803.86 | 43551.06 | 1380.3 | 0 | 100 | 0.0056698 | 0.0037632
71.45 | 0 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 3.33 | 0 | 13.34 | 45589.09 | 1577.01 | 3.33 | 500 | 0.018903 | 0.025722
71.5 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 10 | 33.34 | 43049.46 | 1763.74 | 3.33 | 1000 | 24.4081 | 24.9075
71.55 | 0 | 0 | 0 | 0 | 0 | 10 | 13.34 | 3.33 | 10 | 90.02 | 43733.84 | 1973.82 | 0
71.6 | 0 | 0 | 0 | 3.33 | 13.34 | 10 | 26.67 | 33.34 | 43.34 | 246.71 | 42682 | 1647.04 | 0
71.65 | 0 | 0 | 20 | 3.33 | 10 | 73.35 | 106.69 | 100.02 | 140.02 | 786.82 | 44740.81 | 1406.97 | 0
71.7 | 0 | 0 | 76.68 | 60.01 | 220.04 | 346.73 | 620.13 | 870.19 | 800.15 | 5025.34 | 45502.18 | 23.34 | 0
71.75 | 0 | 60.01 | 300.06 | 590.13 | 1387 | 2560.82 | 3938.21 | 5559.28 | 6246.58 | 36153.64 | 65501.34 | 36.67 | 10
71.8 | 10 | 640.12 | 2904.08 | 6082.26 | 12160.39 | 24564.92 | 36586.6 | 50162.57 | 62703.78 | 322730.24 | 411615.69 | 563.44 | 60.01
71.85 | 6.67 | 1393.63 | 6415.86 | 12697.51 | 25345.6 | 50550.56 | 74525.62 | 100836.33 | 125539.67 | 640052.81 | 1240228.07 | 2910.79 | 213.37
71.9 | 3.33 | 1627.03 | 7029.5 | 14339.27 | 29269.59 | 57336.4 | 86746.64 | 114860.99 | 144527.06 | 735411.38 | 1501662.57 | 3244.28 | 310.05
71.95 | 0 | 1550.35 | 7819.99 | 15632.52 | 30933.92 | 61788.63 | 93042.8 | 126759.3 | 155758.07 | 797287.2 | 1669978.5 | 4138 | 366.73
72 | 3.33 | 1817.07 | 8006.85 | 16234.71 | 31832.47 | 65241.36 | 97790.75 | 130063.71 | 162927.7 | 829617.1 | 1760195.89 | 3991.19 | 333.4
Standard Res
|  | 1 | 5 | 10 | 20 | 40 | 60 | 80 | 100 | 500 |  |  |  |  | Standard
71 | 33.34 | 3397.58 | 16368.53 | 33887.16 | 67583.93 | 134715.86 | 201526.16 | 265917.6 | 338952.55 | 1767621.72 | 3564916.25 | 5345.13 | 216.71 |  | 0.05 AMU | 0.15 AMU
71.05 | 16.67 | 3034.19 | 16585.18 | 33319.81 | 66478.83 | 131812.35 | 197542.7 | 258903.86 | 328928.52 | 1731763.75 | 3466303.59 | 5451.92 | 233.38 | 1 | 0.0032795 | 0.042039
71.1 | 33.34 | 3230.9 | 15207.13 | 32323.52 | 63142.93 | 125734.27 | 188098.41 | 252114.92 | 314922.48 | 1687319.35 | 3383876.85 | 5138.43 | 300.07 | 5 | 0.034423 | 0.16983
71.15 | 13.34 | 3054.2 | 14409.13 | 28879.21 | 58803.23 | 116373.25 | 172254.01 | 228407.85 | 289017.12 | 1569033.62 | 3145203.56 | 5742.16 | 206.72 | 10 | 0.062291 | 0.35314
71.2 | 16.67 | 2580.66 | 13228.04 | 26047.25 | 51925.7 | 104167.17 | 156144.97 | 205376.34 | 258496.42 | 1416147.38 | 2840068.53 | 5718.8 | 226.71 | 20 | 0.15573 | 0.70691
71.25 | 3.33 | 793.48 | 4368.11 | 8178.71 | 17753.98 | 35968.96 | 54802.87 | 73406.69 | 92471.06 | 968131.84 | 1890382.03 | 5141.77 | 43.34 | 40 | 0.28196 | 1.3624
71.3 | 0 | 0 | 16.67 | 36.67 | 76.68 | 90.02 | 136.69 | 243.38 | 333.39 | 50508.7 | 53383.2 | 1807.06 | 0 | 60 | 0.37375 | 1.537
71.35 | 0 | 0 | 0 | 6.67 | 16.67 | 13.34 | 10 | 20 | 50.01 | 50195.56 | 50817.13 | 1957.13 | 0 | 80 | 0.48687 | 2.4159
71.4 | 0 | 0 | 0 | 16.67 | 13.34 | 36.67 | 53.34 | 86.68 | 86.68 | 51154.36 | 52075.37 | 1827.09 | 6.67 | 100 | 0.54261 | 3.1752
71.45 | 0 | 0 | 3.33 | 16.67 | 56.68 | 103.35 | 120.02 | 160.03 | 143.35 | 983.52 | 52325.22 | 1973.8 | 0 | 500 | 2.9367 | 16.3863
71.5 | 0 | 6.67 | 70.01 | 126.69 | 316.73 | 573.45 | 760.15 | 990.22 | 1103.57 | 5972.83 | 54254.15 | 2033.83 | 0 | 1000 | 26.6759 | 33.4318
71.55 | 0 | 243.38 | 936.85 | 1957.16 | 3831.35 | 7427 | 10709.99 | 13219.86 | 17639.59 | 90511.01 | 92276.78 | 1940.48 | 13.34
71.6 | 6.67 | 726.8 | 3324.28 | 6869.46 | 13111.3 | 26194.42 | 38370.32 | 49599.24 | 63343.17 | 310290.63 | 498306.52 | 2687.37 | 106.68
71.65 | 16.67 | 1463.65 | 6489.03 | 12573.84 | 25419.34 | 50246.7 | 73901.77 | 98301.37 | 123019.63 | 622654.06 | 1202471.71 | 4261.31 | 250.04
71.7 | 33.34 | 1507 | 7596.44 | 15313.82 | 30820.35 | 59939.66 | 89573.23 | 116416.43 | 147817.43 | 743196.98 | 1517027.59 | 4788.3 | 360.06
71.75 | 3.33 | 1530.34 | 8036.79 | 16638.53 | 32290.47 | 64198.92 | 96806.27 | 129109.41 | 161959.88 | 817142 | 1692551.93 | 4898.36 | 386.73
71.8 | 3.33 | 1660.37 | 8667.25 | 17060.96 | 34682.04 | 70371.26 | 104871.69 | 136597.71 | 174794.4 | 874867.67 | 1826535.21 | 3854.49 | 440.09
71.85 | 16.67 | 1860.44 | 8964.13 | 18574.65 | 36547.3 | 73341.76 | 109815.33 | 143441.94 | 180859.27 | 915666.04 | 1914030.39 | 3904.53 | 476.75
71.9 | 6.67 | 1973.79 | 9651.34 | 18831.66 | 37336.09 | 75664.21 | 112333.04 | 150397.38 | 187538.06 | 939192.05 | 1996664.78 | 4201.33 | 450.08
71.95 | 20 | 2010.46 | 9938.21 | 19699.45 | 38121.44 | 76659.77 | 115572.12 | 152229.97 | 190807.7 | 972218.32 | 2045267.17 | 4651.56 | 580.11
72 | 10 | 1873.78 | 9521.27 | 19549.28 | 39174.46 | 77254.62 | 115068.48 | 153127.72 | 191774.69 | 973794.55 | 2068479.12 | 4608.18 | 513.43
|  | Narrow |  | Standard
|  | 0.05 AMU | 0.15 AMU | 0.05 AMU | 0.15 AMU
| 1 | 0% | 0% | 0.328% | 4.204%
| 5 | 0% | 0% | 3.442% | 16.98%
| 10 | 0% | 0% | 6.229% | 35.31%
| 20 | 0% | 0% | 15.57% | 70.69%
| 40 | 0.1888% | 0.3763% | 28.2% | 136.2%
| 60 | 0% | 0.3765% | 37.38% | 153.7%
| 80 | 0.567% | 0.3763% | 54.26% | 317.5%
| 100 | 0% | 0.1253% | 48.69% | 241.6%
| 500 | 1.89% | 2.572% | 293.7% | 1639%
| 1000 | 2441% | 2491% | 2668% | 3343%

## 72.5
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Narrow
Narrow
72 | 3.33 | 1817.07 | 8006.85 | 16234.71 | 31832.47 | 65241.36 | 97790.75 | 130063.71 | 162927.7 | 829617.1 | 1760195.89 | 3991.19 | 333.4 |  | 0.05 AMU | 0.15 AMU
72.05 | 10 | 1647.04 | 8105.32 | 15754.23 | 30402.54 | 61513.14 | 93514.53 | 123475.17 | 154696.9 | 790994.44 | 1680890.26 | 4074.6 | 360.06 | 1 | 0 | 0
72.1 | 0 | 1530.33 | 7516.48 | 14746.37 | 29400.69 | 59270.14 | 87768.88 | 116879.31 | 146436 | 746800.43 | 1620773.86 | 4284.71 | 363.4 | 5 | 0 | 0
72.15 | 6.67 | 236.71 | 830.16 | 1637.07 | 3167.64 | 6105.86 | 8581.21 | 9672.08 | 13322.38 | 61059.73 | 672673.21 | 2607.39 | 280.05 | 10 | 0 | 0.00082748
72.2 | 0 | 0 | 10 | 6.67 | 10 | 40.01 | 66.68 | 60.01 | 123.35 | 410.07 | 43303.53 | 20 | 0 | 20 | 0 | 0
72.25 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 10 | 3.33 | 6.67 | 43.34 | 44237.05 | 6.67 | 0 | 40 | 0 | 0.00082748
72.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 41579.15 | 273.39 | 0 | 60 | 0.002296 | 0.00082748
72.35 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 43179.75 | 736.8 | 0 | 80 | 0.0068949 | 0.0057998
72.4 | 0 | 0 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 0 | 0 | 42065.28 | 966.86 | 0 | 100 | 0 | 0
72.45 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 1286.96 | 0 | 500 | 0.018389 | 0.012425
72.5 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 10 | 0 | 26.67 | 30 | 1450.35 | 0 | 1000 | 0.020685 | 0.020709
72.55 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 13.34 | 0 | 20 | 46.67 | 1286.94 | 0
72.6 | 0 | 0 | 0 | 10 | 0 | 3.33 | 13.34 | 6.67 | 20.01 | 53.34 | 116.68 | 1316.93 | 0
72.65 | 0 | 0 | 6.67 | 0 | 6.67 | 3.33 | 13.33 | 33.34 | 40.01 | 233.37 | 446.75 | 1036.89 | 0
72.7 | 0 | 0 | 3.33 | 13.34 | 46.67 | 73.35 | 100.02 | 110.02 | 136.69 | 763.48 | 1443.65 | 43.34 | 0
72.75 | 0 | 40.01 | 250.04 | 503.42 | 1066.88 | 2177.25 | 3104.3 | 4641.76 | 5335.45 | 28636.38 | 61104.66 | 66.68 | 0
72.8 | 0 | 236.7 | 1226.92 | 2353.91 | 4931.61 | 9741.5 | 14629.7 | 19803.45 | 24000.44 | 127084.64 | 256226.67 | 763.49 | 0
72.85 | 0 | 363.39 | 1763.74 | 3407.61 | 6859.42 | 13888.84 | 20757.94 | 27613.38 | 34291.9 | 178143.05 | 357819.77 | 1697.07 | 6.67
72.9 | 0 | 393.4 | 2147.21 | 4294.6 | 8734.03 | 17253.02 | 24928.24 | 33362.57 | 41213.44 | 213293.58 | 428464.21 | 2170.54 | 13.33
72.95 | 0 | 446.75 | 2343.9 | 4451.43 | 9521.22 | 18124.01 | 27205.67 | 36657.46 | 45115.07 | 233488.08 | 473412.42 | 2820.74 | 13.33
73 | 3.33 | 496.76 | 2430.58 | 4331.33 | 9264.36 | 19425.71 | 28091 | 37005.33 | 47510.23 | 239744.84 | 483763.47 | 2797.4 | 13.34
Standard Res
Standard
72 | 10 | 1873.78 | 9521.27 | 19549.28 | 39174.46 | 77254.62 | 115068.48 | 153127.72 | 191774.69 | 973794.55 | 2068479.12 | 4608.18 | 513.43 |  | 0.05 AMU | 0.15 AMU
72.05 | 6.67 | 1870.43 | 9467.89 | 19118.81 | 37739.84 | 75234.55 | 112048.26 | 148284.4 | 185758.51 | 941823.44 | 2004364.97 | 4424.77 | 436.75 | 1 | 0.0046309 | 0.027905
72.1 | 16.67 | 1883.77 | 8877.38 | 18157.52 | 36259.95 | 71948.38 | 108303.14 | 142997.44 | 179173.39 | 911479.64 | 1952287.47 | 4474.81 | 493.43 | 5 | 0.023147 | 0.11317
72.15 | 16.67 | 1783.75 | 8123.57 | 15937.81 | 33563.09 | 66682.56 | 99514.49 | 133243.64 | 164818.56 | 839934.22 | 1814319.1 | 4701.56 | 490.09 | 10 | 0.020829 | 0.23177
72.2 | 0 | 1446.97 | 7393.14 | 15313.67 | 30145.49 | 60121.64 | 90347.82 | 118001.14 | 148180.57 | 759905.08 | 1643377.39 | 4694.88 | 366.73 | 20 | 0.074066 | 0.43643
72.25 | 0 | 233.37 | 970.19 | 2003.83 | 4394.85 | 8711.09 | 13152.11 | 18576.49 | 22262.72 | 116527.26 | 879638.29 | 3784.44 | 266.71 | 40 | 0.090271 | 0.79153
72.3 | 0 | 6.67 | 20.01 | 10 | 26.67 | 26.67 | 73.34 | 93.35 | 120.02 | 616.79 | 49427.18 | 1406.98 | 10 | 60 | 0.12962 | 0.91733
72.35 | 0 | 0 | 0 | 0 | 3.33 | 3.33 | 13.34 | 10 | 13.33 | 60.01 | 50517.08 | 1400.31 | 0 | 80 | 0.16666 | 1.57
72.4 | 0 | 0 | 3.33 | 0 | 0 | 0 | 6.67 | 20 | 26.67 | 80.01 | 51222.89 | 1537.02 | 6.67 | 100 | 0.1898 | 1.9338
72.45 | 0 | 3.33 | 0 | 20 | 3.33 | 16.67 | 46.68 | 46.68 | 66.68 | 213.37 | 503.42 | 1390.3 | 0 | 500 | 0.875 | 10.2535
72.5 | 0 | 6.67 | 33.34 | 30 | 106.68 | 130.02 | 186.7 | 240.05 | 273.38 | 1260.29 | 3197.76 | 1440.33 | 0 | 1000 | 2.2202 | 20.2056
72.55 | 0 | 110.02 | 453.42 | 946.84 | 1767.06 | 3257.66 | 5121.88 | 6465.97 | 7977.03 | 42626.56 | 83203.32 | 1470.37 | 3.33
72.6 | 10 | 360.06 | 1166.91 | 2567.34 | 5238.38 | 9784.8 | 14729.79 | 19913.58 | 24724.87 | 125750.43 | 254189.49 | 2327.25 | 13.34
72.65 | 0 | 330.06 | 1853.8 | 3681.07 | 7149.51 | 14065.65 | 21161.85 | 28612.1 | 34859.66 | 177180.71 | 358775.13 | 3084.21 | 16.67
72.7 | 0 | 340.06 | 2187.18 | 4017.9 | 8000.14 | 16533.55 | 24621.21 | 32189.82 | 41986.21 | 208921.85 | 420221.01 | 3270.95 | 30.01
72.75 | 6.67 | 513.42 | 2487.27 | 4874.9 | 9741.48 | 19282.4 | 27667.15 | 36962.27 | 47281.93 | 238095.85 | 478658.87 | 3641.08 | 16.67
72.8 | 0 | 526.76 | 2827.4 | 5368.58 | 10201.93 | 20523.94 | 30649.88 | 39994.12 | 49527.13 | 255207.02 | 521273.6 | 2533.97 | 13.34
72.85 | 6.67 | 646.78 | 2904.11 | 5642.04 | 11199.27 | 22974.83 | 33218.92 | 44085.97 | 55612.09 | 279012.85 | 560738.51 | 3050.86 | 30.01
72.9 | 6.67 | 650.12 | 2924.1 | 5588.57 | 11122.67 | 23108.38 | 33723.43 | 45182.58 | 57101.88 | 285136.2 | 575556.46 | 3207.56 | 33.34
72.95 | 3.33 | 530.09 | 2947.49 | 6165.66 | 11889.88 | 23151.69 | 34646.1 | 45179.16 | 57040.75 | 293392.62 | 583966.07 | 3370.97 | 26.67
73 | 6.67 | 503.43 | 2880.76 | 5795.42 | 11422.87 | 22891.68 | 34809.7 | 45760.96 | 57232.17 | 291186.66 | 587551.31 | 3280.95 | 33.34
|  | Narrow |  | Standard
|  | 0.05 AMU | 0.15 AMU | 0.05 AMU | 0.15 AMU
| 1 | 0% | 0% | 0.4631% | 2.791%
| 5 | 0% | 0% | 2.315% | 11.32%
| 10 | 0% | 0.08275% | 2.083% | 23.18%
| 20 | 0% | 0% | 7.407% | 43.64%
| 40 | 0% | 0.08275% | 9.027% | 79.15%
| 60 | 0.2296% | 0.08275% | 12.96% | 91.73%
| 80 | 0% | 0% | 18.98% | 193.4%
| 100 | 0.6895% | 0.58% | 16.67% | 157%
| 500 | 1.839% | 1.242% | 87.5% | 1025%
| 1000 | 2.068% | 2.071% | 222% | 2021%

## 67.5
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Narrow
Narrow
67 | 6.67 | 310.06 | 1777.06 | 3337.55 | 6909.44 | 13528.29 | 20053.44 | 26875.04 | 33195.71 | 153143.57 | 323290.12 | 723.48 | 93.35 |  | 0.05 AMU | 0.15 AMU
67.05 | 10 | 430.08 | 1623.67 | 3174.2 | 6265.73 | 13258.15 | 19559.13 | 26467.88 | 32644.06 | 149332.53 | 318184.62 | 673.46 | 103.35 | 1 | 0% | 0%
67.1 | 6.67 | 313.39 | 1390.27 | 3057.44 | 6048.88 | 12350.39 | 16875.77 | 22774.55 | 28628.48 | 132012.03 | 276017.23 | 526.77 | 140.03 | 5 | 0% | 0.1019%
67.15 | 0 | 6.67 | 63.34 | 83.35 | 156.69 | 336.72 | 556.77 | 643.45 | 876.85 | 3154.32 | 6409.61 | 233.38 | 16.67 | 10 | 0.9402% | 1.122%
67.2 | 0 | 0 | 6.67 | 0 | 6.67 | 0 | 6.67 | 3.33 | 30 | 66.68 | 116.69 | 6.67 | 0 | 20 | 0% | 0.2041%
67.25 | 0 | 3.33 | 3.33 | 0 | 0 | 3.33 | 10 | 0 | 0 | 16.67 | 10 | 103.35 | 0 | 40 | 0.6271% | 0.5102%
67.3 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 30.02 | 10 | 390.07 | 0 | 60 | 0% | 0.3061%
67.35 | 0 | 10 | 3.33 | 0 | 6.67 | 0 | 3.33 | 0 | 0 | 10 | 0 | 700.13 | 0 | 80 | 0.3131% | 0.8163%
67.4 | 0 | 0 | 43.39 | 0 | 6.67 | 0 | 0 | 0 | 241.84 | 3.33 | 13.33 | 1016.88 | 0 | 100 | 3.137% | 1.225%
67.45 | 0 | 0 | 0 | 16.67 | 0 | 0 | 6.67 | 6.67 | 3.33 | 3.33 | 53507.02 | 973.53 | 6.67 | 500 | 5.642% | 6.327%
67.5 | 6.67 | 0 | 0 | 10 | 0 | 6.67 | 0 | 3.33 | 33.36 | 60.01 | 38737.31 | 1063.55 | 6.67 | 1000 | 3642% | 4010%
67.55 | 0 | 0 | 3.33 | 10 | 6.67 | 10 | 3.33 | 16.67 | 3.33 | 143.39 | 38771.74 | 1230.25 | 0
67.6 | 33.37 | 3.33 | 6.67 | 23.34 | 20 | 16.67 | 33.34 | 20 | 26.67 | 200.04 | 39654.62 | 1196.92 | 0
67.65 | 0 | 13.34 | 6.67 | 13.34 | 20 | 36.67 | 76.68 | 73.35 | 103.35 | 523.43 | 39197.07 | 893.51 | 0
67.7 | 0 | 13.33 | 46.67 | 96.68 | 246.71 | 646.79 | 940.21 | 1276.97 | 1413.66 | 7183.91 | 44474.19 | 13.34 | 0
67.75 | 0 | 130.02 | 486.75 | 1053.56 | 2237.32 | 4661.93 | 6533.03 | 9496.18 | 11104.98 | 52838.82 | 84721.93 | 70.01 | 33.34
67.8 | 30.01 | 893.51 | 3624.38 | 6992.79 | 14359.66 | 29137.39 | 44061.35 | 58718.02 | 71314.79 | 331755.77 | 535447.76 | 763.49 | 276.72
67.85 | 13.34 | 1393.64 | 6228.91 | 12467.32 | 24440.5 | 48292.94 | 71791.99 | 96363.69 | 119369.43 | 546308.42 | 1118380.64 | 1670.38 | 283.38
67.9 | 50.04 | 1663.71 | 6715.91 | 13942.14 | 28231.41 | 56512.03 | 83231.38 | 108614.53 | 135952.54 | 628287.73 | 1324858.96 | 2033.82 | 476.76
67.95 | 80.03 | 1813.74 | 8050.1 | 15810.85 | 31230.82 | 61442.29 | 92311.18 | 122719.29 | 151743 | 699328.77 | 1505980.55 | 2297.2 | 516.76
68 | 13.33 | 1790.37 | 8467.04 | 16640.3 | 33857.08 | 66531.6 | 98920.54 | 130787.51 | 164781 | 757142.98 | 1626806.82 | 2507.29 | 476.75
Standard Res
Standard
67 | 0 | 450.08 | 2060.47 | 4197.95 | 8585.48 | 16732.37 | 25172.17 | 32363.56 | 40812.26 | 184798.32 | 389953.13 | 663.46 | 166.7 |  | 0.05 AMU | 0.15 AMU
67.05 | 0 | 346.73 | 1883.76 | 3934.49 | 8286.97 | 15874.31 | 24383.93 | 31127.39 | 39495.13 | 179792.93 | 379692.28 | 856.84 | 136.69 | 1 | 3.209% | 12.48%
67.1 | 13.34 | 453.41 | 1897.11 | 4084.51 | 7863.3 | 16041.3 | 23435.37 | 30242.29 | 38783.17 | 171795.02 | 363162 | 1146.88 | 140.02 | 5 | 13.64% | 53.67%
67.15 | 13.34 | 426.74 | 1807.06 | 3454.26 | 6839.36 | 13842.04 | 20718.66 | 27112.61 | 34361.4 | 153628.3 | 326935.53 | 1320.29 | 136.69 | 10 | 18.98% | 95.68%
67.2 | 10 | 283.38 | 1630.37 | 3174.2 | 5922.1 | 12317.13 | 18341.06 | 24317.15 | 30539.44 | 135208.75 | 287649.58 | 1416.96 | 126.69 | 20 | 48.66% | 197.9%
67.25 | 3.33 | 13.34 | 100.02 | 216.7 | 446.75 | 950.19 | 1233.6 | 1857.1 | 2193.91 | 9682.25 | 20835.78 | 1253.61 | 26.67 | 40 | 80.48% | 365.9%
67.3 | 0 | 0 | 0 | 6.67 | 13.34 | 23.34 | 23.34 | 33.34 | 36.68 | 140.02 | 316.72 | 1146.91 | 0 | 60 | 110.4% | 359.6%
67.35 | 0 | 0 | 3.33 | 6.67 | 26.67 | 16.67 | 26.67 | 33.34 | 30 | 150.02 | 373.4 | 1206.92 | 0 | 80 | 163.7% | 696.8%
67.4 | 0 | 3.33 | 10 | 13.34 | 20 | 53.34 | 63.34 | 60.01 | 113.35 | 516.76 | 1006.86 | 1313.6 | 0 | 100 | 177.6% | 864.9%
67.45 | 6.67 | 6.67 | 10 | 13.34 | 16.67 | 66.68 | 43.34 | 110.02 | 176.7 | 693.48 | 71083.9 | 1343.62 | 0 | 500 | 762.7% | 3759%
67.5 | 6.67 | 40.01 | 170.03 | 236.71 | 606.79 | 1003.57 | 1377.02 | 2040.65 | 2214.02 | 9509.77 | 54304.16 | 1246.92 | 6.67 | 1000 | 4355% | 6928%
67.55 | 6.67 | 463.42 | 2013.83 | 3661.09 | 7466.55 | 13886.33 | 19851.46 | 26331.06 | 32960.24 | 143442.41 | 157805.37 | 1496.98 | 163.36
67.6 | 23.34 | 760.15 | 4247.98 | 8287.03 | 16555.16 | 33186.37 | 48691.98 | 63919.82 | 79842.63 | 362177.51 | 635838.67 | 1933.75 | 300.06
67.65 | 26.67 | 1240.25 | 6532.52 | 12290.36 | 24887.81 | 48915.68 | 73000.69 | 96057.64 | 120110.92 | 541476.29 | 1107452.8 | 2794.05 | 516.76
67.7 | 16.67 | 1650.37 | 7139.64 | 14769.77 | 30088.53 | 59123.63 | 88453.32 | 113081.56 | 144978.04 | 646317.5 | 1358009.61 | 2990.8 | 570.11
67.75 | 30.01 | 2010.48 | 8497.19 | 16655.34 | 34264.8 | 67522.65 | 98763.63 | 129142.86 | 162442.13 | 735298.72 | 1576884.24 | 2970.76 | 663.45
67.8 | 50.01 | 2010.45 | 9077.55 | 19141.76 | 38331.97 | 72802.66 | 109664.54 | 143726.69 | 181408.71 | 817791.56 | 1750154.38 | 2357.28 | 676.81
67.85 | 43.34 | 2197.2 | 10125.15 | 19722.7 | 40143.77 | 76993.47 | 115218.89 | 152319.29 | 190210.45 | 862363.89 | 1879472.39 | 2553.99 | 650.12
67.9 | 36.67 | 2017.1 | 9544.66 | 19752.9 | 40438.03 | 80091.02 | 119281.55 | 158408.56 | 196485.43 | 888263.36 | 1932413.79 | 2784.06 | 783.48
67.95 | 53.34 | 2313.93 | 10398.68 | 19996.56 | 42142.66 | 79099.25 | 120692.29 | 159304.32 | 198170.83 | 892194.15 | 1976465.03 | 2653.99 | 673.47
68 | 33.34 | 2340.54 | 9678.03 | 19963.29 | 40758.93 | 80878.61 | 118419.32 | 156616.87 | 197161.18 | 886798.98 | 1962594.78 | 2784.06 | 723.48
| 4.9985 | 58.4989 | 56.9195 | 84.3365 | 67.1714 | 80.5909 | 85.9968 | 76.7485 | 89.0512 | 93.2514 | 36.1408 | 2.2327
|  | Narrow |  | Standard
|  | 0.05 AMU | 0.15 AMU | 0.05 AMU | 0.15 AMU
| 1 | 0% | 0% | 3.209% | 12.48%
| 5 | 0% | 0.1019% | 13.64% | 53.67%
| 10 | 0.9402% | 1.122% | 18.98% | 95.68%
| 20 | 0% | 0.2041% | 48.66% | 197.9%
| 40 | 0.6271% | 0.5102% | 80.48% | 365.9%
| 60 | 0% | 0.3061% | 110.4% | 359.6%
| 80 | 3.137% | 1.225% | 177.6% | 864.9%
| 100 | 0.3131% | 0.8163% | 163.7% | 696.8%
| 500 | 5.642% | 6.327% | 762.7% | 3759%
| 1000 | 3642% | 4010% | 4355% | 6928%

## 68.5
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Narrow
Narrow
68 | 13.33 | 1790.37 | 8467.04 | 16640.3 | 33857.08 | 66531.6 | 98920.54 | 130787.51 | 164781 | 757142.98 | 1626806.82 | 2507.29 | 476.75 |  | 0.05 AMU | 0.15 AMU
68.05 | 43.34 | 1753.72 | 8440.48 | 16411.65 | 32240.23 | 66067.42 | 98175.72 | 130502.73 | 161776.2 | 737512.2 | 1616915.83 | 2570.69 | 490.09 | 1 | 0.1504% | 0.2638%
68.1 | 13.34 | 1793.74 | 7676.54 | 15350.61 | 30258.99 | 59853.37 | 89813.92 | 119061.84 | 149202.18 | 675713.33 | 1544268 | 2640.71 | 586.79 | 5 | 0.1504% | 0.2638%
68.15 | 16.67 | 436.75 | 2037.19 | 3977.89 | 7476.61 | 14787.37 | 21103.29 | 26517.05 | 34131.55 | 141931.49 | 840422.68 | 1366.96 | 86.68 | 10 | 0% | 0.3167%
68.2 | 6.67 | 3.33 | 10 | 20 | 26.67 | 76.68 | 110.02 | 133.36 | 196.7 | 773.47 | 42456.97 | 16.67 | 0 | 20 | 0% | 0.5804%
68.25 | 0 | 0 | 3.33 | 0 | 3.33 | 3.33 | 0 | 20 | 26.67 | 63.34 | 43872.66 | 20 | 0 | 40 | 0.3013% | 0.2111%
68.3 | 0 | 6.67 | 0 | 0 | 10 | 6.67 | 3.33 | 0 | 0 | 13.34 | 42949.46 | 606.78 | 0 | 60 | 0.4517% | 1.055%
68.35 | 13.34 | 0 | 269 | 0 | 0 | 3.33 | 16.67 | 13.34 | 6.67 | 40.01 | 43390.99 | 1343.63 | 0 | 80 | 0.3013% | 0.9498%
68.4 | 0 | 3.33 | 96.97 | 10 | 46.73 | 3.33 | 0 | 16.67 | 0 | 10 | 43861.15 | 1853.76 | 0 | 100 | 0.6026% | 0.9498%
68.45 | 3.33 | 6.67 | 13.34 | 20.01 | 16.67 | 0 | 16.67 | 6.67 | 10 | 63377.36 | 44396.06 | 2030.45 | 3.33 | 500 | 1863% | 2366%
68.5 | 0 | 3.33 | 3.33 | 0 | 0 | 6.67 | 10 | 6.67 | 13.34 | 41253.7 | 41594.83 | 2213.85 | 0 | 1000 | 1879% | 2031%
68.55 | 0 | 6.67 | 0 | 0 | 20 | 6.67 | 40.01 | 46.67 | 36.67 | 44841.35 | 42339.48 | 2073.85 | 0
68.6 | 3.33 | 3.33 | 23.34 | 20 | 36.67 | 76.75 | 100.02 | 83.35 | 110.02 | 47217.47 | 44615.66 | 1967.13 | 0
68.65 | 0 | 3.33 | 26.67 | 40.01 | 70.01 | 180.04 | 200.03 | 296.72 | 477.16 | 45688.29 | 45819.2 | 820.16 | 0
68.7 | 0 | 16.67 | 86.68 | 236.71 | 350.06 | 796.83 | 1453.69 | 1997.24 | 2287.28 | 48383.01 | 48372.69 | 40.01 | 0
68.75 | 0 | 653.45 | 2990.86 | 6015.73 | 12314.83 | 25860.73 | 39260.93 | 52837.2 | 63726.14 | 133529.77 | 242548.98 | 636.79 | 46.68
68.8 | 33.36 | 2485.16 | 11676.53 | 22791.22 | 46075.71 | 92283.05 | 138319.29 | 184631.8 | 229228.22 | 1006987.46 | 2012595.44 | 5902.19 | 76.68
68.85 | 6.67 | 3124.26 | 14940.05 | 30278.76 | 60579.78 | 120065.4 | 180787.88 | 243567.16 | 304626.58 | 1493971.21 | 3009464.18 | 9868.18 | 153.36
68.9 | 6.67 | 3287.77 | 17035.82 | 34051.07 | 67222.15 | 135733.85 | 205476.12 | 273324.36 | 341450.94 | 1739500.05 | 3450356.33 | 11993.62 | 150.03
68.95 | 66.73 | 3831.16 | 18448.05 | 37162.81 | 72601.65 | 147646.49 | 218839.58 | 289521.17 | 369493.36 | 1906373.08 | 3824213.99 | 13608.47 | 216.7
69 | 0 | 3691.11 | 18774.9 | 37145.34 | 74923.1 | 149688.16 | 224166.64 | 300072.36 | 375463.7 | 2000195.49 | 3977346 | 14519.54 | 240.04
Standard Res
Standard
68 | 33.34 | 2340.54 | 9678.03 | 19963.29 | 40758.93 | 80878.61 | 118419.32 | 156616.87 | 197161.18 | 886798.98 | 1962594.78 | 2784.06 | 723.48 |  | 0.05 AMU | 0.15 AMU
68.05 | 16.67 | 2217.2 | 9874.79 | 20063.38 | 40187.15 | 78486.35 | 117082.28 | 155095.5 | 192613.39 | 873197.71 | 1950824.34 | 2634 | 683.47 | 1 | 2.989% | 15.45%
68.1 | 46.69 | 2187.17 | 9658.16 | 19746.5 | 38566.13 | 74875.89 | 113155.53 | 148106.56 | 184441.75 | 840071.85 | 1878082.54 | 3094.14 | 640.12 | 5 | 9.865% | 86.67%
68.15 | 20 | 1937.11 | 8760.68 | 17753.34 | 34763 | 68024.86 | 103059.17 | 135212.69 | 171250.86 | 775195.94 | 1740888.27 | 3691.09 | 620.11 | 10 | 28.55% | 172.8%
68.2 | 20 | 1703.7 | 8003.44 | 16289.89 | 31331.11 | 62310.04 | 93241.28 | 122613.52 | 153680.6 | 694591.63 | 1582396.65 | 3724.41 | 696.78 | 20 | 51.72% | 336.7%
68.25 | 16.67 | 390.06 | 1997.12 | 4304.7 | 8784.14 | 17664.29 | 26914.08 | 36382.86 | 45003.74 | 198467.58 | 976279.3 | 3237.6 | 190.03 | 40 | 91.78% | 658.6%
68.3 | 0 | 6.67 | 20 | 16.67 | 33.34 | 83.35 | 120.02 | 170.03 | 256.71 | 823.49 | 49498.34 | 1890.43 | 0 | 60 | 129.9% | 796.3%
68.35 | 0 | 0 | 3.33 | 0 | 3.33 | 26.67 | 20 | 26.67 | 46.68 | 313.39 | 51901.21 | 2170.53 | 0 | 80 | 164.3% | 1249%
68.4 | 0 | 0 | 6.67 | 20 | 13.34 | 60.01 | 123.36 | 123.36 | 170.03 | 720.13 | 52447.85 | 2273.91 | 0 | 100 | 197.5% | 1554%
68.45 | 0 | 3.33 | 36.67 | 110.02 | 76.68 | 276.72 | 370.07 | 306.72 | 500.09 | 71006.39 | 52423.92 | 2250.54 | 3.33 | 500 | 2470% | 5275%
68.5 | 0 | 66.68 | 220.04 | 636.79 | 1153.61 | 2047.27 | 2897.65 | 3664.85 | 4405.2 | 55084.39 | 60288.5 | 2230.55 | 10 | 1000 | 2703% | 8982%
68.55 | 16.67 | 1273.59 | 7279.68 | 14276.1 | 28051.68 | 54943.75 | 80214.13 | 104677.62 | 130193.42 | 332589.69 | 668326.94 | 4214.65 | 80.01
68.6 | 46.67 | 2517.32 | 11751.52 | 23799.63 | 46457.52 | 92991.19 | 137239 | 181432.93 | 228382.44 | 1064724.24 | 2148734.8 | 8283.67 | 143.36
68.65 | 23.34 | 2867.4 | 15183.56 | 29868.09 | 61268.55 | 120981.91 | 179491.5 | 235519.52 | 300390.74 | 1475600.75 | 2940918.4 | 11105.93 | 203.37
68.7 | 76.7 | 3624.36 | 18087.24 | 35858.65 | 71824.95 | 141830.06 | 211647.19 | 276613.98 | 350755.85 | 1769039.31 | 3555692.58 | 13418.15 | 230.04
68.75 | 76.68 | 4081.19 | 19549.32 | 38710.32 | 76541.99 | 152514.19 | 227869.97 | 299765.48 | 376648.21 | 1946522.19 | 3878361.28 | 13965.63 | 263.38
68.8 | 30.01 | 3991.27 | 20039.89 | 41694.65 | 82107.3 | 163961.05 | 245169.17 | 320385.86 | 404620.81 | 2093577.6 | 4162562.25 | 14539.44 | 243.4
68.85 | 86.71 | 4404.78 | 21652.69 | 44116.18 | 87946.44 | 178352.41 | 264299.43 | 347306.67 | 435654.63 | 2267741.27 | 4521998.64 | 14803.2 | 363.42
68.9 | 100.04 | 4261.31 | 22674.47 | 45232.66 | 91846.39 | 180641.91 | 270646.68 | 357309.9 | 451805.7 | 2346143.18 | 4666851.88 | 15777.61 | 353.4
68.95 | 50.01 | 4281.33 | 22500.72 | 44784.17 | 90817.35 | 179925.45 | 267555.52 | 351694.9 | 446464.74 | 2364069.31 | 4691264.74 | 15707.61 | 310.05
69 | 23.34 | 4601.47 | 22183.7 | 44735.92 | 90820.66 | 179259.22 | 270388.01 | 354417.03 | 445559.81 | 2340859.35 | 4717877.08 | 16381.53 | 256.71

## 73.5
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Narrow
Narrow
73 | 3.33 | 496.76 | 2430.58 | 4331.33 | 9264.36 | 19425.71 | 28091 | 37005.33 | 47510.23 | 239744.84 | 483763.47 | 2797.4 | 13.34 |  | 0.05 AMU | 0.15 AMU
73.05 | 6.67 | 490.09 | 2403.9 | 4601.42 | 9264.28 | 18674.93 | 28404.95 | 37794.24 | 46456.37 | 239497.54 | 481050.29 | 2937.49 | 10 | 1 | 0% | 0.06667%
73.1 | 0 | 450.08 | 2090.49 | 4591.42 | 9014.2 | 16895.78 | 25886.78 | 35183.64 | 43704.18 | 224304.31 | 454647.27 | 2727.37 | 20 | 5 | 0% | 0%
73.15 | 0 | 80.01 | 516.76 | 976.87 | 1887.11 | 3454.38 | 4761.71 | 5869.06 | 8120.51 | 37272.23 | 76073.55 | 2343.9 | 0 | 10 | 0% | 0%
73.2 | 0 | 10 | 0 | 6.67 | 3.33 | 16.67 | 16.67 | 20 | 23.34 | 196.7 | 293.38 | 20 | 6.67 | 20 | 0% | 0%
73.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 46.67 | 3.33 | 20 | 40 | 0.2002% | 0.2%
73.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 26.67 | 3.33 | 236.71 | 60 | 0.4009% | 0.6009%
73.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 6.67 | 3.33 | 1043.56 | 80 | 1.002% | 0.534%
73.4 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 20 | 10 | 0 | 1270.26 | 100 | 0.2002% | 0.2669%
73.45 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 6.67 | 33.34 | 63088.95 | 0 | 1567.01 | 500 | 1.804% | 2.737%
73.5 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 16.67 | 3.33 | 30.01 | 42467.32 | 0 | 1663.71 | 1000 | 2553% | 2974%
73.55 | 3.33 | 3.33 | 0 | 0 | 0 | 3.33 | 23.34 | 10 | 3.33 | 73.34 | 43002.59 | 0 | 1763.77
73.6 | 0 | 0 | 0 | 0 | 10 | 23.34 | 40.01 | 23.34 | 63.34 | 260.05 | 44166.66 | 0 | 1700.4
73.65 | 0 | 0 | 20 | 33.34 | 26.67 | 73.34 | 136.69 | 180.03 | 260.05 | 1303.63 | 43681.94 | 0 | 1223.61
73.7 | 0 | 16.67 | 40.01 | 116.69 | 266.72 | 540.1 | 826.84 | 1040.23 | 1427.03 | 7357.67 | 47228.49 | 6.67 | 20
73.75 | 0 | 170.03 | 916.86 | 1780.43 | 3604.64 | 7100.38 | 11238.17 | 15561.01 | 18790.48 | 96408.56 | 94347.37 | 56.68 | 3.33
73.8 | 10 | 813.5 | 3824.56 | 8257.08 | 16338.83 | 33478.63 | 51167.72 | 68305.06 | 82315.27 | 434246.85 | 607948.23 | 293.38 | 156.7
73.85 | 13.34 | 1613.69 | 7893.37 | 16381.84 | 32196.65 | 64193.34 | 95965.11 | 126752.99 | 159713.68 | 819938.44 | 1527188.99 | 830.16 | 530.1
73.9 | 6.67 | 2163.86 | 9814.76 | 20370.64 | 40394.68 | 80145.26 | 120082.38 | 160803.56 | 202106.66 | 1027470.37 | 2043236.13 | 1310.27 | 1070.22
73.95 | 6.67 | 2477.27 | 11936.67 | 23582.48 | 47934.59 | 95836.12 | 143280.47 | 190855.31 | 238900.74 | 1243585.06 | 2465251.02 | 1787.1 | 1446.99
74 | 3.33 | 2337.24 | 11753.3 | 23896.67 | 47506.42 | 95837.35 | 142433.45 | 190412.55 | 237789.29 | 1270443.33 | 2600225.55 | 1683.72 | 1520.36
Standard Res
Standard
73 | 6.67 | 503.43 | 2880.76 | 5795.42 | 11422.87 | 22891.68 | 34809.7 | 45760.96 | 57232.17 | 291186.66 | 587551.31 | 3280.95 | 33.34 |  | 0.05 AMU | 0.15 AMU
73.05 | 0 | 606.78 | 2700.66 | 5685.46 | 11583.02 | 22417.15 | 33539.74 | 44396.73 | 55228.24 | 279707.29 | 567567.43 | 3037.45 | 40.01 | 1 | 1.629% | 7.047%
73.1 | 3.33 | 563.44 | 2664.06 | 5398.54 | 10995.82 | 21405.41 | 32460.77 | 42018.71 | 53182.78 | 274592.12 | 550562.74 | 3127.51 | 346.74 | 5 | 7.851% | 33.56%
73.15 | 0 | 536.76 | 2457.31 | 4924.94 | 9833.06 | 20016.69 | 29146.63 | 38572.48 | 48619.77 | 248638.01 | 501874.57 | 3004.18 | 1120.23 | 10 | 13.18% | 66.54%
73.2 | 6.67 | 440.08 | 2323.91 | 4458 | 9204.33 | 17423.05 | 26331.12 | 35211.1 | 43407.28 | 221099.65 | 445975.37 | 2490.64 | 1467.01 | 20 | 29.18% | 124.3%
73.25 | 0 | 146.7 | 503.43 | 1053.56 | 2137.23 | 4288.02 | 6872.9 | 8610.76 | 10792.82 | 54284.9 | 109182.53 | 2093.86 | 1797.08 | 40 | 56% | 234.5%
73.3 | 0 | 0 | 0 | 6.67 | 16.67 | 23.34 | 13.33 | 26.67 | 63.34 | 313.39 | 470.09 | 63.34 | 1913.78 | 60 | 83.26% | 345.5%
73.35 | 0 | 0 | 3.33 | 0 | 0 | 10 | 23.34 | 6.67 | 23.34 | 140.03 | 333.4 | 0 | 1930.48 | 80 | 100.6% | 439.9%
73.4 | 0 | 0 | 0 | 3.33 | 3.33 | 30.01 | 20 | 36.67 | 73.35 | 256.71 | 563.44 | 0 | 1997.13 | 100 | 130.1% | 566.1%
73.45 | 0 | 6.67 | 26.67 | 36.67 | 130.02 | 233.37 | 320.06 | 406.74 | 496.76 | 69426.61 | 70684.7 | 0 | 2397.33 | 500 | 2425% | 3204%
73.5 | 0 | 36.67 | 176.7 | 296.72 | 656.81 | 1260.32 | 1873.88 | 2264.09 | 2927.71 | 54586.6 | 56381.37 | 10 | 2250.59 | 1000 | 2505% | 4207%
73.55 | 13.34 | 446.76 | 2130.51 | 4294.81 | 7860.45 | 14817.63 | 21836.15 | 27921.85 | 35949.71 | 98806.17 | 165565.45 | 160.03 | 2307.25
73.6 | 3.33 | 1210.23 | 6175.68 | 13057.99 | 25763.43 | 49511.52 | 75141.78 | 97315.02 | 121629.25 | 512035.22 | 1066730.54 | 623.45 | 2644.06
73.65 | 13.34 | 1937.12 | 8507.15 | 17753.51 | 33580.32 | 67620.89 | 100811.82 | 132889.75 | 167236.09 | 821053.95 | 1646889.85 | 996.88 | 2907.48
73.7 | 16.67 | 2093.82 | 10051.72 | 21609.04 | 41865.57 | 83567.25 | 122989.24 | 162772.77 | 205574.84 | 1029334.27 | 2082007.64 | 1273.6 | 3034.21
73.75 | 16.67 | 2300.54 | 12560.68 | 25282.05 | 50510.3 | 97826.59 | 146779.78 | 192717.05 | 244869.45 | 1250473.23 | 2494628.65 | 1713.71 | 3167.58
73.8 | 46.67 | 2627.36 | 12580.77 | 26070.38 | 52928.84 | 106696.7 | 157253.34 | 205531.83 | 261128.17 | 1369923.21 | 2736475.76 | 1710.39 | 1813.76
73.85 | 13.34 | 2934.12 | 13962.06 | 27222.6 | 54076.62 | 109264.89 | 163473.99 | 213289.38 | 267251.41 | 1415464.64 | 2863582.88 | 1847.09 | 1747.06
73.9 | 30 | 2990.83 | 14192.38 | 28388.35 | 56228.01 | 112646.68 | 167482.23 | 221532.26 | 277695.49 | 1479285.87 | 2984588.56 | 1933.82 | 1813.77
73.95 | 23.34 | 2814.07 | 14476.07 | 28742.21 | 57059.27 | 113656.3 | 170117.54 | 224448.61 | 283463.17 | 1504140.01 | 3009796.53 | 1923.83 | 1753.74
74 | 30.01 | 2517.29 | 14349.28 | 28832.65 | 56653.24 | 115023.9 | 169755.1 | 224421.05 | 279889.51 | 1500044.98 | 3013541.06 | 2017.18 | 1767.07

## 74.5
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Narrow
Narrow
74 | 3.33 | 2337.24 | 11753.3 | 23896.67 | 47506.42 | 95837.35 | 142433.45 | 190412.55 | 237789.29 | 1270443.33 | 2600225.55 | 1683.72 | 1520.36 |  | 0.05 AMU | 0.15 AMU
74.05 | 10 | 2403.92 | 11653.16 | 22951.45 | 46235.86 | 93327.35 | 138704.06 | 186065.76 | 234975.28 | 1245429.05 | 2518032.43 | 1910.46 | 1370.29 | 1 | 0% | 0%
74.1 | 6.67 | 2260.54 | 11032.46 | 21863.15 | 44146.04 | 86776.63 | 131612.66 | 176042.57 | 217386.16 | 1183715.35 | 2409908.49 | 1750.41 | 1233.59 | 5 | 0% | 0.06735%
74.15 | 0 | 483.42 | 2380.62 | 4141.34 | 7666.85 | 15054.54 | 21771.82 | 26792.35 | 36049.51 | 495551.19 | 1170966.82 | 993.54 | 1163.58 | 10 | 0% | 0%
74.2 | 0 | 0 | 13.34 | 3.33 | 56.68 | 83.35 | 140.02 | 126.69 | 213.37 | 40508.63 | 43339.43 | 20 | 10 | 20 | 0.4133% | 0.4047%
74.25 | 0 | 0 | 0 | 0 | 6.67 | 3.33 | 6.67 | 16.67 | 13.34 | 38725.22 | 45257.03 | 0 | 26.67 | 40 | 0% | 0.06735%
74.3 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 6.67 | 6.67 | 39957.52 | 44096.06 | 0 | 270.04 | 60 | 0% | 0%
74.35 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 40779.24 | 43346.41 | 0 | 886.85 | 80 | 0% | 0.1349%
74.4 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 0 | 41376.12 | 43705.68 | 3.33 | 1173.57 | 100 | 0% | 0%
74.45 | 3.33 | 0 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 3.33 | 16.67 | 0 | 1580.36 | 500 | 0.2064% | 0.6067%
74.5 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 3.33 | 43.34 | 0 | 1613.69 | 1000 | 2.686% | 3.169%
74.55 | 0 | 0 | 3.33 | 0 | 13.34 | 3.33 | 0 | 0 | 0 | 23.34 | 96.68 | 0 | 1750.39
74.6 | 0 | 0 | 0 | 0 | 10 | 13.34 | 10 | 30 | 46.68 | 106.68 | 236.7 | 0 | 1833.77
74.65 | 0 | 0 | 3.33 | 6.67 | 10 | 16.67 | 56.68 | 83.35 | 86.68 | 386.73 | 890.19 | 0 | 1450.3
74.7 | 0 | 6.67 | 10 | 53.34 | 80.01 | 223.37 | 276.71 | 383.4 | 373.4 | 2433.97 | 5328.82 | 3.33 | 40.01
74.75 | 0 | 33.34 | 143.35 | 276.71 | 656.81 | 1096.9 | 1973.91 | 2757.6 | 3434.53 | 17128.35 | 36802.36 | 3.33 | 20
74.8 | 0 | 336.73 | 1637.03 | 3137.53 | 6309.15 | 13044.88 | 19193 | 26185.19 | 32208.13 | 164930.63 | 344163.25 | 100.02 | 126.69
74.85 | 3.33 | 866.83 | 3667.72 | 7706.62 | 15454.13 | 30543.05 | 44069.34 | 60211.27 | 74826.47 | 375250.76 | 781849.59 | 670.13 | 680.13
74.9 | 0 | 976.85 | 4594.82 | 9010.8 | 17650.28 | 36082.94 | 53906.47 | 72173.52 | 91163.81 | 451205.52 | 943092.57 | 916.85 | 883.5
74.95 | 6.67 | 1016.87 | 4908.3 | 9421.18 | 18571.49 | 38432.3 | 57784.72 | 76619.98 | 95514.73 | 481756.12 | 1004608.74 | 1070.21 | 1023.53
75 | 0 | 1043.55 | 4708.18 | 9716.47 | 20120.43 | 40003.82 | 59679.84 | 80354.14 | 98789.87 | 497207.36 | 1047394.09 | 1143.57 | 1070.22
Standard Res
Standard
74 | 30.01 | 2517.29 | 14349.28 | 28832.65 | 56653.24 | 115023.9 | 169755.1 | 224421.05 | 279889.51 | 1500044.98 | 3013541.06 | 2017.18 | 1767.07 |  | 0.05 AMU | 0.15 AMU
74.05 | 23.34 | 2850.76 | 14202.35 | 28345.22 | 55204.62 | 111710.86 | 166178.51 | 218117.5 | 273296.46 | 1471543.67 | 2948898.3 | 1860.45 | 1910.46 | 1 | 0.3078% | 1.753%
74.1 | 26.67 | 2647.38 | 13201.22 | 27403.24 | 53775.65 | 106986.77 | 159902.05 | 211145.46 | 265597.2 | 1432817.06 | 2888419.49 | 1770.39 | 2377.25 | 5 | 1.846% | 10.27%
74.15 | 33.34 | 2417.25 | 12387.28 | 25032.04 | 47800.85 | 97752.51 | 146171.06 | 192670.74 | 244434.21 | 1333264.91 | 2671343.97 | 1617.05 | 2533.99 | 10 | 3.846% | 20.84%
74.2 | 36.67 | 2273.89 | 11192.69 | 22243.19 | 44313.06 | 88960.22 | 132590.26 | 174440.35 | 217604.46 | 1210480.5 | 2427526.8 | 1513.67 | 2860.79 | 20 | 8.153% | 39.53%
74.25 | 13.33 | 520.09 | 2990.8 | 6425.79 | 12430.97 | 25354.04 | 38880.93 | 51999.93 | 66940.43 | 752480.08 | 1512605.52 | 1056.89 | 2944.16 | 40 | 14% | 72.69%
74.3 | 0 | 0 | 13.33 | 26.67 | 63.34 | 90.01 | 110.02 | 240.04 | 233.38 | 47905.57 | 52779.2 | 56.68 | 1773.73 | 60 | 18% | 104.6%
74.35 | 3.33 | 0 | 0 | 3.33 | 0 | 13.34 | 26.67 | 23.34 | 36.67 | 51160.16 | 51937.62 | 13.34 | 1890.44 | 80 | 23.69% | 138.6%
74.4 | 0 | 0 | 0 | 10 | 13.34 | 10 | 23.34 | 40.01 | 46.67 | 50537.34 | 49815.99 | 0 | 1963.8 | 100 | 31.08% | 179%
74.45 | 0 | 6.67 | 13.34 | 20 | 23.34 | 70.01 | 60.01 | 100.02 | 143.36 | 690.14 | 17375.63 | 0 | 2377.27 | 500 | 134% | 840.6%
74.5 | 0 | 6.67 | 40.01 | 83.35 | 176.7 | 303.39 | 390.08 | 513.44 | 673.47 | 2904.32 | 20570.98 | 0 | 2167.17 | 1000 | 949.2% | 2164%
74.55 | 0 | 106.68 | 650.13 | 1323.64 | 2507.31 | 4604.99 | 6713 | 8878.12 | 11444.18 | 53975.93 | 110224.23 | 33.34 | 2303.93
74.6 | 13.33 | 500.09 | 2420.62 | 4831.57 | 9404.44 | 18805.36 | 27253.54 | 35532.78 | 44806.34 | 224987.44 | 451440.35 | 283.38 | 2280.59
74.65 | 33.34 | 780.15 | 3807.76 | 7753.34 | 15340.44 | 31641.9 | 46035.5 | 60279.11 | 75380.61 | 376800.19 | 783836.76 | 846.83 | 2824.1
74.7 | 66.68 | 996.86 | 4714.91 | 9581.36 | 18524.42 | 36801.48 | 54803.55 | 72592.74 | 90576.43 | 448194.61 | 940879.75 | 896.86 | 2747.43
74.75 | 30 | 1070.21 | 4921.69 | 10715.46 | 20851.67 | 41150.06 | 61775.15 | 79860.92 | 101165.31 | 502514.18 | 1064517.76 | 1120.22 | 2704.06
74.8 | 53.34 | 1206.94 | 5755.42 | 10765.64 | 21792.64 | 43624.8 | 64559.57 | 83157.58 | 107184.04 | 532306.52 | 1118097.78 | 1136.91 | 1203.58
74.85 | 56.68 | 1183.56 | 5875.48 | 11503.17 | 22804.55 | 46824.3 | 68946.67 | 89413.12 | 112976.07 | 559737.51 | 1189779.65 | 1223.59 | 1340.3
74.9 | 56.68 | 1226.91 | 5805.39 | 12180.32 | 23485.58 | 46964.73 | 71054.81 | 93612.14 | 117672.02 | 580860.41 | 1234164.48 | 1206.93 | 1316.96
74.95 | 33.34 | 1343.62 | 5995.57 | 12250.14 | 24577.57 | 48504.33 | 71858.44 | 94456.75 | 119488.38 | 593301.05 | 1290775.04 | 1510.34 | 1460.33
75 | 33.34 | 1303.59 | 6409.09 | 12025.26 | 24230 | 49119.32 | 71737.78 | 96018.53 | 120669.81 | 589861.46 | 1298202.44 | 1293.62 | 1467

## 77.5
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Narrow
Narrow
77 | 0 | 30 | 160.03 | 386.73 | 783.48 | 1403.64 | 2047.15 | 2747.37 | 3711.08 | 17167.78 | 37646.71 | 36.67 | 3294.3 |  | 0.05 AMU | 0.15 AMU
77.05 | 0 | 46.67 | 233.38 | 330.06 | 706.81 | 1390.28 | 2270.59 | 2610.68 | 3554.4 | 16822.29 | 36738.21 | 33.34 | 3601.08 | 1 | 0% | 0%
77.1 | 0 | 20 | 186.7 | 366.73 | 706.81 | 1400.31 | 2133.87 | 2547.31 | 3410.93 | 15780.9 | 34124.76 | 30 | 3637.77 | 5 | 0% | 0%
77.15 | 0 | 13.33 | 46.67 | 80.01 | 110.02 | 203.37 | 270.05 | 376.73 | 513.43 | 2327.31 | 4718.34 | 6.67 | 2884.11 | 10 | 0% | 0%
77.2 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 3.33 | 0 | 20 | 13.33 | 3.33 | 43.34 | 20 | 0% | 0%
77.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.67 | 40 | 0% | 0%
77.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 133.36 | 60 | 0% | 0.07864%
77.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 710.14 | 80 | 0% | 0%
77.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 996.88 | 100 | 0% | 0%
77.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 1260.26 | 500 | 0% | 0.07864%
77.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1486.99 | 1000 | 0% | 0.2362%
77.55 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 10 | 0 | 1487
77.6 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 6.67 | 0 | 6.67 | 33.34 | 3.33 | 1537.01
77.65 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 13.34 | 6.67 | 50.01 | 146.69 | 3.33 | 1220.26
77.7 | 0 | 0 | 0 | 6.67 | 0 | 13.34 | 10 | 36.67 | 33.34 | 183.36 | 483.43 | 0 | 63.34
77.75 | 0 | 3.33 | 20.01 | 40.01 | 46.67 | 203.37 | 363.42 | 386.74 | 473.42 | 2507.39 | 5472.61 | 3.33 | 16.67
77.8 | 6.67 | 40.01 | 326.73 | 563.44 | 1233.58 | 2423.95 | 3504.3 | 4531.44 | 5665.41 | 27222.96 | 60443.74 | 73.35 | 433.41
77.85 | 6.67 | 110.02 | 380.07 | 963.53 | 1953.79 | 3641.03 | 5245.14 | 6999.46 | 8643.94 | 42872.76 | 94077.3 | 86.68 | 1243.61
77.9 | 26.67 | 116.69 | 503.42 | 1116.89 | 2137.18 | 4341.32 | 6515.83 | 8537.12 | 10385.29 | 51744.98 | 111533.51 | 163.36 | 1870.41
77.95 | 13.34 | 146.69 | 550.1 | 1093.56 | 2217.19 | 4518.12 | 6732.67 | 8934.18 | 11413.02 | 55272.02 | 120483.73 | 153.36 | 2307.26
78 | 26.67 | 186.7 | 566.77 | 1306.97 | 2423.91 | 4728.23 | 7286.29 | 8974.24 | 11833.42 | 57115.04 | 124161.59 | 200.03 | 2420.62
Standard Res
Standard
77 | 20 | 80.01 | 226.71 | 400.07 | 753.47 | 1803.72 | 2320.57 | 3324.29 | 4014.55 | 20463.97 | 43918.11 | 183.37 | 4488.17 |  | 0.05 AMU | 0.15 AMU
77.05 | 13.33 | 73.35 | 226.7 | 460.09 | 900.18 | 1620.36 | 2320.56 | 3404.27 | 4024.53 | 20357.28 | 43686.89 | 93.35 | 4588.19 | 1 | 0% | 0.168%
77.1 | 43.34 | 63.34 | 253.37 | 423.41 | 856.84 | 1553.65 | 2537.32 | 3140.82 | 3921.19 | 19379.32 | 41504.07 | 140.02 | 4975.02 | 5 | 0.1699% | 1.344%
77.15 | 20 | 46.67 | 300.05 | 393.4 | 750.14 | 1553.68 | 2337.26 | 2987.49 | 3721.13 | 17379.69 | 38438.67 | 136.69 | 5258.61 | 10 | 1.36% | 3.361%
77.2 | 16.67 | 73.35 | 153.36 | 316.72 | 673.46 | 1386.96 | 1920.47 | 2730.73 | 3450.99 | 15083.34 | 35100.05 | 96.68 | 4885 | 20 | 1.531% | 5.21%
77.25 | 10 | 20 | 46.67 | 100.01 | 183.37 | 313.39 | 543.43 | 773.49 | 910.17 | 4421.37 | 9401.44 | 33.34 | 4738.33 | 40 | 2.721% | 9.579%
77.3 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 0 | 3.33 | 0 | 13.34 | 23.34 | 0 | 1870.44 | 60 | 2.891% | 11.54%
77.35 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 0 | 3.33 | 6.67 | 0 | 1897.16 | 80 | 3.401% | 15.57%
77.4 | 0 | 0 | 0 | 0 | 0 | 6.67 | 0 | 0 | 3.33 | 23.34 | 20 | 6.67 | 1833.77 | 100 | 4.081% | 22.18%
77.45 | 0 | 0 | 0 | 3.33 | 0 | 10 | 0 | 3.33 | 10 | 116.69 | 166.7 | 0 | 1973.8 | 500 | 19.9% | 93.96%
77.5 | 0 | 0 | 3.33 | 26.67 | 30.01 | 53.34 | 56.68 | 66.68 | 80.01 | 390.08 | 813.54 | 6.67 | 1960.47 | 1000 | 41.5% | 210.8%
77.55 | 23.34 | 10 | 76.68 | 170.03 | 280.05 | 506.76 | 630.12 | 856.83 | 1230.27 | 5085.37 | 11567.94 | 70.01 | 2017.15
77.6 | 43.34 | 96.68 | 410.07 | 923.52 | 1527.01 | 2764.07 | 3864.5 | 5378.54 | 6876.24 | 32327.23 | 69597.91 | 190.03 | 2697.39
77.65 | 80.01 | 216.7 | 466.75 | 1106.9 | 1907.1 | 3601.05 | 5345.13 | 6819.33 | 8537.2 | 41150.16 | 91105.36 | 250.04 | 3207.61
77.7 | 40.01 | 230.04 | 713.47 | 1126.89 | 2097.16 | 4191.32 | 6465.86 | 8417.17 | 10785.77 | 50547.37 | 109201.87 | 310.05 | 3624.38
77.75 | 76.68 | 173.37 | 670.12 | 1310.31 | 2637.35 | 4965.07 | 6966.16 | 9654.8 | 12140.4 | 57512.39 | 124130.84 | 353.39 | 3831.2
77.8 | 103.35 | 176.7 | 706.8 | 1313.6 | 2570.64 | 5018.34 | 7913.44 | 9768.21 | 12143.72 | 60218.66 | 130650.99 | 400.07 | 2717.4
77.85 | 133.36 | 290.06 | 800.16 | 1523.66 | 2954.13 | 5248.48 | 7903.41 | 10372 | 13201.34 | 64368.05 | 139057.58 | 473.42 | 2887.44
77.9 | 123.35 | 223.37 | 866.82 | 1593.69 | 2944.14 | 5938.77 | 8777.41 | 11289.51 | 13815.59 | 66625.72 | 145437.56 | 446.75 | 2780.74
77.95 | 106.68 | 246.71 | 750.14 | 1567.03 | 2884.08 | 5802.18 | 8473.89 | 11115.91 | 13808.76 | 68890.35 | 148491.36 | 470.08 | 3010.8
78 | 106.68 | 196.7 | 826.83 | 1667.02 | 3124.16 | 5675.42 | 8714.06 | 10955.89 | 14202.32 | 67149.2 | 146726.07 | 410.08 | 2850.78
78.05 | 110.02 | 240.05 | 703.46 | 1426.98 | 2844.07 | 5785.37 | 8417.13 | 10678.95 | 13488.3 | 65856.31 | 142518.86 | 473.42 | 2837.41
| #DIV/0! | #DIV/0! | 248.2973 | 62.5054 | 104.104 | 106.4008 | 153.7414 | 164.3055 | 177.5068 | 172.1421 | 180.3551
| 13.8552 | 68.346 | 28.5018 | 23.3017 | 28.5503 | 30.2806 | 37.2817 | 35.3359 | 31.4323 | 36.1035 | 34.8844
| 0.072175 | 0.014631 | 0.035085 | 0.042915 | 0.035026 | 0.033024 | 0.026823 | 0.0283 | 0.031814 | 0.027698 | 0.028666
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1
77.45 | 0 | 0 | 0 | 3.33 | 0 | 10 | 0 | 3.33 | 10 | 116.69 | 166.7 | 0 | 1973.8
77.5 | 0 | 0 | 3.33 | 26.67 | 30.01 | 53.34 | 56.68 | 66.68 | 80.01 | 390.08 | 813.54 | 6.67 | 1960.47
77.55 | 23.34 | 10 | 76.68 | 170.03 | 280.05 | 506.76 | 630.12 | 856.83 | 1230.27 | 5085.37 | 11567.94 | 70.01 | 2017.15
| 7.78 | 3.3333 | 26.67 | 66.6767 | 103.3533 | 190.0333 | 228.9333 | 308.9467 | 440.0933 | 1864.0467 | 4182.7267 | 25.56 | 1983.8067

## 78.5
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1 | Zn, Ga, Ge, As, Se
Narrow
Narrow
78 | 26.67 | 186.7 | 566.77 | 1306.97 | 2423.91 | 4728.23 | 7286.29 | 8974.24 | 11833.42 | 57115.04 | 124161.59 | 200.03 | 2420.62 |  | 0.05 AMU | 0.15 AMU
78.05 | 36.67 | 126.69 | 653.47 | 1250.28 | 2330.56 | 4848.27 | 6942.88 | 9291.04 | 11729.91 | 55010.83 | 122322.55 | 180.03 | 2373.93 | 1 | 0% | 0%
78.1 | 26.67 | 123.35 | 596.78 | 1066.87 | 2217.18 | 4227.99 | 6335.72 | 8480.51 | 10065.03 | 51314.1 | 107788.86 | 173.36 | 2310.59 | 5 | 0% | 0%
78.15 | 0 | 0 | 76.68 | 86.68 | 113.35 | 213.37 | 303.38 | 500.09 | 583.46 | 2624.06 | 5635.74 | 6.67 | 1767.08 | 10 | 0% | 0%
78.2 | 0 | 3.33 | 0 | 0 | 6.67 | 6.67 | 3.33 | 0 | 0 | 16.67 | 26.67 | 0 | 66.68 | 20 | 0% | 0%
78.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 16.67 | 40 | 0% | 0%
78.3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 110.02 | 60 | 0% | 0%
78.35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 706.82 | 80 | 0% | 0%
78.4 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1560.36 | 100 | 0.2125% | 0.06549%
78.45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 1760.42 | 500 | 0% | 0%
78.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 1567.01 | 1000 | 0% | 0%
78.55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1757.07
78.6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1550.36
78.65 | 0 | 0 | 6.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 1356.94
78.7 | 0 | 3.33 | 0 | 3.33 | 6.67 | 0 | 0 | 6.67 | 0 | 6.67 | 3.33 | 0 | 40.01
78.75 | 10 | 3.33 | 13.34 | 13.34 | 10 | 16.67 | 30.01 | 10 | 6.67 | 13.34 | 3.33 | 30.01 | 23.34
78.8 | 66.68 | 93.36 | 80.03 | 90.02 | 73.36 | 50.01 | 93.35 | 83.35 | 43.34 | 96.68 | 80.01 | 150.03 | 333.39
78.85 | 100.02 | 126.69 | 93.35 | 40.01 | 100.02 | 126.69 | 140.03 | 120.02 | 106.7 | 86.68 | 70.01 | 266.73 | 1410.31
78.9 | 116.69 | 160.05 | 93.35 | 113.35 | 150.03 | 166.7 | 203.38 | 170.04 | 163.37 | 120.02 | 116.69 | 280.06 | 2000.48
78.95 | 213.39 | 186.73 | 183.38 | 153.38 | 176.7 | 183.38 | 210.05 | 153.37 | 173.37 | 166.71 | 143.37 | 266.72 | 2387.26
79 | 180.04 | 193.39 | 153.36 | 110.02 | 160.04 | 180.04 | 190.04 | 166.71 | 223.41 | 136.7 | 150.03 | 436.77 | 2960.83
Standard Res
Standard
78 | 106.68 | 196.7 | 826.83 | 1667.02 | 3124.16 | 5675.42 | 8714.06 | 10955.89 | 14202.32 | 67149.2 | 146726.07 | 410.08 | 2850.78 |  | 0.05 AMU | 0.15 AMU
78.05 | 110.02 | 240.05 | 703.46 | 1426.98 | 2844.07 | 5785.37 | 8417.13 | 10678.95 | 13488.3 | 65856.31 | 142518.86 | 473.42 | 2837.41 | 1 | 0% | 0.9008%
78.1 | 96.68 | 223.37 | 796.82 | 1350.29 | 3117.49 | 5615.35 | 8233.75 | 10435.5 | 13651.9 | 63161.58 | 138150.37 | 423.41 | 3404.28 | 5 | 0.2951% | 0.7507%
78.15 | 116.68 | 196.7 | 763.49 | 1430.3 | 2720.74 | 5041.67 | 7593.26 | 9881.65 | 11960.39 | 58634.36 | 126726.33 | 376.73 | 3787.79 | 10 | 0.1473% | 0.4503%
78.2 | 113.35 | 203.37 | 573.44 | 1110.22 | 2163.87 | 4598.1 | 6199.01 | 8417.15 | 10361.96 | 50956.06 | 110378.04 | 373.41 | 3644.39 | 20 | 0.2951% | 0.6506%
78.25 | 30.01 | 10 | 50.01 | 140.02 | 270.04 | 443.42 | 763.49 | 926.85 | 1280.31 | 5492.14 | 12110.93 | 13.33 | 3814.51 | 40 | 0% | 0.9508%
78.3 | 0 | 0 | 0 | 0 | 0 | 3.33 | 10 | 3.33 | 0 | 16.67 | 40.01 | 0 | 2107.16 | 60 | 0.1473% | 0.6005%
78.35 | 0 | 0 | 0 | 0 | 6.67 | 10 | 0 | 0 | 6.67 | 3.33 | 6.67 | 0 | 1807.09 | 80 | 0% | 1.001%
78.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2210.61 | 100 | 0% | 0.5505%
78.45 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 2020.47 | 500 | 0.5901% | 0.5005%
78.5 | 6.67 | 0 | 6.67 | 3.33 | 6.67 | 0 | 3.33 | 0 | 0 | 13.34 | 3.33 | 16.67 | 2260.58 | 1000 | 0.1473% | 0.2502%
78.55 | 40.01 | 60.01 | 43.34 | 26.67 | 33.34 | 63.34 | 36.67 | 66.68 | 36.67 | 16.67 | 13.34 | 103.35 | 2380.59
78.6 | 103.35 | 73.35 | 96.69 | 73.35 | 56.68 | 116.69 | 70.01 | 76.68 | 123.36 | 63.35 | 40.01 | 140.03 | 2824.07
78.65 | 163.38 | 170.05 | 116.69 | 153.37 | 136.7 | 126.7 | 216.72 | 113.36 | 106.69 | 86.68 | 80.01 | 290.07 | 3627.72
78.7 | 186.71 | 146.7 | 173.38 | 176.71 | 143.36 | 126.7 | 210.04 | 170.05 | 130.03 | 140.03 | 76.68 | 266.73 | 4151.26
78.75 | 126.7 | 200.05 | 196.73 | 120.02 | 156.7 | 173.38 | 173.37 | 176.71 | 160.03 | 136.7 | 120.02 | 326.73 | 4154.62
78.8 | 230.07 | 213.39 | 173.38 | 170.04 | 160.03 | 190.04 | 203.38 | 203.38 | 193.39 | 140.02 | 176.71 | 383.43 | 3317.66
78.85 | 270.07 | 200.06 | 220.09 | 216.72 | 223.4 | 240.05 | 226.73 | 223.38 | 243.4 | 153.36 | 103.35 | 400.09 | 3694.45
78.9 | 213.4 | 176.72 | 213.39 | 193.38 | 213.4 | 186.72 | 223.38 | 170.04 | 163.37 | 136.7 | 106.69 | 383.41 | 3907.87
78.95 | 183.39 | 230.05 | 226.73 | 260.09 | 193.38 | 200.05 | 233.39 | 233.41 | 263.41 | 163.36 | 156.7 | 496.78 | 3777.82
79 | 230.06 | 243.4 | 193.38 | 166.71 | 220.06 | 183.37 | 256.74 | 283.43 | 186.72 | 196.71 | 193.37 | 326.74 | 3631.09

## Gd Plot
| 77.5 |  | 78.5
| Narrow | Standard | Narrow | Standard
1 | 0% | 0.168% | 0% | 0.9008%
5 | 0% | 1.344% | 0% | 0.7507%
10 | 0% | 3.361% | 0% | 0.4503%
20 | 0% | 5.21% | 0% | 0.6506%
40 | 0% | 9.579% | 0% | 0.9508%
60 | 0.07864% | 11.54% | 0% | 0.6005%
80 | 0% | 15.57% | 0% | 1.001%
100 | 0% | 22.18% | 0.06549% | 0.5505%
500 | 0.07864% | 93.96% | 0% | 0.5005%
1000 | 0.2362% | 210.8% | 0% | 0.2502%

## Nd Plot
| 71.5 |  | 72.5
| Narrow | Standard | Narrow | Standard
1 | 0% | 4.204% | 0% | 2.791%
5 | 0% | 16.98% | 0% | 11.32%
10 | 0% | 35.31% | 0.08275% | 23.18%
20 | 0% | 70.69% | 0% | 43.64%
40 | 0.3763% | 136.2% | 0.08275% | 79.15%
60 | 0.3765% | 153.7% | 0.08275% | 91.73%
80 | 0.1253% | 241.6% | 0.58% | 157%
100 | 0.3763% | 317.5% | 0% | 193.4%
500 | 2.572% | 1639% | 1.242% | 1025%
1000 | 2491% | 3343% | 2.071% | 2021%

## Sm Plot
| 73.5 |  | 74.5
| Narrow | Standard | Narrow | Standard
1 | 0.06667% | 7.047% | 0% | 1.753%
5 | 0% | 33.56% | 0.06735% | 10.27%
10 | 0% | 66.54% | 0% | 20.84%
20 | 0% | 124.3% | 0.4047% | 39.53%
40 | 0.2% | 234.5% | 0.06735% | 72.69%
60 | 0.6009% | 345.5% | 0% | 104.6%
80 | 0.534% | 439.9% | 0.1349% | 138.6%
100 | 0.2669% | 566.1% | 0% | 179%
500 | 2.737% | 3204% | 0.6067% | 840.6%
1000 | 2974% | 4207% | 3.169% | 2164%

## Ba Plot (2)
| 67.5 |  | 68.5
| Narrow | Standard | Narrow | Standard
1 | 0% | 12.48% | 0.2638% | 15.45%
5 | 0.1019% | 53.67% | 0.2638% | 86.67%
10 | 1.122% | 95.68% | 0.3167% | 172.8%
20 | 0.2041% | 197.9% | 0.5804% | 336.7%
40 | 0.5102% | 365.9% | 0.2111% | 658.6%
60 | 0.3061% | 359.6% | 1.055% | 796.3%
80 | 0.8163% | 696.8% | 0.9498% | 1249%
100 | 1.225% | 864.9% | 0.9498% | 1554%
500 | 6.327% | 3759% | 2366% | 5275%
1000 | 4010% | 6928% | 2031% | 8982%

## .55 plots
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1
68.55 | 0 | 6.67 | 0 | 0 | 20 | 6.67 | 40.01 | 46.67 | 36.67 | 44841.35 | 42339.48 | 2073.85 | 0 | H
68.55 | 16.67 | 1273.59 | 7279.68 | 14276.1 | 28051.68 | 54943.75 | 80214.13 | 104677.62 | 130193.42 | 332589.69 | 668326.94 | 4214.65 | 80.01 | L
69 | 0 | 3691.11 | 18774.9 | 37145.34 | 74923.1 | 149688.16 | 224166.64 | 300072.36 | 375463.7 | 2000195.49 | 3977346 | 14519.54 | 240.04 | H
69 | 23.34 | 4601.47 | 22183.7 | 44735.92 | 90820.66 | 179259.22 | 270388.01 | 354417.03 | 445559.81 | 2340859.35 | 4717877.08 | 16381.53 | 256.71 | L
| 0.71422 | 0.27678 | 0.32815 | 0.31912 | 0.30887 | 0.3065 | 0.29666 | 0.29535 | 0.2922 | 0.14208 | 0.14166
71.55 | 0 | 0 | 0 | 0 | 0 | 10 | 13.34 | 3.33 | 10 | 90.02 | 43733.84 | 1973.82 | 0 | H
71.55 | 0 | 243.38 | 936.85 | 1957.16 | 3831.35 | 7427 | 10709.99 | 13219.86 | 17639.59 | 90511.01 | 92276.78 | 1940.48 | 13.34 | L
72 | 3.33 | 1817.07 | 8006.85 | 16234.71 | 31832.47 | 65241.36 | 97790.75 | 130063.71 | 162927.7 | 829617.1 | 1760195.89 | 3991.19 | 333.4 | H
72 | 10 | 1873.78 | 9521.27 | 19549.28 | 39174.46 | 77254.62 | 115068.48 | 153127.72 | 191774.69 | 973794.55 | 2068479.12 | 4608.18 | 513.43 | L
| 0 | 0.12989 | 0.098395 | 0.10011 | 0.097802 | 0.096137 | 0.093075 | 0.086332 | 0.091981 | 0.092947 | 0.044611
77.55 | 0 | 0 | 0 | 0 | 0 | 0 | 3.33 | 0 | 0 | 3.33 | 10 | 0 | 1487 | H
77.55 | 23.34 | 10 | 76.68 | 170.03 | 280.05 | 506.76 | 630.12 | 856.83 | 1230.27 | 5085.37 | 11567.94 | 70.01 | 2017.15 | L
78 | 26.67 | 186.7 | 566.77 | 1306.97 | 2423.91 | 4728.23 | 7286.29 | 8974.24 | 11833.42 | 57115.04 | 124161.59 | 200.03 | 2420.62 | H
78 | 106.68 | 196.7 | 826.83 | 1667.02 | 3124.16 | 5675.42 | 8714.06 | 10955.89 | 14202.32 | 67149.2 | 146726.07 | 410.08 | 2850.78 | L
| 0.21879 | 0.050839 | 0.09274 | 0.102 | 0.08964 | 0.08929 | 0.072311 | 0.078207 | 0.086625 | 0.075732 | 0.07884
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm
68.55 | 16.67 | 1273.59 | 7279.68 | 14276.1 | 28051.68 | 54943.75 | 80214.13 | 104677.62 | 130193.42 | 332589.69 | 668326.94
71.55 | 0 | 243.38 | 936.85 | 1957.16 | 3831.35 | 7427 | 10709.99 | 13219.86 | 17639.59 | 90511.01 | 92276.78
77.55 | 23.34 | 10 | 76.68 | 170.03 | 280.05 | 506.76 | 630.12 | 856.83 | 1230.27 | 5085.37 | 11567.94

## 0.15 AMU Plot
|  | 1 | 5 | 10 | 20
68.45 | 0 | 3.33 | 36.67 | 110.02 | 76.68 | 276.72 | 370.07 | 306.72 | 500.09 | 71006.39 | 52423.92 | 2250.54 | 3.33
68.5 | 0 | 66.68 | 220.04 | 636.79 | 1153.61 | 2047.27 | 2897.65 | 3664.85 | 4405.2 | 55084.39 | 60288.5 | 2230.55 | 10
68.55 | 16.67 | 1273.59 | 7279.68 | 14276.1 | 28051.68 | 54943.75 | 80214.13 | 104677.62 | 130193.42 | 332589.69 | 668326.94 | 4214.65 | 80.01
| 5.5567 | 447.8667 | 2512.13 | 5007.6367 | 9760.6567 | 19089.2467 | 27827.2833 | 36216.3967 | 45032.9033 | 152893.49 | 260346.4533 | 2898.58 | 31.1133
|  | 15.45% | 86.67% | 172.8% | 336.7% | 658.6% | 960% | 1249% | 1554% | 5275% | 8982% | 1
71.45 | 0 | 0 | 3.33 | 16.67 | 56.68 | 103.35 | 120.02 | 160.03 | 143.35 | 983.52 | 52325.22 | 1973.8 | 0
71.5 | 0 | 6.67 | 70.01 | 126.69 | 316.73 | 573.45 | 760.15 | 990.22 | 1103.57 | 5972.83 | 54254.15 | 2033.83 | 0
71.55 | 0 | 243.38 | 936.85 | 1957.16 | 3831.35 | 7427 | 10709.99 | 13219.86 | 17639.59 | 90511.01 | 92276.78 | 1940.48 | 13.34
| 0 | 83.35 | 336.73 | 700.1733 | 1401.5867 | 2701.2667 | 3863.3867 | 4790.0367 | 6295.5033 | 32489.12 | 66285.3833 | 1982.7033 | 4.4467
|  | 4.204% | 16.98% | 35.31% | 70.69% | 136.2% | 194.9% | 241.6% | 317.5% | 1639% | 3343% | 1
| Blank | 1 ppb | 5 ppb | 10 ppb | 20 ppb | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | Solution 2 | Solution 1
77.45 | 0 | 0 | 0 | 3.33 | 0 | 10 | 0 | 3.33 | 10 | 116.69 | 166.7 | 0 | 1973.8
77.5 | 0 | 0 | 3.33 | 26.67 | 30.01 | 53.34 | 56.68 | 66.68 | 80.01 | 390.08 | 813.54 | 6.67 | 1960.47
77.55 | 23.34 | 10 | 76.68 | 170.03 | 280.05 | 506.76 | 630.12 | 856.83 | 1230.27 | 5085.37 | 11567.94 | 70.01 | 2017.15
| 7.78 | 3.3333 | 26.67 | 66.6767 | 103.3533 | 190.0333 | 228.9333 | 308.9467 | 440.0933 | 1864.0467 | 4182.7267 | 25.56 | 1983.8067
|  | 0.168% | 1.344% | 3.361% | 5.21% | 9.579% | 11.54% | 15.57% | 22.18% | 93.96% | 210.8%
|  | 1 | 5 | 10 | 20
Low res
| Blank | 1 ppb Ge | 5 ppb Ge | 10 ppb Ge | 20 ppb Ge | 40 ppb | 60 ppb | 80 ppb | 100 ppb | 500 ppb | 1 ppm | 50 ppb Nd | Solution 1
71.45 | 0 | 0 | 3.33 | 16.67 | 56.68 | 103.35 | 120.02 | 160.03 | 143.35 | 983.52 | 52325.22 | 1973.8 | 0
71.5 | 0 | 6.67 | 70.01 | 126.69 | 316.73 | 573.45 | 760.15 | 990.22 | 1103.57 | 5972.83 | 54254.15 | 2033.83 | 0
71.55 | 0 | 243.38 | 936.85 | 1957.16 | 3831.35 | 7427 | 10709.99 | 13219.86 | 17639.59 | 90511.01 | 92276.78 | 1940.48 | 13.34
|  | 83.35 | 336.73 | 700.1733 | 1401.5867 | 2701.2667 | 3863.3867 | 4790.0367 | 6295.5033 | 32489.12 | 66285.3833 | 1982.7033