**Rapid Viability Polymerase Chain Reaction Method for Detection of *Francisella tularensis***

Staci R. Kane,a Sanjiv R. Shah,b# and Teneile M. Alfaroa

aLawrence Livermore National Laboratory, Livermore, California, USA

bNational Homeland Security Research Center, Office of Research and Development, U.S. Environmental Protection Agency, Washington, District of Columbia, USA

Running Head: Rapid Viability PCR Method for *Francisella tularensis*

**Declarations of interest: none**

#Address correspondence to Sanjiv R. Shah, shah.sanjiv@epa.gov.

1200 Pennsylvania Ave, NW, 8801R,

US Environmental Protection Agency, National Homeland Security Research Center, Washington, DC, 20460, USA

**Abstract**

*Francisella tularensis,* which causes potentially fatal tularemia, has been considered an attractive agent of bioterrorism and biological warfare due to its low infectious dose, reported environmental persistence, and ability to be transmitted to humans via multiple exposure routes. Due to slow growth on even selective culture media, detection of viable *F. tularensis* from environmental and drinking water samples, usually takes more than 3 days. Therefore, a rapid viability polymerase chain reaction (RV-PCR) method was developed to detect and identify viable *F. tularensis* cells in environmental samples. The method uses a change in PCR response during high throughput (48-well) sample incubation in Brain Heart Infusion/Vitox/Fildes/Histidine growth medium to detect viable *F. tularensis* presence, which is 2-3 times faster than the current plate culture-based method. Using the method, < 102 live *F. tularensis* cells were detected in complex sample matrices containing chemical and biological interferences.

**Highlights**

* Rapid detection of viable *F. tularensis* is critical during a high-consequence tularemia incident.
* Pre- and post-enrichment differential PCR analysis-based method provides more timely results than current culture-based methods.
* The RV-PCR method features high-throughput 48-well sample processing, a small foot-print, reduced labor requirements, and less biohazardous waste generated.
* Even with typical inhibitors found in water, the method detects low levels of viable pathogen cells.

**Key Words**

*Francisella tularensis*, tularemia, bioterrorism, detection, RV-PCR, contamination

*Francisella tularensis* (*F. tularensis*),the pathogen that causes tularemia in humans and animals, could be introduced into water infrastructure due to a natural outbreak, laboratory accident, or intentional contamination. It is known that *F. tularensis* and several other vegetative bacterial pathogens can remain viable and infectious for some time in certain environments including water (Anda et al., 2001; Berrada et al., 2011; Gilbert et al., 2012). As reviewed by Rice (2015), drinking water outbreaks of tularemia have been reported in the U.S. and several countries including Bulgaria, Georgia, Germany, Italy, Kosovo, Norway, Russia, Sweden, and Turkey.

*F. tularensis* is a category A select agent because it can readily be weaponized, has a low infective dose (1 – 10 cells; Jones et al., 2005; Saslaw et al., 1961), and high morbidity and mortality. Therefore, national security concerns have been raised regarding this potential bioterrorism agent (Oyston et al., 2004).  As a result, consensus-based recommendations have been developed for civilian defense if *F. tularensis* is used as a biological weapon (CDC, 2000; Dennis et al., 2001).

*Francisella tularensis* is a Gram-negative coccobacillus whose main transmission vectors are arthropods (ticks, deer flies), while small mammals (e.g., rabbits, muskrats) serve as reservoir hosts. Protozoa have also been suggested as an important environmental reservoir (Abd et al., 2003). A security concern surrounding this bacterium is its potential persistence in the environment. Due to its historical usage as a biological weapon (Gürcan, 2014) and the occurrence of natural tularemia outbreaks, there is a need for rapid and sensitive analytical methods for detection of viable *F. tularensis* in environmental samples.

Methods to more rapidly determine pathogen presence and viability are needed to ensure public safety and to help mitigate impacts of facility and infrastructure closures following a biological agent release. The current plate culture-based methods used for detectionof *F. tularensis* and other biothreat agents are labor-intensive and have a low throughput (~30–40 samples processed per laboratory per 8-hr shift), with confirmed results obtained only after several days. In fact, detection via culture-based methods and subsequent confirmation takes 3 days or longer (Stewart et al., 1995). A review of methods for soil and water sample processing and analysis for *F. tularensis* described challenges with using traditional culture with these complex sample types (US EPA, 2015). In fact, multiple instances were reported where culture identification was unsuccessful.

We recently reported the development of a RV-PCR method for *Yersinia pestis* detection from sample matrices containing similar types of chemical and biological challenge material (Kane et al., 2018). In the current effort, a similar approach was used including high-throughput sample incubation with 48-well plates (E&K Scientific, Santa Clara, CA; Cat. No. EK-2044); however, differences included 1) use of optimized Brain Heart Infusion (BHI)/Vitox/Fildes/Histidine (BVFH) growth medium (Morris et al., 2017), which was concentrated 6-fold rather than 10-fold as described for the *Y. pestis* growth medium, 2) addition of 0.6-mL of 6-fold concentrated BVFH medium to 3-mL sample to yield 1X concentration, 3) incubation at 37°C instead of 30°C, and 4) use of a 30-h incubation period instead of 24-h, since *F. tularensis* has a longer doubling time than *Y. pestis*.

The 3-mL samples were prepared with phosphate buffered saline (PBS; Teknova, Inc., Hollister, CA; Cat. No. P0261) with or without challenge materials including 1) 4 mg/mL Arizona Test Dust (ATD; ISO 12103-1, A3 Medium Test Dust; Powder Technology, Arden Hills, MN) shown to contain ~5 × 104 colony-forming units (CFU) per 10 mg including fungi and bacterial spores (Rose et al., 2011), 2) 10 μg/mL Fe2+ as iron sulfate heptahydrate (Sigma-Aldrich, St. Louis, MO; Cat. No. 215422), and 3) 50 μg/mL humic acid (Sigma-Aldrich, Cat. No. 53680-10G) as a surrogate for natural organic matter. These iron and humic acid concentrations were at the upper end of the range of values expected for drinking water samples (Farrah et al., 1978; NRC, 1979; WHO, 1996; US EPA, 2005). PBS was used as a substitute for water samples because it maintained cell viability and represented a reproducible matrix; in addition, its use is consistent with the EPA protocol for collection, concentration and resuspension of vegetative pathogens from large volume water samples (US EPA, 2017). *F. tularensis* Schu S4 cells were grown in BVFH medium at 37°C (180 rpm) to an OD600 (optical density at 600 nm) of ~ 1.5 which was diluted 10-fold in BVFH medium (to ~1 × 109 CFU/mL), followed by centrifugation (3,000 × *g* for 10 min at 4°C), washing one time in PBS buffer, and repeated centrifugation. The cell pellet was then suspended and diluted in PBS to obtain cell densities from 12-2500 CFU/3-mL based on serial dilution and plating of the inoculum onto Chocolate Agar (Hardy Diagnostics, Santa Maria, CA; Cat. No. E14) with incubation at 37°C for 3 days.

The 6X BVFH medium was prepared as follows: 1) 22.2 g BactoTM Brain Heart Infusion Broth Base powder (BD Biosciences, Franklin Lakes, NJ; Cat. No. 237500) and 0.6 g L-histidine (Sigma-Aldrich, Cat. No. H8000-5G) were dissolved in 28 mL Millipore water and brought to boiling for 1 minute followed by cooling the solution; 2) 12 mL of sterile Vitox Supplement (Thermo Fisher Scientific, Waltham, MA; Cat. No. SR0090A) was added, 3) 60 mL of sterile Fildes Enrichment (Thermo Fisher Scientific, Cat. No. R45037) was added, and 4) the solution was filter-sterilized with a disposable filtration system. The 6X solution was stored at 4°C and brought to room temperature before use. The 6X BVFH was mixed with the water sample (PBS) by pipettor, 0.5-mL aliquots were removed before (T0 aliquot) and after sample incubation at 37°C with shaking at 180 rpm for 30-h (T30 aliquot). The 48-well plate was sealed with a sterile AeraSealTM breathable adhesive seal (Excel Scientific, Victorville, CA; Cat. No. BS-25) during incubation. The aliquots were processed as previously described (Kane et al., 2019) using a modified Promega paramagnetic particle (PMP)-based DNA extraction protocol (MagneSil® Blood Genomic, Max Yield System; Promega, Madison, WI; Cat. No. MD1360) yielding T0 and T30 DNA extracts for PCR analysis.

In some cases, a duplicate aliquot was removed and processed using an automated Roche MagNA Pure Compact instrument with Nucleic Acid Isolation Kit I (Roche, Indianapolis, IN; Cat. No. 03730964001) protocol for comparison purposes. Briefly, the sample aliquots were treated by heat lysis at 70°C for 10 min, cooling for 3 min, addition of 300 μL Roche chemical lysis buffer, followed by vortexing tubes for 10 sec at 1,800-2,000 rpm, and incubation at 22–25°C for 30 min (with 3-5 sec vortexing every 5 min). These crude sample lysates were then processed on the Roche instrument using the pre-programmed DNA-Bacteria Purification Protocol yielding 100 μL purified DNA extracts. A validated inactivation procedure per the Centers for Disease Control and Prevention (CDC) select agent regulations (CDC and USDA, 2018) was approved via LLNL Institutional Biosafety Committee review prior to transfer of crude sample lysates to the robot (within the same laboratory) for subsequent extraction and DNA purification.

Two real-time PCR assays targeting the *pdp*D, pathogenicity determinant protein D unique to Type A *F. tularensis* were used (Table 1) including the F4 assay designed at LLNL and an assay designed at the CDC (Kugeler et al., 2006), refered to here as F5. Real-time PCR analysis was conducted using an Applied Biosystems® 7500 Fast Real-Time PCR System (Foster City, CA) as previously described (Kane et al., 2019). In addition, Platinum™ Taq DNA Polymerase (Plat Taq; Invitrogen, Carlsbad, CA; Cat. No. 10966026) was included at 0.25 µL per reaction (1.25 U) to determine its effect on assay sensitivity. Three replicate samples were analyzed per experimental condition, and three replicate PCR analyses were conducted per sample replicate. DNA extracts from T0 and T30 from the same samples were analyzed on the same plate to standardize the analysis conditions. Both undiluted and 10-fold-diluted DNA extracts (prepared in PCR water) for both T0 and T30 were analyzed to check for PCR inhibition (i.e., if the difference between CT values for 10-fold diluted and undiluted extracts is negative and/or significantly less than three).

For cases where no PCR response was obtained (non-detect [NDT] results), the CT values were set to 45 (since 45 PCR cycles were used), to calculate ∆CT. A ∆CT ≥ 6 represented an increase in DNA concentration of approximately 2-log (100-fold) because of the presence of viable cells in the original sample that propagated during incubation. For individual replicates within an experiment, the RV-PCR result was considered positive when at least two of three replicates met the algorithm requirement. Specifically, if two or three of three PCR replicates were NDT, then a CT value of 45 was assigned as the average T0 CT or T30 CT (as appropriate) to calculate ΔCT; whereas, if two or three of three replicates showed positive CT values, the average of the positive CT values was used to calculate ΔCT. The RV-PCR method sensitivity of detection was equivalent to the *F. tularensis* cell level where 100% of the spiked samples had ∆CT ≥ 6. This value was essentially an analytical sensitivity of detection for the RV-PCR method and did not account for any losses that could occur from sampling and sample handling prior to RV-PCR analysis.

**Table 1. Nucleotide Sequences of the Primer/Probe Sets Used for *Francisella tularensis* RV-PCR Analysis**

| **Assay ID** | **Forward Primer\*** | **Reverse Primer\*** | **Probe\*** | **Amplicon Length (bp)** | **Reference** |
| --- | --- | --- | --- | --- | --- |
| **F4** | TTGCTCCAGTAGCTGCAAGATT | CCAAGTGCTTGGTGGTGGTA | FAM-TGCTGCCGAGATGTTTTCATTATTAACTGATGC-BHQ | 125 | This study |
| **F5** | GAGACATCAATTAAAAGAAGCAATACCTT | CCAAGAGTACTATTTCCGGTTGGT | FAM-AAAATTCTGCTCAGCAGGATTTTGATTTGGTT-BHQ | 105 | Kugeler et al. (2006) |

\* Sequences are listed in 5’ to 3’ orientation. Acronyms: bp, base pair; FAM, fluorescein; BHQ, Black Hole Quencher

Real-time PCR assay optimization showed that addition of 1.25 U Plat Taq per reaction resulted in better sensitivity, with average CT values reduced with the enzyme addition by an average 9.8 and 8.9 CT units for F4 and F5 assays, respectively. Both assays showed detection of 5 fg Schu S4 genomic DNA with Plat Taq added; therefore, it was included along with the AmpliTaq Gold® DNA polymerase from the MasterMix. In addition, DNA extracts from the automated Roche MagNA Pure® Compact showed an average 1.8 ± 0.9 lower CT values than the manual Promega MagneSil® DNA kit extracts. Although both assays consistently demonstrated < 30 genome equivalent sensitivity, other *F. tularensis* assays could readily be integrated into the RV-PCR method, as well.

Concentration of the growth medium enabled use of a larger volume water sample (3-mL plus 0.6 mL 6X BVFH medium per 5-mL well) and demonstrated reproducible growth even at low inoculum levels, 101 CFU/3-mL (data not shown). The resulting RV-PCR method showed good sensitivity, using the criterion of ΔCT values ≥ 6 for positive detection even in the presence of chemical and biological backgrounds (Table 2). For samples containing Fe and humic acid, 101 Schu S4 CFU-level were detected, whereas, 102 CFU-level were detected for samples with ATD (containing background microbes). In general, PCR inhibition did not appear to contribute to higher CT values since the difference between T30 CT values for undiluted and 10-fold diluted extracts from ATD-spiked samples showed roughly a three CT difference (data not shown), as expected for a 10-fold lower DNA concentration. It is suggested that both undiluted and 10-fold diluted DNA extracts (T0 and T30) are analyzed for samples with complex matrices to assess PCR inhibition and provide more data, potentially reducing false negative rates.

Growth of *F. tularensis* cells may have been enhanced by iron and/or nutrients associated with the humic acid formulation. These data suggest that the BVFH growth medium could be further optimized, although this could also enhance growth of competing microbes. Conversely, in the presence of ATD, *F. tularensis* growth appeared to be inhibited for low cell levels (101 CFU/3-mL sample) leading to non-detect results (although growth was not monitored in this case), which could not be attributed to PCR inhibition (i.e., both undiluted and 10-fold diluted DNA extracts were non-detect). Additional testing is needed to assess RV-PCR method performance with a range of real-world water samples, concentrated from larger volumes to improve detection limits, as well as testing with a dead *F. tularensis* cell background for natural decay or post-decontamination scenarios.

**Table 2. RV-PCR results for different starting cell levels per sample with and without chemical and biological backgrounds**

| **Treatment** | **Starting CFU /  3-mL Sample\*** | **ΔCT Range** | **Positive RV-PCR Samples\*\*\*** |
| --- | --- | --- | --- |
| **Control** | 2500 | 7.2–9.3 | 3 of 3 |
| 1200 | 7.4–9.3 | 3 of 3 |
| 250 | 6.8–8.1 | 3 of 3 |
| 120 | 8.9–9.3 | 3 of 3 |
| 25 | 10.6–12.3 | 3 of 3 |
| 12 | 10.3–12.7 | 3 of 3 |
| **Fe/Humic acid** | 2500 | 11.5–17.2 | 3 of 3 |
| 1200 | 14.5–22.7 | 3 of 3 |
| 250 | 12.5–13.9 | 3 of 3 |
| 120 | 15.4–18.3 | 3 of 3 |
| 25 | 16.4–16.8 | 3 of 3 |
| 12 | 15.9–17.8 | 3 of 3 |
| **ATD\*\*** | 2500 | 9.2–10.7 | 3 of 3 |
| 1200 | 5.5–8.3 | 2 of 3 |
| 250 | 6.4–7.5 | 3 of 3 |
| 120 | 7.2–8.2 | 3 of 3 |
| 25 | 0–4.9 | 0 of 3 |
| 12 | 0 - 0 | 0 of 3 |

\* CFU values are averages corrected for dilution, expressed with 2 significant figure accuracy.

\*\* The ΔCT range includes the highest ΔCT value per ATD sample replicate, either from undiluted or 10-fold diluted DNA extracts.

\*\*\* The number of positive samples per three replicate samples.

Acronyms: ATD, Arizona Test Dust; CT, cycle threshold.

While traditional culture methods are considered the gold standard for viability analysis, this effort demonstrated that RV-PCR could detect down to the 102-CFU level per 3-mL sample with background microbes in <36 h (sample processing and analysis time), compared to more than 72 h required for confirmed results from plate culture analysis. Additionally, the RV-PCR method uses a single 48-well plate for 48 samples (and controls) for growth compared with the culture method that uses 11 Chocolate Agar plates, dilution tubes, and an enrichment culture tube per sample, with additional plates to re-streak from the enrichment culture for presumptive *F. tularensis* colony isolation. Therefore, more incubator space (larger laboratory footprint) would be required for plates and tubes compared with 48-well plates used for RV-PCR analysis. Furthermore, by incorporating automated DNA extraction, less consumable waste is generated and DNA extracts can be obtained more rapidly with reduced labor required.

With proper training and proficiency demonstration, the RV-PCR method for detection of viable *F. tularensis* from water samples could help enhance the laboratory capability for rapid, reliable, and high-throughput sample analysis to effectively respond to an intentional, accidental, or natural outbreak incident resulting in water infrastructure contamination. Additionally, subsequent to following sample-type-specific processing and concentration steps, this RV-PCR method could be used for other sample types during a response to a wide-area tularemia incident, although further verification testing would be required (i.e., for different sample types and potential interferences).

**Acknowledgments**

We thank Worth Calfee (EPA) and John Archer (EPA) for technical review of this manuscript.

The United States Environmental Protection Agency through its Office of Research and Development funded and managed the research described here under an Interagency Agreement (EPA IA DW-89-92454101-0). The DOE contractor role did not include establishing Agency policy. It has been subjected to the Agency’s review and has been approved for publication. Note that approval does not signify that the contents necessarily reflect the views of the Agency. Mention of trade names, products, or services does not convey official EPA approval, endorsement, or recommendation.

Lawrence Livermore National Laboratory is operated by Lawrence Livermore National Security (LLNS), LLC, for the U.S. Department of Energy, National Nuclear Security Administration under Contract DE-AC52-07NA27344.

**References**

Abd, H., Johansson, T., Golovliov, I., Sandstrom, G., Forsman, M., 2003. Survival and growth of *Francisella tularensis* in *Acanthamoeba castellanii*. Appl. Environ. Microbiol. 69, 600–606. doi: [10.1128/AEM.69.1.600-606.2003](https://dx.doi.org/10.1128%2FAEM.69.1.600-606.2003)

Anda, P., Segura del Poz, J., J.M., Diaz García, Escudero, R., García Peña, F.J., Lopez Velasco, M.C., Sellek, R.E., Jiménez Chillarón, M.R., Sánchez Serrano, L.P., Martínez Navarro, J.F., 2001. Waterborne outbreak of tularemia associated with crayfish fishing. Emerg. Infect. Dis. 7, 575-582. doi: 10.3201/eid0707.010740

Berrada, Z.L. and Telford S.R, III. 2011. Survival of *Francisella tularensis* Type A in brackish water. Arch. Microbiol. 193, 223-226. doi: 10.1007/s00203-010-0655-8

CDC, 2000. Biological and chemical terrorism: strategic plan for preparedness and response, recommendations of the CDC Strategic Planning Workgroup. MMWR Morb. Mortal. Wkly. Rep. 49 (RR-4), 1–14. <https://www.cdc.gov/mmwr/PDF/rr/rr4904.pdf>.

CDC and USDA, 2018. Guidance on the Inactivation or Removal of Select Agents and Toxins for Future Use. 7 CFR Part 331, 9 CFR Part 121.3, 42 CFR Part 73.3. September.

Dennis, D.T., Inglesby, T.V., Henderson, D.A., Barlett, J.G., Ascher, M.S., Eitzen, E., Fine, A.D., Friedlander, A.M., Hauer, J., Layton, M., Lillibridge, S.R., McDade, J.E., Osterholm, M.T., O’Toole, T., Parker, G., Perl, T.M. Russell, P.K., Tonat, K., 2001. Tularemia as a biological weapon: medical and public health management. JAMA 285, 2763–2773. DOI: [10.1001/jama.285.21.2763](https://doi.org/10.1001/jama.285.21.2763)

Farrah, S.R., Goyal, S.M., Gerba, C.P., Mahajan, V.K.,Wallis, C., Melnick, J.L., 1978. Concentration of humic acid from tapwater. Water Research 12, 303-306. <https://doi.org/10.1016/0043-1354(78)90117-3>.

Gilbert, S.E., Rose, L.J., 2012. Survival and persistence of nonspore-forming biothreat agents in water. Lett. Appl. Microbiol. 55, 189–194. <https://doi.org/10.1111/j.1472-765X.2012.03277.x>

Jones, R.M., Nicas, M., Hubbard, A., Sylvester, M.D., Reingold, A., 2005. The infectious dose of *Francisella tularensis* (Tularemia). Appl. Biosafety 10, 227-239. <https://doi.org/10.1177/153567600501000405>

Kane, S.R., Shah, S.R., Alfaro, T.M., 2019. Development of a Rapid Viability Polymerase Chain Reaction method for detection of *Yersinia pestis.* J. Microbiol. Meth. 162, 21-27. <https://doi.org/10.1016/j.mimet.2019.05.005>

Kugeler, K.J., Pappert, R., Zhou, Y., Petersen, J.M., 2006. Real-time PCR assays for *Francisella tularensis* Types A and B. Emerg. Infect. Dis. 12, 1799-1801. doi: [10.3201/eid1211.060629](https://dx.doi.org/10.3201%2Feid1211.060629)

Morris, B.J., Buse, H.Y., Adcock, N.J., Rice, E.W., 2017. A novel broth medium for enhanced growth of *Francisella tularensis*. Lett. Appl. Microbiol. 64, 394-400. doi: [10.1111/lam.12725](https://dx.doi.org/10.1111%2Flam.12725)

National Research Council (NRC). 1979. Iron. A Report of the Subcommittee on Iron, Committee on Medical and Biologic Effects of Environmental Pollutants, Division of Medical Sciences, Assembly of Life Sciences. University Park Press, Baltimore. 248 pp.

Oyston, P.C., Sjostedt, A., Titball, R.W., 2004. Tularaemia: bioterrorism defense renews interest in *Francisella tularensis*. Nat. Rev. Microbiol. 2, 967-978. doi: 10.1038/nrmicro1045

Rice, E.W., 2015. Occurrence and control of tularemia in drinking water. J. Am. Water Works Assoc. 107, E486-E496. <https://doi.org/10.5942/jawwa.2015.107.0130>

Rose, L.J., Hodges, L., O'Connell, H., Noble-Wang, J., 2011. National validation study of a cellulose sponge wipe-processing method for use after sampling *Bacillus anthracis* spores from surfaces. Appl. Environ. Microbiol. 77, 8355–8359. <https://doi.org/10.1128/AEM.05377-11>.

Saslaw, S., Eigelsbach, H.T., Prior, J.A., Wilson, H.E., Carhart S., 1961. Tularemia vaccine study, II: respiratory challenge. Arch. Intern. Med. 107, 134-146. doi:10.1001/archinte.1961.03620050068007

Stewart, S.J., 1995. Francisella. In: Murray, P.A,. Baron, E.J., Pfaller, M.A., Tenover, F.C., Yolken, R.H., eds. Manual of clinical microbiology. 6th ed. Washington, DC: American Society for Microbiology, 1995:545–8.

US EPA, 2005. Method 415.3 Determination of Total Organic Carbon and Specific UV Absorbance at 254 nm in Source Water. Revision 1.1. EPA/600/R-05/055.   
https://cfpub.epa.gov/si/si\_public\_file\_download.cfm?p\_download\_id=525079

US EPA, 2015. Literature Review on Processing and Analytical Methods for *Francisella tularensis* in Soil and Water. EPA 600/R-15/118 U.S. Environmental Protection Agency, Washington, DC.

US EPA, 2017. Sampling Guidance for Unknown Contaminants in Drinking Water. EPA-817-R-08-003. U.S. Environmental Protection Agency, Washington, DC. <https://www.epa.gov/waterlabnetwork/sampling-guidanceunknown-contaminants-drinking-water>.

World Health Organization (WHO). 1996. Guidelines for Drinking-Water Quality, 2nd ed. Vol. 2. Health Criteria and Other Supporting Information. Geneva: World Health Organization. WHO/SDE/WSH/03.04/08.