## Abbreviations
Abbreviations
ASR | Acoustic Startle Reflex
ANOVA | analysis of variance
AAALAC | Association for Acredidation of Laboratory Animal Care
BL | baseline
CXP | Collision cell exit potential
CE | collision energy
CS | conditioned stimulus
dB | decibel
dl | deciliter
DP | depolarizing potential
EP | Entrance Potential
fT4 | free T4
GD | Gestational Day
g or gm | gram
IACUC | Institutional Animal Care and Use Committee
ITI | intertrial interval
kg | kilogram
LC/MS/MS | liquid chromatography mass spectroscoscopy
LE | Long Evans
LLOQ | lower limit of quantification
MaxG | maximal grams of force=Startle amplitude
ME | matrix effects
MMI | methimazole
MDL | method detection limit
MQL | method quantification limit
µ | micro
mm | milimeter
mg | milligram
ml | milliliter
mM | millimolar
msec | millisecond
ng | nanogram
NeuN | neuronal nuclei
# | number
OECD | Organization for Economic Cooperation and Development
% | percent
ClO4 | perchlorate
PFHxS | perflourohexane sulfonate
PVH | periventricular heterotopia
pg | picogram
PN or PND | Postnatal Day
PP | prepulse
PPI | Prepulse Inhibition
PTU | propylthiouracil
qual | qualitative
quant | quantitative
qRT-PCR | quantitative real time-polymerase chain reaction
RSD | relatiave standard deviation
rt | retention time
rT3 | reverse T3
NIS | sodium-iodine symporter
SPE | solid phase extraction
std | standard
SEM or SE | standard error
SEM | standard error of mean
T4 | tetra-iodothyronine, thyroxine
TSH | thyroid stimulating hormone
TPO | thyroperoxidase
TFC | trace fear conditioning
T3 | tri-iodothyronine
UCS | unconditioned stimulus
v:v | volume:volume
ACN
qtrap
m/z

## Table 1-Optimized LC settings
Table 1. Optimized ion transitions, retention time (Rt) declustering potential (DP), entrance potential (EP), collision energy (CE), and collision cell exit potential (CXP) for AB Sciex 6500+ Qtrap.
Analyte | Q1 Mass | Q2 Mass | DP | EP | CE | CXP | Analyte Status
T3.1 | 651.8 | 605.8 | 120 | 10 | 30 | 38 | Quantitative
T3.2 | 651.8 | 478.8 | 120 | 10 | 50 | 30 | Qualitative
T3-13C6 | 657.8 | 611.7 | 95 | 10 | 33 | 38 | Surrogate
T3-13C12 | 663.8 | 617.8 | 100 | 10 | 32 | 40 | Internal Std
rT3 Quant | 651.8 | 605.8 | 120 | 10 | 30 | 35 | Quantitative
rT3 Qual | 651.8 | 507.8 | 120 | 10 | 30 | 35 | Qualitative
rT3-13C6 | 657.8 | 611.7 | 120 | 10 | 33 | 40 | Surrogate
T4 Quant | 777.6 | 731.6 | 120 | 10 | 31 | 50 | Quantitative
T4 Qual | 777.6 | 350.9 | 120 | 10 | 65 | 35 | Qualitative
T4-13C6 | 783.5 | 737.7 | 105 | 10 | 35 | 50 | Surrogate
T4-13C12 | 789.7 | 743.6 | 100 | 10 | 36 | 40 | Internal Std

## Table 2 SPE comparisons
Table 2. Comparison of solid phase extraction media for the concentration of thyroid hormones from brain tissue and the removal of matrix interference.
Step | Evolute Express CX | Evolute Express AX | Oasis PRiME HLB | Strata X Pro
Pre-Treatment | Dilute volume of 2% Formic Acid:ACN (60:40) v/v | Dilute with equal volume of 5% NH4OH:ACN (1:1) v/v | Dilute with equal volume of 2% Formic Acid | Dilute with equal volume of 2% Formic Acid
Conditioning | 3 mL MeOH | 3 mL MeOH | 3 mL MeOH | -
Equilibration | 3 mL 2% Formic Acid | 3 mL H2O | 3 mL H2O | -
Load | Pre-treaded Sample | Pre-treaded Sample | Pre-treaded Sample | Pre-treaded Sample
Wash 1 | 3 mL 50mM NH4OAC, pH 6 | 3 mL 50mM NH4OAC, pH 9 | 3 mL 30% MeOH | 3 mL 30% MeOH
Wash 2 | 3 mL 2% Formic Acid | 2 mL MeOH | -
Wash 3 | 3 mL MeOH | 3 mL 2% Formic Acid in DCM | -
Elution | 1 mL 2.5% NH4OH in MeOH | 1 mL MeOH | 1 mL MeOH | 1 mL ACN/MeOH (90:10)

## Table 4-optimizing SPE
A. Evaluation of Evolute Express AX 60 mg cartridges using the manufacturer’s recommended method. Recovery was poor, <50% for all 3 hormone analytes.
| T3-13C6 | rT3-13C6 | T4-13C6
Recovery % | 46 | 41 | 35
Matrix Effects % (ME) | 99 | 110 | 112
Process Efficiency (%) | 45 | 45 | 37
B. Recovery and process efficiency increases using Evolute Express AX 60 mg cartridges with an optimized wash step (2.0% formic acid in dichloromethane) and nitrogen evaporation in salinized glass tubes.
| T3-13C6 | rT3-13C6 | T4-13C6
Recovery % | 62 | 58 | 63
Matrix Effects % (ME) | 98 | 91 | 100
Process Efficiency (%) | 70 | 62 | 73
C. Further enhancement of recovery using Evolute Express AX 60 mg cartridges with an optimized wash step (2.0% formic acid in dichloromethane) and nitrogen evaporation in low retention polypropylene tubes versus salinized glass tubes used in B. Recovery improved relative to conditions tallied in B.
| T3-13C6 | rT3-13C6 | T4-13C6
Recovery % | 80 | 78 | 82
Matrix Effects % (ME) | 98 | 91 | 100
Process Efficiency (%) | 88 | 82 | 92

## Table 5-Precision Accuracy
Table 5. Precision and Accuracy Results using samples fortified at 5.00 ng/g of 13C6 labelled thyroid hormones. The samples were processed using the Evolute Express AX 60 mg cartridges with the additional wash step. High recovery with low intra- and inter-day variability was evident. RSD=Relative Standard Deviation of averaged concentrations of replicates within and across days.
| Intra-day (n=5) |  | Inter-day (n=15)
Analyte | Measured (ng/g) | %RSD | Measured (ng/g) | %RSD | % Recovery
T3-13C6 | 4.75 ± 0.088 | 1.85 | 4.74 ± 0.220 | 4.65 | 94.9 ± 4.4
rT3-13C6 | 5.02 ± 0.589 | 11.7 | 5.14 ± 0.744 | 14.5 | 102 ± 4.9
T4-13C6 | 5.09 ± 0.146 | 2.87 | 5.00 ± 0.293 | 5.85 | 100 ± 5.9

## Figure 3A-3C Calibration curves
Calibration Curves: Solvent vs Matric-Matched Samples
Figure 3A | Conc. (ng/g) | T3 Solvent | 13C6-T3 Solvent | 13C6-T3 Matrix Matched
| 0.05 | 0.007843 | 0.006885 | 0.010012
| 0.1 | 0.012386 | 0.011128 | 0.014658
| 0.25 | 0.031859 | 0.029018 | 0.038212
| 0.5 | 0.074207 | 0.068688 | 0.056717
| 1 | 0.12915 | 0.12566 | 0.11405
| 2.5 | 0.34004 | 0.31668 | 0.30948
| 5 | 0.60704 | 0.55275 | 0.59919
| 10 | 1.4676 | 1.31 | 1.2991
| 25 | 3.317 | 3.1 | 2.9701
Figure 3b | Conc. (ng/g) | rT3 Solvent | 13C6-rT3 Solvent | 13C6-rT3 Matrix Matched
0.05
0.1
| 0.25 |  |  |  | `
0.5
1
2.5
5
10
25
Figure 3C | Conc. (ng/g) | T4 Solvent | 13C6-T4 Solvent | 13C6-T4 Matrix-Matched
| 0.05 | 0.010323 | 0.013028 | 0.030153
| 0.1 | 0.023043 | 0.02117 | 0.051835
| 0.25 | 0.053966 | 0.050557 | 0.067167
| 0.5 | 0.12478 | 0.12347 | 0.13951
| 1 | 0.23361 | 0.21893 | 0.27003
| 2.5 | 0.58609 | 0.58911 | 0.6158
| 5 | 1.0545 | 1.0461 | 1.2424
| 10 | 2.5445 | 2.3285 | 2.5354
| 25 | 5.4873 | 5.2872 | 6.0272

## Figure 4A PN2 forebrain TH
Figure 4A PN2 Forebrain Thyroid Hormones
Mean +/- SEM n=6
| T3 (ng/g) | rT3 (ng/g) | T4 (ng/g)
| 1.516 | 0.147 | 1.021
| 1.267 | 0.192 | 1.068
| 0.971 | 0.124 | 0.881
| 1.36 | 0.105 | 0.807
| 1.644 | 0.156 | 1.509
| 1.524 | 0.084 | 0.779

## Figure 4B 4C reproducibility
Figure 4 B and C - % Recovery over Time
Recovrey Spiked at 5 ng/mL
Run | 13C12-T3 (% Rec) | `13C12-T4 (%Rec)
1 | 98.5868 | 98.5868
2 | 105.781 | 105.781
3 | 98.7832 | 98.7832
4 | 106.2436 | 106.2436
5 | 105.8218 | 105.8218
6 | 104.3196 | 104.3196
7 | 107.9942 | 107.9942
8 | 95.0662 | 95.0662
9 | 101.0076 | 101.0076
10 | 104.3302 | 104.3302
11 | 71.7288 | 71.7288
12 | 98.7824 | 98.7824
13 | 98.0866 | 98.0866
14 | 101.0626 | 101.0626
15 | 72.581 | 72.581
16 | 97.5254 | 97.5254
17 | 104.7754 | 104.7754
18 | 104.1144 | 104.1144
19 | 104.1134 | 104.1134
20 | 104.5746 | 104.5746
21 | 106.566 | 106.566
22 | 103.8532 | 103.8532
23 | 97.9862 | 97.9862
24 | 106.9632 | 106.9632
25 | 100.4336 | 100.4336
26 | 101.529 | 101.529
27 | 95.757 | 95.757
28 | 100.2512 | 100.2512
29 | 106.1736 | 106.1736
30 | 101.1332 | 101.1332
31 | 104.8642 | 104.8642
32 | 105.3022 | 105.3022
33 | 100.0912 | 100.0912
34 | 100.199 | 100.199
35 | 102.832 | 102.832
36 | 104.9646 | 104.9646
37 | 92.7248 | 92.7248
38 | 104.1462 | 104.1462
39 | 101.3396 | 101.3396
40 | 102.013 | 102.013
41 | 101.1232 | 101.1232
42 | 102.0608 | 102.0608
43 | 98.1138 | 98.1138
44 | 99.8862 | 99.8862

## Figure 5A serum TH
M0918 Brain Hormone Methods - Serum Hormone Ontogeny
Mean
T3 ng/ml | 23 | 26 | 35 | 66 | (n=5/age)
se | 3 | 3.9 | 1.3 | 11.4
Mean
T4 ng/ml | 4.484 | 7.626 | 17.7056 | 44.3588 | (n=5/age)
se | 0.39318 | 0.87795 | 1.674 | 7.2148

## Figure 5B Brain Perfuse
M0918 Brain Hormone Methods Perfused vs NonPerfused Brains June 3 2020
MEANS | Age-PND | PERFUSE? | n | wt in mg | T3 ng/gm | T4 ng/gm
| 0 | NP | 5 | 214.2 | 0.89 | 0.5864
| 0 | P | 5 | 211.2 | 0.9086 | 0.5546
| 2 | NP | 7 | 300.714 | 0.933 | 0.64871
| 2 | P | 7 | 326 | 0.88671 | 0.59957
| 6 | NP | 6 | 287.833 | 1.7827 | 1.3917
| 6 | P | 6 | 296.5 | 1.8213 | 1.5082
| 14 | NP | 7 | 607 | 4.075 | 3.0676
| 14 | P | 7 | 589.857 | 4.1883 | 3.0317
SEMS | Age -PND | PERFUSE? | n | wt in mg | T3 ng/gm | T4 ng/gm
| 0 | NP | 5 | 4.6195 | 0.08856 | 0.08281
| 0 | P | 5 | 4.3977 | 0.08195 | 0.0959
| 2 | NP | 7 | 7.0228 | 0.036 | 0.04349
| 2 | P | 7 | 7.5246 | 0.05952 | 0.05559
| 6 | NP | 6 | 12.051 | 0.04237 | 0.08845
| 6 | P | 6 | 11.8793 | 0.03761 | 0.16216
| 14 | NP | 7 | 32.4661 | 0.25321 | 0.3721
| 14 | P | 7 | 39.7866 | 0.18046 | 0.41591
P=Perfused
NP=Non-perfused

## Figure 5C Brain TH by Sex
M1019 Brain Hormone August 19 2022 Means and SEMs
MEANS
AGE | SEX | n | Wt in mg | T3 ng/gm | T4 ng/gm
0 | Female | 6 | 207.333 | 0.97783 | 0.64417
0 | Male | 4 | 220.75 | 0.7815 | 0.46
2 | Female | 7 | 316.857 | 0.83286 | 0.59543
2 | Male | 7 | 309.857 | 0.98686 | 0.65286
6 | Female | 6 | 291.667 | 1.8072 | 1.4292
6 | Male | 6 | 292.667 | 1.7968 | 1.4707
14 | Female | 8 | 627.125 | 4.1891 | 3.2576
14 | Male | 6 | 560.167 | 4.055 | 2.7723
SEMS
AGE | SEX | n | Wt in mg | T3 ng/gm | T4 ng/gm
0 | Female | 6 | 2.2311 | 0.08063 | 0.0849
0 | Male | 4 | 4.6075 | 0.01935 | 0.04655
2 | Female | 7 | 6.731 | 0.04734 | 0.04915
2 | Male | 7 | 10.48 | 0.02828 | 0.04993
6 | Female | 6 | 14.3798 | 0.0414 | 0.10861
6 | Male | 6 | 9.3262 | 0.04043 | 0.15332
14 | Female | 8 | 18.6848 | 0.14592 | 0.38181
14 | Male | 6 | 50.6165 | 0.30831 | 0.37031