## Sheet1
Sample Name | Solvent | Sample Volume | Preparation Date | Measurement Date and Results + Qualifier Flags
|  |  |  | 4/14/2023 (4-day after) | 4/17/2023 (7-day after) | 4/24/2023 (14-day after) | Also done on 4/24/2023 | Method
tap water | tap water | 500 mL | 2023-04-10 | 2.36 ppt B | - | - | - | 3887
100 ppq PFOS |  | 1 L | 2023-04-10 | - | - | 25.3 ppt | - | 3887
1 ppt PFOS |  | 500 mL | 2023-04-10 | 5.06 ppt B | 6.34 ppt B | 48.9 ppt | 25.1 ppt | 3887
10 ppt PFOS |  | 125 mL | 2023-04-10 | - | - | 22.7 ppt | - | 3887
100 ppt PFOS |  | 125 mL | 2023-04-10 | 108 ppt | 116 ppt | 101 ppt | 116 ppt | 3887
1 ppm PFOS |  | 5 mL | 2023-04-10 | - | - | 1.65 ppm | - | 4424
DI water | DI water | 5 mL | 2023-04-10 | - | - | < 35 ppt U | - | 4424
100 ppm PFOS |  | 5 mL | 2023-04-10 | 119 ppm | 116 ppm | 123 ppm | 159 ppm | 4424
500 ppm PFOS |  | 5 mL | 2023-04-10 | 762 ppm | 793 ppm | 986 ppm | 820 ppm | 4424
1000 ppm PFOS |  | 5 mL | 2023-04-10 | 957 ppm | 981 ppm | 1096 ppm | 1161 ppm | 4424
1 ppt PFOS in TW 125 mL bottle (extra) | tap water | 125 mL | 2023-04-10 | - | - | 5.53 ppt | - | 3887
Flag | Description
U | The analyte was analyzed for, but not detected above the reported method detection limit (MDL).
B | Laboratory Reagent Blank (LRB) result is >50% of method reporting limit (MRL).
Sample Preparation Overview
For solid phase extraction (SPE), the glass manifold, WAX cartridges, and sample reservoirs were prepared according to K-LRTD-SOP-3887-1. Before SPE, the internal standard, MPFAC-24ES, was aliquoted to the sample containers as received to obtain an expected concentration of 40 ng L-1. The volume of MPFAC-24ES added to each sample was determined by the sample volume as received, which varied by sample. SPE was performed on the experimental samples corresponding QC samples (Laboratory Blank, Laboratory Fortified Blank, and Laboratory Fortified Sample Matrix). 4 mL of 1% methanolic ammonium was passed through the WAX cartridges for final extraction and captured into respective 15 mL falcon tubes. Each sample was diluted to 4 mL final volumes using 0.3% methanolic ammonium hydroxide. 10 µL of acetic acid was aliquoted to each sample for pH adjustment. It was assumed that all PFOS and internal standard reached the 4 mL final volume extraction.
Sample LC/MS/MS Analysis Overview
Calibration standards and QC samples for LC/MS/MS analysis were prepared according to K-LRTD-SOP-3887-1 with the native standard solutions (PFAC30-PAR) and mass-labeled isotopic standard solutions (internal standards; MPFAC-24ES, M3HFPO-DA) from Wellington Laboratories, Guelph, Ontario, Canada. PFAC30-PAR contains 36 PFAS compounds at 1.0 µg mL-1 each. MPFAC-24ES contains 19 mass-labeled PFAS compounds at 1.0 µg mL-1 each, and M3HFPO-DA contains one mass-labeled PFAS compound at 50 µg mL-1. Calibration concentrations ranged from 5 to 10,000 ng L-1 (ppt) for each analyte present in the native standard solutions.
Aliquots (200 uL) from the calibration standards, QC samples (calibration blanks and calibration controls), and all experimental samples (4 mL extracts) were transferred to LC vials directly prior to the injection sequence. All samples were vortexed before aliquoted to vials. Samples were analyzed on a Waters LC/MS Micro with a gradient of 9.5-minute run time with a starting and ending eluent composition of 98% Eluent B (2 mM ammonium acetate in 95:5 water: acetonitrile) and 2% Eluent A (HPLC grade acetonitrile).
Calibration was calculated for all analytes with compound analysis being accepted for calibration curves of R2 (>0.99) for each analyte. The Method Detection Limit (MDL) – as outlined in CESER SOP K-IO-SOP-3706 – for the analyte with these instruments has not yet been determined. Therefore, the Minimum Reporting Limit (MRL) for the analyte was established by the lowest concentration calibration standard that agrees within 50% of the true concentration. All collected data resulted in concentrations above the MRL, satisfying the QC requirement.
Final Concentrations of QC samples and experimental samples were calculated following the method in the SOP K-LRTD-SOP-3887-1 and passed QC requirements. Results are reported as Total PFOS (TPFOS).