## Read Me
Figure 4: Focusing Trap Extended Purge Study
Standards with and without 4% CO2 added were prepared at 0.5 pbbv (5 ppbv CF4) and 20 ppbv (200 ppbv CF4) VFC concentrations.
The standards with and without CO2 added were analyzed on System 2 using focusing trap purge times of 1, 3, 6, and 9 min.
Recoveries were determined by comparison to the 1 minute purge calibration curve (200 ml injection) except for tetrafluoromethane.
Tetrafluoromethane comparative recovery data were generated using a 1 minute focusing trap purge tetrafluoromethane calibration curve prepared using 20 ml injections.
1,4 Difluorobenzene was used as the internal standard.
Abbreviations
cal | calibration
CAS | chemical abstract services
CO2 | carbon dioxide
Conc | concentration
d | detected
ID | identification
min | minute
mL | milliliter
N.D. | not detected
ppb | part-per-billion
ppbv | parts-per billion by volume
PQL | practical quantitation limit
Qion | qualifier Ion
R.T. | retention time
VFC | volatile fluorinated compound

## 4% CO2, 0.5 ppbv VFC Recoveries
Recovery of 0.5 ppbv VFC calibration standard with 4% CO2 injected at 200 mL (20 mL for tetrafluoromethane)
Internal Standard 20% low due to matrix
| Canister ID | 45688 | 45688 | 45688 | 45688
Description | CAS # | 1 minute trap purge | 3 minute trap purge | 6 minute trap purge | 9 minute trap purge
Tetrafluoromethane | 75-73-0 | No comparable level | Not Relevant | Not Relevant | Not Relevant
Hexafluoroethane | 76-16-4 | 0% | 24.32% | 91.89% | 108.1%
Chlorotrifluoromethane | 75-72-9 | 0% | 97.22% | 100% | 100%
Tetrafluoroethene | 116-14-3 | 2.381% | 90.48% | 95.24% | 95.24%
Trifluoromethane | 75-46-7 | 63.04% | 54.35% | 57.61% | 55.43%
Difluoromethane | 75-10-5 | 154% | 110% | 102% | 98%
Octafluoropropane | 76-19-7 | 159.5% | 97.62% | 97.62% | 95.24%
Fluoromethane | 593-53-3 | 163.8% | 74.14% | 55.17% | 43.1%
Pentafluoroethane | 354-33-6 | 123.9% | 95.65% | 93.48% | 93.48%
1,1,1-Trifluoroethane | 420-46-2 | 138.9% | 108.3% | 97.22% | 94.44%
Hexafluoropropene | 116-15-4 | 166.7% | 88.89% | 91.67% | 91.67%
Hexafluoropropene Oxide | 428-59-1 | 0% | 81.25% | 97.92% | 81.25%
Chlorodifluoromethane | 75-45-6 | 130.4% | 97.83% | 100% | 93.48%
Octafluorocyclobutane | 115-25-3 | 179.5% | 100% | 97.44% | 94.87%
1,1,1,2-Tetrafluoroethane | 811-97-2 | 135.3% | 98.04% | 100% | 96.08%
Decafluorobutane | 355-25-9 | 134.1% | 100% | 95.12% | 90.24%
1H-Heptafluoropropane | 2252-84-8 | 152.2% | 100% | 95.65% | 86.96%
Dodecafluoropentane | 678-26-2 | 86.96% | 106.5% | 100% | 97.83%
Trichlorofluoromethane | 75-69-4 | 91.89% | 102.7% | 100% | 100%
Octafluorocyclopentene | 559-40-0 | 87.18% | 107.7% | 100% | 97.44%
1H-Nonafluorobutane | 375-17-7 | 86.27% | 103.9% | 98.04% | 92.16%
Tetradecafluorohexane | 355-42-0 | 86.67% | 102.2% | 88.89% | 91.11%
1H-Perfluoropentane | 375-61-1 | 90.2% | 103.9% | 100% | 98.04%
E1 | 3330-15-2 | 92.86% | 107.1% | 104.8% | 102.4%
Hexadecafluoroheptane | 335-57-9 | 120.5% | 100% | 100% | 97.73%
1H-Perfluorohexane | 355-37-3 | 104.2% | 102.1% | 100% | 97.92%
Octadecafluorooctane | 307-34-6 | 186.5% | 108.1% | 102.7% | 102.7%
1H-Perfluoroheptane | 375-83-7 | 137.5% | 100% | 97.92% | 93.75%
E2 | 3330-14-1 | 202.9% | 111.4% | 105.7% | 105.7%
1H-Perfluorooctane | 335-65-9 | 215.9% | 111.4% | 102.3% | 100%

## 4% CO2, 20 ppbv VFC Recoveries
Recovery of 20 ppbv VFC calibration standard with 4% CO2 injected at 200 mL (20 mL for tetrafluoromethane)
Internal Standard 20% low due to matrix
| Canister ID | 42881 | 42881 | 42881 | 42881
Description | CAS # | 1 minute trap purge | 3 minute trap purge | 6 minute trap purge | 9 minute trap purge
Tetrafluoromethane | 75-73-0 | 81.48% | Not Relevant | Not Relevant | Not Relevant
Hexafluoroethane | 76-16-4 | 1.444% | 28.04% | 86.04% | 97.47%
Chlorotrifluoromethane | 75-72-9 | 38.81% | 94.53% | 96.9% | 97.26%
Tetrafluoroethene | 116-14-3 | 15.29% | 92.38% | 96.34% | 95.21%
Trifluoromethane | 75-46-7 | 130.1% | 104.5% | 97.42% | 96.42%
Difluoromethane | 75-10-5 | 144.2% | 105.7% | 97.65% | 95.16%
Octafluoropropane | 76-19-7 | 169.4% | 97.5% | 96.05% | 93.7%
Fluoromethane | 593-53-3 | 130.4% | 69.69% | 51.27% | 42.53%
Pentafluoroethane | 354-33-6 | 124% | 95.57% | 96.16% | 92.92%
1,1,1-Trifluoroethane | 420-46-2 | 153.6% | 99.52% | 100.7% | 94.61%
Hexafluoropropene | 116-15-4 | 165.5% | 98.23% | 97.31% | 97.25%
Hexafluoropropene Oxide | 428-59-1 | 0% | 83.33% | 72.18% | 87.7%
Chlorodifluoromethane | 75-45-6 | 0% | 97.67% | 96.01% | 92.52%
Octafluorocyclobutane | 115-25-3 | 189.2% | 99.22% | 94.85% | 94.08%
1,1,1,2-Tetrafluoroethane | 811-97-2 | 140.1% | 96% | 94.35% | 89.9%
Decafluorobutane | 355-25-9 | 138% | 99.19% | 90.48% | 88.17%
1H-Heptafluoropropane | 2252-84-8 | 192.8% | 99.08% | 90.65% | 85.87%
Dodecafluoropentane | 678-26-2 | 79.87% | 99.57% | 92.29% | 90.51%
Trichlorofluoromethane | 75-69-4 | 83.69% | 97.76% | 93.23% | 92.34%
Octafluorocyclopentene | 559-40-0 | 78.43% | 99.22% | 91.44% | 90.11%
1H-Nonafluorobutane | 375-17-7 | 77.18% | 99.23% | 89.8% | 86.53%
Tetradecafluorohexane | 355-42-0 | 88.92% | 98.8% | 94.94% | 91.98%
1H-Perfluoropentane | 375-61-1 | 83.45% | 91.6% | 85.38% | 88.43%
E1 | 3330-15-2 | 81.22% | 93.84% | 91.7% | 84.53%
Hexadecafluoroheptane | 335-57-9 | 117.3% | 99.26% | 97.49% | 95.08%
1H-Perfluorohexane | 355-37-3 | 96.62% | 99.57% | 95.15% | 92.2%
Octadecafluorooctane | 307-34-6 | 196.5% | 102.8% | 98.03% | 96.69%
1H-Perfluoroheptane | 375-83-7 | 133.2% | 99.04% | 95.91% | 92.29%
E2 | 3330-14-1 | 198.7% | 100.8% | 96.32% | 96.14%
1H-Perfluorooctane | 335-65-9 | NA | 101.4% | 92.49% | 88.56%

## 4% CO2, 0.5 ppb VFC, 200mL Data
092623/092623_23.D/epatemp.txt (QT Reviewed): Riemer cal 0.5 ppbv (200ml), No CO2 | [5 empty] | 092123/092123_03.D/epatemp.txt (QT Reviewed): 4%CO2_0.5ppbv_1 mintrap | [6 empty] | 092123/092123_05.D/epatemp.txt (QT Reviewed): 4%CO2_0.5ppbv_3 mintrap | [6 empty] | 092123/092123_13.D/epatemp.txt (QT Reviewed): 4%CO2_0.5ppbv_6 mintrap | [6 empty] | 092523/092523_04.D/epatemp.txt (QT Reviewed): 4%CO2_0.5ppbv_9 mintrap
Compound | R.T. | QIon | Response | Conc | Units | Compound | R.T. | QIon | Response | Conc | Units | % Recovery | Compound | R.T. | QIon | Response | Conc | Units | % Recovery | Compound | R.T. | QIon | Response | Conc | Units | % Recovery | Compound | R.T. | QIon | Response | Conc | Units | % Recovery
d5-Chlorobenzene | 33.899 | 117 | 724041 | 5 | ppb | d5-Chlorobenzene | 33.909 | 117 | 811244 | 5 | ppb |  | d5-Chlorobenzene | 33.899 | 117 | 873676 | 5 | ppb |  | d5-Chlorobenzene | 33.899 | 117 | 836380 | 5 | ppb |  | d5-Chlorobenzene | 33.909 | 117 | 834670 | 5 | ppb
1,4-Difluorobenzene | 30 | 114 | 854245 | 5 | ppb | 1,4-Difluorobenzene | 29.987 | 114 | 780849 | 5 | ppb |  | 1,4-Difluorobenzene | 30 | 114 | 1033133 | 5 | ppb |  | 1,4-Difluorobenzene | 29.989 | 114 | 992726 | 5 | ppb |  | 1,4-Difluorobenzene | 29.987 | 114 | 995510 | 5 | ppb
Tetrafluoromethane | 7.273 | 69 | 6593 | 42.62 | ppb | Tetrafluoromethane | 7.147 | 69 | 1528m | 11.85 | ppb | NA | Tetrafluoromethane | 0 |  | 0 | N.D. |  | NA | Tetrafluoromethane | 0 | 0 | N.D. |  |  | NA | Tetrafluoromethane | 0 |  | 0 | N.D. |  | NA
Hexafluoroethane | 8.261 | 119 | 8260 | 0.37 | ppb | Hexafluoroethane | 0 |  | 0 | N.D. |  | 0 | Hexafluoroethane | 8.009 | 119 | 2396m | 0.09 | ppb | 24.3243 | Hexafluoroethane | 8.26 | 119 | 8851m | 0.34 | ppb | 91.8919 | Hexafluoroethane | 8.28 | 119 | 10468 | 0.4 | ppb | 108.1081
Chlorotrifluoromethane | 8.667 | 85 | 2950 | 0.36 | ppb | Chlorotrifluoromethane | 0 |  | 0 | N.D. |  | 0 | Chlorotrifluoromethane | 8.58 | 85 | 3447 | 0.35 | ppb | 97.2222 | Chlorotrifluoromethane | 8.667 | 85 | 3361 | 0.36 | ppb | 100 | Chlorotrifluoromethane | 8.686 | 85 | 3401 | 0.36 | ppb | 100
Tetrafluoroethene | 8.658 | 81 | 7711 | 0.42 | ppb | Tetrafluoroethene | 8.512 | 81 | 247m | 0.01 | ppb | 2.381 | Tetrafluoroethene | 8.561 | 81 | 8622 | 0.38 | ppb | 90.4762 | Tetrafluoroethene | 8.648 | 81 | 8669 | 0.4 | ppb | 95.2381 | Tetrafluoroethene | 8.667 | 81 | 8653 | 0.4 | ppb | 95.2381
Trifluoromethane | 8.87 | 69 | 9654m | 0.92 | ppb | Trifluoromethane | 8.551 | 69 | 5546m | 0.58 | ppb | 63.0435 | Trifluoromethane | 8.803 | 69 | 6372 | 0.5 | ppb | 54.3478 | Trifluoromethane | 8.861 | 69 | 6478 | 0.53 | ppb | 57.6087 | Trifluoromethane | 8.87 | 69 | 6263 | 0.51 | ppb | 55.4348
Difluoromethane | 10.932 | 51 | 3900 | 0.5 | ppb | Difluoromethane | 10.788 | 51 | 5477 | 0.77 | ppb | 154 | Difluoromethane | 10.922 | 51 | 5107 | 0.55 | ppb | 110 | Difluoromethane | 10.932 | 51 | 4590 | 0.51 | ppb | 102 | Difluoromethane | 10.942 | 51 | 4426 | 0.49 | ppb | 98
Octafluoropropane | 11.056 | 69 | 38832 | 0.42 | ppb | Octafluoropropane | 10.901 | 69 | 56751 | 0.67 | ppb | 159.5238 | Octafluoropropane | 11.035 | 69 | 46101 | 0.41 | ppb | 97.619 | Octafluoropropane | 11.045 | 69 | 44415 | 0.41 | ppb | 97.619 | Octafluoropropane | 11.055 | 69 | 43544 | 0.4 | ppb | 95.2381
Fluoromethane | 11.354 | 33 | 1841 | 0.58 | ppb | Fluoromethane | 11.292 | 33 | 2755m | 0.95 | ppb | 163.7931 | Fluoromethane | 11.364 | 33 | 1650 | 0.43 | ppb | 74.1379 | Fluoromethane | 11.364 | 33 | 1190m | 0.32 | ppb | 55.1724 | Fluoromethane | 11.375 | 33 | 936 | 0.25 | ppb | 43.1034
Pentafluoroethane | 12.672 | 51 | 12592 | 0.46 | ppb | Pentafluoroethane | 12.579 | 51 | 14076 | 0.57 | ppb | 123.913 | Pentafluoroethane | 12.672 | 51 | 14605 | 0.44 | ppb | 95.6522 | Pentafluoroethane | 12.672 | 51 | 13751 | 0.43 | ppb | 93.4783 | Pentafluoroethane | 12.682 | 51 | 13708 | 0.43 | ppb | 93.4783
1,1,1-Trifluoroethane | 13.034 | 64 | 789 | 0.36 | ppb | 1,1,1-Trifluoroethane | 12.946 | 64 | 1004m | 0.5 | ppb | 138.8889 | 1,1,1-Trifluoroethane | 13.033 | 64 | 1034 | 0.39 | ppb | 108.3333 | 1,1,1-Trifluoroethane | 13.033 | 64 | 880 | 0.35 | ppb | 97.2222 | 1,1,1-Trifluoroethane | 13.044 | 64 | 874 | 0.34 | ppb | 94.4444
Hexafluoropropene | 13.034 | 131 | 11861 | 0.36 | ppb | Hexafluoropropene | 12.936 | 131 | 18261 | 0.6 | ppb | 166.6667 | Hexafluoropropene | 13.023 | 131 | 12911 | 0.32 | ppb | 88.8889 | Hexafluoropropene | 13.023 | 131 | 12677 | 0.33 | ppb | 91.6667 | Hexafluoropropene | 13.033 | 131 | 12881 | 0.33 | ppb | 91.6667
Hexafluoropropene_oxide | 13.121 | 69 | 3473m | 0.48 | ppb | Hexafluoropropene_oxide | 0 |  | 0 | N.D. | d | 0 | Hexafluoropropene_oxide | 13.12 | 69 | 3373m | 0.39 | ppb | 81.25 | Hexafluoropropene_oxide | 13.12 | 69 | 3901m | 0.47 | ppb | 97.9167 | Hexafluoropropene_oxide | 13.121 | 69 | 3273m | 0.39 | ppb | 81.25
Chlorodifluoromethane | 13.567 | 51 | 12857 | 0.46 | ppb | Chlorodifluoromethane | 13.491 | 51 | 15218 | 0.6 | ppb | 130.4348 | Chlorodifluoromethane | 13.567 | 51 | 15132 | 0.45 | ppb | 97.8261 | Chlorodifluoromethane | 13.567 | 51 | 14715 | 0.46 | ppb | 100 | Chlorodifluoromethane | 13.578 | 51 | 13921 | 0.43 | ppb | 93.4783
Octafluorocyclobutane | 15.483 | 100 | 32124 | 0.39 | ppb | Octafluorocyclobutane | 15.456 | 100 | 51798 | 0.7 | ppb | 179.4872 | Octafluorocyclobutane | 15.483 | 100 | 38359 | 0.39 | ppb | 100 | Octafluorocyclobutane | 15.483 | 100 | 35990 | 0.38 | ppb | 97.4359 | Octafluorocyclobutane | 15.483 | 100 | 35403 | 0.37 | ppb | 94.8718
1,1,1,2-Tetrafluoroethane | 15.61 | 33 | 4690 | 0.51 | ppb | 1,1,1,2-Tetrafluoroethane | 15.574 | 33 | 5799 | 0.69 | ppb | 135.2941 | 1,1,1,2-Tetrafluoroethane | 15.601 | 33 | 5544 | 0.5 | ppb | 98.0392 | 1,1,1,2-Tetrafluoroethane | 15.601 | 33 | 5498 | 0.51 | ppb | 100 | 1,1,1,2-Tetrafluoroethane | 15.61 | 33 | 5231 | 0.49 | ppb | 96.0784
Decafluorobutane | 15.701 | 119 | 16184 | 0.41 | ppb | Decafluorobutane | 15.665 | 119 | 19952 | 0.55 | ppb | 134.1463 | Decafluorobutane | 15.701 | 119 | 19744 | 0.41 | ppb | 100 | Decafluorobutane | 15.701 | 119 | 18082 | 0.39 | ppb | 95.122 | Decafluorobutane | 15.71 | 119 | 17200 | 0.37 | ppb | 90.2439
1H-Heptafluoropropane | 17.673 | 51 | 18797 | 0.46 | ppb | 1H-Heptafluoropropane | 17.66 | 51 | 26095 | 0.7 | ppb | 152.1739 | 1H-Heptafluoropropane | 17.68 | 51 | 22750 | 0.46 | ppb | 100 | 1H-Heptafluoropropane | 17.672 | 51 | 20628 | 0.44 | ppb | 95.6522 | 1H-Heptafluoropropane | 17.682 | 51 | 19192 | 0.4 | ppb | 86.9565
Dodecafluoropentane | 20.478 | 69 | 70122 | 0.46 | ppb | Dodecafluoropentane | 20.468 | 69 | 56351 | 0.4 | ppb | 86.9565 | Dodecafluoropentane | 20.477 | 69 | 89528 | 0.49 | ppb | 106.5217 | Dodecafluoropentane | 20.477 | 69 | 81014 | 0.46 | ppb | 100 | Dodecafluoropentane | 20.479 | 69 | 79387 | 0.45 | ppb | 97.8261
Trichlorofluoromethane | 20.88 | 101 | 21052 | 0.37 | ppb | Trichlorofluoromethane | 20.859 | 101 | 17595 | 0.34 | ppb | 91.8919 | Trichlorofluoromethane | 20.88 | 101 | 26331 | 0.38 | ppb | 102.7027 | Trichlorofluoromethane | 20.868 | 101 | 24469 | 0.37 | ppb | 100 | Trichlorofluoromethane | 20.881 | 101 | 24245 | 0.37 | ppb | 100
Octafluorocyclopentene | 21.443 | 93 | 20553 | 0.39 | ppb | Octafluorocyclopentene | 21.433 | 93 | 16463 | 0.34 | ppb | 87.1795 | Octafluorocyclopentene | 21.443 | 93 | 26681 | 0.42 | ppb | 107.6923 | Octafluorocyclopentene | 21.443 | 93 | 23940 | 0.39 | ppb | 100 | Octafluorocyclopentene | 21.445 | 93 | 23637 | 0.38 | ppb | 97.4359
1H-Nonafluorobutane | 22.326 | 51 | 34429 | 0.51 | ppb | 1H-Nonafluorobutane | 22.324 | 51 | 26893 | 0.44 | ppb | 86.2745 | 1H-Nonafluorobutane | 22.334 | 51 | 43345 | 0.53 | ppb | 103.9216 | 1H-Nonafluorobutane | 22.334 | 51 | 38703 | 0.5 | ppb | 98.0392 | 1H-Nonafluorobutane | 22.332 | 51 | 36829 | 0.47 | ppb | 92.1569
Tetradecafluorohexane | 24.049 | 69 | 65068 | 0.45 | ppb | Tetradecafluorohexane | 23.992 | 69 | 50853m | 0.39 | ppb | 86.6667 | Tetradecafluorohexane | 24.041 | 69 | 78951 | 0.46 | ppb | 102.2222 | Tetradecafluorohexane | 24.033 | 69 | 66850m | 0.4 | ppb | 88.8889 | Tetradecafluorohexane | 24.016 | 69 | 68703m | 0.41 | ppb | 91.1111
1H-Perfluoropentane | 25.683 | 51 | 29592m | 0.51 | ppb | 1H-Perfluoropentane | 25.639 | 51 | 24755m | 0.46 | ppb | 90.1961 | 1H-Perfluoropentane | 25.683 | 51 | 37626m | 0.53 | ppb | 103.9216 | 1H-Perfluoropentane | 25.674 | 51 | 34589m | 0.51 | ppb | 100 | 1H-Perfluoropentane | 25.656 | 51 | 34009m | 0.5 | ppb | 98.0392
E1 | 25.751 | 101 | 48586m | 0.42 | ppb | E1 | 25.715 | 101 | 41512m | 0.39 | ppb | 92.8571 | E1 | 25.751 | 101 | 63524m | 0.45 | ppb | 107.1429 | E1 | 25.742 | 101 | 58823m | 0.44 | ppb | 104.7619 | E1 | 25.724 | 101 | 58176m | 0.43 | ppb | 102.381
Hexadecafluoroheptane | 26.873 | 69 | 83724 | 0.44 | ppb | Hexadecafluoroheptane | 26.837 | 69 | 93694 | 0.53 | ppb | 120.4545 | Hexadecafluoroheptane | 26.873 | 69 | 102429 | 0.44 | ppb | 100 | Hexadecafluoroheptane | 26.864 | 69 | 97672 | 0.44 | ppb | 100 | Hexadecafluoroheptane | 26.854 | 69 | 96885 | 0.43 | ppb | 97.7273
1H-Perfluorohexane | 28.333 | 51 | 46599 | 0.48 | ppb | 1H-Perfluorohexane | 28.319 | 51 | 44653 | 0.5 | ppb | 104.1667 | 1H-Perfluorohexane | 28.333 | 51 | 57475 | 0.49 | ppb | 102.0833 | 1H-Perfluorohexane | 28.333 | 51 | 54338 | 0.48 | ppb | 100 | 1H-Perfluorohexane | 28.319 | 51 | 52775 | 0.47 | ppb | 97.9167
Octadecafluorooctane | 29.126 | 119 | 39055 | 0.37 | ppb | Octadecafluorooctane | 29.113 | 119 | 66069 | 0.69 | ppb | 186.4865 | Octadecafluorooctane | 29.126 | 119 | 50506 | 0.4 | ppb | 108.1081 | Octadecafluorooctane | 29.115 | 119 | 46378 | 0.38 | ppb | 102.7027 | Octadecafluorooctane | 29.113 | 119 | 45844 | 0.38 | ppb | 102.7027
1H-Perfluoroheptane | 30.564 | 51 | 62918 | 0.48 | ppb | 1H-Perfluoroheptane | 30.562 | 51 | 79714 | 0.66 | ppb | 137.5 | 1H-Perfluoroheptane | 30.564 | 51 | 76439 | 0.48 | ppb | 100 | 1H-Perfluoroheptane | 30.564 | 51 | 72098 | 0.47 | ppb | 97.9167 | 1H-Perfluoroheptane | 30.562 | 51 | 69635 | 0.45 | ppb | 93.75
E2 | 32.404 | 169 | 72878 | 0.35 | ppb | E2 | 32.402 | 169 | 132734 | 0.71 | ppb | 202.8571 | E2 | 32.404 | 169 | 96457 | 0.39 | ppb | 111.4286 | E2 | 32.404 | 169 | 87156 | 0.37 | ppb | 105.7143 | E2 | 32.402 | 169 | 88259 | 0.37 | ppb | 105.7143
1H-Perfluorooctane | 32.761 | 51 | 47194 | 0.44 | ppb | 1H-Perfluorooctane | 32.759 | 51 | 92021 | 0.95 | ppb | 215.9091 | 1H-Perfluorooctane | 32.761 | 51 | 63425 | 0.49 | ppb | 111.3636 | 1H-Perfluorooctane | 32.761 | 51 | 55546 | 0.45 | ppb | 102.2727 | 1H-Perfluorooctane | 32.77 | 51 | 54149 | 0.44 | ppb | 100

## 4% CO2, 20 ppb VFC, 200mL Data
092623/092623_24.D/epatemp.txt (QT Reviewed): Riemer cal 20 ppb (200ml), No CO2 | [5 empty] | 092123/092123_06.D/epatemp.txt (QT Reviewed): 4%CO2_20ppbv_1 mintrap | [6 empty] | 092123/092123_08.D/epatemp.txt (QT Reviewed): 4%CO2_20ppbv_3 mintrap | [6 empty] | 092123/092123_14.D/epatemp.txt (Not Reviewed): 4%CO2_20ppbv_6 mintrap | [6 empty] | 092523/092523_05.D/epatemp.txt (QT Reviewed): 4%CO2_20ppbv_9 mintrap
Compound | R.T. | QIon | Response | Conc | Units | Compound | R.T. | QIon | Response | Conc | Units | % Recovery | Compound | R.T. | QIon | Response | Conc | Units | % Recovery | Compound | R.T. | QIon | Response | Conc | Units | % Recovery | Compound | R.T. | QIon | Response | Conc | Units | % Recovery
d5-Chlorobenzene | 33.899 | 117 | 698325 | 5 | ppb | d5-Chlorobenzene | 33.899 | 117 | 776927 | 5 | ppb |  | d5-Chlorobenzene | 33.899 | 117 | 813867 | 5 | ppb |  | d5-Chlorobenzene | 33.899 | 117 | 823572 | 5 | ppb |  | d5-Chlorobenzene | 33.911 | 117 | 826507 | 5 | ppb
1,4-Difluorobenzene | 29.966 | 114 | 826811 | 5 | ppb | 1,4-Difluorobenzene | 29.977 | 114 | 744235 | 5 | ppb |  | 1,4-Difluorobenzene | 30 | 114 | 934790 | 5 | ppb |  | 1,4-Difluorobenzene | 29.989 | 114 | 963310 | 5 | ppb |  | 1,4-Difluorobenzene | 29.989 | 114 | 965798 | 5 | ppb
Tetrafluoromethane | 7.273 | 69 | 180293 | 512.65 | ppb | Tetrafluoromethane | 7.147 | 69 | 66893 | 290.18 | ppb |  | Tetrafluoromethane | 7.273 | 69 | 3752 | 23.45 | ppb |  | Tetrafluoromethane | 7.283 | 69 | 2074 | 12.99 | ppb |  | Tetrafluoromethane | 7.283 | 69 | 1597 | 10.07 | ppb
Hexafluoroethane | 8.27 | 119 | 386984 | 16.62 | ppb | Hexafluoroethane | 8.561 | 119 | 4635 | 0.24 | ppb | 1.444 | Hexafluoroethane | 8.019 | 119 | 117331m | 4.66 | ppb | 28.0385 | Hexafluoroethane | 8.27 | 119 | 384646 | 14.3 | ppb | 86.0409 | Hexafluoroethane | 8.28 | 119 | 439785 | 16.2 | ppb | 97.4729
Chlorotrifluoromethane | 8.667 | 85 | 119213 | 16.44 | ppb | Chlorotrifluoromethane | 8.502 | 85 | 43786m | 6.38 | ppb | 38.8078 | Chlorotrifluoromethane | 8.58 | 85 | 128018 | 15.54 | ppb | 94.5255 | Chlorotrifluoromethane | 8.677 | 85 | 134990 | 15.93 | ppb | 96.8978 | Chlorotrifluoromethane | 8.687 | 85 | 135765 | 15.99 | ppb | 97.2628
Tetrafluoroethene | 8.658 | 81 | 311143 | 19.42 | ppb | Tetrafluoroethene | 8.018 | 81 | 47364m | 2.97 | ppb | 15.2935 | Tetrafluoroethene | 8.561 | 81 | 328037 | 17.94 | ppb | 92.379 | Tetrafluoroethene | 8.667 | 81 | 350838 | 18.71 | ppb | 96.344 | Tetrafluoroethene | 8.667 | 81 | 348038 | 18.49 | ppb | 95.2111
Trifluoromethane | 8.861 | 69 | 216012 | 22.89 | ppb | Trifluoromethane | 8.512 | 69 | 246831m | 29.79 | ppb | 130.1442 | Trifluoromethane | 8.803 | 69 | 254420 | 23.93 | ppb | 104.5435 | Trifluoromethane | 8.87 | 69 | 245765 | 22.3 | ppb | 97.4225 | Trifluoromethane | 8.87 | 69 | 244015 | 22.07 | ppb | 96.4176
Difluoromethane | 10.932 | 51 | 167777 | 23.78 | ppb | Difluoromethane | 10.798 | 51 | 211787 | 34.3 | ppb | 144.2389 | Difluoromethane | 10.922 | 51 | 199737 | 25.13 | ppb | 105.677 | Difluoromethane | 10.942 | 51 | 191146 | 23.22 | ppb | 97.6451 | Difluoromethane | 10.942 | 51 | 187106 | 22.63 | ppb | 95.164
Octafluoropropane | 11.056 | 69 | 1558257 | 19.99 | ppb | Octafluoropropane | 10.911 | 69 | 2115044 | 33.87 | ppb | 169.4347 | Octafluoropropane | 11.035 | 69 | 1724827 | 19.49 | ppb | 97.4987 | Octafluoropropane | 11.055 | 69 | 1755278 | 19.2 | ppb | 96.048 | Octafluoropropane | 11.056 | 69 | 1722990 | 18.73 | ppb | 93.6968
Fluoromethane | 11.354 | 33 | 79382m | 27.25 | ppb | Fluoromethane | 11.323 | 33 | 91452 | 35.54 | ppb | 130.422 | Fluoromethane | 11.364 | 33 | 63703 | 18.99 | ppb | 69.6881 | Fluoromethane | 11.375 | 33 | 48811 | 13.97 | ppb | 51.2661 | Fluoromethane | 11.375 | 33 | 40806 | 11.59 | ppb | 42.5321
Pentafluoroethane | 12.672 | 51 | 500101 | 21.88 | ppb | Pentafluoroethane | 12.589 | 51 | 537181 | 27.14 | ppb | 124.0402 | Pentafluoroethane | 12.672 | 51 | 544113 | 20.91 | ppb | 95.5667 | Pentafluoroethane | 12.682 | 51 | 563719 | 21.04 | ppb | 96.1609 | Pentafluoroethane | 12.682 | 51 | 548798 | 20.33 | ppb | 92.9159
1,1,1-Trifluoroethane | 13.044 | 64 | 31689 | 16.52 | ppb | 1,1,1-Trifluoroethane | 12.946 | 64 | 41416 | 25.38 | ppb | 153.632 | 1,1,1-Trifluoroethane | 13.033 | 64 | 35668 | 16.44 | ppb | 99.5157 | 1,1,1-Trifluoroethane | 13.044 | 64 | 37139 | 16.63 | ppb | 100.6659 | 1,1,1-Trifluoroethane | 13.044 | 64 | 35205 | 15.63 | ppb | 94.6126
Hexafluoropropene | 13.034 | 131 | 437193 | 15.26 | ppb | Hexafluoropropene | 12.946 | 131 | 595114 | 25.26 | ppb | 165.5308 | Hexafluoropropene | 13.023 | 131 | 486747 | 14.99 | ppb | 98.2307 | Hexafluoropropene | 13.033 | 131 | 497435 | 14.85 | ppb | 97.3132 | Hexafluoropropene | 13.034 | 131 | 498392 | 14.84 | ppb | 97.2477
Hexafluoropropene_oxide | 13.164 | 69 | 149605m | 21.96 | ppb | Hexafluoropropene_oxide | 0 |  | 0 | N.D. | d | NA | Hexafluoropropene_oxide | 13.11 | 69 | 141448m | 18.3 | ppb | 83.3333 | Hexafluoropropene_oxide | 13.044 | 69 | 126506m | 197.09 | ppb | 897.4954 | Hexafluoropropene_oxide | 13.121 | 69 | 153638m | 19.26 | ppb | 87.7049
Chlorodifluoromethane | 13.567 | 51 | 516260 | 22.33 | ppb | Chlorodifluoromethane | 0 |  | 0 | N.D. | d | NA | Chlorodifluoromethane | 13.567 | 51 | 572284 | 21.81 | ppb | 97.6713 | Chlorodifluoromethane | 13.567 | 51 | 581154 | 21.44 | ppb | 96.0143 | Chlorodifluoromethane | 13.578 | 51 | 564752 | 20.66 | ppb | 92.5213
Octafluorocyclobutane | 15.474 | 100 | 1224468 | 18.06 | ppb | Octafluorocyclobutane | 15.446 | 100 | 1769923 | 34.17 | ppb | 189.2027 | Octafluorocyclobutane | 15.474 | 100 | 1375180 | 17.92 | ppb | 99.2248 | Octafluorocyclobutane | 15.474 | 100 | 1364775 | 17.13 | ppb | 94.8505 | Octafluorocyclobutane | 15.483 | 100 | 1358684 | 16.99 | ppb | 94.0753
1,1,1,2-Tetrafluoroethane | 15.61 | 33 | 186744 | 26.53 | ppb | 1,1,1,2-Tetrafluoroethane | 15.564 | 33 | 210083 | 37.18 | ppb | 140.1432 | 1,1,1,2-Tetrafluoroethane | 15.61 | 33 | 204903 | 25.47 | ppb | 96.0045 | 1,1,1,2-Tetrafluoroethane | 15.61 | 33 | 208411 | 25.03 | ppb | 94.346 | 1,1,1,2-Tetrafluoroethane | 15.61 | 33 | 201422 | 23.85 | ppb | 89.8982
Decafluorobutane | 15.701 | 119 | 615548 | 18.6 | ppb | Decafluorobutane | 15.674 | 119 | 719101 | 25.66 | ppb | 137.957 | Decafluorobutane | 15.701 | 119 | 691334 | 18.45 | ppb | 99.1935 | Decafluorobutane | 15.701 | 119 | 658676 | 16.83 | ppb | 90.4839 | Decafluorobutane | 15.701 | 119 | 645900 | 16.4 | ppb | 88.172
1H-Heptafluoropropane | 17.672 | 51 | 719885 | 22.79 | ppb | 1H-Heptafluoropropane | 17.658 | 51 | 958393 | 43.95 | ppb | 192.8477 | 1H-Heptafluoropropane | 17.672 | 51 | 808440 | 22.58 | ppb | 99.0785 | 1H-Heptafluoropropane | 17.68 | 51 | 778276 | 20.66 | ppb | 90.6538 | 1H-Heptafluoropropane | 17.68 | 51 | 747864 | 19.57 | ppb | 85.871
Dodecafluoropentane | 20.466 | 69 | 2690894 | 20.76 | ppb | Dodecafluoropentane | 20.466 | 69 | 1988834 | 16.58 | ppb | 79.8651 | Dodecafluoropentane | 20.466 | 69 | 3031428 | 20.67 | ppb | 99.5665 | Dodecafluoropentane | 20.466 | 69 | 2925539 | 19.16 | ppb | 92.2929 | Dodecafluoropentane | 20.466 | 69 | 2883637 | 18.79 | ppb | 90.5106
Trichlorofluoromethane | 20.868 | 101 | 841557 | 16.98 | ppb | Trichlorofluoromethane | 20.857 | 101 | 645976 | 14.21 | ppb | 83.6867 | Trichlorofluoromethane | 20.868 | 101 | 932775 | 16.6 | ppb | 97.7621 | Trichlorofluoromethane | 20.868 | 101 | 921471 | 15.83 | ppb | 93.2273 | Trichlorofluoromethane | 20.88 | 101 | 915933 | 15.68 | ppb | 92.3439
Octafluorocyclopentene | 21.431 | 93 | 837644 | 17.99 | ppb | Octafluorocyclopentene | 21.431 | 93 | 603779 | 14.11 | ppb | 78.4325 | Octafluorocyclopentene | 21.431 | 93 | 940196 | 17.85 | ppb | 99.2218 | Octafluorocyclopentene | 21.431 | 93 | 899744 | 16.45 | ppb | 91.4397 | Octafluorocyclopentene | 21.431 | 93 | 890048 | 16.21 | ppb | 90.1056
1H-Nonafluorobutane | 22.326 | 51 | 1323025 | 23.23 | ppb | 1H-Nonafluorobutane | 22.326 | 51 | 949477 | 17.93 | ppb | 77.1847 | 1H-Nonafluorobutane | 22.326 | 51 | 1485770 | 23.05 | ppb | 99.2251 | 1H-Nonafluorobutane | 22.326 | 51 | 1404502 | 20.86 | ppb | 89.7977 | 1H-Nonafluorobutane | 22.326 | 51 | 1363461 | 20.1 | ppb | 86.526
Tetradecafluorohexane | 23.939 | 69 | 2160261m | 19.95 | ppb | Tetradecafluorohexane | 23.986 | 69 | 1783910m | 17.74 | ppb | 88.9223 | Tetradecafluorohexane | 24.033 | 69 | 2421733m | 19.71 | ppb | 98.797 | Tetradecafluorohexane | 24.025 | 69 | 2424836m | 18.94 | ppb | 94.9373 | Tetradecafluorohexane | 24.002 | 69 | 2375031m | 18.35 | ppb | 91.9799
1H-Perfluoropentane | 25.606 | 51 | 1070312m | 24.29 | ppb | 1H-Perfluoropentane | 25.64 | 51 | 842671m | 20.27 | ppb | 83.45 | 1H-Perfluoropentane | 25.674 | 51 | 1135485m | 22.25 | ppb | 91.6015 | 1H-Perfluoropentane | 25.666 | 51 | 1110176m | 20.74 | ppb | 85.3849 | 1H-Perfluoropentane | 25.649 | 51 | 1142913m | 21.48 | ppb | 88.4315
E1 | 25.674 | 101 | 1802945m | 19.65 | ppb | E1 | 25.708 | 101 | 1373829m | 15.96 | ppb | 81.2214 | E1 | 25.742 | 101 | 1938944m | 18.44 | ppb | 93.8422 | E1 | 25.734 | 101 | 1962338m | 18.02 | ppb | 91.7048 | E1 | 25.717 | 101 | 1841624m | 16.61 | ppb | 84.5293
Hexadecafluoroheptane | 26.797 | 69 | 3056622 | 20.33 | ppb | Hexadecafluoroheptane | 26.822 | 69 | 3093229 | 23.85 | ppb | 117.3143 | Hexadecafluoroheptane | 26.856 | 69 | 3436346 | 20.18 | ppb | 99.2622 | Hexadecafluoroheptane | 26.848 | 69 | 3492567 | 19.82 | ppb | 97.4914 | Hexadecafluoroheptane | 26.831 | 69 | 3435036 | 19.33 | ppb | 95.0812
1H-Perfluorohexane | 28.298 | 51 | 1759999 | 23.08 | ppb | 1H-Perfluorohexane | 28.31 | 51 | 1542935 | 22.3 | ppb | 96.6205 | 1H-Perfluorohexane | 28.333 | 51 | 1983310 | 22.98 | ppb | 99.5667 | 1H-Perfluorohexane | 28.321 | 51 | 1973731 | 21.96 | ppb | 95.1473 | 1H-Perfluorohexane | 28.321 | 51 | 1930929 | 21.28 | ppb | 92.201
Octadecafluorooctane | 29.08 | 119 | 1443452 | 17.24 | ppb | Octadecafluorooctane | 29.092 | 119 | 2030283 | 33.88 | ppb | 196.5197 | Octadecafluorooctane | 29.103 | 119 | 1667967 | 17.73 | ppb | 102.8422 | Octadecafluorooctane | 29.103 | 119 | 1655123 | 16.9 | ppb | 98.0278 | Octadecafluorooctane | 29.103 | 119 | 1641646 | 16.67 | ppb | 96.6937
1H-Perfluoroheptane | 30.552 | 51 | 2319102 | 22.96 | ppb | 1H-Perfluoroheptane | 30.552 | 51 | 2530836 | 30.58 | ppb | 133.1882 | 1H-Perfluoroheptane | 30.552 | 51 | 2603754 | 22.74 | ppb | 99.0418 | 1H-Perfluoroheptane | 30.552 | 51 | 2620007 | 22.02 | ppb | 95.9059 | 1H-Perfluoroheptane | 30.552 | 51 | 2553033 | 21.19 | ppb | 92.2909
E2 | 32.393 | 169 | 2863774 | 17.1 | ppb | E2 | 32.392 | 169 | 4151736 | 33.98 | ppb | 198.7135 | E2 | 32.393 | 169 | 3258191 | 17.24 | ppb | 100.8187 | E2 | 32.393 | 169 | 3235573 | 16.47 | ppb | 96.3158 | E2 | 32.393 | 169 | 3239760 | 16.44 | ppb | 96.1404
1H-Perfluorooctane | 32.749 | 51 | 1845139 | 22.91 | ppb | 1H-Perfluorooctane | 32.749 | 51 | 2835988 | Below Cal |  | NA | 1H-Perfluorooctane | 32.749 | 51 | 2106550 | 23.23 | ppb | 101.3968 | 1H-Perfluorooctane | 32.749 | 51 | 2030182 | 21.19 | ppb | 92.4924 | 1H-Perfluorooctane | 32.749 | 51 | 1970472 | 20.29 | ppb | 88.5639

## 4% CO2, 20 ppb VFC, 20 mL Data
091323/091323_04.D/epatemp.txt (QT Reviewed): No CO2 Riemer cal 20 | [5 empty] | 091323/091323_03.D/epatemp.txt (QT Reviewed): 4% CO2 Riemer cal 20
Compound | R.T. | QIon | Response | Conc | Units | Compound | R.T. | QIon | Response | Conc | Units | % Recovery
d5-Chlorobenzene | 33.908 | 117 | 798359 | 5 | ppb | d5-Chlorobenzene | 33.908 | 117 | 969225 | 5 | ppb
1,4-Difluorobenzene | 30.032 | 114 | 923319 | 5 | ppb | 1,4-Difluorobenzene | 30.033 | 114 | 933992 | 5 | ppb
Tetrafluoromethane | 7.273 | 69 | 234691m | 135.35 | ppb | Tetrafluoromethane | 7.225 | 69 | 212489 | 110.28 | ppb | 81.4777
Hexafluoroethane | 8.261 | 119 | 56961 | 2.31 | ppb | Hexafluoroethane | 8.406 | 119 | 1142 | 0.05 | ppb | NA
Chlorotrifluoromethane | 8.667 | 85 | 16448 | 1.89 | ppb | Chlorotrifluoromethane | 8.348 | 85 | 12523 | 1.42 | ppb | NA
Tetrafluoroethene | 8.657 | 81 | 37717 | 1.89 | ppb | Tetrafluoroethene | 8.319 | 81 | 22157 | 1.1 | ppb | NA
Trifluoromethane | 8.851 | 69 | 20603 | 1.82 | ppb | Trifluoromethane | 8.629 | 69 | 20232 | 1.77 | ppb | NA
Difluoromethane | 10.931 | 51 | 14290 | 1.71 | ppb | Difluoromethane | 10.88 | 51 | 12396 | 1.47 | ppb | NA
Octafluoropropane | 11.055 | 69 | 180358 | 1.81 | ppb | Octafluoropropane | 11.014 | 69 | 206643 | 2.05 | ppb | NA
Fluoromethane | 11.353 | 33 | 5811 | 1.69 | ppb | Fluoromethane | 11.312 | 33 | 5477 | 1.57 | ppb | NA
Pentafluoroethane | 12.692 | 51 | 55578 | 1.9 | ppb | Pentafluoroethane | 12.661 | 51 | 63379 | 2.15 | ppb | NA
1,1,1-Trifluoroethane | 13.055 | 64 | 4508 | 1.93 | ppb | 1,1,1-Trifluoroethane | 13.022 | 64 | 5153 | 2.18 | ppb | NA
Hexafluoropropene | 13.044 | 131 | 62290 | 1.74 | ppb | Hexafluoropropene | 13.011 | 131 | 61419 | 1.7 | ppb | NA
Hexafluoropropene_oxide | 13.131 | 69 | 14041m | 1.81 | ppb | Hexafluoropropene_oxide | 13.109 | 69 | 12734m | 1.62 | ppb | NA
Chlorodifluoromethane | 13.577 | 51 | 58281 | 1.96 | ppb | Chlorodifluoromethane | 13.555 | 51 | 69473 | 2.32 | ppb | NA
Octafluorocyclobutane | 15.484 | 100 | 158569 | 1.82 | ppb | Octafluorocyclobutane | 15.464 | 100 | 116954 | 1.32 | ppb | NA
1,1,1,2-Tetrafluoroethane | 15.62 | 33 | 18934 | 1.93 | ppb | 1,1,1,2-Tetrafluoroethane | 15.61 | 33 | 22716 | 2.29 | ppb | NA
Decafluorobutane | 15.702 | 119 | 75175 | 1.78 | ppb | Decafluorobutane | 15.682 | 119 | 56027 | 1.31 | ppb | NA
1H-Heptafluoropropane | 17.681 | 51 | 79062 | 1.82 | ppb | 1H-Heptafluoropropane | 17.674 | 51 | 56680 | 1.28 | ppb | NA
Dodecafluoropentane | 20.479 | 69 | 317232 | 1.94 | ppb | Dodecafluoropentane | 20.479 | 69 | 270487 | 1.64 | ppb | NA
Trichlorofluoromethane | 20.881 | 101 | 107863 | 1.77 | ppb | Trichlorofluoromethane | 20.881 | 101 | 83661 | 1.35 | ppb | NA
Octafluorocyclopentene | 21.444 | 93 | 105575 | 1.86 | ppb | Octafluorocyclopentene | 21.444 | 93 | 87492 | 1.53 | ppb | NA
1H-Nonafluorobutane | 22.34 | 51 | 137334 | 1.91 | ppb | 1H-Nonafluorobutane | 22.34 | 51 | 110477 | 1.51 | ppb | NA
Tetradecafluorohexane | 24.11 | 69 | 333314 | 2.2 | ppb | Tetradecafluorohexane | 24.11 | 69 | 332252 | 2.16 | ppb | NA
1H-Perfluoropentane | 25.748 | 51 | 130617m | 2.1 | ppb | 1H-Perfluoropentane | 25.74 | 51 | 131324m | 2.08 | ppb | NA
E1 | 25.808 | 101 | 246886m | 2.01 | ppb | E1 | 25.808 | 101 | 218813m | 1.75 | ppb | NA
Hexadecafluoroheptane | 26.913 | 69 | 443276 | 2.17 | ppb | Hexadecafluoroheptane | 26.913 | 69 | 564240 | 2.75 | ppb | NA
1H-Perfluorohexane | 28.365 | 51 | 218746 | 2.11 | ppb | 1H-Perfluorohexane | 28.365 | 51 | 238810 | 2.28 | ppb | NA
Octadecafluorooctane | 29.135 | 119 | 236844 | 2.14 | ppb | Octadecafluorooctane | 29.136 | 119 | 333952 | 3.01 | ppb | NA
1H-Perfluoroheptane | 30.573 | 51 | 294112 | 2.09 | ppb | 1H-Perfluoroheptane | 30.573 | 51 | 390091 | 2.76 | ppb | NA
E2 | 32.402 | 169 | 406201 | 1.86 | ppb | E2 | 32.402 | 169 | 514159 | 2.33 | ppb | NA
1H-Perfluorooctane | 32.77 | 51 | 214665 | 1.9 | ppb | 1H-Perfluorooctane | 32.77 | 51 | 262853 | 2.3 | ppb | NA