## Read Me
Figure S1: OTM-50 Calibration Curves
Calibration curves were generated by analyzing variable volume VFC standards prepared in 6 L canisters using System 2.
CF4 calibration curves were generated by injecting 20 mL volumes of 10, 50, 100, 150, 200, and 300 ppbv CF4 standards.
Calibration curves for the remaining 29 target compounds were generated by injection 200 mL volumes of 1, 5, 10, 15, 20, and 30 ppbv VFC standards.
Relative response factors (RRFs) were calculated using 1,4-difluorobenzene as the internal standard.
Abbreviations
BDL | below detection limit
C | Celsius
Cal | calibration
CF4 | tetrafluoromethane
Conc | concentration
Dil | dilution
ID | identification
L | liter
mL | milliliter
OTM-50 | Other Test Method 50
ppb | parts-per-billion
ppbv | parts per billion by volume
RRF | relative response factor
RT | retention time
VFC | volatile fluorinated compound

## CF4_20mL
050824/050824_03.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 10 ppbv_Can 36850_vol 20 ml | [8 empty] | 050824/050824_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 50 ppbv_Can 37123_vol 20 ml | [8 empty] | 050824/050824_05.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 100 ppbv_Can 40857_vol 20 ml | [8 empty] | 050824/050824_06.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 150 ppbv_Can 40872_vol 20 ml | [8 empty] | 050824/050824_07.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal200 ppbv_Can 42881_vol 20 ml | [8 empty] | 050824/050824_08.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal300 ppbv_Can 45674_vol 20 ml
Compound | R.T. | QIon | Response | Conc | num Conc | Units | Vol (mL) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Vol (mL) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Vol (mL) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Vol (mL) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Vol (mL) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Vol (mL) | RRF
d5-Chlorobenzene | 35.182 | 117 | 898062 | 50 |  | ppb |  |  | d5-Chlorobenzene | 35.181 | 117 | 898407 | 50 |  | ppb |  |  | d5-Chlorobenzene | 35.182 | 117 | 886970 | 50 |  | ppb |  |  | d5-Chlorobenzene | 35.181 | 117 | 882485 | 50 |  | ppb |  |  | d5-Chlorobenzene | 35.179 | 117 | 881529 | 50 |  | ppb |  |  | d5-Chlorobenzene | 35.181 | 117 | 868247 | 50 |  | ppb
1,4-difluorobenzene | 30.846 | 114 | 958904 | 50 |  | ppb |  |  | 1,4-difluorobenzene | 30.845 | 114 | 965962 | 50 |  | ppb |  |  | 1,4-difluorobenzene | 30.846 | 114 | 947756 | 50 |  | ppb |  |  | 1,4-difluorobenzene | 30.846 | 114 | 942737 | 50 |  | ppb |  |  | 1,4-difluorobenzene | 30.843 | 114 | 943860 | 50 |  | ppb |  |  | 1,4-difluorobenzene | 30.845 | 114 | 917547 | 50 |  | ppb
tetrafluoromethane | 7.03 | 69 | 3703 | 7.987 |  | ppb | 20 | 0.024175 | tetrafluoromethane | 7.052 | 69 | 22239 | 48.209 |  | ppb | 20 | 0.023878 | tetrafluoromethane | 7.03 | 69 | 44304 | 99.46 |  | ppb | 20 | 0.0235 | tetrafluoromethane | 7.03 | 69 | 62234 | 142.374 |  | ppb | 20 | 0.023183 | tetrafluoromethane | 7.051 | 69 | 79795 | 184.828 |  | ppb | 20 | 0.02287 | tetrafluoromethane | 7.03 | 69 | 123361 | 305.876 |  | ppb | 20 | 0.021977

## VFCs_200mL
050824/050824_18.D/epatemp.txt (QT Reviewed): Riemer PFAS_Cal 1 ppbv_Can 36850_vol 200 ml | [8 empty] | 050824/050824_19.D/epatemp.txt (QT Reviewed): Riemer PFAS_Cal 2.5 ppbv_Can 36862_vol 200 ml | [8 empty] | 050824/050824_20.D/epatemp.txt (QT Reviewed): Riemer PFAS_Cal 5.0 ppbv_Can 37123_vol 200 ml | [8 empty] | 050824/050824_21.D/epatemp.txt (QT Reviewed): Riemer PFAS_Cal 10 ppbv_Can 40857_vol 200 ml | [8 empty] | 050824/050824_22.D/epatemp.txt (QT Reviewed): Riemer PFAS_Cal 20 ppbv_Can 42881_vol 200 ml | [8 empty] | 050824/050824_23.D/epatemp.txt (QT Reviewed): Riemer PFAS_Cal 30 ppbv_Can 45674_vol 200 ml
Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | RRF | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | RRF
d5-Chlorobenzene | 35.136 | 117 | 833888 | 5 |  | ppb |  |  | d5-Chlorobenzene | 35.136 | 117 | 846284 | 5 |  | ppb |  |  | d5-Chlorobenzene | 35.135 | 117 | 827025 | 5 |  | ppb |  |  | d5-Chlorobenzene | 35.124 | 117 | 825511 | 5 |  | ppb |  |  | d5-Chlorobenzene | 35.124 | 117 | 825907 | 5 |  | ppb |  |  | d5-Chlorobenzene | 35.124 | 117 | 828636 | 5 |  | ppb
1,4-difluorobenzene | 30.765 | 114 | 925713 | 5 |  | ppb |  |  | 1,4-difluorobenzene | 30.765 | 114 | 941949 | 5 |  | ppb |  |  | 1,4-difluorobenzene | 30.753 | 114 | 912525 | 5 |  | ppb |  |  | 1,4-difluorobenzene | 30.765 | 114 | 898098 | 5 |  | ppb |  |  | 1,4-difluorobenzene | 30.765 | 114 | 904952 | 5 |  | ppb |  |  | 1,4-difluorobenzene | 30.765 | 114 | 893266 | 5 |  | ppb
tetrafluoromethane | 0 | 0 | N.D. | [5 empty] | tetrafluoromethane | 0 | 0 | N.D. | [5 empty] | tetrafluoromethane | 0 | 0 | N.D. | [5 empty] | tetrafluoromethane | 7.03 | 69 | 414 | No |  | Calib |  |  | tetrafluoromethane | 7.03 | 69 | 861 | No |  | Calib | # |  | tetrafluoromethane | 7.051 | 69 | 1313 | No |  | Calib | #
hexafluoroethane | 8.182 | 119 | 8647 | 0.808 |  | ppb | 200 | 0.057803 | hexafluoroethane | 8.193 | 119 | 23780 | 2.19 |  | ppb | 200 | 0.057638 | hexafluoroethane | 8.183 | 119 | 50197 | 4.812 |  | ppb | 200 | 0.057158 | hexafluoroethane | 8.182 | 119 | 99337 | 9.892 |  | ppb | 200 | 0.055908 | hexafluoroethane | 8.182 | 119 | 183960 | 18.435 |  | ppb | 200 | 0.055135 | hexafluoroethane | 8.193 | 119 | 276153 | 29.023 |  | ppb | 200 | 0.053259
chlorotrifluoromethane | 8.606 | 85 | 2512 | 0.881 |  | ppb | 200 | 0.015401 | chlorotrifluoromethane | 8.617 | 85 | 6717 | 2.327 |  | ppb | 200 | 0.015322 | chlorotrifluoromethane | 8.596 | 85 | 14119 | 5.104 |  | ppb | 200 | 0.015157 | chlorotrifluoromethane | 8.596 | 85 | 27606 | 10.344 |  | ppb | 200 | 0.014858 | chlorotrifluoromethane | 8.596 | 85 | 50790 | 19.245 |  | ppb | 200 | 0.014582 | chlorotrifluoromethane | 8.606 | 85 | 75990 | 30.901 |  | ppb | 200 | 0.013765
trifluoromethane | 8.824 | 69 | 3787 | 0.902 |  | ppb | 200 | 0.022677 | trifluoromethane | 8.835 | 69 | 9733 | 2.288 |  | ppb | 200 | 0.02258 | trifluoromethane | 8.824 | 69 | 21212 | 5.202 |  | ppb | 200 | 0.022343 | trifluoromethane | 8.824 | 69 | 41410 | 10.495 |  | ppb | 200 | 0.021967 | trifluoromethane | 8.824 | 69 | 74710 | 19.047 |  | ppb | 200 | 0.021672 | trifluoromethane | 8.835 | 69 | 114393 | 31.496 |  | ppb | 200 | 0.02033
octafluoropropane | 11.411 | 69 | 30246 | 0.86 |  | ppb | 200 | 0.18996 | octafluoropropane | 11.411 | 69 | 82863 | 2.326 |  | ppb | 200 | 0.1891 | octafluoropropane | 11.397 | 69 | 173291 | 5.077 |  | ppb | 200 | 0.18702 | octafluoropropane | 11.397 | 69 | 335260 | 10.184 |  | ppb | 200 | 0.18328 | octafluoropropane | 11.397 | 69 | 626740 | 19.322 |  | ppb | 200 | 0.17922 | octafluoropropane | 11.411 | 69 | 932205 | 30.724 |  | ppb | 200 | 0.16983
difluoromethane | 11.18 | 51 | 3050 | 0.895 |  | ppb | 200 | 0.018406 | difluoromethane | 11.194 | 51 | 8133 | 2.354 |  | ppb | 200 | 0.018339 | difluoromethane | 11.18 | 51 | 17195 | 5.175 |  | ppb | 200 | 0.018206 | difluoromethane | 11.18 | 51 | 33793 | 10.437 |  | ppb | 200 | 0.018026 | difluoromethane | 11.18 | 51 | 62463 | 19.233 |  | ppb | 200 | 0.017944 | difluoromethane | 11.18 | 51 | 95937 | 31.723 |  | ppb | 200 | 0.016928
pentafluoroethane | 13.133 | 51 | 9629 | 0.91 |  | ppb | 200 | 0.057152 | pentafluoroethane | 13.134 | 51 | 26578 | 2.475 |  | ppb | 200 | 0.057002 | pentafluoroethane | 13.134 | 51 | 54510 | 5.27 |  | ppb | 200 | 0.056675 | pentafluoroethane | 13.134 | 51 | 106570 | 10.572 |  | ppb | 200 | 0.056121 | pentafluoroethane | 13.134 | 51 | 202776 | 20.017 |  | ppb | 200 | 0.055971 | pentafluoroethane | 13.134 | 51 | 301742 | 31.288 |  | ppb | 200 | 0.053982
octafluorocyclobutane | 16.313 | 100 | 25447 | 0.889 |  | ppb | 200 | 0.15461 | octafluorocyclobutane | 16.313 | 100 | 68355 | 2.361 |  | ppb | 200 | 0.15368 | octafluorocyclobutane | 16.313 | 100 | 142483 | 5.144 |  | ppb | 200 | 0.15177 | octafluorocyclobutane | 16.313 | 100 | 273474 | 10.249 |  | ppb | 200 | 0.14855 | octafluorocyclobutane | 16.3 | 100 | 506292 | 19.347 |  | ppb | 200 | 0.14459 | octafluorocyclobutane | 16.299 | 100 | 753873 | 31.383 |  | ppb | 200 | 0.13446
fluoromethane | 11.628 | 33 | 1664 | 1.053 |  | ppb | 200 | 0.0085353 | fluoromethane | 11.643 | 33 | 4060 | 2.535 |  | ppb | 200 | 0.0085014 | fluoromethane | 11.628 | 33 | 9141 | 5.946 |  | ppb | 200 | 0.0084235 | fluoromethane | 11.628 | 33 | 17386 | 11.447 |  | ppb | 200 | 0.0084558 | fluoromethane | 11.628 | 33 | 32435 | 21.432 |  | ppb | 200 | 0.0083617 | fluoromethane | 11.643 | 33 | 47802 | 37.514 |  | ppb | 200 | 0.0071325
tetrafluoroethylene | 8.585 | 81 | 6085 | 0.87 |  | ppb | 200 | 0.037778 | tetrafluoroethylene | 8.606 | 81 | 16281 | 2.295 |  | ppb | 200 | 0.037657 | tetrafluoroethylene | 8.585 | 81 | 34846 | 5.117 |  | ppb | 200 | 0.037313 | tetrafluoroethylene | 8.585 | 81 | 67209 | 10.214 |  | ppb | 200 | 0.036633 | tetrafluoroethylene | 8.585 | 81 | 122587 | 18.649 |  | ppb | 200 | 0.036319 | tetrafluoroethylene | 8.596 | 81 | 188590 | 30.511 |  | ppb | 200 | 0.034598
hexafluoropropene | 13.539 | 131 | 8622 | 0.837 |  | ppb | 200 | 0.055639 | hexafluoropropene | 13.553 | 131 | 23319 | 2.237 |  | ppb | 200 | 0.055333 | hexafluoropropene | 13.539 | 131 | 48806 | 4.888 |  | ppb | 200 | 0.05471 | hexafluoropropene | 13.539 | 131 | 92791 | 9.696 |  | ppb | 200 | 0.053279 | hexafluoropropene | 13.539 | 131 | 175284 | 18.718 |  | ppb | 200 | 0.05174 | hexafluoropropene | 13.539 | 131 | 261316 | 29.488 |  | ppb | 200 | 0.049603
1,1,1-trifluoroethane | 13.524 | 64 | 696 | 0.879 |  | ppb | 200 | 0.0042768 | 1,1,1-trifluoroethane | 13.524 | 64 | 1852 | 2.312 |  | ppb | 200 | 0.004252 | 1,1,1-trifluoroethane | 13.524 | 64 | 3830 | 4.992 |  | ppb | 200 | 0.0042039 | 1,1,1-trifluoroethane | 13.524 | 64 | 7484 | 10.168 |  | ppb | 200 | 0.0040977 | 1,1,1-trifluoroethane | 13.51 | 64 | 13684 | 18.922 |  | ppb | 200 | 0.0039957 | 1,1,1-trifluoroethane | 13.524 | 64 | 20797 | 30.941 |  | ppb | 200 | 0.0037623
hexafluoropropene_oxide | 13.756 | 69 | 657 | 0.754 |  | ppb | 200 | 0.0047064 | hexafluoropropene_oxide | 13.756 | 69 | 2171 | 2.358 |  | ppb | 200 | 0.0048872 | hexafluoropropene_oxide | 13.742 | 69 | 5247 | 5.483 |  | ppb | 200 | 0.0052435 | hexafluoropropene_oxide | 13.742 | 69 | 12396 | 11.495 |  | ppb | 200 | 0.0060037 | hexafluoropropene_oxide | 13.742 | 69 | 28789 | 21.492 |  | ppb | 200 | 0.0074011 | hexafluoropropene_oxide | 13.756 | 69 | 46965 | 32.318 |  | ppb | 200 | 0.0081343
chlorodifluoromethane | 14.002 | 51 | 9936 | 0.865 |  | ppb | 200 | 0.062042 | chlorodifluoromethane | 14.016 | 51 | 27587 | 2.369 |  | ppb | 200 | 0.061813 | chlorodifluoromethane | 14.002 | 51 | 56196 | 5.019 |  | ppb | 200 | 0.06135 | chlorodifluoromethane | 14.002 | 51 | 109204 | 10.096 |  | ppb | 200 | 0.060219 | chlorodifluoromethane | 14.002 | 51 | 206745 | 19.177 |  | ppb | 200 | 0.059566 | chlorodifluoromethane | 14.002 | 51 | 309533 | 29.948 |  | ppb | 200 | 0.057853
1,1,1,2-tetrafluoroethane | 16.299 | 33 | 3977 | 0.989 |  | ppb | 200 | 0.02172 | 1,1,1,2-tetrafluoroethane | 16.299 | 33 | 11791 | 2.902 |  | ppb | 200 | 0.021567 | 1,1,1,2-tetrafluoroethane | 16.3 | 33 | 25718 | 6.631 |  | ppb | 200 | 0.021251 | 1,1,1,2-tetrafluoroethane | 16.3 | 33 | 49440 | 12.888 |  | ppb | 200 | 0.021357 | 1,1,1,2-tetrafluoroethane | 16.3 | 33 | 90752 | 24.58 |  | ppb | 200 | 0.020399 | 1,1,1,2-tetrafluoroethane | 16.299 | 33 | 137144 | Below Cal |  | ppb | 200 | #VALUE!
decafluorobutane | 16.541 | 119 | 11497 | 0.844 |  | ppb | 200 | 0.073576 | decafluorobutane | 16.541 | 119 | 30765 | 2.235 |  | ppb | 200 | 0.073067 | decafluorobutane | 16.541 | 119 | 61551 | 4.675 |  | ppb | 200 | 0.07214 | decafluorobutane | 16.541 | 119 | 124551 | 10.018 |  | ppb | 200 | 0.069217 | decafluorobutane | 16.541 | 119 | 226648 | 18.676 |  | ppb | 200 | 0.067052 | decafluorobutane | 16.541 | 119 | 337002 | 29.645 |  | ppb | 200 | 0.063631
1H-Heptafluoropropane | 18.497 | 69 | 20236 | 0.904 |  | ppb | 200 | 0.12091 | 1H-Heptafluoropropane | 18.498 | 69 | 53863 | 2.377 |  | ppb | 200 | 0.12028 | 1H-Heptafluoropropane | 18.498 | 69 | 112206 | 5.163 |  | ppb | 200 | 0.11908 | 1H-Heptafluoropropane | 18.498 | 69 | 215189 | 10.242 |  | ppb | 200 | 0.11697 | 1H-Heptafluoropropane | 18.498 | 69 | 402744 | 19.412 |  | ppb | 200 | 0.11463 | 1H-Heptafluoropropane | 18.498 | 69 | 601576 | 30.884 |  | ppb | 200 | 0.10903
octafluorocyclopentene | 22.327 | 93 | 16859 | 0.87 |  | ppb | 200 | 0.10467 | octafluorocyclopentene | 22.327 | 93 | 46038 | 2.352 |  | ppb | 200 | 0.1039 | octafluorocyclopentene | 22.322 | 93 | 93743 | 5.015 |  | ppb | 200 | 0.10242 | octafluorocyclopentene | 22.315 | 93 | 177925 | 9.954 |  | ppb | 200 | 0.099514 | octafluorocyclopentene | 22.315 | 93 | 342559 | 19.894 |  | ppb | 200 | 0.095139 | octafluorocyclopentene | 22.315 | 93 | 487118 | 30.389 |  | ppb | 200 | 0.089724
trichlorofluoromethane | 21.553 | 101 | 17938 | 0.945 |  | ppb | 200 | 0.10253 | trichlorofluoromethane | 21.553 | 101 | 47966 | 2.497 |  | ppb | 200 | 0.10197 | trichlorofluoromethane | 21.548 | 101 | 99294 | 5.398 |  | ppb | 200 | 0.10079 | trichlorofluoromethane | 21.553 | 101 | 189272 | 10.612 |  | ppb | 200 | 0.099297 | trichlorofluoromethane | 21.553 | 101 | 348551 | 19.783 |  | ppb | 200 | 0.097346 | trichlorofluoromethane | 21.553 | 101 | 517364 | 32.103 |  | ppb | 200 | 0.090207
dodecafluoro-n-pentane | 21.415 | 69 | 54200 | 0.888 |  | ppb | 200 | 0.32967 | dodecafluoro-n-pentane | 21.415 | 69 | 143261 | 2.323 |  | ppb | 200 | 0.32736 | dodecafluoro-n-pentane | 21.397 | 69 | 313846 | 5.341 |  | ppb | 200 | 0.32197 | dodecafluoro-n-pentane | 21.397 | 69 | 553387 | 9.781 |  | ppb | 200 | 0.31499 | dodecafluoro-n-pentane | 21.397 | 69 | 1023088 | 18.822 |  | ppb | 200 | 0.30033 | dodecafluoro-n-pentane | 21.397 | 69 | 1462478 | 28.97 |  | ppb | 200 | 0.28257
1H-Nonafluorobutane | 23.173 | 51 | 25485 | 0.844 |  | ppb | 200 | 0.16309 | 1H-Nonafluorobutane | 23.173 | 51 | 65582 | 2.148 |  | ppb | 200 | 0.16207 | 1H-Nonafluorobutane | 23.168 | 51 | 132682 | 4.546 |  | ppb | 200 | 0.15992 | 1H-Nonafluorobutane | 23.161 | 51 | 255454 | 9.273 |  | ppb | 200 | 0.15337 | 1H-Nonafluorobutane | 23.161 | 51 | 478286 | 18.182 |  | ppb | 200 | 0.14534 | 1H-Nonafluorobutane | 23.161 | 51 | 676748 | 27.111 |  | ppb | 200 | 0.13972
tetradecafluorohexane | 24.904 | 69 | 60607 | 0.89 |  | ppb | 200 | 0.36781 | tetradecafluorohexane | 24.897 | 69 | 159055 | 2.315 |  | ppb | 200 | 0.3647 | tetradecafluorohexane | 24.888 | 69 | 322681 | 4.933 |  | ppb | 200 | 0.35842 | tetradecafluorohexane | 24.905 | 69 | 597611 | 9.626 |  | ppb | 200 | 0.34564 | tetradecafluorohexane | 24.897 | 69 | 1113259 | 18.901 |  | ppb | 200 | 0.32543 | tetradecafluorohexane | 24.897 | 69 | 1586111 | 29.044 |  | ppb | 200 | 0.30568
1H-Perfluoropentane | 26.489 | 69 | 19301 | 0.879 |  | ppb | 200 | 0.1186 | 1H-Perfluoropentane | 26.489 | 69 | 52777 | 2.381 |  | ppb | 200 | 0.11766 | 1H-Perfluoropentane | 26.488 | 69 | 100367 | 4.744 |  | ppb | 200 | 0.11592 | 1H-Perfluoropentane | 26.489 | 69 | 204082 | 10.182 |  | ppb | 200 | 0.11159 | 1H-Perfluoropentane | 26.489 | 69 | 359736 | 18.752 |  | ppb | 200 | 0.10599 | 1H-Perfluoropentane | 26.479 | 69 | 510745 | 28.631 |  | ppb | 200 | 0.099852
E1 | 26.567 | 101 | 43136 | 0.932 |  | ppb | 200 | 0.24999 | E1 | 26.567 | 101 | 108675 | 2.326 |  | ppb | 200 | 0.24801 | E1 | 26.566 | 101 | 213313 | 4.788 |  | ppb | 200 | 0.24411 | E1 | 26.567 | 101 | 396264 | 9.43 |  | ppb | 200 | 0.23395 | E1 | 26.567 | 101 | 760129 | 19.127 |  | ppb | 200 | 0.21958 | E1 | 26.567 | 101 | 1053363 | 28.135 |  | ppb | 200 | 0.20957
hexadecafluoroheptane | 27.668 | 69 | 67952 | 0.899 |  | ppb | 200 | 0.40826 | hexadecafluoroheptane | 27.669 | 69 | 186406 | 2.444 |  | ppb | 200 | 0.40486 | hexadecafluoroheptane | 27.658 | 69 | 363369 | 4.996 |  | ppb | 200 | 0.39852 | hexadecafluoroheptane | 27.669 | 69 | 696385 | 10.092 |  | ppb | 200 | 0.38417 | hexadecafluoroheptane | 27.659 | 69 | 1299234 | 19.594 |  | ppb | 200 | 0.36636 | hexadecafluoroheptane | 27.659 | 69 | 1892989 | 31.155 |  | ppb | 200 | 0.3401
1H-Perfluorohexane | 29.121 | 51 | 40437 | 0.837 |  | ppb | 200 | 0.26094 | 1H-Perfluorohexane | 29.121 | 51 | 112285 | 2.304 |  | ppb | 200 | 0.25869 | 1H-Perfluorohexane | 29.11 | 51 | 220393 | 4.742 |  | ppb | 200 | 0.25466 | 1H-Perfluorohexane | 29.112 | 51 | 416152 | 9.494 |  | ppb | 200 | 0.24403 | 1H-Perfluorohexane | 29.112 | 51 | 776714 | 18.607 |  | ppb | 200 | 0.23064 | 1H-Perfluorohexane | 29.111 | 51 | 1115296 | 28.561 |  | ppb | 200 | 0.21858
octadecafluorooctane | 29.926 | 119 | 32999 | 0.546 |  | ppb | 200 | 0.32644 | octadecafluorooctane | 29.926 | 119 | 89851 | 1.475 |  | ppb | 200 | 0.32335 | octadecafluorooctane | 29.913 | 119 | 164127 | 2.839 |  | ppb | 200 | 0.31677 | octadecafluorooctane | 29.914 | 119 | 329125 | 6.528 |  | ppb | 200 | 0.28069 | octadecafluorooctane | 29.914 | 119 | 602701 | 14.347 |  | ppb | 200 | 0.23211 | octadecafluorooctane | 29.914 | 119 | 913318 | 24.906 |  | ppb | 200 | 0.20526
1H-Perfluoroheptane | 31.467 | 51 | 117463 | 0.914 |  | ppb | 200 | 0.69414 | 1H-Perfluoroheptane | 31.467 | 51 | 317778 | 2.455 |  | ppb | 200 | 0.68709 | 1H-Perfluoroheptane | 31.466 | 51 | 601505 | 4.885 |  | ppb | 200 | 0.67468 | 1H-Perfluoroheptane | 31.467 | 51 | 1144512 | 9.873 |  | ppb | 200 | 0.64538 | 1H-Perfluoroheptane | 31.455 | 51 | 2100962 | 19.162 |  | ppb | 200 | 0.60579 | 1H-Perfluoroheptane | 31.455 | 51 | 3082800 | 31.376 |  | ppb | 200 | 0.54997
1H-Perfluorooctane | 33.997 | 51 | 93905 | 0.882 |  | ppb | 200 | 0.57506 | 1H-Perfluorooctane | 33.997 | 51 | 249959 | 2.339 |  | ppb | 200 | 0.56726 | 1H-Perfluorooctane | 33.996 | 51 | 441083 | 4.347 |  | ppb | 200 | 0.55598 | 1H-Perfluorooctane | 33.997 | 51 | 862256 | 9.23 |  | ppb | 200 | 0.52009 | 1H-Perfluorooctane | 33.986 | 51 | 1559133 | 18.178 |  | ppb | 200 | 0.47389 | 1H-Perfluorooctane | 33.986 | 51 | 2351752 | 31.544 |  | ppb | 200 | 0.41731
E2 | 33.479 | 169 | 138864 | 0.856 |  | ppb | 200 | 0.87621 | E2 | 33.48 | 169 | 374037 | 2.284 |  | ppb | 200 | 0.86928 | E2 | 33.479 | 169 | 702237 | 4.494 |  | ppb | 200 | 0.8562 | E2 | 33.48 | 169 | 1341038 | 9.156 |  | ppb | 200 | 0.81542 | E2 | 33.48 | 169 | 2397981 | 17.273 |  | ppb | 200 | 0.76705 | E2 | 33.48 | 169 | 3624101 | 28.292 |  | ppb | 200 | 0.71701