## Cover Letter
MEMORANDUM | (LABORATORY DATA REPORT)
EPA Metals Laboratory
To: |  | Mark Sievers, TetraTech | Analyst: | Lisa Costantino
|  |  | Lab: | Metals
QA Reviewer: |  | Rick Wilkin | Date: | 2024-07-15
Technical Directive No.: |  | EPA MA140 | Originator: | R. Wilkin
Task No.: |  | SHC 402.4.1 | Copies: | L. Costantino
R. Ellison
T. Lee
E. Barth
M. Noerpel
Sample Site/Project: |  | Geochemical Characterization of Acid Mine Drainage, Shamokin ROAR R3
Date Collected: |  | 5/20/2024 - 5/21/2024 | Sample Set No.: | 8044
Date Received: |  | 2024-05-23 | Sample Matrix: | Surface water preserved with HNO3, filtered
Date Analyzed: |  | 6/12/224 | Analysis Type: | HR-ICP-MS
No. Samples Analyzed: |  | 14 | Sample Preparation: | NA
Method(s) Used : |  | K-GCRD-1158-1 Standard Operating Procedure for the Determination of
Rare Earth Elements using HR-ICP-MS
Comments: The MDLs and QLs reflect those determined in August 2023. See notes on the Data and QA Data tabs.

## Data
EPA Metals Laboratory
Results Report
Laboratory: | Metals |  | Report Date: | 2024-07-15
Technical Directive: | EPA MA140 |  |  |  |  | Sample Results
Analysts: | LMC |  |  | Field Sample ID |  | BM-SW-1 |  | BM-SW-1 FD |  | HC-SW-1 |  | HC-SW-1 FD |  | Cam-SW-1 |  | Cam-SW-1 FD |  | SR-SW-2 |  | SR-SW-2 FD |  | SR-SW-5 |  | SR-SW-9 |  | EM-SW-1 |  | EM-SW-1 FD |  | EM-SW-5 |  | EM-SW-9
|  |  |  | Sample Lab ID |  | 8044-1 |  | 8044-2 |  | 8044-9 |  | 8044-10 |  | 8044-21 |  | 8044-22 |  | 8044-31 |  | 8044-32 |  | 8044-35 |  | 8044-38 |  | 8044-45 |  | 8044-46 |  | 8044-51 |  | 8044-54
Method: | EPA 200.8 |  |  | Date Collected |  | 2024-05-20 |  | 2024-05-20 |  | 2024-05-20 |  | 2024-05-20 |  | 2024-05-21 |  | 2024-05-21 |  | 2024-05-21 |  | 2024-05-21 |  | 2024-05-21 |  | 2024-05-21 |  | 2024-05-21 |  | 2024-05-21 |  | 2024-05-21 |  | 2024-05-21
|  |  |  | Date Analyzed |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12
Analytes |  |  | Unit | MDL | QL | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF | Data | DF
Names |  | Codes
Y89(MR) |  | 7440-65-5 | ug/L | 0.002 | 0.02 | 8.965 | 1 | 9.055 | 1 | 1.555 | 1 | 1.581 | 1 | 1.001 | 1 | 1.007 | 1 | 1.391 | 1 | 1.401 | 1 | 0.889 | 1 | 1.508 | 1 | 1.721 | 1 | 1.684 | 1 | 1.121 | 1 | 1.499 | 1
La139(HR) |  | 7439-91-0 | ug/L | 0.002 | 0.02 | 6.36 | 1 | 6.516 | 1 | 0.944 | 1 | 0.932 | 1 | 0.628 | 1 | 0.651 | 1 | 0.904 | 1 | 0.893 | 1 | 0.459 | 1 | 0.999 | 1 | 1.377 | 1 | 1.322 | 1 | 0.957 | 1 | 1.251 | 1
Ce140(HR) |  | 7440-45-1 | ug/L | 0.002 | 0.02 | 14.549 | 1 | 14.834 | 1 | 1.858 | 1 | 1.83 | 1 | 1.07 | 1 | 1.105 | 1 | 1.734 | 1 | 1.732 | 1 | 0.913 | 1 | 1.947 | 1 | 2.664 | 1 | 2.554 | 1 | 1.637 | 1 | 2.249 | 1
Pr141(HR) |  | 7440-10-0 | ug/L | 0.002 | 0.02 | 1.759 | 1 | 1.819 | 1 | 0.211 | 1 | 0.21 | 1 | 0.115 | 1 | 0.122 | 1 | 0.196 | 1 | 0.198 | 1 | 0.108 | 1 | 0.225 | 1 | 0.301 | 1 | 0.287 | 1 | 0.169 | 1 | 0.23 | 1
Nd146(MR) |  | 7440-00-08 | ug/L | 0.002 | 0.02 | 7.597 | 1 | 7.63 | 1 | 0.835 | 1 | 0.851 | 1 | 0.453 | 1 | 0.478 | 1 | 0.799 | 1 | 0.828 | 1 | 0.469 | 1 | 0.891 | 1 | 1.139 | 1 | 1.114 | 1 | 0.626 | 1 | 0.877 | 1
Sm147(MR) |  | 7440-19-9 | ug/L | 0.002 | 0.02 | 1.684 | 1 | 1.72 | 1 | 0.174 | 1 | 0.179 | 1 | 0.095 | 1 | 0.091 | 1 | 0.161 | 1 | 0.155 | 1 | 0.097 | 1 | 0.164 | 1 | 0.225 | 1 | 0.217 | 1 | 0.113 | 1 | 0.153 | 1
Eu151(MR) |  | 7440-53-1 | ug/L | 0.002 | 0.02 | 0.344 | 1 | 0.348 | 1 | 0.034 | 1 | 0.035 | 1 | 0.015 (BQL) | 1 | 0.017 (BQL) | 1 | 0.032 | 1 | 0.035 | 1 | 0.021 | 1 | 0.036 | 1 | 0.041 | 1 | 0.043 | 1 | 0.019 (BQL) | 1 | 0.032 | 1
Gd157(HR) |  | 7440-54-2 | ug/L | 0.002 | 0.02 | 1.885 | 1 | 1.941 | 1 | 0.25 | 1 | 0.24 | 1 | 0.159 | 1 | 0.114 | 1 | 0.237 | 1 | 0.217 | 1 | 0.133 | 1 | 0.225 | 1 | 0.284 | 1 | 0.312 | 1 | 0.148 | 1 | 0.212 | 1
Tb159(HR) |  | 7440-27-9 | ug/L | 0.002 | 0.02 | 0.279 | 1 | 0.283 | 1 | 0.036 | 1 | 0.036 | 1 | 0.019 (BQL) | 1 | 0.016 (BQL) | 1 | 0.033 | 1 | 0.035 | 1 | 0.02 | 1 | 0.035 | 1 | 0.045 | 1 | 0.041 | 1 | 0.024 | 1 | 0.03 | 1
Dy163(MR) |  | 7429-91-6 | ug/L | 0.002 | 0.02 | 1.728 | 1 | 1.758 | 1 | 0.22 | 1 | 0.216 | 1 | 0.12 | 1 | 0.12 | 1 | 0.206 | 1 | 0.215 | 1 | 0.125 | 1 | 0.227 | 1 | 0.267 | 1 | 0.263 | 1 | 0.141 | 1 | 0.211 | 1
Ho165(MR) |  | 7440-60-0 | ug/L | 0.002 | 0.02 | 0.34 | 1 | 0.342 | 1 | 0.045 | 1 | 0.046 | 1 | 0.026 | 1 | 0.026 | 1 | 0.046 | 1 | 0.046 | 1 | 0.028 | 1 | 0.048 | 1 | 0.055 | 1 | 0.054 | 1 | 0.033 | 1 | 0.046 | 1
Er166(HR) |  | 7440-52-0 | ug/L | 0.002 | 0.02 | 0.922 | 1 | 0.931 | 1 | 0.118 | 1 | 0.12 | 1 | 0.07 | 1 | 0.074 | 1 | 0.128 | 1 | 0.134 | 1 | 0.09 | 1 | 0.141 | 1 | 0.156 | 1 | 0.148 | 1 | 0.094 | 1 | 0.123 | 1
Tm169(HR) |  | 7440-30-4 | ug/L | 0.002 | 0.02 | 0.13 | 1 | 0.136 | 1 | 0.015 (BQL) | 1 | 0.014 (BQL) | 1 | 0.009 (BQL) | 1 | 0.009 (BQL) | 1 | 0.018 (BQL) | 1 | 0.02 | 1 | 0.013 | 1 | 0.022 | 1 | 0.022 | 1 | 0.022 | 1 | 0.012 (BQL) | 1 | 0.016 (BQL) | 1
Yb172(MR) |  | 7440-64-4 | ug/L | 0.002 | 0.02 | 0.725 | 1 | 0.763 | 1 | 0.08 | 1 | 0.077 | 1 | 0.042 | 1 | 0.046 | 1 | 0.128 | 1 | 0.127 | 1 | 0.093 | 1 | 0.137 | 1 | 0.122 | 1 | 0.117 | 1 | 0.067 | 1 | 0.086 | 1
Lu175(HR) |  | 7439-94-3 | ug/L | 0.002 | 0.02 | 0.111 | 1 | 0.115 | 1 | 0.012 (BQL) | 1 | 0.013 (BQL) | 1 | 0.007 (BQL) | 1 | 0.007 (BQL) | 1 | 0.022 | 1 | 0.021 | 1 | 0.018 | 1 | 0.025 | 1 | 0.02 | 1 | 0.018 (BQL) | 1 | 0.012 (BQL) | 1 | 0.015 (BQL) | 1
Comments: All ICV, CCVs, Method Blanks, Low Level QL Checks, Sample Duplicates, Serial
Dilutions, Matrix Spikes, and Interference Checks were within specified ranges.
See Comments on QA Data tab.
Notes: If the parameter was detected above the quantitation limit (QL), the numeric result is reported. BQL denotes that the parameter was not detected at or above the quantitation limit; (BQL) denotes that the parameter was detected above the method detection limit (MDL) but below QL and the estimated numeric result is reported. "<QL value" denotes that the parameter was not detected at all. All of the results are corrected with dilution factors (DF), if applicable. Note that the applicable MDL is given only for the case where DF=1.

## QA Data
EPA Metals Laboratory
Results Report
Laboratory: | Metals |  | Report Date: | 2024-07-15
Technical Directive: | EPA MA140 |  |  |  |  | Quality Control Data Summary (1) | [26 empty] | Quality Control Data Summary (1) | [20 empty] | Quality Control Data Summary (1)
Analysts: | LMC |  |  | QC Sample ID |  | ICV |  |  | CCV-1 |  |  | CCV-2 |  |  | CCV-3 |  |  | CCV-4 |  |  | ICB |  | CCB-1 |  | CCB-2 |  | CCB-3 |  | CCB-4 |  | ICS1 |  |  | ICS2 |  |  |  | ICS3 |  |  |  | Serial Dilution 1 |  |  | Matrix Spike 1 |  |  |  | Matrix Spike 2 |  |  |  | Lab Dup 1 |  |  | Lab Dup 2 |  |  | LLQLS
|  |  |  | Additional ID |  | ICV |  |  | CCV-1 |  |  | CCV-2 |  |  | CCV-3 |  |  | CCV-4 |  |  | ICB |  | CCB-1 |  | CCB-2 |  | CCB-3 |  | CCB-4 |  | ICS1 |  |  | ICS2 |  |  |  | ICS3 |  |  |  | Serial Dilution |  |  | Matrix Spike |  |  |  | 7439-94-3 |  |  |  | Lab Dup 1 |  |  | Lab Dup 2 |  |  | LLQL 1
Method: | EPA 200.8 |  |  | Date Prepared |  | 2024-04-09 |  |  | 2024-04-09 |  |  | 2024-04-09 |  |  | 2024-04-09 |  |  | 2024-04-09 |  |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | NA |  | 2024-04-09 |  |  | 2024-04-09 |  |  |  | 2024-04-09 |  |  |  | 2024-05-10 |  |  | 2024-05-10 |  |  |  | 2024-05-10 |  |  |  | 2024-06-12 |  |  | 2024-06-12 |  |  | 2024-04-09
|  |  |  | Date Analyzed |  | 2024-06-12 |  |  | 2024-06-12 |  |  | 2024-06-12 |  |  | 2024-06-12 |  |  | NA |  |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | 2024-06-12 |  | NA |  | 2024-06-12 |  |  | 2024-06-12 |  |  |  | 2024-06-12 |  |  |  | 2024-06-12 |  |  | 2024-06-12 |  |  |  | 2024-06-12 |  |  |  | 2024-06-12 |  |  | 2024-06-12 |  |  | 2024-06-12
Analytes |  |  | Unit | MDL | QL | Data | True Values | % Rec. | Data | True Values | % Rec. | Data | True Values | % Rec. | Data | True Values | % Rec. | Data | True Values | % Rec. | Data | Pass/ Fail | Data | Pass/ Fail | Data | Pass/ Fail | Data | Pass/ Fail | Data | Pass/ Fail | Data | True Values | % Rec. | Data | True Values | % Rec. | Pass/Fail | Data | True Values | % Rec. | Pass/Fail | Initial Result | Dilution Result | RPD | Data | Spike | True Values | % Rec. | Data | Spike | True Values | % Rec. | Data | Dup Data | Pass/Fail | Data | Dup Data | Pass/Fail | Data | True Values | % Rec.
Names |  | Codes
Y89(MR) | Y89(MR) | 7440-65-5 | ug/L | 0.002 | 0.02 | 1.939 | 2 | 97 | 1.991 | 2 | 99.6 | 1.998 | 2 | 99.9 | 1.993 | 2 | 99.7 | NA | 2 | NA | 0.001 | Pass | -0.001 | Pass | 0.001 | Pass | 0.002 | Pass | NA | NA | -0.001 | ----- | ----- | 0.006 | ----- | ----- | ----- | 0.004 | ----- | ----- | ----- | 8.965 | 1.738 | 3.1 | 8.965 | 9.774 | 1 | 80.9 | 1.721 | 2.746 | 1 | 102.5 | 8.965 | 8.97 | Pass | 1.721 | 1.683 | Pass | 0.023 | 0.02 | 115
La139(HR) | La139(HR) | 7439-91-0 | ug/L | 0.002 | 0.02 | 1.933 | 2 | 96.7 | 2.028 | 2 | 101.4 | 2.044 | 2 | 102.2 | 2.047 | 2 | 102.4 | NA | 2 | NA | 0 | Pass | -0.001 | Pass | 0 | Pass | -0.001 | Pass | NA | NA | 0.005 | ----- | ----- | 0.005 | ----- | ----- | ----- | 0.003 | ----- | ----- | ----- | 6.36 | 1.252 | 1.6 | 6.36 | 7.35 | 1 | 99 | 1.322 | 2.334 | 1 | 101.2 | 6.36 | 6.531 | Pass | 1.377 | 1.337 | Pass | 0.019 | 0.02 | 95
Ce140(HR) | Ce140(HR) | 7440-45-1 | ug/L | 0.002 | 0.02 | 1.908 | 2 | 95.4 | 2.023 | 2 | 101.2 | 2.023 | 2 | 101.2 | 2.046 | 2 | 102.3 | NA | 2 | NA | -0.001 | Pass | 0 | Pass | 0.001 | Pass | 0 | Pass | NA | NA | 0.002 | ----- | ----- | 0.002 | ----- | ----- | ----- | 0.003 | ----- | ----- | ----- | 14.549 | 2.884 | 0.9 | 14.549 | 15.547 | 1 | 99.8 | 2.664 | 3.6 | 1 | 93.6 | 14.549 | 14.892 | Pass | 2.664 | 2.585 | Pass | 0.021 | 0.02 | 105
Pr141(HR) | Pr141(HR) | 7440-10-0 | ug/L | 0.002 | 0.02 | 1.898 | 2 | 94.9 | 2.032 | 2 | 101.6 | 2.013 | 2 | 100.7 | 2.021 | 2 | 101.1 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0 | ----- | ----- | 1.994 | 2 | 99.7 | Pass | 0 | ----- | ----- | ----- | 1.759 | 0.356 | 1.2 | 1.759 | 2.799 | 1 | 104 | 0.301 | 1.26 | 1 | 95.9 | 1.759 | 1.833 | Pass | 0.301 | 0.282 | Pass | 0.021 | 0.02 | 105
Nd146(MR) | Nd146(MR) | 7440-00-08 | ug/L | 0.002 | 0.02 | 1.935 | 2 | 96.8 | 1.984 | 2 | 99.2 | 1.984 | 2 | 99.2 | 2.041 | 2 | 102.1 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | -0.001 | Pass | NA | NA | 0.005 | ----- | ----- | 2.031 | 2 | 101.6 | Pass | 0 | ----- | ----- | ----- | 7.597 | 1.506 | 0.9 | 7.597 | 8.616 | 1 | 101.9 | 1.139 | 2.152 | 1 | 101.3 | 7.597 | 7.476 | Pass | 1.139 | 1.135 | Pass | 0.017 | 0.02 | 85
Sm147(MR) | Sm147(MR) | 7440-19-9 | ug/L | 0.002 | 0.02 | 1.94 | 2 | 97 | 1.996 | 2 | 99.8 | 2.045 | 2 | 102.3 | 2.062 | 2 | 103.1 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0.004 | ----- | ----- | 0 | ----- | ----- | ----- | 2.125 | 2 | 106.3 | Pass | 1.684 | 0.348 | 3.3 | 1.684 | 2.718 | 1 | 103.4 | 0.225 | 1.241 | 1 | 101.6 | 1.684 | 1.695 | Pass | 0.225 | 0.216 | Pass | 0.021 | 0.02 | 105
Eu151(MR) | Eu151(MR) | 7440-53-1 | ug/L | 0.002 | 0.02 | 1.932 | 2 | 96.6 | 1.992 | 2 | 99.6 | 2.011 | 2 | 100.6 | 2.034 | 2 | 101.7 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0.016 | ----- | ----- | 2.039 | 2 | 102 | Pass | 0 | ----- | ----- | ----- | 0.344 | 0.07 | 1.7 | 0.344 | 1.364 | 1 | 102 | 0.041 | 1.034 | 1 | 99.3 | 0.344 | 0.346 | Pass | 0.041 | 0.042 | NC | 0.019 | 0.02 | 95
Gd157(HR) | Gd157(HR) | 7440-54-2 | ug/L | 0.002 | 0.02 | 1.839 | 2 | 92 | 2.058 | 2 | 102.9 | 2.047 | 2 | 102.4 | 2.035 | 2 | 100.7 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0 | ----- | ----- | -0.016 | ----- | ----- | Pass | 2.073 | 2 | 103.7 | Pass | 1.885 | 0.366 | 3 | 1.885 | 2.944 | 1 | 105.9 | 0.312 | 1.279 | 1 | 96.7 | 1.885 | 1.74 | Pass | 0.284 | 0.294 | Pass | 0.021 | 0.02 | 105
Tb159(HR) | Tb159(HR) | 7440-27-9 | ug/L | 0.002 | 0.02 | 1.878 | 2 | 93.9 | 2.005 | 2 | 100.3 | 2.05 | 2 | 102.5 | 2.049 | 2 | 102.5 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0 | ----- | ----- | 0 | ----- | ----- | Pass | 2.062 | 2 | 103.1 | Pass | 0.279 | 0.056 | 0.4 | 0.279 | 1.309 | 1 | 103 | 0.045 | 1.046 | 1 | 100.1 | 0.279 | 0.263 | Pass | 0.045 | 0.046 | NC | 0.018 | 0.02 | 90
Dy163(MR) | Dy163(MR) | 7429-91-6 | ug/L | 0.002 | 0.02 | 1.938 | 2 | 96.9 | 1.978 | 2 | 98.9 | 2.014 | 2 | 100.7 | 2.021 | 2 | 101.1 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0 | ----- | ----- | 0 | ----- | ----- | ----- | -0.002 | ----- | ----- | Pass | 1.728 | 0.353 | 2.1 | 1.728 | 2.737 | 1 | 100.9 | 0.267 | 1.294 | 1 | 102.7 | 1.728 | 1.739 | Pass | 0.267 | 0.271 | Pass | 0.019 | 0.02 | 95
Ho165(MR) | Ho165(MR) | 7440-60-0 | ug/L | 0.002 | 0.02 | 1.937 | 2 | 96.9 | 1.971 | 2 | 98.6 | 1.992 | 2 | 99.6 | 2.014 | 2 | 100.7 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0 | ----- | ----- | 0 | ----- | ----- | ----- | 0 | ----- | ----- | Pass | 0.34 | 0.069 | 1.5 | 0.34 | 1.356 | 1 | 101.6 | 0.054 | 1.069 | 1 | 101.5 | 0.34 | 0.34 | Pass | 0.055 | 0.057 | NC | 0.019 | 0.02 | 95
Er166(HR) | Er166(HR) | 7440-52-0 | ug/L | 0.002 | 0.02 | 1.893 | 2 | 94.7 | 2.005 | 2 | 100.3 | 2.015 | 2 | 100.8 | 2.009 | 2 | 100.5 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0 | ----- | ----- | 0 | ----- | ----- | Pass | 0 | ----- | ----- | Pass | 0.922 | 0.189 | 2.5 | 0.922 | 1.936 | 1 | 101.4 | 0.156 | 1.131 | 1 | 97.5 | 0.922 | 0.938 | Pass | 0.156 | 0.153 | Pass | 0.019 | 0.02 | 95
Tm169(HR) | Tm169(HR) | 7440-30-4 | ug/L | 0.002 | 0.02 | 1.902 | 2 | 95.1 | 2.009 | 2 | 100.5 | 2.078 | 2 | 103.9 | 2.07 | 2 | 103.5 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0 | ----- | ----- | 0 | ----- | ----- | Pass | 0 | ----- | ----- | ----- | 0.13 | 0.026 | 0 | 0.13 | 1.163 | 1 | 103.3 | 0.022 | 0.999 | 1 | 97.7 | 0.13 | 0.134 | Pass | 0.022 | 0.02 | NC | 0.02 | 0.02 | 100
Yb172(MR) | Yb172(MR) | 7440-64-4 | ug/L | 0.002 | 0.02 | 1.897 | 2 | 94.9 | 1.955 | 2 | 97.8 | 2.004 | 2 | 100.2 | 1.974 | 2 | 98.7 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0 | ----- | ----- | -0.006 | ----- | ----- | ----- | -0.027 | ----- | ----- | ----- | 0.725 | 0.147 | 1.4 | 0.725 | 1.691 | 1 | 96.6 | 0.122 | 1.105 | 1 | 98.3 | 0.725 | 0.747 | Pass | 0.122 | 0.114 | Pass | 0.018 | 0.02 | 90
Lu175(HR) | Lu175(HR) | 7439-94-3 | ug/L | 0.002 | 0.02 | 1.93 | 2 | 96.5 | 2.034 | 2 | 101.7 | 2.077 | 2 | 103.9 | 2.03 | 2 | 101.5 | NA | 2 | NA | 0 | Pass | 0 | Pass | 0 | Pass | 0 | Pass | NA | NA | 0 | ----- | ----- | 0 | ----- | ----- | ----- | -0.013 | ----- | ----- | Pass | 0.111 | 0.022 | 0.9 | 0.111 | 1.109 | 1 | 99.8 | 0.02 | 1.001 | 1 | 98.1 | 0.111 | 0.121 | Pass | 0.02 | 0.019 | NC | 0.019 | 0.02 | 95
Comments: All QC checks for the ICV, CCVs, Method Blanks, Low Level QL Check, Sample
Duplicates, Serial Dilution, Matrix Spikes, and Interference Checks met the criteria established in
K-GCRD-1158. NA = not analyzed. NC = not calculated.

## summary
| NASC | PAAS | BM-SW-1 | BM-SW-1 FD | HC-SW-1 | HC-SW-1 FD | Cam-SW-1 | Cam-SW-1 FD | SR-SW-2 | SR-SW-2 FD | SR-SW-5 | SR-SW-9 | EM-SW-1 | EM-SW-1 FD | EM-SW-5 | EM-SW-9 | BM-SW-1 | BM-SW-1 FD | HC-SW-1 | HC-SW-1 FD | Cam-SW-1 | Cam-SW-1 FD | SR-SW-2 | SR-SW-2 FD | SR-SW-5 | SR-SW-9 | EM-SW-1 | EM-SW-1 FD | EM-SW-5 | EM-SW-9 | 10 | 11 | 12 | 13 | 14 | 15
La | 31.1 | 38.2 | 6.36 | 6.516 | 0.944 | 0.932 | 0.628 | 0.651 | 0.904 | 0.893 | 0.459 | 0.999 | 1.377 | 1.322 | 0.957 | 1.251 | 0.0002045 | 0.00017058 | 2.4712e-05 | 2.4398e-05 | 1.644e-05 | 1.7042e-05 | 2.3665e-05 | 2.3377e-05 | 1.2016e-05 | 2.6152e-05 | 3.6047e-05 | 3.4607e-05 | 2.5052e-05 | 3.2749e-05 | 0 | 0 | 0 | 0 | 0 | 0
Ce | 66.7 | 79.6 | 14.549 | 14.834 | 1.858 | 1.83 | 1.07 | 1.105 | 1.734 | 1.732 | 0.913 | 1.947 | 2.664 | 2.554 | 1.637 | 2.249 | 0.00018278 | 0.00018636 | 2.3342e-05 | 2.299e-05 | 1.3442e-05 | 1.3882e-05 | 2.1784e-05 | 2.1759e-05 | 1.147e-05 | 2.446e-05 | 3.3467e-05 | 3.2085e-05 | 2.0565e-05 | 2.8254e-05 | 0 | 0 | 0 | 0 | 0 | 0
Pr | 7.7 | 8.83 | 1.759 | 1.819 | 0.211 | 0.21 | 0.115 | 0.122 | 0.196 | 0.198 | 0.108 | 0.225 | 0.301 | 0.287 | 0.169 | 0.23 | 0.00019921 | 0.000206 | 2.3896e-05 | 2.3783e-05 | 1.3024e-05 | 1.3817e-05 | 2.2197e-05 | 2.2424e-05 | 1.2231e-05 | 2.5481e-05 | 3.4088e-05 | 3.2503e-05 | 1.9139e-05 | 2.6048e-05 | 0 | 0 | 0 | 0 | 0 | 0
Nd | 27.4 | 33.9 | 7.597 | 7.63 | 0.835 | 0.851 | 0.453 | 0.478 | 0.799 | 0.828 | 0.469 | 0.891 | 1.139 | 1.114 | 0.626 | 0.877 | 0.0002241 | 0.00022507 | 2.4631e-05 | 2.5103e-05 | 1.3363e-05 | 1.41e-05 | 2.3569e-05 | 2.4425e-05 | 1.3835e-05 | 2.6283e-05 | 3.3599e-05 | 3.2861e-05 | 1.8466e-05 | 2.587e-05 | 0 | 0 | 0 | 0 | 0 | 0
Sm | 5.59 | 5.55 | 1.684 | 1.72 | 0.174 | 0.179 | 0.095 | 0.091 | 0.161 | 0.155 | 0.097 | 0.164 | 0.225 | 0.217 | 0.113 | 0.153 | 0.00030342 | 0.00030991 | 3.1351e-05 | 3.2252e-05 | 1.7117e-05 | 1.6396e-05 | 2.9009e-05 | 2.7928e-05 | 1.7477e-05 | 2.955e-05 | 4.0541e-05 | 3.9099e-05 | 2.036e-05 | 2.7568e-05 | 0 | 0 | 0 | 0 | 0 | 0
Eu | 1.18 | 1.08 | 0.344 | 0.348 | 0.034 | 0.035 | 0.015 | 0.017 | 0.032 | 0.035 | 0.021 | 0.036 | 0.041 | 0.043 | 0.019 | 0.032 | 0.00031852 | 0.00032222 | 3.1481e-05 | 3.2407e-05 | 1.3889e-05 | 1.5741e-05 | 2.963e-05 | 3.2407e-05 | 1.9444e-05 | 3.3333e-05 | 3.7963e-05 | 3.9815e-05 | 1.7593e-05 | 2.963e-05 | 0 | 0 | 0 | 0 | 0 | 0
Gd | 4.9 | 4.66 | 1.885 | 1.941 | 0.25 | 0.24 | 0.159 | 0.114 | 0.237 | 0.217 | 0.133 | 0.225 | 0.284 | 0.312 | 0.148 | 0.212 | 0.00040451 | 0.00041652 | 5.3648e-05 | 5.1502e-05 | 3.412e-05 | 2.4464e-05 | 5.0858e-05 | 4.6567e-05 | 2.8541e-05 | 4.8283e-05 | 6.0944e-05 | 6.6953e-05 | 3.176e-05 | 4.5494e-05 | 0 | 0 | 0 | 0 | 0 | 0
Tb | 0.85 | 0.774 | 0.279 | 0.283 | 0.036 | 0.036 | 0.019 | 0.016 | 0.033 | 0.035 | 0.02 | 0.035 | 0.045 | 0.041 | 0.024 | 0.03 | 0.00036047 | 0.00036563 | 4.6512e-05 | 4.6512e-05 | 2.4548e-05 | 2.0672e-05 | 4.2636e-05 | 4.522e-05 | 2.584e-05 | 4.522e-05 | 5.814e-05 | 5.2972e-05 | 3.1008e-05 | 3.876e-05 | 0 | 0 | 0 | 0 | 0 | 0
Dy | 4.17 | 4.68 | 1.728 | 1.758 | 0.22 | 0.216 | 0.12 | 0.12 | 0.206 | 0.215 | 0.125 | 0.227 | 0.267 | 0.263 | 0.141 | 0.211 | 0.00036923 | 0.00037564 | 4.7009e-05 | 4.6154e-05 | 2.5641e-05 | 2.5641e-05 | 4.4017e-05 | 4.594e-05 | 2.6709e-05 | 4.8504e-05 | 5.7051e-05 | 5.6197e-05 | 3.0128e-05 | 4.5085e-05 | 0 | 0 | 0 | 0 | 0 | 0
Y | 21 | 27.3 | 8.965 | 9.055 | 1.555 | 1.581 | 1.001 | 1.007 | 1.391 | 1.401 | 0.889 | 1.508 | 1.721 | 1.684 | 1.121 | 1.499 | 0.0004269 | 0.00043119 | 7.4048e-05 | 7.5286e-05 | 4.7667e-05 | 4.7952e-05 | 6.6238e-05 | 6.6714e-05 | 4.2333e-05 | 7.181e-05 | 8.1952e-05 | 8.019e-05 | 5.3381e-05 | 7.1381e-05 | 0 | 0 | 0 | 0 | 0 | 0
Ho | 1.02 | 0.991 | 0.34 | 0.342 | 0.045 | 0.046 | 0.026 | 0.026 | 0.046 | 0.046 | 0.028 | 0.048 | 0.055 | 0.054 | 0.033 | 0.046 | 0.00034309 | 0.00034511 | 4.5409e-05 | 4.6418e-05 | 2.6236e-05 | 2.6236e-05 | 4.6418e-05 | 4.6418e-05 | 2.8254e-05 | 4.8436e-05 | 5.5499e-05 | 5.449e-05 | 3.33e-05 | 4.6418e-05 | 0 | 0 | 0 | 0 | 0 | 0
Er | 2.84 | 2.85 | 0.922 | 0.931 | 0.118 | 0.12 | 0.07 | 0.074 | 0.128 | 0.134 | 0.09 | 0.141 | 0.156 | 0.148 | 0.094 | 0.123 | 0.00032351 | 0.00032667 | 4.1404e-05 | 4.2105e-05 | 2.4561e-05 | 2.5965e-05 | 4.4912e-05 | 4.7018e-05 | 3.1579e-05 | 4.9474e-05 | 5.4737e-05 | 5.193e-05 | 3.2982e-05 | 4.3158e-05 | 0 | 0 | 0 | 0 | 0 | 0
Tm | 0.48 | 0.405 | 0.13 | 0.136 | 0.015 | 0.014 | 0.009 | 0.009 | 0.018 | 0.02 | 0.013 | 0.022 | 0.022 | 0.022 | 0.012 | 0.016 | 0.00032099 | 0.0003358 | 3.7037e-05 | 3.4568e-05 | 2.2222e-05 | 2.2222e-05 | 4.4444e-05 | 4.9383e-05 | 3.2099e-05 | 5.4321e-05 | 5.4321e-05 | 5.4321e-05 | 2.963e-05 | 3.9506e-05 | 0 | 0 | 0 | 0 | 0 | 0
Yb | 3.06 | 2.82 | 0.725 | 0.763 | 0.08 | 0.077 | 0.042 | 0.046 | 0.128 | 0.127 | 0.093 | 0.137 | 0.122 | 0.117 | 0.067 | 0.086 | 0.00025709 | 0.00027057 | 2.8369e-05 | 2.7305e-05 | 1.4894e-05 | 1.6312e-05 | 4.539e-05 | 4.5035e-05 | 3.2979e-05 | 4.8582e-05 | 4.3262e-05 | 4.1489e-05 | 2.3759e-05 | 3.0496e-05 | 0 | 0 | 0 | 0 | 0 | 0
Lu | 0.46 | 0.433 | 0.111 | 0.115 | 0.012 | 0.013 | 0.007 | 0.007 | 0.022 | 0.021 | 0.018 | 0.025 | 0.02 | 0.018 | 0.012 | 0.015 | 0.00025635 | 0.00026559 | 2.7714e-05 | 3.0023e-05 | 1.6166e-05 | 1.6166e-05 | 5.0808e-05 | 4.8499e-05 | 4.157e-05 | 5.7737e-05 | 4.6189e-05 | 4.157e-05 | 2.7714e-05 | 3.4642e-05 | 0 | 0 | 0 | 0 | 0 | 0
Sum, ug/L |  | 212.073 | 47.378 | 48.191 | 6.387 | 6.38 | 3.829 | 3.883 | 6.035 | 6.057 | 3.476 | 6.63 | 8.439 | 8.196 | 5.173 | 7.03
Ce* linear |  |  | 0.90549 | 0.98974 | 0.96041 | 0.95433 | 0.91246 | 0.89972 | 0.94998 | 0.95015 | 0.94609 | 0.94745 | 0.95436 | 0.9562 | 0.93073 | 0.96107
Ce* geometric |  |  | 0.90556 | 0.99415 | 0.96054 | 0.9544 | 0.91866 | 0.90467 | 0.95046 | 0.95036 | 0.94613 | 0.94753 | 0.95474 | 0.95667 | 0.93918 | 0.96738
H/L |  | 1 | 1.5543 | 1.5672 | 1.4904 | 1.4992 | 1.4798 | 1.4534 | 2.0345 | 2.0556 | 2.6878 | 2.0204 | 1.4811 | 1.4769 | 1.4168 | 1.376
M/L+H |  | 1 | 1.4239 | 1.4157 | 1.3969 | 1.3887 | 1.3223 | 1.1406 | 1.1338 | 1.1275 | 1.0333 | 1.0602 | 1.1969 | 1.2475 | 1.0409 | 1.1124
Field Dups
Blanks
Normalized
Extra Lab Dups

## locations
Site | Samples ID | Location ID | Latitude | Longitude
Big Mountain Mine | BM-SW-1 | BM-SW-A | 40.7725 | -76.5385
Big Mountain Mine | BM-SW-1-FD | BM-SW-A | 40.7725 | -76.5385
Henry Clay Stirling Mine Pump | HC-SW-1 | HC-SW-A | 40.7774 | -76.5681
Henry Clay Stirling Mine Pump | HC-SW-1-FD | HC-SW-A | 40.7774 | -76.5681
Cameron/Glen Burn | Cam-SW-1 | Cam-SW-A | 40.7938 | -76.565
Cameron/Glen Burn | Cam-SW-1-FD | Cam-SW-A | 40.7938 | -76.565
Scott Ridge Mine Tunnel - Downstream Weir | SR-SW-1 | SR-SW-A | 40.7916 | -76.4896
Scott Ridge Mine Tunnel - Downstream Weir | SR-SW-1-FD | SR-SW-A | 40.7916 | -76.4896
Scott Ridge Mine Tunnel - Left Branch Scott AMD | SR-SW-5 | SR-SW-B | 40.7921 | -76.4895
Scott Ridge Mine Tunnel - Right Branch Scott AMD | SR-SW-9 | SR-SW-C | 40.7922 | -76.4892
Excelsior Mine Strip Pit Overflow - Weir | EM-SW-1 | EM-SW-A | 40.7739 | -76.4936
Excelsior Mine Strip Pit Overflow - Weir | EM-SW-1-FD | EM-SW-A | 40.7739 | -76.4936
Excelsior Mine Strip Pit Overflow - East End Pond - Blue Water | EM-SW-5 | EM-SW-B | 40.7733 | -76.4902
Excelsior Mine Strip Pit Overflow - Center Pond | EM-SW-9 | EM-SW-C | 40.7732 | -76.4912