## Summary
Summary
Test Name: | Rainbow Fall 2024 PFAS Assessment Data
Purpose: | OTM-50 Quantification of Rainbow Fall 2024
Reporter: | Bill Preston
Sample Medium: | Silonite treated 6-liter canisters
Analysis Technique: | Thermal Desorption - Gas Chromatography/Time-of-Flight Mass Spectrometry (TD-GC/TOF)
Instrument: | Markes CIA Advantage/Kori-xr Water Condenser/Unity-xr Thermal Desorber
Agilent 7890B Gas Chromatograph
Agilent 5977 Mass Selective Detector (Extractor Source-6 mm)
Markes Acquisition Method: | PFAS TD Plus istd_1minpurge_10 min bake/PFAS TD Plus istd_9minpurge_25min trap bake
Agilent Acquisition Method: | Agilent_112524_PFASvolatilesAlkenesMixRT_d.M
QAPP Title: | Incineration of PFAS using a Pilot-scale Tunnel Furnace
QAPP ID: | J-AMCD-0032499-QP-1-1
Instrument Operator: | Jose Gomez/Bill Preston
Test Date: | NA
Analysis Dates:
Review Status: | Initial: | 11/25/2024-12/03/2024 | Jose Gomez
| 2nd Electronic GCMS data Review: | 11/25/2024-12/03/2024 | Bill Preston
QA:
Revision: | 0

## Data Analysis
Data Reduction Methods used for PFAS Volatiles Analysis
6 level 1 min_CF4_Nov26th, 2024_CC_RF.M
6 level 9 min Nov26th, 2024_CC_RF.m
QA
Sample/cal check standard internal standard response were checked to be within 60% and 140%. A singular, continuing calibration (CC) point was used for each method (CC levels were analyzed November 25th) for sample internal standard assessment using 1,4-difluorobenzene-see worksheets Nov 26th-29th:Dec 3rd for each individual sample.
Sample/cal retention time adherence check to be within 5 seconds of the average RT of the CC levels ( CC levels were analyzed November 25th) for each individual sample.
Sample duplicate precision assessed among common analytes greater than 3x the MDL were within 25 % average Relative Percent Deviation (RPD)-see dupicate precision sheets.
Calibration curve check and daily cal check validations were assessed when compared to the November 25th CC level for each 1 and 9 minute purge method (within 20% per analyte).
Daily laboratory blanks were tabulated. All values greater than 0.015 ppbv were flagged.
Post cal blanks all had values well below 50% of the practical quantification limit (PQL-which is the lowest calibration level).
Interferences
NA

## Method Parameters
Analytical Method for Air Samples Collected using 1, 1.4 and 6-liter Canisters
Instrumentation
Markes International, Inc. | CIA Advantage
Kori-xr Water Condenser
Unity-xr Thermal Desorber
Agilent Technologies | 7890B Gas Chromatograph
Agilent Technologies | 5977 Extractor Source 6 mm
Agilent Acquisition Method: | Agilent_112524_PFASvolatilesAlkenesMixRT_d.M
Markes Acquisition Method (1 minute trap bake) | PFAS TD Plus istd_1minpurge_10 min bake
Markes Acquisition Method (9 minute trap bake) | PFAS TD Plus istd_9minpurge_25min trap bake
Markes Thermal Desorption Method Parameters
MFC Sampling
Split On: | √
Split Flow: | 10 | mL/min
Flow Path Temp.: | 120 | ⁰C
GC Cycle Time: | 0 | min
Minimum Carrier Pressure: | 5 | psi
PreSampling
Sample Purge Time: | 2 | min
Purge Flow: | 50 | mL/min
Add Internal Standard: | √
Loop Fill Time: | 1 | min
Loop Equilibration Time: | 0 | min
Loop Inject Time: | 1 | min
Flow: | 50 | mL/min
Kori Trap Low Temp.: | -30 | ⁰C
Kori Trap High Temp.: | 300 | ⁰C
Sampling
Sample by Volume: | √
Sample Quantity: | 200 | mL
Flow: | 25 | mL/min
Use dedicated purge channel: | o
Post Sampling Purge Time: | 5 | min
Purge Flow: | 50 | mL/min
Trap Settings
Trap type | Markes ODS proprietary
Trap Purge Time: | 1 min CF4/9 minute rest | min | (CO2 ambient for study)
Purge Flow: | 50 | mL/min
Trap Low Temp.: | -30 | ⁰C
Trap High Temp.: | 280 | ⁰C
Trap Heating Rate: | 40 | ⁰C/second
Trap Hold Time: | 10 and 25 | min
Split On: | √
Split Flow: | 3 | mL/min
Agilent Gas Chromatograph Parameters
Column: | Agilent GS-GasPro (cat. # 113-4362)
Dimensions: | 60 m x 0.320 mm
Film Thickness | 0 | μm
Inlet Parameters
Mode: | Splitless
Temp.: | 200 | ⁰C
Pressure: | 24.5 | psi
Column Flow | 1.8 | mL/min
Purge Flow to Split Vent Time: | 999.99 | min
Purge Flow to Vent: | 15 | mL/min
Septum Purge Flow: | 3 | mL/min
Total Flow: | 19.8 | mL/min
Oven Parameters
Initial Temp.: | 50 | ⁰C
Initial Time: | 1 | min
Ramp Rate: | 5 | ⁰C/min
Temp.: | 130 | ⁰C
Hold Time: | 0 | min
Ramp Rate: | 10 | ⁰C/min
Final Temp.: | 240 | ⁰C
Final Time: | 30 | min
Extractor Source
Measurement: | SIM/Scan
Use External Trigger: | √
Electron Multipler Voltage: | set to tuning voltage +100V (typically 1100-1300V) | V
SIMScan paramters
Mass Range Low: | 29 | m/z
Mass Range High: | 450 | m/z
Threshold | 150
Filament delay | 6.6 | min
Gain | 1.5
Temperatures/Voltages
Transfer Line Temp.: | 250 | ⁰C
Quad temp: | 150 | ⁰C
Ion Source Temp.: | 230 | ⁰C
Ionization: | -70 | V

## Rainbow ppbv composite
PFAS VOC Rainbow Fall 2024
Sample Name |  | N2 Blank | Combustion Blank | Run 1 Port 4 PFBA inj | Run 2 Port 6 PFBA inj | Run 3 Port 8 PFBA inj | Run 4 Port 10 PFBA inj | Run 5 Port 4 PFOA inj | Run 6 Port 6 PFOA inj | Run 7 Port 8 PFOA inj | Run 8 Port 10 PFOA inj | Post Test N2 Blank* | Post Test N2 Blank
Silonite Canister ID |  | 36834 | 40882 | 40880 | 40864 | 40860 | 40879 | 42891 | 40877 | 40869 | 42896 | 1750* | 40865
Description | CAS # | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv
Tetrafluoromethane | 75-73-0 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
Hexafluoroethane (FC-116) | 76-16-4 | 0.028438 | BDL | 1.3205 | 1.7711 | 1.3091 | 0.66758 | 0.99803 | 2.253 | 2.0096 | 1.4723 | BDL | BDL
Chlorotrifluoromethane (CFC-13) | 75-72-9 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
Trifluoromethane (Fluoroform) | 75-46-7 | BDL | BDL | BDL | 4.3219 | 64.3736 | 166.3949 | BDL | 1.683 | 33.5067 | 115.1308 | BDL | BDL
Octafluoropropane | 76-19-7 | 0.15641 | BDL | BDL | BDL | 0.06042 | 0.061219 | BDL | 0.072 | 0.29367 | 0.4539 | 0.034843 | 0.057383
Difluoromethane (HFC-32) | 75-10-5 | BDL | BDL | BDL | BDL | BDL | 0.19532 | BDL | BDL | BDL | BDL | BDL | BDL
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | 354-33-6 | BDL | 0.066794 | 0.058045 | 0.078264 | 6.5225 | 82.3983 | 0.11025 | 0.057 | 0.16987 | 2.0746 | 0.056619 | BDL
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | 115-25-3 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 0.093108 | BDL | BDL
Fluoromethane (HFC-41) | 593-53-3 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
Tetrafluoroethylene | 116-14-3 | BDL | BDL | BDL | BDL | 0.55817 | 6.5854 | BDL | BDL | BDL | 0.066921 | BDL | BDL
Hexafluoropropene | 116-15-4 | BDL | BDL | BDL | BDL | 0.031649 | 1.7929 | BDL | BDL | BDL | 0.037825 | BDL | BDL
1,1,1-trifluoroethane | 420-46-2 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
Hexafluoropropene Oxide | 428-59-1 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
Chlorodifluoromethane (HCFC-22) | 75-45-6 | BDL | 0.1452 | 0.10738 | 0.15653 | 0.1381 | 0.14576 | 0.15087 | 0.135 | 0.14107 | 0.16003 | 0.2047 | BDL
1,1,1,2-Tetrafluoroethane | 811-97-2 | BDL | 0.18005 | BDL | BDL | 0.14674 | 0.11369 | BDL | 0.108 | 0.13244 | BDL | 0.1655 | BDL
Decafluorobutane (Perfluorobutane) | 355-25-9 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | 2252-84-8 | BDL | BDL | BDL | BDL | 0.48624 | 22.9631 | 0.11025 | 0.057 | BDL | 0.3026 | BDL | BDL
Octafluorocyclopentene (FC-C1418) | 559-40-0 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 0.19203 | BDL | BDL
Trichlorofluoromethane (CFC-11) | 75-69-4 | BDL | 0.11326 | 0.095773 | 0.10725 | 0.11796 | 0.11078 | 0.12766 | 0.111 | 0.11516 | 0.12511 | 0.18728 | BDL
Dodecafluoro-n-pentane | 678-26-2 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | 375-17-7 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
Tetradecafluorohexane | 355-42-0 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
1,1,1,2,2,3,3,4,4,5,5-Undecafluoropentane (1H-Perfluoropentane) | 375-61-1 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | 3330-15-2 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
Hexadecafluoroheptane | 335-57-9 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | 355-37-3 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
Octadecafluorooctane (Perfluorooctane) | 307-34-6 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | 375-83-7 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 0.084379 | BDL | BDL
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | 335-65-9 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | 3330-14-1 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL
BDL=Below Detection Limit
Values between MDL and lowest cal (PQL) are denoted in Italics (J values)
< ppbv CF4 PQL used for J italics | 10*dil factor | 27.9768 | 28.5702 | 28.5521 | 28.5169 | 28.3055 | 28.6798 | 28.5426 | 29.514 | 28.3244 | 28.6249 | 21.4239 | 28.2267
<Nominal ppbv remaining targets PQL used for J italics | 1*dil factor | 2.8438 | 2.9041 | 2.9022 | 2.8986 | 2.8772 | 2.9152 | 2.9013 | 3 | 2.8791 | 2.9096 | 2.1777 | 2.8691
| Aggregate targeted | 0.18484 | 0.50531 | 1.5817 | 6.435 | 73.7445 | 281.4289 | 1.497 | 4.476 | 36.3685 | 120.1936 | 0.64894 | 0.057383
|  | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL | 9 Minute Method ran at 200 mL
|  | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL | 1 Minute Method ran at 20 mL

## Rainbow UG per m3
PFAS VOC Rainbow 2024
Sample Name |  | N2 Blank | Combustion Blank | Run 1 Port 4 PFBA inj | Run 2 Port 6 PFBA inj | Run 3 Port 8 PFBA inj | Run 4 Port 10 PFBA inj | Run 5 Port 4 PFOA inj | Run 6 Port 6 PFOA inj | Run 7 Port 8 PFOA inj | Run 8 Port 10 PFOA inj | Post Test N2 Blank* | Post Test N2 Blank
Silonite Canister ID |  | 36834 | 40882 | 40880 | 40864 | 40860 | 40879 | 42891 | 40877 | 40869 | 42896 | 1750* | 40865 | Conversion Factors
Description | CAS # | ug/m3 | ug/m3 | ug/m3 | ug/m3 | ug/m3 | ug/m3 | ug/m3 | ug/m3 | ug/m3 | ug/m3 | ug/m3 | ug/m3 | Molecular wt. (ng/nmole) | ng/l=µg/m3 per 1 ppb
Tetrafluoromethane | 75-73-0 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 88.0043 | 3.6583
Hexafluoroethane (FC-116) | 76-16-4 | 0.16315 | BDL | 7.5759 | 10.1608 | 7.5104 | 3.83 | 5.7258 | 12.9256 | 11.5292 | 8.4465 | BDL | BDL | 138.01 | 5.7371
Chlorotrifluoromethane (CFC-13) | 75-72-9 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 104.46 | 4.3424
Trifluoromethane (Fluoroform) | 75-46-7 | BDL | BDL | BDL | 12.578 | 187.3468 | 484.2602 | BDL | 4.898 | 97.5148 | 335.0659 | BDL | BDL | 70.01 | 2.9103
Octafluoropropane | 76-19-7 | 1.2225 | BDL | BDL | BDL | 0.47224 | 0.47849 | BDL | 0.56275 | 2.2953 | 3.5477 | 0.27233 | 0.4485 | 188.02 | 7.816
Difluoromethane (HFC-32) | 75-10-5 | BDL | BDL | BDL | BDL | BDL | 0.4224 | BDL | BDL | BDL | BDL | BDL | BDL | 52.024 | 2.1626
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | 354-33-6 | BDL | 0.33325 | 0.2896 | 0.39047 | 32.5423 | 411.1026 | 0.55005 | 0.28439 | 0.8475 | 10.3504 | 0.28249 | BDL | 120.02 | 4.9892
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | 115-25-3 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 0.77421 | BDL | BDL | 200.03 | 8.3152
Fluoromethane (HFC-41) | 593-53-3 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 34.03 | 1.4146
Tetrafluoroethylene | 116-14-3 | BDL | BDL | BDL | BDL | 2.3208 | 27.3811 | BDL | BDL | BDL | 0.27825 | BDL | BDL | 100.02 | 4.1578
Hexafluoropropene | 116-15-4 | BDL | BDL | BDL | BDL | 0.19737 | 11.1808 | BDL | BDL | BDL | 0.23589 | BDL | BDL | 150.02 | 6.2363
1,1,1-trifluoroethane | 420-46-2 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 84.04 | 3.4935
Hexafluoropropene Oxide | 428-59-1 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 166.02 | 6.9014
Chlorodifluoromethane (HCFC-22) | 75-45-6 | BDL | 0.52194 | 0.38599 | 0.56264 | 0.49642 | 0.52394 | 0.54229 | 0.48526 | 0.5071 | 0.57523 | 0.73581 | BDL | 86.47 | 3.5945
1,1,1,2-Tetrafluoroethane | 811-97-2 | BDL | 0.76367 | BDL | BDL | 0.62236 | 0.48221 | BDL | 0.45807 | 0.56172 | BDL | 0.70196 | BDL | 102.03 | 4.2414
Decafluorobutane (Perfluorobutane) | 355-25-9 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 238.03 | 9.8949
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | 2252-84-8 | BDL | BDL | BDL | BDL | 3.4368 | 162.3058 | 0.77925 | 0.40288 | BDL | 2.1388 | BDL | BDL | 170.03 | 7.0681
Octafluorocyclopentene (FC-C1418) | 559-40-0 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 1.6927 | BDL | BDL | 212.042 | 8.8146
Trichlorofluoromethane (CFC-11) | 75-69-4 | BDL | 0.64676 | 0.54691 | 0.61245 | 0.67363 | 0.63259 | 0.72897 | 0.63386 | 0.65763 | 0.71446 | 1.0694 | BDL | 137.37 | 5.7105
Dodecafluoro-n-pentane | 678-26-2 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 288.036 | 11.9736
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | 375-17-7 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 220.04 | 9.147
Tetradecafluorohexane | 355-42-0 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 338.04 | 14.0523
1,1,1,2,2,3,3,4,4,5,5-Undecafluoropentane (1H-Perfluoropentane) | 375-61-1 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 270.04 | 11.2255
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | 3330-15-2 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 286.04 | 11.8906
Hexadecafluoroheptane | 335-57-9 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 388.051 | 16.1312
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | 355-37-3 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 320.05 | 13.3044
Octadecafluorooctane (Perfluorooctane) | 307-34-6 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 438.06 | 18.2101
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | 375-83-7 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 1.298 | BDL | BDL | 370.06 | 15.3833
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | 335-65-9 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 420.07 | 17.4623
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | 3330-14-1 | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | BDL | 452.07 | 18.7925
BDL=Below Detection Limit
| Aggregate targeted | 1.3856 | 2.2656 | 8.7984 | 24.3043 | 235.619 | 1102.6 | 8.3263 | 20.6508 | 113.9132 | 365.1181 | 3.062 | 0.4485

## Samples Daily Blanks
PFAS VOC Rainbow Fall 2024
Sample Name |  | 42879 | 42879 | 42879 | 42879 | 42879 | 42879
Silonite Canister ID |  | Batch Blank 200 mL | Batch Blank 200 mL | Batch Blank 20 mL | Batch Blank 200 mL | Batch Blank 20 mL | Batch Blank 200 mL
Date Analyzed |  | 2024-11-26 | 2024-11-27 | 2024-11-29 | 2024-11-29 | 2024-12-03 | 2024-12-03 | Conversion Factors
Description | CAS # | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | Molecular wt. (ng/nmole) | ng/l=µg/m3 per 1 ppb
Tetrafluoromethane | 75-73-0 | NA | NA | BDL | NA | BDL | NA | 88.0043 | 3.6583
Hexafluoroethane (FC-116) | 76-16-4 | BDL | BDL | NA | BDL | NA | BDL | 138.01 | 5.7371
Chlorotrifluoromethane (CFC-13) | 75-72-9 | BDL | BDL | NA | BDL | NA | BDL | 104.46 | 4.3424
Trifluoromethane (Fluoroform) | 75-46-7 | BDL | BDL | NA | BDL | NA | BDL | 70.01 | 2.9103
Octafluoropropane | 76-19-7 | BDL | BDL | NA | BDL | NA | BDL | 188.02 | 7.816
Difluoromethane (HFC-32) | 75-10-5 | BDL | BDL | NA | BDL | NA | BDL | 52.024 | 2.1626
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | 354-33-6 | 0.022 | 0.016 | NA | 0.014 | NA | BDL | 120.02 | 4.9892
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | 115-25-3 | BDL | 0.016 | NA | BDL | NA | BDL | 200.03 | 8.3152
Fluoromethane (HFC-41) | 593-53-3 | BDL | BDL | NA | BDL | NA | BDL | 34.03 | 1.4146
Tetrafluoroethylene | 116-14-3 | BDL | BDL | NA | BDL | NA | BDL | 100.02 | 4.1578
Hexafluoropropene | 116-15-4 | BDL | BDL | NA | BDL | NA | BDL | 150.02 | 6.2363
1,1,1-trifluoroethane | 420-46-2 | BDL | BDL | NA | BDL | NA | BDL | 84.04 | 3.4935
Hexafluoropropene Oxide | 428-59-1 | BDL | BDL | NA | BDL | NA | BDL | 166.02 | 6.9014
Chlorodifluoromethane (HCFC-22) | 75-45-6 | BDL | 0.029 | NA | BDL | NA | BDL | 86.47 | 3.5945
1,1,1,2-Tetrafluoroethane | 811-97-2 | BDL | BDL | NA | BDL | NA | BDL | 102.03 | 4.2414
Decafluorobutane (Perfluorobutane) | 355-25-9 | BDL | BDL | NA | BDL | NA | BDL | 238.03 | 9.8949
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | 2252-84-8 | BDL | 0.02 | NA | BDL | NA | BDL | 170.03 | 7.0681
Octafluorocyclopentene (FC-C1418) | 559-40-0 | BDL | BDL | NA | BDL | NA | BDL | 212.042 | 8.8146
Trichlorofluoromethane (CFC-11) | 75-69-4 | BDL | BDL | NA | BDL | NA | BDL | 137.37 | 5.7105
Dodecafluoro-n-pentane | 678-26-2 | 0.013 | BDL | NA | BDL | NA | BDL | 288.036 | 11.9736
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | 375-17-7 | BDL | BDL | NA | BDL | NA | BDL | 220.04 | 9.147
Tetradecafluorohexane | 355-42-0 | BDL | BDL | NA | BDL | NA | BDL | 338.04 | 14.0523
1,1,1,2,2,3,3,4,4,5,5-Undecafluoropentane (1H-Perfluoropentane) | 375-61-1 | BDL | BDL | NA | BDL | NA | BDL | 270.04 | 11.2255
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | 3330-15-2 | BDL | BDL | NA | BDL | NA | BDL | 286.04 | 11.8906
Hexadecafluoroheptane | 335-57-9 | BDL | BDL | NA | BDL | NA | BDL | 388.051 | 16.1312
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | 355-37-3 | BDL | BDL | NA | BDL | NA | BDL | 320.05 | 13.3044
Octadecafluorooctane (Perfluorooctane) | 307-34-6 | BDL | BDL | NA | BDL | NA | BDL | 438.06 | 18.2101
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | 375-83-7 | 0.014 | 0.014 | NA | 0.011 | NA | 0.011 | 370.06 | 15.3833
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | 335-65-9 | 0.054 | 0.082 | NA | 0.096 | NA | 0.127 | 420.07 | 17.4623
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | 3330-14-1 | BDL | BDL | NA | BDL | NA | BDL | 452.07 | 18.7925
Batch Blanks > 0.015 ppbv are flagged

## Post cal blanks
PFAS VOC Rainbow Fall 2024
Acceptance criteria is being less than 50% of the lowest cal level (PQL) | 1 minute purge/20 ml method | 1 minute purge/20 ml method | 9 minute purge/200 ml method | 9 minute purge/200 ml method
Can ID: 42879 | 112624 | 112724 | 112924 | 120324
Compound | ppbv | ppbv | ppbv | ppbv | Nominal
Description | Post Cal Blank_CF4 Post 10 ppbv Cal | Post Cal Blank_CF4 Post 10 ppbv Cal | Post Cal Blank_Rest Post 10 ppbv Cal | Post Cal Blank_Rest Post 10 ppbv Cal | 9 minute
Date | 2024-11-26 | 2024-11-27 | 2024-11-29 | 2024-12-03 | PQL | 0.5*PQL
Tetrafluoromethane | BDL | BDL | NA | NA | 10 | 5
Hexafluoroethane (FC-116) | NA | NA | BDL | BDL | 1 | 0.5
Chlorotrifluoromethane (CFC-13) | NA | NA | BDL | BDL | 1 | 0.5
Trifluoromethane (Fluoroform) | NA | NA | BDL | BDL | 1 | 0.5
Octafluoropropane | NA | NA | BDL | BDL | 1 | 0.5
Difluoromethane (HFC-32) | NA | NA | BDL | BDL | 1 | 0.5
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | NA | NA | BDL | BDL | 1 | 0.5
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | NA | NA | BDL | BDL | 1 | 0.5
Fluoromethane (HFC-41) | NA | NA | BDL | BDL | 1 | 0.5
Tetrafluoroethylene | NA | NA | BDL | BDL | 1 | 0.5
Hexafluoropropene | NA | NA | BDL | BDL | 1 | 0.5
1,1,1-trifluoroethane | NA | NA | BDL | BDL | 1 | 0.5
Hexafluoropropene Oxide | NA | NA | BDL | BDL | 1 | 0.5
Chlorodifluoromethane (HCFC-22) | NA | NA | BDL | BDL | 1 | 0.5
1,1,1,2-Tetrafluoroethane | NA | NA | BDL | BDL | 1 | 0.5
Decafluorobutane (Perfluorobutane) | NA | NA | BDL | BDL | 1 | 0.5
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | NA | NA | BDL | BDL | 1 | 0.5
Octafluorocyclopentene (FC-C1418) | NA | NA | BDL | BDL | 1 | 0.5
Trichlorofluoromethane (CFC-11) | NA | NA | BDL | BDL | 1 | 0.5
Dodecafluoro-n-pentane | NA | NA | BDL | BDL | 1 | 0.5
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | NA | NA | BDL | BDL | 1 | 0.5
Tetradecafluorohexane | NA | NA | BDL | BDL | 1 | 0.5
1,1,1,2,2,3,3,4,4,5,5-Undecafluoropentane (1H-Perfluoropentane) | NA | NA | BDL | BDL | 1 | 0.5
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | NA | NA | BDL | BDL | 1 | 0.5
Hexadecafluoroheptane | NA | NA | BDL | BDL | 1 | 0.5
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | NA | NA | BDL | BDL | 1 | 0.5
Octadecafluorooctane (Perfluorooctane) | NA | NA | BDL | BDL | 1 | 0.5
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | NA | NA | 0.011 | 0.012 | 1 | 0.5
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | NA | NA | 0.14 | 0.16 | 1 | 0.5
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | NA | NA | BDL | BDL | 1 | 0.5
above 50% of the PQL (lowest calibration level)
detected but below 50% of PQL

## Ical Bias Assessments Cal
Referenced ICAL-SCWO 2024 Initial Calibration Curve Intrinsic Assessments % of Expected-Nov.14th,2024 |  |  | (used continuing cal) | [9 empty] | 111424/111424_recalcRest_L1_10.D/epatemp.txt (QT Reviewed): Riemer PFAS Cal 1_1.0 ppbv_vol 200 mL | [12 empty] | 111424/111424_recalcRest_L2_11.D/epatemp.txt (QT Reviewed): Riemer PFAS Cal 2_2.5 ppbv_vol 200 ml | [12 empty] | 111424/111424_recalcRest_L3_12.D/epatemp.txt (QT Reviewed): Riemer PFAS Cal 3_5.0 ppbv_vol 200 ml | [12 empty] | 111424/111424_recalcRest_L4_13.D/epatemp.txt (QT Reviewed): Riemer PFAS Cal 4_10 ppbv_vol 200 ml | [12 empty] | 111424/111424_recalcRest_L5_14.D/epatemp.txt (QT Reviewed): Riemer PFAS Cal 5_20 ppbv_vol 200 ml | [12 empty] | 111424/111424_recalcRest_L6_15.D/epatemp.txt (QT Reviewed): Riemer PFAS Cal 6_30 ppbv_vol 200 ml | [12 empty] | 111424/111424_recalcCF4_L10_03.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 10_10ppbv_vol 20 ml | [12 empty] | 111424/111424_recalcCF4_L50_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 50_50 ppbv_vol 20 ml | [12 empty] | 111424/111424_recalcCF4_L100_05.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 100_100 ppbv_vol 20 ml | [12 empty] | 111424/111424_recalcCF4_L150_06.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 150_150 ppbv_vol 20 ml | [12 empty] | 111424/111424_recalcCF4_L200_07.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 200_200 ppbv_vol 20 ml | [12 empty] | 111424/111424_recalcCF4_L300_08.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal 300_300 ppbv_vol 20 ml
Test Description | 9 min_L1 | 9 min_L2 | 9 min_L3 | 9 min_L4 | 9 min_L5 | 9 min_L6 | 1 min_L1 | 1 min_L2 | 1 min_L3 | 1 min_L4 | 1 min_L5 | 1min_L6 | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb
Can ID | Cal level 1 | Cal level 2 | Cal level 3 | Cal level 4 | Cal level 5 | Cal Level 6 | Cal level 1 | CF4 Cal level 2 | CF4 Cal level 3 | CF4 Cal level 4 | CF4 Cal level 5 | CF4 Cal Level 6 | 1,4-difluorobenzene | 30.445 | 114 | 861499 | 5 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.434 | 114 | 866893 | 5 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.446 | 114 | 874185 | 5 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.446 | 114 | 880585 | 5 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.445 | 114 | 890285 | 5 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.446 | 114 | 928261 | 5 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.537 | 114 | 888887 | 50 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.538 | 114 | 870940 | 50 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.538 | 114 | 836391 | 50 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.537 | 114 | 824023 | 50 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.537 | 114 | 849673 | 50 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.537 | 114 | 833790 | 50 |  | ppb
ppbv Conc | 1 | 2.5 | 5 | 10 | 20 | 30 | 10 | 50 | 100 | 150 | 200 | 300 | hexafluoroethane | 8.093 | 119 | 12285 | 0.837 |  | ppb | [5 empty] | 85.5% | hexafluoroethane | 8.105 | 119 | 33301 | 2.262 |  | ppb |  |  | 92.42% |  |  |  | hexafluoroethane | 8.105 | 119 | 69966 | 4.746 |  | ppb | [5 empty] | 96.96% | hexafluoroethane | 8.093 | 119 | 144297 | 9.858 |  | ppb | [5 empty] | 100.7% | hexafluoroethane | 8.105 | 119 | 283398 | 19.696 |  | ppb | [5 empty] | 100.6% | hexafluoroethane | 8.105 | 119 | 427410 | 29.307 |  | ppb | [5 empty] | 99.79% | tetrafluoromethane | 6.994 | 69 | 6506 | 9.088 |  | ppb |  |  |  | 0.089699 |  | 92.38% | tetrafluoromethane | 7.007 | 69 | 32508 | 46.904 |  | ppb |  |  |  | 0.333 |  | 95.35% | tetrafluoromethane | 6.995 | 69 | 64762 | 98.944 |  | ppb |  |  |  | 0.089699 |  | 100.6% | tetrafluoromethane | 6.994 | 69 | 96257 | 151.877 |  | ppb |  |  |  | 0.089699 |  | 102.9% | tetrafluoromethane | 7.019 | 69 | 124349 | 192.889 |  | ppb |  |  |  | 0.02815 |  | 98.03% | tetrafluoromethane | 6.994 | 69 | 180691 | 295.808 |  | ppb |  |  |  | 0.089699 |  | 100.2%
Tetrafluoromethane | Not used | Not used | Not used | Not used | Not used | Not used | 92.38% | 95.35% | 100.6% | 102.9% | 98.03% | 100.2% | chlorotrifluoromethane | 8.515 | 85 | 3208 | 0.856 |  | ppb | [5 empty] | 86.99% | chlorotrifluoromethane | 8.515 | 85 | 8689 | 2.315 |  | ppb |  |  | 94.11% |  |  |  | chlorotrifluoromethane | 8.515 | 85 | 18059 | 4.807 |  | ppb | [5 empty] | 97.7% | chlorotrifluoromethane | 8.516 | 85 | 36911 | 9.909 |  | ppb | [5 empty] | 100.7% | chlorotrifluoromethane | 8.515 | 85 | 72195 | 19.778 |  | ppb | [5 empty] | 100.5% | chlorotrifluoromethane | 8.515 | 85 | 108735 | 29.491 |  | ppb | [5 empty] | 99.9%
Hexafluoroethane (FC-116) | 85.5% | 92.42% | 96.96% | 100.7% | 100.6% | 99.79% | Not used | Not used | Not used | Not used | Not used | Not used | trifluoromethane | 8.72 | 69 | 2398 | 0.861 |  | ppb | [5 empty] | 87.32% | trifluoromethane | 8.721 | 69 | 6798 | 2.431 |  | ppb |  |  | 98.62% |  |  |  | trifluoromethane | 8.721 | 69 | 14063 | 5.008 |  | ppb | [5 empty] | 101.6% | trifluoromethane | 8.721 | 69 | 28308 | 10.09 |  | ppb | [5 empty] | 102.3% | trifluoromethane | 8.72 | 69 | 54374 | 19.472 |  | ppb | [5 empty] | 98.74% | trifluoromethane | 8.72 | 69 | 84952 | 29.685 |  | ppb | [5 empty] | 100.4%
Chlorotrifluoromethane (CFC-13) | 86.99% | 94.11% | 97.7% | 100.7% | 100.5% | 99.9% | Not used | Not used | Not used | Not used | Not used | Not used | octafluoropropane | 11.178 | 69 | 22914 | 0.866 |  | ppb | [5 empty] | 88.46% | octafluoropropane | 11.178 | 69 | 61099 | 2.304 |  | ppb |  |  | 94.14% |  |  |  | octafluoropropane | 11.178 | 69 | 126996 | 4.786 |  | ppb | [5 empty] | 97.77% | octafluoropropane | 11.178 | 69 | 261062 | 9.925 |  | ppb | [5 empty] | 101.4% | octafluoropropane | 11.178 | 69 | 505478 | 19.606 |  | ppb | [5 empty] | 100.1% | octafluoropropane | 11.178 | 69 | 764354 | 29.367 |  | ppb | [5 empty] | 99.99%
Trifluoromethane (Fluoroform) | 87.32% | 98.62% | 101.6% | 102.3% | 98.74% | 100.4% | Not used | Not used | Not used | Not used | Not used | Not used | difluoromethane | 10.99 | 51 | 1732 | 0.898 |  | ppb | [5 empty] | 92.39% | difluoromethane | 10.99 | 51 | 4520 | 2.338 |  | ppb |  |  | 96.21% |  |  |  | difluoromethane | 10.99 | 51 | 9368 | 4.836 |  | ppb | [5 empty] | 99.51% | difluoromethane | 10.99 | 51 | 18709 | 9.712 |  | ppb | [5 empty] | 99.92% | difluoromethane | 10.99 | 51 | 37052 | 19.528 |  | ppb | [5 empty] | 100.5% | difluoromethane | 10.99 | 51 | 56158 | 29.142 |  | ppb | [5 empty] | 99.94%
Octafluoropropane | 88.46% | 94.14% | 97.77% | 101.4% | 100.1% | 99.99% | Not used | Not used | Not used | Not used | Not used | Not used | 1,1,1,2,2-pentafluoroe... | 12.871 | 51 | 6305 | 0.865 |  | ppb | [5 empty] | 87.46% | 1,1,1,2,2-pentafluoroe... | 12.871 | 51 | 16911 | 2.311 |  | ppb |  |  | 93.47% |  |  |  | 1,1,1,2,2-pentafluoroe... | 12.871 | 51 | 35717 | 4.865 |  | ppb | [5 empty] | 98.38% | 1,1,1,2,2-pentafluoroe... | 12.871 | 51 | 73741 | 10.073 |  | ppb | [5 empty] | 101.9% | 1,1,1,2,2-pentafluoroe... | 12.871 | 51 | 143363 | 19.741 |  | ppb | [5 empty] | 99.8% | 1,1,1,2,2-pentafluoroe... | 12.871 | 51 | 220372 | 29.691 |  | ppb | [5 empty] | 100.1%
*Difluoromethane (HFC-32) | 92.39% | 96.21% | 99.51% | 99.92% | 100.5% | 99.94% | Not used | Not used | Not used | Not used | Not used | Not used | 1,1,2,2,3,3,4,4-octafl... | 15.953 | 100 | 24377 | 0.868 |  | ppb | [5 empty] | 89.3% | 1,1,2,2,3,3,4,4-octafl... | 15.953 | 100 | 65048 | 2.314 |  | ppb |  |  | 95.23% |  |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.953 | 100 | 133315 | 4.743 |  | ppb | [5 empty] | 97.59% | 1,1,2,2,3,3,4,4-octafl... | 15.953 | 100 | 273035 | 9.814 |  | ppb | [5 empty] | 101% | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 529898 | 19.5 |  | ppb | [5 empty] | 100.3% | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 796978 | 29.146 |  | ppb | [5 empty] | 99.95%
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | 87.46% | 93.47% | 98.38% | 101.9% | 99.8% | 100.1% | Not used | Not used | Not used | Not used | Not used | Not used | fluoromethane | 11.438 | 33 | 875 | 0.974 |  | ppb | [5 empty] | 98.48% | fluoromethane | 11.438 | 33 | 2093 | 2.327 |  | ppb |  |  | 94.12% |  |  |  | fluoromethane | 11.438 | 33 | 4178m | 4.641 |  | ppb | [5 empty] | 93.85% | fluoromethane | 11.439 | 33 | 8772 | 9.846 |  | ppb | [5 empty] | 99.56% | fluoromethane | 11.438 | 33 | 17142m | 19.692 |  | ppb | [5 empty] | 99.56% | fluoromethane | 11.438 | 33 | 25916 | 29.587 |  | ppb | [5 empty] | 99.72%
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | 89.3% | 95.23% | 97.59% | 101% | 100.3% | 99.95% | Not used | Not used | Not used | Not used | Not used | Not used | tetrafluoroethylene | 8.503 | 81 | 5178 | 0.889 |  | ppb | [5 empty] | 90.35% | tetrafluoroethylene | 8.503 | 81 | 13411 | 2.296 |  | ppb |  |  | 93.33% |  |  |  | tetrafluoroethylene | 8.503 | 81 | 28632 | 4.888 |  | ppb | [5 empty] | 99.35% | tetrafluoroethylene | 8.503 | 81 | 58404 | 10.007 |  | ppb | [5 empty] | 101.7% | tetrafluoroethylene | 8.503 | 81 | 113386 | 19.619 |  | ppb | [5 empty] | 99.69% | tetrafluoroethylene | 8.503 | 81 | 174208 | 29.552 |  | ppb | [5 empty] | 100.1%
Fluoromethane (HFC-41) | 98.48% | 94.12% | 93.85% | 99.56% | 99.56% | 99.72% | Not used | Not used | Not used | Not used | Not used | Not used | hexafluoropropene | 13.276 | 131 | 12749 | 0.868 |  | ppb | [5 empty] | 88.84% | hexafluoropropene | 13.277 | 131 | 33729 | 2.291 |  | ppb |  |  | 93.8% |  |  |  | hexafluoropropene | 13.276 | 131 | 70348 | 4.778 |  | ppb | [5 empty] | 97.81% | hexafluoropropene | 13.262 | 131 | 144042 | 9.877 |  | ppb | [5 empty] | 101.1% | hexafluoropropene | 13.262 | 131 | 279263 | 19.573 |  | ppb | [5 empty] | 100.2% | hexafluoropropene | 13.262 | 131 | 421148 | 29.289 |  | ppb | [5 empty] | 99.93%
Tetrafluoroethylene | 90.35% | 93.33% | 99.35% | 101.7% | 99.69% | 100.1% | Not used | Not used | Not used | Not used | Not used | Not used | 1,1,1-trifluoroethane | 13.262 | 64 | 821 | 0.996 |  | ppb | [5 empty] | 102.3% | 1,1,1-trifluoroethane | 13.262 | 64 | 1825 | 2.206 |  | ppb |  |  | 90.6% |  |  |  | 1,1,1-trifluoroethane | 13.262 | 64 | 3957 | 4.777 |  | ppb | [5 empty] | 98.09% | 1,1,1-trifluoroethane | 13.262 | 64 | 7997 | 9.713 |  | ppb | [5 empty] | 99.72% | 1,1,1-trifluoroethane | 13.262 | 64 | 15931 | 19.674 |  | ppb | [5 empty] | 101% | 1,1,1-trifluoroethane | 13.262 | 64 | 23962 | 29.158 |  | ppb | [5 empty] | 99.79%
Hexafluoropropene | 88.84% | 93.8% | 97.81% | 101.1% | 100.2% | 99.93% | Not used | Not used | Not used | Not used | Not used | Not used | hexafluoropropene_oxide | 13.392 | 69 | 686m | 0.993 |  | ppb | [5 empty] | 100.9% | hexafluoropropene_oxide | 13.45 | 69 | 1567m | 2.205 |  | ppb |  |  | 89.63% |  |  |  | hexafluoropropene_oxide | 13.45 | 69 | 3365m | 4.505 |  | ppb | [5 empty] | 91.57% | hexafluoropropene_oxide | 13.45 | 69 | 8092 | 9.832 |  | ppb | [5 empty] | 99.92% | hexafluoropropene_oxide | 13.45 | 69 | 18574 | 19.355 |  | ppb | [5 empty] | 98.35% | hexafluoropropene_oxide | 13.45 | 69 | 33431 | 29.329 |  | ppb | [5 empty] | 99.35%
1,1,1-trifluoroethane | 102.3% | 90.6% | 98.09% | 99.72% | 101% | 99.79% | Not used | Not used | Not used | Not used | Not used | Not used | chlorodifluoromethane | 13.768 | 51 | 6605 | 0.879 |  | ppb | [5 empty] | 88.88% | chlorodifluoromethane | 13.769 | 51 | 17763 | 2.356 |  | ppb |  |  | 95.29% |  |  |  | chlorodifluoromethane | 13.769 | 51 | 36935 | 4.889 |  | ppb | [5 empty] | 98.87% | chlorodifluoromethane | 13.769 | 51 | 75257 | 10.015 |  | ppb | [5 empty] | 101.3% | chlorodifluoromethane | 13.769 | 51 | 146534 | 19.767 |  | ppb | [5 empty] | 99.93% | chlorodifluoromethane | 13.769 | 51 | 223657 | 29.685 |  | ppb | [5 empty] | 100.1%
Hexafluoropropene Oxide | 100.9% | 89.63% | 91.57% | 99.92% | 98.35% | 99.35% | Not used | Not used | Not used | Not used | Not used | Not used | 1,1,1,2-tetrafluoroethane | 15.967 | 33 | 2338 | 0.885 |  | ppb | [5 empty] | 89.48% | 1,1,1,2-tetrafluoroethane | 15.967 | 33 | 6365 | 2.408 |  | ppb |  |  | 97.39% |  |  |  | 1,1,1,2-tetrafluoroethane | 15.981 | 33 | 12862 | 4.868 |  | ppb | [5 empty] | 98.44% | 1,1,1,2-tetrafluoroethane | 15.967 | 33 | 26122 | 9.997 |  | ppb | [5 empty] | 101.1% | 1,1,1,2-tetrafluoroethane | 15.967 | 33 | 50416 | 19.782 |  | ppb | [5 empty] | 100% | 1,1,1,2-tetrafluoroethane | 15.967 | 33 | 75952 | 29.681 |  | ppb | [5 empty] | 100%
Chlorodifluoromethane (HCFC-22) | 88.88% | 95.29% | 98.87% | 101.3% | 99.93% | 100.1% | Not used | Not used | Not used | Not used | Not used | Not used | decafluorobutane | 16.175 | 119 | 13499 | 0.91 |  | ppb | [5 empty] | 93.43% | decafluorobutane | 16.175 | 119 | 35478 | 2.393 |  | ppb |  |  | 98.28% |  |  |  | decafluorobutane | 16.175 | 119 | 72155 | 4.879 |  | ppb | [5 empty] | 100.2% | decafluorobutane | 16.175 | 119 | 145386 | 9.981 |  | ppb | [5 empty] | 102.5% | decafluorobutane | 16.175 | 119 | 278743 | 19.794 |  | ppb | [5 empty] | 101.6% | decafluorobutane | 16.161 | 119 | 415675 | 29.677 |  | ppb | [5 empty] | 101.6%
1,1,1,2-Tetrafluoroethane | 89.48% | 97.39% | 98.44% | 101.1% | 100% | 100% | Not used | Not used | Not used | Not used | Not used | Not used | 1H-Heptafluoropropane | 18.147 | 51 | 9543 | 0.886 |  | ppb | [5 empty] | 90.04% | 1H-Heptafluoropropane | 18.148 | 51 | 25607 | 2.373 |  | ppb |  |  | 96.46% |  |  |  | 1H-Heptafluoropropane | 18.148 | 51 | 52065 | 4.82 |  | ppb | [5 empty] | 97.97% | 1H-Heptafluoropropane | 18.148 | 51 | 106773 | 9.97 |  | ppb | [5 empty] | 101.3% | 1H-Heptafluoropropane | 18.148 | 51 | 206220 | 19.639 |  | ppb | [5 empty] | 99.79% | 1H-Heptafluoropropane | 18.148 | 51 | 312340 | 29.459 |  | ppb | [5 empty] | 99.79%
Decafluorobutane (Perfluorobutane) | 93.43% | 98.28% | 100.2% | 102.5% | 101.6% | 101.6% | Not used | Not used | Not used | Not used | Not used | Not used | octafluorocyclopentene | 21.967 | 93 | 17863 | 0.906 |  | ppb | [5 empty] | 93.02% | octafluorocyclopentene | 21.968 | 93 | 47265 | 2.388 |  | ppb |  |  | 98.07% |  |  |  | octafluorocyclopentene | 21.968 | 93 | 97093 | 4.884 |  | ppb | [5 empty] | 100.3% | octafluorocyclopentene | 21.968 | 93 | 196912 | 9.916 |  | ppb | [5 empty] | 101.8% | octafluorocyclopentene | 21.955 | 93 | 380945 | 19.278 |  | ppb | [5 empty] | 98.96% | octafluorocyclopentene | 21.955 | 93 | 592255 | 29.241 |  | ppb | [5 empty] | 100.1%
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | 90.04% | 96.46% | 97.97% | 101.3% | 99.79% | 99.79% | Not used | Not used | Not used | Not used | Not used | Not used | trichlorofluoromethane | 21.274 | 101 | 19675 | 0.94 |  | ppb | [5 empty] | 95.33% | trichlorofluoromethane | 21.274 | 101 | 53363 | 2.542 |  | ppb |  |  | 103.1% |  |  |  | trichlorofluoromethane | 21.257 | 101 | 103884 | 4.931 |  | ppb | [5 empty] | 100% | trichlorofluoromethane | 21.275 | 101 | 210730 | 10.033 |  | ppb | [5 empty] | 101.8% | trichlorofluoromethane | 21.257 | 101 | 408908 | 19.642 |  | ppb | [5 empty] | 99.6% | trichlorofluoromethane | 21.257 | 101 | 632129 | 29.747 |  | ppb | [5 empty] | 100.6%
Octafluorocyclopentene (FC-C1418) | 93.02% | 98.07% | 100.3% | 101.8% | 98.96% | 100.1% | Not used | Not used | Not used | Not used | Not used | Not used | dodecafluoro-n-pentane | 21.03 | 69 | 39766 | 0.934 |  | ppb | [5 empty] | 94.73% | dodecafluoro-n-pentane | 21.012 | 69 | 103722 | 2.432 |  | ppb |  |  | 98.66% |  |  |  | dodecafluoro-n-pentane | 21.012 | 69 | 209879 | 4.911 |  | ppb | [5 empty] | 99.61% | dodecafluoro-n-pentane | 21.013 | 69 | 425214 | 10.009 |  | ppb | [5 empty] | 101.5% | dodecafluoro-n-pentane | 21.012 | 69 | 821295 | 19.617 |  | ppb | [5 empty] | 99.48% | dodecafluoro-n-pentane | 21.012 | 69 | 1258639 | 29.633 |  | ppb | [5 empty] | 100.2%
Trichlorofluoromethane (CFC-11) | 95.33% | 103.1% | 100% | 101.8% | 99.6% | 100.6% | Not used | Not used | Not used | Not used | Not used | Not used | 1H-Nonafluorobutane | 22.826 | 51 | 17324 | 1.001 |  | ppb | [5 empty] | 101% | 1H-Nonafluorobutane | 22.826 | 51 | 43650 | 2.521 |  | ppb |  |  | 101.8% |  |  |  | 1H-Nonafluorobutane | 22.826 | 51 | 89815 | 5.198 |  | ppb | [5 empty] | 104.9% | 1H-Nonafluorobutane | 22.826 | 51 | 177218 | 10.398 |  | ppb | [5 empty] | 104.9% | 1H-Nonafluorobutane | 22.826 | 51 | 322148 | 19.408 |  | ppb | [5 empty] | 97.92% | 1H-Nonafluorobutane | 22.814 | 51 | 503017 | 30.494 |  | ppb | [5 empty] | 102.6%
Dodecafluoro-n-pentane | 94.73% | 98.66% | 99.61% | 101.5% | 99.48% | 100.2% | Not used | Not used | Not used | Not used | Not used | Not used | tetradecafluorohexane | 24.495 | 69 | 40433m | 0.86 |  | ppb | [5 empty] | 87.49% | tetradecafluorohexane | 24.487 | 69 | 109956m | 2.338 |  | ppb |  |  | 95.14% |  |  |  | tetradecafluorohexane | 24.495 | 69 | 223145m | 4.751 |  | ppb | [5 empty] | 96.66% | tetradecafluorohexane | 24.495 | 69 | 450248m | 9.71 |  | ppb | [5 empty] | 98.78% | tetradecafluorohexane | 24.487 | 69 | 859470m | 19.069 |  | ppb | [5 empty] | 96.99% | tetradecafluorohexane | 24.582 | 69 | 1322710m | 29.458 |  | ppb | [5 empty] | 99.89%
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | 101% | 101.8% | 104.9% | 104.9% | 97.92% | 102.6% | Not used | Not used | Not used | Not used | Not used | Not used | 1H-Perfluoropentane | 26.144 | 69 | 14141 | 0.842 |  | ppb | [5 empty] | 85.05% | 1H-Perfluoropentane | 26.144 | 69 | 37799m | 2.259 |  | ppb |  |  | 91.27% |  |  |  | 1H-Perfluoropentane | 26.144 | 69 | 76623 | 4.617 |  | ppb | [5 empty] | 93.27% | 1H-Perfluoropentane | 26.144 | 69 | 154467 | 9.575 |  | ppb | [5 empty] | 96.72% | 1H-Perfluoropentane | 26.134 | 69 | 310403 | 20.725 |  | ppb | [5 empty] | 104.7% | 1H-Perfluoropentane | 26.134 | 69 | 426222 | 29.299 |  | ppb | [5 empty] | 98.65%
Tetradecafluorohexane | 87.49% | 95.14% | 96.66% | 98.78% | 96.99% | 99.89% | Not used | Not used | Not used | Not used | Not used | Not used | heptafluoropropyl 1,2,... | 26.212 | 101 | 38791m | 0.961 |  | ppb | [5 empty] | 97.76% | heptafluoropropyl 1,2,... | 26.212 | 101 | 101832m | 2.517 |  | ppb |  |  | 102.4% |  |  |  | heptafluoropropyl 1,2,... | 26.212 | 101 | 203502m | 5.019 |  | ppb | [5 empty] | 102.1% | heptafluoropropyl 1,2,... | 26.212 | 101 | 403808m | 10.016 |  | ppb | [5 empty] | 101.9% | heptafluoropropyl 1,2,... | 26.212 | 101 | 785557m | 19.777 |  | ppb | [5 empty] | 100.6% | heptafluoropropyl 1,2,... | 26.212 | 101 | 1187796m | 29.442 |  | ppb | [5 empty] | 99.84%
1,1,1,2,2,3,3,4,4,5,5-Undecafluoropentane (1H-Perfluoropentane) | 85.05% | 91.27% | 93.27% | 96.72% | 104.7% | 98.65% | Not used | Not used | Not used | Not used | Not used | Not used | hexadecafluoroheptane | 27.304 | 69 | 50349 | 0.915 |  | ppb | [5 empty] | 92.61% | hexadecafluoroheptane | 27.304 | 69 | 135133 | 2.448 |  | ppb |  |  | 99.11% |  |  |  | hexadecafluoroheptane | 27.304 | 69 | 267001 | 4.824 |  | ppb | [5 empty] | 97.65% | hexadecafluoroheptane | 27.304 | 69 | 549352 | 9.978 |  | ppb | [5 empty] | 101% | hexadecafluoroheptane | 27.294 | 69 | 1074715 | 19.779 |  | ppb | [5 empty] | 100.1% | hexadecafluoroheptane | 27.294 | 69 | 1638828 | 29.657 |  | ppb | [5 empty] | 100.1%
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | 97.76% | 102.4% | 102.1% | 101.9% | 100.6% | 99.84% | Not used | Not used | Not used | Not used | Not used | Not used | 1H-Perfluorohexane | 28.776 | 51 | 23882 | 0.91 |  | ppb | [5 empty] | 92.29% | 1H-Perfluorohexane | 28.776 | 51 | 61876 | 2.351 |  | ppb |  |  | 95.38% |  |  |  | 1H-Perfluorohexane | 28.776 | 51 | 127861 | 4.845 |  | ppb | [5 empty] | 98.28% | 1H-Perfluorohexane | 28.776 | 51 | 260851 | 9.928 |  | ppb | [5 empty] | 100.7% | 1H-Perfluorohexane | 28.766 | 51 | 505774 | 19.466 |  | ppb | [5 empty] | 98.71% | 1H-Perfluorohexane | 28.766 | 51 | 775379 | 29.297 |  | ppb | [5 empty] | 99.04%
Hexadecafluoroheptane | 92.61% | 99.11% | 97.65% | 101% | 100.1% | 100.1% | Not used | Not used | Not used | Not used | Not used | Not used | octadecafluorooctane | 29.56 | 119 | 34223 | 0.93 |  | ppb | [5 empty] | 94.8% | octadecafluorooctane | 29.548 | 119 | 91547 | 2.482 |  | ppb |  |  | 101.2% |  |  |  | octadecafluorooctane | 29.548 | 119 | 169838 | 4.59 |  | ppb | [5 empty] | 93.58% | octadecafluorooctane | 29.549 | 119 | 361305 | 9.82 |  | ppb | [5 empty] | 100.1% | octadecafluorooctane | 29.548 | 119 | 719682 | 19.838 |  | ppb | [5 empty] | 101.1% | octadecafluorooctane | 29.537 | 119 | 1083617 | 29.358 |  | ppb | [5 empty] | 99.76%
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | 92.29% | 95.38% | 98.28% | 100.7% | 98.71% | 99.04% | Not used | Not used | Not used | Not used | Not used | Not used | 1H-Perfluoroheptane | 31.089 | 51 | 67827 | 0.921 |  | ppb | [5 empty] | 93.22% | 1H-Perfluoroheptane | 31.09 | 51 | 178130 | 2.413 |  | ppb |  |  | 97.69% |  |  |  | 1H-Perfluoroheptane | 31.09 | 51 | 349150 | 4.716 |  | ppb | [5 empty] | 95.47% | 1H-Perfluoroheptane | 31.078 | 51 | 732782 | 9.952 |  | ppb | [5 empty] | 100.7% | 1H-Perfluoroheptane | 31.078 | 51 | 1441738 | 19.846 |  | ppb | [5 empty] | 100.4% | 1H-Perfluoroheptane | 31.078 | 51 | 2185650 | 29.574 |  | ppb | [5 empty] | 99.78%
Octadecafluorooctane (Perfluorooctane) | 94.8% | 101.2% | 93.58% | 100.1% | 101.1% | 99.76% | Not used | Not used | Not used | Not used | Not used | Not used | 1H-Perfluorooctane | 33.505 | 51 | 55396 | 1.002 |  | ppb | [5 empty] | 102.1% | 1H-Perfluorooctane | 33.505 | 51 | 142911 | 2.571 |  | ppb |  |  | 104.8% |  |  |  | 1H-Perfluorooctane | 33.493 | 51 | 246225 | 4.396 |  | ppb | [5 empty] | 89.62% | 1H-Perfluorooctane | 33.494 | 51 | 556507 | 9.883 |  | ppb | [5 empty] | 100.7% | 1H-Perfluorooctane | 33.493 | 51 | 1135588 | 20.022 |  | ppb | [5 empty] | 102% | 1H-Perfluorooctane | 33.482 | 51 | 1727104 | 29.305 |  | ppb | [5 empty] | 99.58%
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | 93.22% | 97.69% | 95.47% | 100.7% | 100.4% | 99.78% | Not used | Not used | Not used | Not used | Not used | Not used | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 169708 | 0.935 |  | ppb | [5 empty] | 94.64% | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 444163 | 2.443 |  | ppb |  |  | 98.91% |  |  |  | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 853538 | 4.688 |  | ppb | [5 empty] | 94.9% | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 1798514 | 9.968 |  | ppb | [5 empty] | 100.9% | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 3513022 | 19.876 |  | ppb | [5 empty] | 100.6% | 2H-peffluoro-5-methyl-... | 33.033 | 169 | 5283919 | 29.604 |  | ppb | [5 empty] | 99.88%
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | 102.1% | 104.8% | 89.62% | 100.7% | 102% | 99.58% | Not used | Not used | Not used | Not used | Not used | Not used
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | 94.64% | 98.91% | 94.9% | 100.9% | 100.6% | 99.88% | Not used | Not used | Not used | Not used | Not used | Not used

## Duplicate QC Checks
Duplicate Relative Percent Deviation (must be greater than 3x MDL to qualify for valid comparison )
Sample Name | Rainbow Run 1 Port 4 PFBA 1-minute Method |  |  |  | Rainbow Run 1 Port 4 PFBA 9-minute Method |  |  |  | Rainbow Run 7 Port 8 PFOA 1-minute Method |  |  |  | Rainbow Run 7 Port 8 PFOA 9-minute Method
Can ID |  |  | 20 ml dup % |  |  |  | 200 ml dup % | Comments |  |  | 20 ml dup % |  |  |  | 200 ml dup % | Comments
Date Analyzed | 2024-11-26 |  |  |  | 2024-11-26 |  |  |  | 2024-11-29 |  |  |  | 2024-11-29
Compound | ppbv |  | %RPD |  | ppbv |  | %RPD |  | ppbv |  | %RPD |  | ppbv |  | %RPD
1,4 difluorobenzene (IS) | Original | Duplicate |  |  | Original | Duplicate |  |  | Original | Duplicate |  |  | Original | Duplicate
Tetrafluoromethane | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA
Hexafluoroethane (FC-116) | NA | NA | NA |  | 1.3205 | 1.3466 | 1.959% |  | NA | NA | NA |  | 2.0096 | 2.0067 | 0.1434%
Chlorotrifluoromethane (CFC-13) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Trifluoromethane (Fluoroform) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | 33.5067 | 33.2217 | 0.8543%
Octafluoropropane | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | 0.29367 | 0.29655 | 0.9756%
Difluoromethane (HFC-32) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | NA | NA | NA |  | 0.058045 | BDL | NA |  | NA | NA | NA |  | 0.16987 | 0.14971 | 12.61%
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Fluoromethane (HFC-41) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Tetrafluoroethylene | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Hexafluoropropene | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
1,1,1-trifluoroethane | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Hexafluoropropene Oxide | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Chlorodifluoromethane (HCFC-22) | NA | NA | NA |  | 0.10738 | 0.095773 | 11.43% |  | NA | NA | NA |  | 0.14107 | 0.1382 | 2.062%
1,1,1,2-Tetrafluoroethane | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | 0.13244 | 0.10077 | 27.16%
Decafluorobutane (Perfluorobutane) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Octafluorocyclopentene (FC-C1418) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Trichlorofluoromethane (CFC-11) | NA | NA | NA |  | 0.095773 | 0.098676 | 2.985% |  | NA | NA | NA |  | 0.11516 | 0.11516 | 0%
Dodecafluoro-n-pentane | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Tetradecafluorohexane | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
1,1,1,2,2,3,3,4,4,5,5-Undecfluoropentane (1H-Perfluoropentane) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
hexadecafluoroheptane | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
Octadecafluorooctane (Perfluorooctane) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | NA | NA | NA |  | BDL | BDL | NA |  | NA | NA | NA |  | BDL | BDL | NA
|  |  | NA | Average RPD |  |  | 5.457% |  |  |  | NA | Average RPD |  |  | 6.258%
[8 empty] | Average RPD | [7 empty] | Average RPD

## Daily cal curve checks
Daily Calibration Checks | 2024-11-26 | 2024-11-26 | 2024-11-27 | 2024-11-27 | 2024-11-29 | 2024-11-29 | 2024-12-03 | 2024-12-03 | 112624/112624_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_initial_100ppbv_vol 20 ml | [14 empty] | Opening Cal Check | 112624/112624_05.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | Opening cal check | 112624/112624_19.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [14 empty] | Closing Cal Check | 112624/112624_20.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [14 empty] | Closing cal check | 112724/112724_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_initial_100ppbv_vol 20 ml | [14 empty] | Opening Cal Check | 112724/112724_05.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | Opening cal check | 112724/112724_17.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [14 empty] | Closing Cal Check | 112724/112724_18.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [14 empty] | closing cal check | 112924/112924_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_initial_100ppbv_vol 20 ml | [14 empty] | Opening Cal Check | 112924/112924_06.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | Opening cal check | 112924/112924_20.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [14 empty] | Closing Cal Check | 112924/112924_21.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [14 empty] | Opening cal check | 120324/120324_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_initial_100ppbv_vol 20 ml | [14 empty] | Closing Cal Check | 120324/120324_06.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | Opening cal check | 120324/120324_10.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [14 empty] | Closing Cal Check | 120324/120324_11.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [14 empty] | Opening cal check | CF4 Method |  | Lower Limit | Uppper Limit | Rest Method |  | Lower Limit | Uppper Limit
| Opening | Closing | Opening | Closing | Opening | Closing | Opening | Closing | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | % Recovery | Average IS area | 1053174 | 631904.4 | 1474443.6 | Average IS area | 1032225 | 619335 | 1445115
| Recovery % | Recovery % | Recovery % | Recovery % | Recovery % | Recovery % | Recovery % | Recovery % | 1,4-difluorobenzene | 30.48 | 114 | 1082936 | 50 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.399 | 114 | 1038373 | 5 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.48 | 114 | 1072880 | 50 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.4 | 114 | 1043367 | 5 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.48 | 114 | 1063940 | 50 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.388 | 114 | 1049943 | 5 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.469 | 114 | 1094387 | 50 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.388 | 114 | 1059081 | 5 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.491 | 114 | 1035814 | 50 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.365 | 114 | 1023085 | 5 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.457 | 114 | 1048353 | 50 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.354 | 114 | 1019948 | 5 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.48 | 114 | 940577 | 50 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.342 | 114 | 947736 | 5 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.457 | 114 | 961607 | 50 |  | ppb | [6 empty] | Pass | Pass |  | 1,4-difluorobenzene | 30.319 | 114 | 957751 | 5 |  | ppb | [6 empty] | Pass | Pass |  | Average RT | 30.503 | 30.4197 | 30.5863 | Average RT | 30.411 | 30.3277 | 30.4943
Tetrafluoromethane | 96.12% | 99.86% | 102.2% | 97.37% | 94.97% | 91.06% | 85.78% | 83.06% | tetrafluoromethane | 6.994 | 69 | 78707 | 94.565 | 94.565 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 94.565 | Pass | Pass | 96.12% | hexafluoroethane | 8.093 | 119 | 161695 | 9.636 | 9.636 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.636 | Pass | Pass | 98.43% | tetrafluoromethane | 6.995 | 69 | 81010 | 98.244 | 98.244 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 98.244 | Pass | Pass | 99.86% | hexafluoroethane | 8.093 | 119 | 162256 | 9.623 | 9.623 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.623 | Pass | Pass | 98.29% | tetrafluoromethane | 6.994 | 69 | 82247 | 100.583 | 100.583 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 100.583 | Pass | Pass | 102.2% | hexafluoroethane | 8.093 | 119 | 166556 | 9.816 | 9.816 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.816 | Pass | Pass | 100.3% | tetrafluoromethane | 6.995 | 69 | 80568 | 95.788 | 95.788 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 95.788 | Pass | Pass | 97.37% | hexafluoroethane | 8.093 | 119 | 168067 | 9.82 | 9.82 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.82 | Pass | Pass | 100.3% | tetrafluoromethane | 6.994 | 69 | 74377 | 93.428 | 93.428 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 93.428 | Pass | Pass | 94.97% | hexafluoroethane | 8.093 | 119 | 163708 | 9.902 | 9.902 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.902 | Pass | Pass | 101.1% | tetrafluoromethane | 6.995 | 69 | 72180 | 89.584 | 89.584 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 89.584 | Pass | Pass | 91.06% | hexafluoroethane | 8.093 | 119 | 158786 | 9.634 | 9.634 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.634 | Pass | Pass | 98.41% | tetrafluoromethane | 6.982 | 69 | 61007 | 84.393 | 84.393 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 84.393 | Pass | Pass | 85.78% | hexafluoroethane | 8.093 | 119 | 143478m | 9.368 | 9.368 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.368 | Pass | Pass | 95.69% | tetrafluoromethane | 6.982 | 69 | 60393 | 81.716 | 81.716 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 81.716 | Pass | Pass | 83.06% | hexafluoroethane | 8.093 | 119 | 151992 | 9.82 | 9.82 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.82 | Pass | Fail | 100.3%
Hexafluoroethane (FC-116) | 98.43% | 98.29% | 100.3% | 100.3% | 101.1% | 98.41% | 95.69% | 100.3% | [16 empty] | chlorotrifluoromethane | 8.515 | 85 | 42364 | 9.714 | 9.714 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.714 |  |  | 98.72% | [16 empty] | chlorotrifluoromethane | 8.515 | 85 | 43052 | 9.825 | 9.825 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.825 |  |  | 99.85% | [16 empty] | chlorotrifluoromethane | 8.515 | 85 | 44539 | 10.101 | 10.101 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.101 |  |  | 102.7% | [16 empty] | chlorotrifluoromethane | 8.515 | 85 | 45132 | 10.147 | 10.147 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.147 |  |  | 103.1% | [16 empty] | chlorotrifluoromethane | 8.515 | 85 | 44000 | 10.24 | 10.24 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.24 |  |  | 104.1% | [16 empty] | chlorotrifluoromethane | 8.515 | 85 | 42686 | 9.965 | 9.965 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.965 |  |  | 101.3% | [16 empty] | chlorotrifluoromethane | 8.503 | 85 | 39853 | 10.013 | 10.013 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.013 |  |  | 101.8% | [16 empty] | chlorotrifluoromethane | 8.503 | 85 | 40944 | 10.179 | 10.179 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.179 |  |  | 103.4%
Chlorotrifluoromethane (CFC-13) | 98.72% | 99.85% | 102.7% | 103.1% | 104.1% | 101.3% | 101.8% | 103.4% | [16 empty] | trifluoromethane | 8.708 | 69 | 37287 | 9.874 | 9.874 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 9.874 |  |  | 100.1% | [16 empty] | trifluoromethane | 8.708 | 69 | 38023 | 10.021 | 10.021 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.021 |  |  | 101.6% | [16 empty] | trifluoromethane | 8.708 | 69 | 39839 | 10.434 | 10.434 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.434 |  |  | 105.8% | [16 empty] | trifluoromethane | 8.708 | 69 | 39903 | 10.36 | 10.36 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.36 |  |  | 105.1% | [16 empty] | trifluoromethane | 8.709 | 69 | 39204 | 10.537 | 10.537 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.537 |  |  | 106.9% | [16 empty] | trifluoromethane | 8.709 | 69 | 38963 | 10.504 | 10.504 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.504 |  |  | 106.5% | [16 empty] | trifluoromethane | 8.708 | 69 | 35994 | 10.443 | 10.443 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.443 |  |  | 105.9% | [16 empty] | trifluoromethane | 8.708 | 69 | 38590 | 11.079 | 11.079 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 11.079 |  |  | 112.4%
Trifluoromethane (Fluoroform) | 100.1% | 101.6% | 105.8% | 105.1% | 106.9% | 106.5% | 105.9% | 112.4% | [16 empty] | octafluoropropane | 11.178 | 69 | 328369 | 9.721 | 9.721 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 9.721 |  |  | 99.3% | [16 empty] | octafluoropropane | 11.178 | 69 | 339379 | 9.998 | 9.998 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 9.998 |  |  | 102.1% | [16 empty] | octafluoropropane | 11.178 | 69 | 350883 | 10.273 | 10.273 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.273 |  |  | 104.9% | [16 empty] | octafluoropropane | 11.178 | 69 | 355402 | 10.315 | 10.315 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.315 |  |  | 105.4% | [16 empty] | octafluoropropane | 11.178 | 69 | 344476 | 10.35 | 10.35 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.35 |  |  | 105.7% | [16 empty] | octafluoropropane | 11.178 | 69 | 341069 | 10.279 | 10.279 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.279 |  |  | 105% | [16 empty] | octafluoropropane | 11.178 | 69 | 310451 | 10.069 | 10.069 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.069 |  |  | 102.8% | [16 empty] | octafluoropropane | 11.178 | 69 | 324804 | 10.424 | 10.424 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.424 |  |  | 106.5%
Octafluoropropane | 99.3% | 102.1% | 104.9% | 105.4% | 105.7% | 105% | 102.8% | 106.5% | [16 empty] | difluoromethane | 10.975 | 51 | 25054 | 9.846 | 9.846 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 9.846 |  |  | 101.3% | [16 empty] | difluoromethane | 10.975 | 51 | 26160 | 10.231 | 10.231 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.231 |  |  | 105.3% | [16 empty] | difluoromethane | 10.975 | 51 | 26313 | 10.226 | 10.226 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.226 |  |  | 105.2% | [16 empty] | difluoromethane | 10.961 | 51 | 26740 | 10.303 | 10.303 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.303 |  |  | 106% | [16 empty] | difluoromethane | 10.961 | 51 | 26298 | 10.489 | 10.489 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.489 |  |  | 107.9% | [16 empty] | difluoromethane | 10.961 | 51 | 25987 | 10.397 | 10.397 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.397 |  |  | 107% | [16 empty] | difluoromethane | 10.961 | 51 | 23511 | 10.123 | 10.123 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.123 |  |  | 104.1% | [16 empty] | difluoromethane | 10.961 | 51 | 24618 | 10.489 | 10.489 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.489 |  |  | 107.9%
Difluoromethane (HFC-32) | 101.3% | 105.3% | 105.2% | 106% | 107.9% | 107% | 104.1% | 107.9% | [16 empty] | 1,1,1,2,2-pentafluoroe... | 12.856 | 51 | 94143 | 9.936 | 9.936 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 9.936 |  |  | 100.5% | [16 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 97677 | 10.26 | 10.26 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.26 |  |  | 103.7% | [16 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 101499 | 10.594 | 10.594 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.594 |  |  | 107.1% | [16 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 103038 | 10.662 | 10.662 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.662 |  |  | 107.8% | [16 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 100169 | 10.73 | 10.73 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.73 |  |  | 108.5% | [16 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 97951 | 10.525 | 10.525 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.525 |  |  | 106.4% | [16 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 89435 | 10.342 | 10.342 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.342 |  |  | 104.6% | [16 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 93653 | 10.716 | 10.716 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.716 |  |  | 108.4%
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | 100.5% | 103.7% | 107.1% | 107.8% | 108.5% | 106.4% | 104.6% | 108.4% | [16 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 325747 | 10.037 | 10.037 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.037 |  |  | 103.3% | [16 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 329640 | 10.108 | 10.108 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.108 |  |  | 104% | [16 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 339554 | 10.347 | 10.347 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.347 |  |  | 106.5% | [16 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 342381 | 10.343 | 10.343 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.343 |  |  | 106.4% | [16 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 331990 | 10.382 | 10.382 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.382 |  |  | 106.8% | [16 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 326371 | 10.238 | 10.238 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.238 |  |  | 105.3% | [16 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 301091 | 10.164 | 10.164 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.164 |  |  | 104.6% | [16 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 309157 | 10.327 | 10.327 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.327 |  |  | 106.2%
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | 103.3% | 104% | 106.5% | 106.4% | 106.8% | 105.3% | 104.6% | 106.2% | [16 empty] | fluoromethane | 11.409 | 33 | 11147 | 10.097 | 10.097 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.097 |  |  | 102.1% | [16 empty] | fluoromethane | 11.409 | 33 | 11540m | 10.403 | 10.403 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.403 |  |  | 105.2% | [16 empty] | fluoromethane | 11.409 | 33 | 11952 | 10.707 | 10.707 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.707 |  |  | 108.3% | [16 empty] | fluoromethane | 11.409 | 33 | 12430 | 11.039 | 11.039 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 11.039 |  |  | 111.6% | [16 empty] | fluoromethane | 11.41 | 33 | 12348 | 11.352 | 11.352 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 11.352 |  |  | 114.8% | [16 empty] | fluoromethane | 11.41 | 33 | 11918 | 10.99 | 10.99 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.99 |  |  | 111.1% | [16 empty] | fluoromethane | 11.409 | 33 | 10542 | 10.462 | 10.462 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.462 |  |  | 105.8% | [16 empty] | fluoromethane | 11.409 | 33 | 11197 | 10.996 | 10.996 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.996 |  |  | 111.2%
Fluoromethane (HFC-41) | 102.1% | 105.2% | 108.3% | 111.6% | 114.8% | 111.1% | 105.8% | 111.2% | [16 empty] | tetrafluoroethylene | 8.491 | 81 | 71716 | 9.834 | 9.834 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 9.834 |  |  | 99.94% | [16 empty] | tetrafluoroethylene | 8.491 | 81 | 72540 | 9.899 | 9.899 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 9.899 |  |  | 100.6% | [16 empty] | tetrafluoroethylene | 8.503 | 81 | 74938 | 10.162 | 10.162 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.162 |  |  | 103.3% | [16 empty] | tetrafluoroethylene | 8.491 | 81 | 75945 | 10.21 | 10.21 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.21 |  |  | 103.8% | [16 empty] | tetrafluoroethylene | 8.491 | 81 | 73373 | 10.211 | 10.211 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.211 |  |  | 103.8% | [16 empty] | tetrafluoroethylene | 8.491 | 81 | 72484 | 10.118 | 10.118 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.118 |  |  | 102.8% | [16 empty] | tetrafluoroethylene | 8.491 | 81 | 65773 | 9.881 | 9.881 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 9.881 |  |  | 100.4% | [16 empty] | tetrafluoroethylene | 8.491 | 81 | 67931 | 10.099 | 10.099 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.099 |  |  | 102.6%
Tetrafluoroethylene | 99.94% | 100.6% | 103.3% | 103.8% | 103.8% | 102.8% | 100.4% | 102.6% | [16 empty] | hexafluoropropene | 13.247 | 131 | 154996 | 9.594 | 9.594 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.594 |  |  | 98.2% | [16 empty] | hexafluoropropene | 13.248 | 131 | 154380 | 9.51 | 9.51 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.51 |  |  | 97.34% | [16 empty] | hexafluoropropene | 13.248 | 131 | 158162 | 9.682 | 9.682 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.682 |  |  | 99.1% | [16 empty] | hexafluoropropene | 13.248 | 131 | 157747 | 9.573 | 9.573 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.573 |  |  | 97.98% | [16 empty] | hexafluoropropene | 13.248 | 131 | 153597 | 9.649 | 9.649 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.649 |  |  | 98.76% | [16 empty] | hexafluoropropene | 13.248 | 131 | 147789 | 9.313 | 9.313 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.313 |  |  | 95.32% | [16 empty] | hexafluoropropene | 13.247 | 131 | 140498 | 9.528 | 9.528 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.528 |  |  | 97.52% | [16 empty] | hexafluoropropene | 13.248 | 131 | 141708 | 9.51 | 9.51 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.51 |  |  | 97.34%
Hexafluoropropene | 98.2% | 97.34% | 99.1% | 97.98% | 98.76% | 95.32% | 97.52% | 97.34% | [16 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9640 | 9.788 | 9.788 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.788 |  |  | 100.5% | [16 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9732 | 9.834 | 9.834 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.834 |  |  | 101% | [16 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 10109 | 10.151 | 10.151 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 10.151 |  |  | 104.2% | [16 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9975 | 9.93 | 9.93 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.93 |  |  | 102% | [16 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9641 | 9.935 | 9.935 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.935 |  |  | 102% | [16 empty] | 1,1,1-trifluoroethane | 13.219 | 64 | 9578 | 9.901 | 9.901 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.901 |  |  | 101.7% | [16 empty] | 1,1,1-trifluoroethane | 13.218 | 64 | 8909 | 9.911 | 9.911 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.911 |  |  | 101.8% | [16 empty] | 1,1,1-trifluoroethane | 13.219 | 64 | 9070 | 9.984 | 9.984 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.984 |  |  | 102.5%
1,1,1-trifluoroethane | 100.5% | 101% | 104.2% | 102% | 102% | 101.7% | 101.8% | 102.5% | [16 empty] | hexafluoropropene_oxide | 13.435 | 69 | 10939 | 10.52 | 10.52 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 10.52 |  |  | 106.9% | [16 empty] | hexafluoropropene_oxide | 13.436 | 69 | 10233 | 9.794 | 9.794 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.794 |  |  | 99.53% | [16 empty] | hexafluoropropene_oxide | 13.436 | 69 | 11453 | 10.893 | 10.893 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 10.893 |  |  | 110.7% | [16 empty] | hexafluoropropene_oxide | 13.436 | 69 | 11018 | 10.389 | 10.389 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 10.389 |  |  | 105.6% | [16 empty] | hexafluoropropene_oxide | 13.436 | 69 | 11964 | 11.678 | 11.678 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 11.678 |  |  | 118.7% | [16 empty] | hexafluoropropene_oxide | 13.421 | 69 | 9060m | 8.871 | 8.871 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 8.871 |  |  | 90.15% | [16 empty] | hexafluoropropene_oxide | 13.421 | 69 | 9101m | 9.59 | 9.59 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.59 |  |  | 97.46% | [16 empty] | hexafluoropropene_oxide | 13.421 | 69 | 9160 | 9.551 | 9.551 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.551 |  |  | 97.06%
Hexafluoropropene Oxide | 106.9% | 99.53% | 110.7% | 105.6% | 118.7% | 90.15% | 97.46% | 97.06% | [16 empty] | chlorodifluoromethane | 13.739 | 51 | 95589 | 9.911 | 9.911 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 9.911 |  |  | 100.2% | [16 empty] | chlorodifluoromethane | 13.74 | 51 | 99702 | 10.288 | 10.288 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.288 |  |  | 104% | [16 empty] | chlorodifluoromethane | 13.74 | 51 | 102717 | 10.533 | 10.533 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.533 |  |  | 106.5% | [16 empty] | chlorodifluoromethane | 13.74 | 51 | 104406 | 10.614 | 10.614 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.614 |  |  | 107.3% | [16 empty] | chlorodifluoromethane | 13.725 | 51 | 101004 | 10.629 | 10.629 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.629 |  |  | 107.5% | [16 empty] | chlorodifluoromethane | 13.725 | 51 | 99617 | 10.515 | 10.515 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.515 |  |  | 106.3% | [16 empty] | chlorodifluoromethane | 13.725 | 51 | 91033 | 10.341 | 10.341 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.341 |  |  | 104.6% | [16 empty] | chlorodifluoromethane | 13.725 | 51 | 94482 | 10.621 | 10.621 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.621 |  |  | 107.4%
Chlorodifluoromethane (HCFC-22) | 100.2% | 104% | 106.5% | 107.3% | 107.5% | 106.3% | 104.6% | 107.4% | [16 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 31829 | 9.982 | 9.982 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 9.982 |  |  | 100.9% | [16 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 33493 | 10.454 | 10.454 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.454 |  |  | 105.7% | [16 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 35066 | 10.876 | 10.876 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.876 |  |  | 110% | [16 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 35875 | 11.031 | 11.031 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 11.031 |  |  | 111.5% | [16 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 34201 | 10.886 | 10.886 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.886 |  |  | 110.1% | [16 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 33713 | 10.764 | 10.764 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.764 |  |  | 108.8% | [16 empty] | 1,1,1,2-tetrafluoroethane | 15.911 | 33 | 30445 | 10.461 | 10.461 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.461 |  |  | 105.8% | [16 empty] | 1,1,1,2-tetrafluoroethane | 15.911 | 33 | 32270 | 10.972 | 10.972 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.972 |  |  | 110.9%
1,1,1,2-Tetrafluoroethane | 100.9% | 105.7% | 110% | 111.5% | 110.1% | 108.8% | 105.8% | 110.9% | [16 empty] | decafluorobutane | 16.161 | 119 | 169526 | 9.801 | 9.801 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 9.801 |  |  | 100.6% | [16 empty] | decafluorobutane | 16.175 | 119 | 173271 | 9.97 | 9.97 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 9.97 |  |  | 102.4% | [16 empty] | decafluorobutane | 16.175 | 119 | 176708 | 10.104 | 10.104 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.104 |  |  | 103.7% | [16 empty] | decafluorobutane | 16.161 | 119 | 178800 | 10.135 | 10.135 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.135 |  |  | 104.1% | [16 empty] | decafluorobutane | 16.161 | 119 | 174318 | 10.229 | 10.229 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.229 |  |  | 105% | [16 empty] | decafluorobutane | 16.161 | 119 | 170042 | 10.008 | 10.008 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.008 |  |  | 102.8% | [16 empty] | decafluorobutane | 16.161 | 119 | 160648 | 10.176 | 10.176 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.176 |  |  | 104.5% | [16 empty] | decafluorobutane | 16.161 | 119 | 164011 | 10.28 | 10.28 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.28 |  |  | 105.5%
Decafluorobutane (Perfluorobutane) | 100.6% | 102.4% | 103.7% | 104.1% | 105% | 102.8% | 104.5% | 105.5% | [16 empty] | 1H-Heptafluoropropane | 18.112 | 51 | 140153 | 9.828 | 9.828 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 9.828 |  |  | 99.88% | [16 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 147074 | 10.264 | 10.264 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.264 |  |  | 104.3% | [16 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 150915 | 10.467 | 10.467 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.467 |  |  | 106.4% | [16 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 157604 | 10.836 | 10.836 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.836 |  |  | 110.1% | [16 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 150470 | 10.71 | 10.71 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.71 |  |  | 108.8% | [16 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 151320 | 10.803 | 10.803 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.803 |  |  | 109.8% | [16 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 135104 | 10.38 | 10.38 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.38 |  |  | 105.5% | [16 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 142107 | 10.804 | 10.804 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.804 |  |  | 109.8%
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | 99.88% | 104.3% | 106.4% | 110.1% | 108.8% | 109.8% | 105.5% | 109.8% | [16 empty] | octafluorocyclopentene | 21.943 | 93 | 235023 | 9.957 | 9.957 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 9.957 |  |  | 102.2% | [16 empty] | octafluorocyclopentene | 21.943 | 93 | 240403 | 10.136 | 10.136 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.136 |  |  | 104.1% | [16 empty] | octafluorocyclopentene | 21.943 | 93 | 244195 | 10.232 | 10.232 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.232 |  |  | 105.1% | [16 empty] | octafluorocyclopentene | 21.943 | 93 | 249672 | 10.371 | 10.371 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.371 |  |  | 106.5% | [16 empty] | octafluorocyclopentene | 21.943 | 93 | 245432 | 10.553 | 10.553 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.553 |  |  | 108.3% | [16 empty] | octafluorocyclopentene | 21.931 | 93 | 236496 | 10.2 | 10.2 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.2 |  |  | 104.7% | [16 empty] | octafluorocyclopentene | 21.931 | 93 | 223215 | 10.361 | 10.361 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.361 |  |  | 106.4% | [16 empty] | octafluorocyclopentene | 21.931 | 93 | 228756 | 10.507 | 10.507 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.507 |  |  | 107.9%
Octafluorocyclopentene (FC-C1418) | 102.2% | 104.1% | 105.1% | 106.5% | 108.3% | 104.7% | 106.4% | 107.9% | [16 empty] | trichlorofluoromethane | 21.239 | 101 | 235289 | 9.971 | 9.971 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 9.971 |  |  | 101.1% | [16 empty] | trichlorofluoromethane | 21.239 | 101 | 240278 | 10.134 | 10.134 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.134 |  |  | 102.8% | [16 empty] | trichlorofluoromethane | 21.239 | 101 | 246979 | 10.352 | 10.352 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.352 |  |  | 105% | [16 empty] | trichlorofluoromethane | 21.239 | 101 | 248487 | 10.325 | 10.325 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.325 |  |  | 104.7% | [16 empty] | trichlorofluoromethane | 21.24 | 101 | 245935 | 10.578 | 10.578 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.578 |  |  | 107.3% | [16 empty] | trichlorofluoromethane | 21.24 | 101 | 237303 | 10.239 | 10.239 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.239 |  |  | 103.8% | [16 empty] | trichlorofluoromethane | 21.239 | 101 | 227323 | 10.555 | 10.555 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.555 |  |  | 107% | [16 empty] | trichlorofluoromethane | 21.239 | 101 | 232544 | 10.685 | 10.685 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.685 |  |  | 108.4%
Trichlorofluoromethane (CFC-11) | 101.1% | 102.8% | 105% | 104.7% | 107.3% | 103.8% | 107% | 108.4% | [16 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 549309 | 10.213 | 10.213 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.213 |  |  | 103.6% | [16 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 571942 | 10.583 | 10.583 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.583 |  |  | 107.3% | [16 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 581564 | 10.694 | 10.694 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.694 |  |  | 108.5% | [16 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 587590m | 10.712 | 10.712 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.712 |  |  | 108.6% | [16 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 591803 | 11.168 | 11.168 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 11.168 |  |  | 113.3% | [16 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 584950 | 11.073 | 11.073 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 11.073 |  |  | 112.3% | [16 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 529884 | 10.794 | 10.794 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.794 |  |  | 109.5% | [16 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 557251 | 11.233 | 11.233 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 11.233 |  |  | 113.9%
Dodecafluoro-n-pentane | 103.6% | 107.3% | 108.5% | 108.6% | 113.3% | 112.3% | 109.5% | 113.9% | [16 empty] | 1H-Nonafluorobutane | 22.801 | 51 | 222716 | 10.484 | 10.484 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.484 |  |  | 105.8% | [16 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 233362 | 10.933 | 10.933 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.933 |  |  | 110.3% | [16 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 231808 | 10.792 | 10.792 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.792 |  |  | 108.9% | [16 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 241608 | 11.151 | 11.151 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.151 |  |  | 112.5% | [16 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 236240 | 11.287 | 11.287 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.287 |  |  | 113.9% | [16 empty] | 1H-Nonafluorobutane | 22.79 | 51 | 229369 | 10.993 | 10.993 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.993 |  |  | 110.9% | [16 empty] | 1H-Nonafluorobutane | 22.789 | 51 | 214533 | 11.065 | 11.065 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.065 |  |  | 111.7% | [16 empty] | 1H-Nonafluorobutane | 22.79 | 51 | 220921 | 11.275 | 11.275 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.275 |  |  | 113.8%
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | 105.8% | 110.3% | 108.9% | 112.5% | 113.9% | 110.9% | 111.7% | 113.8% | [16 empty] | tetradecafluorohexane | 24.455 | 69 | 563397m | 9.847 | 9.847 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 9.847 |  |  | 100.2% | [16 empty] | tetradecafluorohexane | 24.456 | 69 | 580547m | 10.098 | 10.098 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.098 |  |  | 102.7% | [16 empty] | tetradecafluorohexane | 24.456 | 69 | 596577m | 10.312 | 10.312 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.312 |  |  | 104.9% | [16 empty] | tetradecafluorohexane | 24.416 | 69 | 601255m | 10.303 | 10.303 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.303 |  |  | 104.8% | [16 empty] | tetradecafluorohexane | 24.408 | 69 | 609180m | 10.807 | 10.807 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.807 |  |  | 109.9% | [16 empty] | tetradecafluorohexane | 24.408 | 69 | 595209m | 10.591 | 10.591 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.591 |  |  | 107.7% | [16 empty] | tetradecafluorohexane | 24.392 | 69 | 543626m | 10.41 | 10.41 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.41 |  |  | 105.9% | [16 empty] | tetradecafluorohexane | 24.353 | 69 | 556165m | 10.539 | 10.539 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.539 |  |  | 107.2%
Tetradecafluorohexane | 100.2% | 102.7% | 104.9% | 104.8% | 109.9% | 107.7% | 105.9% | 107.2% | [16 empty] | 1H-Perfluoropentane | 26.105 | 69 | 186847 | 10.134 | 10.134 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.134 |  |  | 102.4% | [16 empty] | 1H-Perfluoropentane | 26.105 | 69 | 196289 | 10.595 | 10.595 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.595 |  |  | 107% | [16 empty] | 1H-Perfluoropentane | 26.095 | 69 | 190013m | 10.192 | 10.192 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.192 |  |  | 102.9% | [16 empty] | 1H-Perfluoropentane | 26.076 | 69 | 204588m | 10.879 | 10.879 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.879 |  |  | 109.9% | [16 empty] | 1H-Perfluoropentane | 26.066 | 69 | 208076m | 11.454 | 11.454 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 11.454 |  |  | 115.7% | [16 empty] | 1H-Perfluoropentane | 26.066 | 69 | 194714m | 10.751 | 10.751 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.751 |  |  | 108.6% | [16 empty] | 1H-Perfluoropentane | 26.056 | 69 | 180190m | 10.707 | 10.707 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.707 |  |  | 108.2% | [16 empty] | 1H-Perfluoropentane | 26.027 | 69 | 185143m | 10.887 | 10.887 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.887 |  |  | 110%
1,1,1,2,2,3,3,4,4,5,5-Undecafluoropentane (1H-Perfluoropentane) | 102.4% | 107% | 102.9% | 109.9% | 115.7% | 108.6% | 108.2% | 110% | [16 empty] | heptafluoropropyl 1,2,... | 26.173 | 101 | 466707m | 9.955 | 9.955 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 9.955 |  |  | 101.3% | [16 empty] | heptafluoropropyl 1,2,... | 26.173 | 101 | 467901m | 9.933 | 9.933 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 9.933 |  |  | 101% | [16 empty] | heptafluoropropyl 1,2,... | 26.173 | 101 | 489920m | 10.335 | 10.335 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.335 |  |  | 105.1% | [16 empty] | heptafluoropropyl 1,2,... | 26.154 | 101 | 489443m | 10.236 | 10.236 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.236 |  |  | 104.1% | [16 empty] | heptafluoropropyl 1,2,... | 26.144 | 101 | 499494m | 10.814 | 10.814 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.814 |  |  | 110% | [16 empty] | heptafluoropropyl 1,2,... | 26.144 | 101 | 496995m | 10.793 | 10.793 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.793 |  |  | 109.8% | [16 empty] | heptafluoropropyl 1,2,... | 26.124 | 101 | 455105m | 10.636 | 10.636 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.636 |  |  | 108.2% | [16 empty] | heptafluoropropyl 1,2,... | 26.105 | 101 | 464944m | 10.753 | 10.753 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.753 |  |  | 109.4%
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | 101.3% | 101% | 105.1% | 104.1% | 110% | 109.8% | 108.2% | 109.4% | [16 empty] | hexadecafluoroheptane | 27.275 | 69 | 673979 | 10.006 | 10.006 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.006 |  |  | 101.3% | [16 empty] | hexadecafluoroheptane | 27.285 | 69 | 691987 | 10.225 | 10.225 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.225 |  |  | 103.5% | [16 empty] | hexadecafluoroheptane | 27.275 | 69 | 711695 | 10.45 | 10.45 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.45 |  |  | 105.8% | [16 empty] | hexadecafluoroheptane | 27.265 | 69 | 714489 | 10.4 | 10.4 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.4 |  |  | 105.3% | [16 empty] | hexadecafluoroheptane | 27.255 | 69 | 704131 | 10.61 | 10.61 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.61 |  |  | 107.4% | [16 empty] | hexadecafluoroheptane | 27.255 | 69 | 686926 | 10.383 | 10.383 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.383 |  |  | 105.1% | [16 empty] | hexadecafluoroheptane | 27.245 | 69 | 640340 | 10.416 | 10.416 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.416 |  |  | 105.4% | [16 empty] | hexadecafluoroheptane | 27.226 | 69 | 661489 | 10.648 | 10.648 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.648 |  |  | 107.8%
Hexadecafluoroheptane | 101.3% | 103.5% | 105.8% | 105.3% | 107.4% | 105.1% | 105.4% | 107.8% | [16 empty] | 1H-Perfluorohexane | 28.747 | 51 | 325895 | 10.136 | 10.136 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.136 |  |  | 102.8% | [16 empty] | 1H-Perfluorohexane | 28.757 | 51 | 337248 | 10.439 | 10.439 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.439 |  |  | 105.9% | [16 empty] | 1H-Perfluorohexane | 28.737 | 51 | 347360 | 10.685 | 10.685 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.685 |  |  | 108.4% | [16 empty] | 1H-Perfluorohexane | 28.737 | 51 | 350739 | 10.696 | 10.696 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.696 |  |  | 108.5% | [16 empty] | 1H-Perfluorohexane | 28.718 | 51 | 343655 | 10.848 | 10.848 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.848 |  |  | 110% | [16 empty] | 1H-Perfluorohexane | 28.718 | 51 | 337809 | 10.696 | 10.696 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.696 |  |  | 108.5% | [16 empty] | 1H-Perfluorohexane | 28.708 | 51 | 309545 | 10.548 | 10.548 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.548 |  |  | 107% | [16 empty] | 1H-Perfluorohexane | 28.688 | 51 | 320709 | 10.814 | 10.814 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.814 |  |  | 109.7%
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | 102.8% | 105.9% | 108.4% | 108.5% | 110% | 108.5% | 107% | 109.7% | [16 empty] | octadecafluorooctane | 29.537 | 119 | 409773 | 9.813 | 9.813 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.813 |  |  | 100% | [16 empty] | octadecafluorooctane | 29.548 | 119 | 407237 | 9.706 | 9.706 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.706 |  |  | 98.94% | [16 empty] | octadecafluorooctane | 29.537 | 119 | 417927 | 9.898 | 9.898 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.898 |  |  | 100.9% | [16 empty] | octadecafluorooctane | 29.537 | 119 | 417977 | 9.814 | 9.814 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.814 |  |  | 100% | [16 empty] | octadecafluorooctane | 29.514 | 119 | 408181 | 9.921 | 9.921 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.921 |  |  | 101.1% | [16 empty] | octadecafluorooctane | 29.514 | 119 | 400089 | 9.754 | 9.754 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.754 |  |  | 99.43% | [16 empty] | octadecafluorooctane | 29.514 | 119 | 372714 | 9.779 | 9.779 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.779 |  |  | 99.68% | [16 empty] | octadecafluorooctane | 29.502 | 119 | 377638 | 9.805 | 9.805 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.805 |  |  | 99.95%
Octadecafluorooctane (Perfluorooctane) | 100% | 98.94% | 100.9% | 100% | 101.1% | 99.43% | 99.68% | 99.95% | [16 empty] | 1H-Perfluoroheptane | 31.066 | 51 | 894988 | 9.994 | 9.994 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 9.994 |  |  | 101.2% | [16 empty] | 1H-Perfluoroheptane | 31.09 | 51 | 923485 | 10.263 | 10.263 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.263 |  |  | 103.9% | [16 empty] | 1H-Perfluoroheptane | 31.067 | 51 | 951905 | 10.513 | 10.513 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.513 |  |  | 106.4% | [16 empty] | 1H-Perfluoroheptane | 31.078 | 51 | 961481 | 10.527 | 10.527 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.527 |  |  | 106.5% | [16 empty] | 1H-Perfluoroheptane | 31.055 | 51 | 933689 | 10.582 | 10.582 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.582 |  |  | 107.1% | [16 empty] | 1H-Perfluoroheptane | 31.055 | 51 | 921534 | 10.477 | 10.477 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.477 |  |  | 106% | [16 empty] | 1H-Perfluoroheptane | 31.055 | 51 | 825974 | 10.106 | 10.106 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.106 |  |  | 102.3% | [16 empty] | 1H-Perfluoroheptane | 31.044 | 51 | 855374 | 10.356 | 10.356 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.356 |  |  | 104.8%
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | 101.2% | 103.9% | 106.4% | 106.5% | 107.1% | 106% | 102.3% | 104.8% | [16 empty] | 1H-Perfluorooctane | 33.482 | 51 | 688094 | 10.024 | 10.024 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.024 |  |  | 102.2% | [16 empty] | 1H-Perfluorooctane | 33.505 | 51 | 704508 | 10.214 | 10.214 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.214 |  |  | 104.1% | [16 empty] | 1H-Perfluorooctane | 33.482 | 51 | 722734 | 10.413 | 10.413 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.413 |  |  | 106.1% | [16 empty] | 1H-Perfluorooctane | 33.505 | 51 | 726063 | 10.371 | 10.371 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.371 |  |  | 105.7% | [16 empty] | 1H-Perfluorooctane | 33.482 | 51 | 698436 | 10.327 | 10.327 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.327 |  |  | 105.3% | [16 empty] | 1H-Perfluorooctane | 33.47 | 51 | 687014 | 10.189 | 10.189 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.189 |  |  | 103.9% | [16 empty] | 1H-Perfluorooctane | 33.47 | 51 | 612025 | 9.769 | 9.769 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 9.769 |  |  | 99.58% | [16 empty] | 1H-Perfluorooctane | 33.47 | 51 | 627515 | 9.911 | 9.911 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 9.911 |  |  | 101%
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | 102.2% | 104.1% | 106.1% | 105.7% | 105.3% | 103.9% | 99.58% | 101% | [16 empty] | 2H-peffluoro-5-methyl-... | 33.033 | 169 | 1860293 | 9.797 | 9.797 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.797 |  |  | 99.16% | [16 empty] | 2H-peffluoro-5-methyl-... | 33.056 | 169 | 1829586 | 9.589 | 9.589 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.589 |  |  | 97.05% | [16 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 1848463 | 9.628 | 9.628 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.628 |  |  | 97.45% | [16 empty] | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 1820840 | 9.402 | 9.402 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.402 |  |  | 95.16% | [16 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 1765997 | 9.44 | 9.44 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.44 |  |  | 95.55% | [16 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 1681240 | 9.014 | 9.014 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.014 |  |  | 91.23% | [16 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 1589566 | 9.172 | 9.172 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.172 |  |  | 92.83% | [16 empty] | 2H-peffluoro-5-methyl-... | 33.01 | 169 | 1561311 | 8.915 | 8.915 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 8.915 |  |  | 90.23%
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | 99.16% | 97.05% | 97.45% | 95.16% | 95.55% | 91.23% | 92.83% | 90.23%

## 112524
112524/112524_01.D/epatemp.txt (QT Reviewed): AirBlank_Vol 200 ml | [14 empty] | 112524/112524_02.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | 112524/112524_03.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | 112524/112524_04.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | 112524/112524_05.D/epatemp.txt (QT Reviewed): AirBlank_Vol 200 ml | [14 empty] | 112524/112524_06.D/epatemp.txt (QT eviewed): Riemer PFASCF4t_Calck_100 ppbv_vol 20 ml | [14 empty] | CF4 Method |  | Lower Limit | Uppper Limit | Rest Method |  | Lower Limit | Uppper Limit
Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Average IS area | 1053174 | 631904.4 | 1474443.6 | Average IS area | 1032225 | 619335 | 1445115
1,4-difluorobenzene | 30.308 | 114 | 1426220 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.411 | 114 | 1345627 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.411 | 114 | 1149560 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.411 | 114 | 1032225 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.331 | 114 | 1075612 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.503 | 114 | 1053174 | 50 |  | ppb | [6 empty] | Pass | Pass | Average RT | 30.503 | 30.4197 | 30.5863 | Average RT | 30.411 | 30.3277 | 30.4943
hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Fail | hexafluoroethane | 8.093 | 119 | 207856 | 9.733 | 9.733 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.733 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 177155 | 9.711 | 9.711 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.711 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 163304 | 9.969 | 9.969 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.969 | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 6.994 | 69 | 79567 | 99.425 | 99.425 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 99.425 | Pass | Pass
chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL |  |  | chlorotrifluoromethane | 8.515 | 85 | 53880 | 9.843 | 9.843 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.843 |  |  | chlorotrifluoromethane | 8.503 | 85 | 45796 | 9.793 | 9.793 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.793 |  |  | chlorotrifluoromethane | 8.515 | 85 | 42658 | 10.159 | 10.159 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.159 |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL
trifluoromethane | 8.721 | 69 | 319 | 0.069 | 0.069 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 0.069 |  |  | trifluoromethane | 8.708 | 69 | 46880 | 11.172 | 11.172 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 11.172 |  |  | trifluoromethane | 8.708 | 69 | 39430 | 11 | 11 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 11 |  |  | trifluoromethane | 8.708 | 69 | 37013 | 11.499 | 11.499 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 11.499 |  |  | trifluoromethane | 8.697 | 69 | 218m | 0.065 | 0.065 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 0.065
octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL |  |  | octafluoropropane | 11.178 | 69 | 415828 | 10.732 | 10.732 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.732 |  |  | octafluoropropane | 11.178 | 69 | 358919 | 10.843 | 10.843 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.843 |  |  | octafluoropropane | 11.178 | 69 | 328755 | 11.06 | 11.06 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 11.06 |  |  | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL
difluoromethane | 10.99 | 51 | 792 | 0.248 | 0.248 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 0.248 |  |  | difluoromethane | 10.975 | 51 | 30844 | 10.73 | 10.73 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.73 |  |  | difluoromethane | 10.961 | 51 | 25767 | 10.493 | 10.493 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.493 |  |  | difluoromethane | 10.975 | 51 | 24588 | 11.151 | 11.151 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 11.151 |  |  | difluoromethane | 10.932 | 51 | 258 | 0.112 | 0.112 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 0.112
1,1,1,2,2-pentafluoroe... | 12.857 | 51 | 1578 | 0.131 | 0.131 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.131 |  |  | 1,1,1,2,2-pentafluoroe... | 12.857 | 51 | 117683 | 10.903 | 10.903 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.903 |  |  | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 102172 | 11.081 | 11.081 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 11.081 |  |  | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 93151 | 11.251 | 11.251 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 11.251 |  |  | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 773 | 0.09 | 0.09 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.09
1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 414701 | 10.089 | 10.089 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.089 |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 349924 | 9.965 | 9.965 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 9.965 |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 321666 | 10.201 | 10.201 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.201 |  |  | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL
fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL |  |  | fluoromethane | 11.409 | 33 | 13879 | 10.456 | 10.456 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.456 |  |  | fluoromethane | 11.409 | 33 | 11797 | 10.403 | 10.403 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.403 |  |  | fluoromethane | 11.409 | 33 | 10854 | 10.659 | 10.659 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.659 |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL
tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL |  |  | tetrafluoroethylene | 8.491 | 81 | 89953 | 10.384 | 10.384 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.384 |  |  | tetrafluoroethylene | 8.491 | 81 | 77001 | 10.405 | 10.405 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.405 |  |  | tetrafluoroethylene | 8.491 | 81 | 71338 | 10.735 | 10.735 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.735 |  |  | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL
hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL |  |  | hexafluoropropene | 13.248 | 131 | 200848 | 9.353 | 9.353 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.353 |  |  | hexafluoropropene | 13.247 | 131 | 172475 | 9.401 | 9.401 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.401 |  |  | hexafluoropropene | 13.247 | 131 | 156908 | 9.525 | 9.525 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.525 |  |  | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL
1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL |  |  | 1,1,1-trifluoroethane | 13.233 | 64 | 12043 | 9.753 | 9.753 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.753 |  |  | 1,1,1-trifluoroethane | 13.233 | 64 | 10254 | 9.72 | 9.72 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.72 |  |  | 1,1,1-trifluoroethane | 13.233 | 64 | 9536 | 10.067 | 10.067 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 10.067 |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL
hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL |  |  | hexafluoropropene_oxide | 13.436 | 69 | 11684 | 9.074 | 9.074 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.074 |  |  | hexafluoropropene_oxide | 13.435 | 69 | 11511 | 10.464 | 10.464 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 10.464 |  |  | hexafluoropropene_oxide | 13.435 | 69 | 10171 | 10.297 | 10.297 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 10.297 |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL
chlorodifluoromethane | 13.74 | 51 | 5513 | 0.442 | 0.442 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 0.442 |  |  | chlorodifluoromethane | 13.74 | 51 | 120369 | 10.814 | 10.814 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.814 |  |  | chlorodifluoromethane | 13.739 | 51 | 103332 | 10.867 | 10.867 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.867 |  |  | chlorodifluoromethane | 13.739 | 51 | 94820 | 11.105 | 11.105 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 11.105 |  |  | chlorodifluoromethane | 13.711 | 51 | 2737 | 0.308 | 0.308 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 0.308
1,1,1,2-tetrafluoroethane | 15.953 | 33 | 1145 | 0.261 | 0.261 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 0.261 |  |  | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 39097 | 10.079 | 10.079 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.079 |  |  | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 34344 | 10.364 | 10.364 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.364 |  |  | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 31349 | 10.535 | 10.535 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.535 |  |  | 1,1,1,2-tetrafluoroethane | 15.911 | 33 | 863 | 0.278 | 0.278 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 0.278
decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL |  |  | decafluorobutane | 16.161 | 119 | 217661 | 10.044 | 10.044 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.044 |  |  | decafluorobutane | 16.161 | 119 | 187807 | 10.144 | 10.144 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.144 |  |  | decafluorobutane | 16.161 | 119 | 170053 | 10.229 | 10.229 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.229 |  |  | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL
1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL |  |  | 1H-Heptafluoropropane | 18.113 | 51 | 177111 | 11.136 | 11.136 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 11.136 |  |  | 1H-Heptafluoropropane | 18.112 | 51 | 154140 | 11.344 | 11.344 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 11.344 |  |  | 1H-Heptafluoropropane | 18.112 | 51 | 139487 | 11.433 | 11.433 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 11.433 |  |  | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL
octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL |  |  | octafluorocyclopentene | 21.943 | 93 | 293734 | 9.822 | 9.822 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 9.822 |  |  | octafluorocyclopentene | 21.943 | 93 | 255837 | 10.014 | 10.014 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.014 |  |  | octafluorocyclopentene | 21.943 | 93 | 228540 | 9.963 | 9.963 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 9.963 |  |  | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL
trichlorofluoromethane | 21.257 | 101 | 8900 | 0.257 | 0.257 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 0.257 |  |  | trichlorofluoromethane | 21.239 | 101 | 296707 | 9.283 | 9.283 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 9.283 |  |  | trichlorofluoromethane | 21.239 | 101 | 259519 | 9.505 | 9.505 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 9.505 |  |  | trichlorofluoromethane | 21.239 | 101 | 231281 | 9.433 | 9.433 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 9.433 |  |  | trichlorofluoromethane | 21.222 | 101 | 6526 | 0.255 | 0.255 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 0.255
dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL |  |  | dodecafluoro-n-pentane | 21.012 | 69 | 687016 | 10.741 | 10.741 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.741 |  |  | dodecafluoro-n-pentane | 21.012 | 69 | 596115 | 10.91 | 10.91 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.91 |  |  | dodecafluoro-n-pentane | 21.012 | 69 | 527160 | 10.744 | 10.744 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.744 |  |  | dodecafluoro-n-pentane | 20.978 | 69 | 1017m | 0.02 | 0.02 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 0.02
1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL |  |  | 1H-Nonafluorobutane | 22.802 | 51 | 267945 | 10.163 | 10.163 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.163 |  |  | 1H-Nonafluorobutane | 22.801 | 51 | 234513 | 10.412 | 10.412 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.412 |  |  | 1H-Nonafluorobutane | 22.801 | 51 | 209269 | 10.347 | 10.347 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.347 |  |  | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL
tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL |  |  | tetradecafluorohexane | 24.487 | 69 | 702066m | 10.226 | 10.226 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.226 |  |  | tetradecafluorohexane | 24.495 | 69 | 609531m | 10.393 | 10.393 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.393 |  |  | tetradecafluorohexane | 24.495 | 69 | 559082m | 10.616 | 10.616 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.616 |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL
1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL |  |  | 1H-Perfluoropentane | 26.115 | 69 | 245121 | 10.878 | 10.878 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.878 |  |  | 1H-Perfluoropentane | 26.124 | 69 | 202829 | 10.536 | 10.536 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.536 |  |  | 1H-Perfluoropentane | 26.124 | 69 | 181455 | 10.498 | 10.498 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.498 |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL
heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL |  |  | heptafluoropropyl 1,2,... | 26.193 | 101 | 572073m | 9.394 | 9.394 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 9.394 |  |  | heptafluoropropyl 1,2,... | 26.192 | 101 | 522464m | 10.042 | 10.042 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.042 |  |  | heptafluoropropyl 1,2,... | 26.192 | 101 | 458097m | 9.806 | 9.806 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 9.806 |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL
hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL |  |  | hexadecafluoroheptane | 27.294 | 69 | 848478 | 10.294 | 10.294 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.294 |  |  | hexadecafluoroheptane | 27.294 | 69 | 730649 | 10.376 | 10.376 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.376 |  |  | hexadecafluoroheptane | 27.294 | 69 | 661523 | 10.462 | 10.462 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.462 |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL
1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL |  |  | 1H-Perfluorohexane | 28.757 | 51 | 402330 | 10.284 | 10.284 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.284 |  |  | 1H-Perfluorohexane | 28.756 | 51 | 352464 | 10.546 | 10.546 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.546 |  |  | 1H-Perfluorohexane | 28.756 | 51 | 315140 | 10.501 | 10.501 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.501 |  |  | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL
octadecafluorooctane | 29.48 | 119 | 212m | 0.003 | 0.003 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL |  |  | octadecafluorooctane | 29.549 | 119 | 525704 | 9.538 | 9.538 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.538 |  |  | octadecafluorooctane | 29.548 | 119 | 447326 | 9.5 | 9.5 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.5 |  |  | octadecafluorooctane | 29.548 | 119 | 407222 | 9.632 | 9.632 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.632 |  |  | octadecafluorooctane | 29.491 | 119 | 296 | 0.007 | 0.007 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL
1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL |  |  | 1H-Perfluoroheptane | 31.067 | 51 | 1125832 | 10.26 | 10.26 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.26 |  |  | 1H-Perfluoroheptane | 31.078 | 51 | 960464 | 10.246 | 10.246 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.246 |  |  | 1H-Perfluoroheptane | 31.078 | 51 | 879515 | 10.449 | 10.449 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.449 |  |  | 1H-Perfluoroheptane | 31.067 | 51 | 881 | 0.01 | 0.01 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL
1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | BDL |  |  | 1H-Perfluorooctane | 33.482 | 51 | 858545 | 10.018 | 10.018 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.018 |  |  | 1H-Perfluorooctane | 33.493 | 51 | 733631 | 10.021 | 10.021 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.021 |  |  | 1H-Perfluorooctane | 33.482 | 51 | 674178 | 10.256 | 10.256 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.256 |  |  | 1H-Perfluorooctane | 33.505 | 51 | 4648 | 0.068 | 0.068 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.068
2H-peffluoro-5-methyl-... | 33.022 | 169 | 1205m | 0.004 | 0.004 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL |  |  | 2H-peffluoro-5-methyl-... | 33.033 | 169 | 2439449 | 9.05 | 9.05 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.05 |  |  | 2H-peffluoro-5-methyl-... | 33.033 | 169 | 2050460 | 8.905 | 8.905 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 8.905 |  |  | 2H-peffluoro-5-methyl-... | 33.033 | 169 | 1864919 | 9.02 | 9.02 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.02 |  |  | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 3703 | 0.017 | 0.017 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL

## 112624
112624/112624_01.D/epatemp.txt (QT Reviewed): AirBlank_Vol 200 ml | [14 empty] | 112624/112624_02.D/epatemp.txt (QT Reviewed): Rainbow BatchBlank_CF4_Vol 20 ml | [14 empty] | 112624/112624_03.D/epatemp.txt (QT Reviewed): Rainbow BatchBlank_Rest_Vol 200 ml | [14 empty] | 112624/112624_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_initial_100ppbv_vol 20 ml | [14 empty] | 112624/112624_05.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | 112624/112624_06.D/epatemp.txt (QT Reviewed): Rainbow Post_BatchBlankCK_Rest_vol 200 ml | [14 empty] | 112624/112624_07.D/epatemp.txt (QT Reviewed): Rainbow N2_Blank_vol 20 mL | [14 empty] | 112624/112624_08.D/epatemp.txt (QT Reviewed): Rainbow N2_Blank_vol 200 mL | [14 empty] | 112624/112624_09.D/epatemp.txt (QT Reviewed): Rainbow Combustion_Blank_vol 20 mL | [14 empty] | 112624/112624_10.D/epatemp.txt (QT Reviewed): Rainbow Combustion_Blank_vol 200 mL | [14 empty] | 112624/112624_11.D/epatemp.txt (QT Reviewed): Rainbow Run1_Port4_PFBA_Inj_20 mL | [14 empty] | 112624/112624_12.D/epatemp.txt (QT Reviewed): Rainbow Run1_Port4_PFBA_Inj_200 mL | [14 empty] | 112624/112624_13.D/epatemp.txt (QT Reviewed): Rainbow Run1_Port4_PFBA_Injj_dup_20 mL | [14 empty] | 112624/112624_14.D/epatemp.txt (QT Reviewed): Rainbow Run1_Port4_PFBA_Inj_dup_200 mL | [14 empty] | 112624/112624_15.D/epatemp.txt (QT Reviewed): 2.5% CO2_10ppbv_CF4_20 mL | [14 empty] | 112624/112624_16.D/epatemp.txt (QT Reviewed): 2.5% CO2_1ppbv_Rest_200 mL | [14 empty] | 112624/112624_17.D/epatemp.txt (QT Reviewed): 2.5% CO2_300ppbv_CF4_20 mL | [14 empty] | 112624/112624_18.D/epatemp.txt (QT Reviewed): 2.5% CO2_30ppbv_CF4_200 mL | [14 empty] | 112624/112624_19.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [14 empty] | 112624/112624_20.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [14 empty] | 112624/112624_21.D/epatemp.txt (QT Reviewed): Post BatchBlankCK_CF4 20 mL | [14 empty] | CF4 Method |  | Lower Limit | Uppper Limit | Rest Method |  | Lower Limit | Uppper Limit
Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Average IS area | 1053174 | 631904.4 | 1474443.6 | Average IS area | 1032225 | 619335 | 1445115
1,4-difluorobenzene | 30.331 | 114 | 1075517 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1060554 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.377 | 114 | 1069145 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1082936 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.399 | 114 | 1038373 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.365 | 114 | 1090229 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1084097 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.411 | 114 | 1079019 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1196013 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.411 | 114 | 1032577 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1246889 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.411 | 114 | 1010945 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1198180 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.397 | 114 | 1002742 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1251881 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.423 | 114 | 1063954 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1204137 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.423 | 114 | 1032864 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1072880 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.4 | 114 | 1043367 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.503 | 114 | 1090499 | 50 |  | ppb | [6 empty] | Pass | Pass | Average RT | 30.503 | 30.4197 | 30.5863 | Average RT | 30.411 | 30.3277 | 30.4943
hexafluoroethane | 8.093 | 119 | 95m | 0.005 | 0.005 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | BDL | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 6.994 | 69 | 78707 | 94.565 | 94.565 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 94.565 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 161695 | 9.636 | 9.636 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.636 | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.8438 | 2.8438 | 0.089699 | 0.25508 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 172m | 0.01 | 0.01 | ppb | 200 | 2.8438 | 2.8438 | 0.0084458 | 0.024018 | 0.028438 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.9041 | 2.9041 | 0.089699 | 0.26049 | BDL | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.0084458 | 0.024527 | BDL | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.9022 | 2.9022 | 0.089699 | 0.26033 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 7432m | 0.455 | 0.455 | ppb | 200 | 2.9022 | 2.9022 | 0.0084458 | 0.024512 | 1.3205 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.9022 | 2.9022 | 0.089699 | 0.26033 | BDL | Pass | Pass | hexafluoroethane | 8.105 | 119 | 7517m | 0.464 | 0.464 | ppb | 200 | 2.9022 | 2.9022 | 0.0084458 | 0.024512 | 1.3466 | Pass | Pass | tetrafluoromethane | 6.958 | 69 | 8080 | 8.398 | 8.398 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 8.398 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 15860 | 0.922 | 0.922 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 0.922 | Pass | Pass | tetrafluoromethane | 6.958 | 69 | 225778 | 243.964 | 243.964 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 243.964 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 433248 | 25.957 | 25.957 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 25.957 | Pass | Pass | tetrafluoromethane | 6.995 | 69 | 81010 | 98.244 | 98.244 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 98.244 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 162256 | 9.623 | 9.623 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.623 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | BDL | Pass | Pass
chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 8.515 | 85 | 42364 | 9.714 | 9.714 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.714 |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.029884 | 0.084984 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.029884 | 0.086786 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.029884 | 0.086731 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.029884 | 0.086731 | BDL | [17 empty] | chlorotrifluoromethane | 8.515 | 85 | 4531 | 1.014 | 1.014 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 1.014 | [17 empty] | chlorotrifluoromethane | 8.515 | 85 | 117921 | 27.184 | 27.184 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 27.184 | [17 empty] | chlorotrifluoromethane | 8.515 | 85 | 43052 | 9.825 | 9.825 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.825
trifluoromethane | 8.72 | 69 | 215m | 0.055 | 0.055 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 0.055 | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL | [17 empty] | trifluoromethane | 8.708 | 69 | 37287 | 9.874 | 9.874 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 9.874 |  |  | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.0356 | 0.10124 | BDL | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.0356 | 0.10338 | BDL | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0356 | 0.10332 | BDL | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0356 | 0.10332 | BDL | [17 empty] | trifluoromethane | 8.721 | 69 | 3888 | 1.005 | 1.005 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 1.005 | [17 empty] | trifluoromethane | 8.721 | 69 | 110549 | 29.431 | 29.431 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 29.431 | [17 empty] | trifluoromethane | 8.708 | 69 | 38023 | 10.021 | 10.021 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.021
octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 11.178 | 69 | 328369 | 9.721 | 9.721 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 9.721 |  |  | octafluoropropane | 11.178 | 69 | 534m | 0.015 | 0.015 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 11.178 | 69 | 1933m | 0.055 | 0.055 | ppb | 200 | 2.8438 | 2.8438 | 0.015822 | 0.044993 | 0.15641 | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.015822 | 0.045947 | BDL | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.015822 | 0.045918 | BDL | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.015822 | 0.045918 | BDL | [17 empty] | octafluoropropane | 11.178 | 69 | 35880 | 1.037 | 1.037 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 1.037 | [17 empty] | octafluoropropane | 11.178 | 69 | 931465 | 27.721 | 27.721 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 27.721 | [17 empty] | octafluoropropane | 11.178 | 69 | 339379 | 9.998 | 9.998 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 9.998
difluoromethane | 10.975 | 51 | 230m | 0.087 | 0.087 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 0.087 | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL | [17 empty] | difluoromethane | 10.975 | 51 | 25054 | 9.846 | 9.846 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 9.846 |  |  | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.007355 | 0.020916 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.007355 | 0.021359 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.007355 | 0.021346 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.007355 | 0.021346 | BDL | [17 empty] | difluoromethane | 10.975 | 51 | 2796 | 1.072 | 1.072 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 1.072 | [17 empty] | difluoromethane | 10.975 | 51 | 74339 | 29.369 | 29.369 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 29.369 | [17 empty] | difluoromethane | 10.975 | 51 | 26160 | 10.231 | 10.231 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.231
1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 693m | 0.071 | 0.071 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.071 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | BDL | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.856 | 51 | 94143 | 9.936 | 9.936 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 9.936 |  |  | 1,1,1,2,2-pentafluoroe... | 12.857 | 51 | 221m | 0.022 | 0.022 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.022 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.013843 | 0.039366 | BDL | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.857 | 51 | 212m | 0.023 | 0.023 | ppb | 200 | 2.9041 | 2.9041 | 0.013843 | 0.040201 | 0.066794 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 182m | 0.02 | 0.02 | ppb | 200 | 2.9022 | 2.9022 | 0.013843 | 0.040175 | 0.058045 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.013843 | 0.040175 | BDL | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 10047 | 1.035 | 1.035 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 1.035 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 269462 | 28.592 | 28.592 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 28.592 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 97677 | 10.26 | 10.26 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.26
1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 325747 | 10.037 | 10.037 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.037 |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 407m | 0.012 | 0.012 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.013563 | 0.038568 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.013563 | 0.039387 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.013563 | 0.039362 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.013563 | 0.039362 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 35251 | 1.06 | 1.06 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 1.06 | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 904368 | 28.013 | 28.013 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 28.013 | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 329640 | 10.108 | 10.108 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.108
fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 11.409 | 33 | 11147 | 10.097 | 10.097 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.097 |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.333 | 0.94697 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.333 | 0.96706 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.333 | 0.96644 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.333 | 0.96644 | BDL | [17 empty] | fluoromethane | 11.424 | 33 | 902m | 0.797 | 0.797 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 0.797 | [17 empty] | fluoromethane | 11.424 | 33 | 23254m | 21.176 | 21.176 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 21.176 | [17 empty] | fluoromethane | 11.409 | 33 | 11540m | 10.403 | 10.403 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.403
tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 8.491 | 81 | 71716 | 9.834 | 9.834 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 9.834 |  |  | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.019104 | 0.054326 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.019104 | 0.055478 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.019104 | 0.055443 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.019104 | 0.055443 | BDL | [17 empty] | tetrafluoroethylene | 8.503 | 81 | 7407 | 0.991 | 0.991 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 0.991 | [17 empty] | tetrafluoroethylene | 8.503 | 81 | 196399 | 27.074 | 27.074 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 27.074 | [17 empty] | tetrafluoroethylene | 8.491 | 81 | 72540 | 9.899 | 9.899 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 9.899
hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 13.247 | 131 | 154996 | 9.594 | 9.594 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.594 |  |  | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.0086848 | 0.024697 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.0086848 | 0.025221 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0086848 | 0.025205 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0086848 | 0.025205 | BDL | [17 empty] | hexafluoropropene | 13.248 | 131 | 16284 | 0.984 | 0.984 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 0.984 | [17 empty] | hexafluoropropene | 13.248 | 131 | 423903m | 26.378 | 26.378 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 26.378 | [17 empty] | hexafluoropropene | 13.248 | 131 | 154380 | 9.51 | 9.51 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.51
1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9640 | 9.788 | 9.788 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.788 |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.034251 | 0.097402 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.034251 | 0.099468 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.034251 | 0.099405 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.034251 | 0.099405 | BDL | [17 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 1018 | 1.009 | 1.009 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 1.009 | [17 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 26780 | 27.336 | 27.336 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 27.336 | [17 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9732 | 9.834 | 9.834 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.834
hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 13.435 | 69 | 10939 | 10.52 | 10.52 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 10.52 |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.33918 | 0.96455 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.33918 | 0.98501 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.33918 | 0.98439 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.33918 | 0.98439 | BDL | [17 empty] | hexafluoropropene_oxide | 13.45 | 69 | 1177m | 1.105 | 1.105 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 1.105 | [17 empty] | hexafluoropropene_oxide | 13.436 | 69 | 28219m | 27.284 | 27.284 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 27.284 | [17 empty] | hexafluoropropene_oxide | 13.436 | 69 | 10233 | 9.794 | 9.794 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.794
chlorodifluoromethane | 13.74 | 51 | 2673m | 0.268 | 0.268 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 0.268 | [17 empty] | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | BDL | [17 empty] | chlorodifluoromethane | 13.739 | 51 | 95589 | 9.911 | 9.911 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 9.911 |  |  | chlorodifluoromethane | 13.74 | 51 | 228m | 0.023 | 0.023 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | BDL | [17 empty] | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.024044 | 0.068374 | BDL | [17 empty] | chlorodifluoromethane | 13.74 | 51 | 476m | 0.05 | 0.05 | ppb | 200 | 2.9041 | 2.9041 | 0.024044 | 0.069825 | 0.1452 | [17 empty] | chlorodifluoromethane | 13.739 | 51 | 343m | 0.037 | 0.037 | ppb | 200 | 2.9022 | 2.9022 | 0.024044 | 0.06978 | 0.10738 | [17 empty] | chlorodifluoromethane | 13.739 | 51 | 311m | 0.033 | 0.033 | ppb | 200 | 2.9022 | 2.9022 | 0.024044 | 0.06978 | 0.095773 | [17 empty] | chlorodifluoromethane | 13.74 | 51 | 10506 | 1.063 | 1.063 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 1.063 | [17 empty] | chlorodifluoromethane | 13.74 | 51 | 273630 | 28.523 | 28.523 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 28.523 | [17 empty] | chlorodifluoromethane | 13.74 | 51 | 99702 | 10.288 | 10.288 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.288
1,1,1,2-tetrafluoroethane | 15.939 | 33 | 645m | 0.195 | 0.195 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 0.195 | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 31829 | 9.982 | 9.982 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 9.982 |  |  | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.02815 | 0.080053 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 198m | 0.062 | 0.062 | ppb | 200 | 2.9041 | 2.9041 | 0.02815 | 0.081751 | 0.18005 | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.02815 | 0.081699 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.02815 | 0.081699 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 3396 | 1.039 | 1.039 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 1.039 | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 93306 | 29.418 | 29.418 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 29.418 | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 33493 | 10.454 | 10.454 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.454
decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 16.161 | 119 | 169526 | 9.801 | 9.801 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 9.801 |  |  | decafluorobutane | 16.147 | 119 | 141m | 0.008 | 0.008 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.023141 | 0.065808 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.023141 | 0.067204 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.023141 | 0.067161 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.023141 | 0.067161 | BDL | [17 empty] | decafluorobutane | 16.161 | 119 | 19102 | 1.078 | 1.078 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 1.078 | [17 empty] | decafluorobutane | 16.161 | 119 | 458240m | 26.634 | 26.634 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 26.634 | [17 empty] | decafluorobutane | 16.175 | 119 | 173271 | 9.97 | 9.97 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 9.97
1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 18.112 | 51 | 140153 | 9.828 | 9.828 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 9.828 |  |  | 1H-Heptafluoropropane | 18.13 | 51 | 237m | 0.016 | 0.016 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.016117 | 0.045833 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.016117 | 0.046805 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.016117 | 0.046775 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.016117 | 0.046775 | BDL | [17 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 15628 | 1.07 | 1.07 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 1.07 | [17 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 394456 | 27.809 | 27.809 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 27.809 | [17 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 147074 | 10.264 | 10.264 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.264
octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 21.943 | 93 | 235023 | 9.957 | 9.957 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 9.957 |  |  | octafluorocyclopentene | 21.943 | 93 | 102m | 0.004 | 0.004 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.0051927 | 0.014767 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.0051927 | 0.01508 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0051927 | 0.01507 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0051927 | 0.01507 | BDL | [17 empty] | octafluorocyclopentene | 21.943 | 93 | 24941 | 1.031 | 1.031 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 1.031 | [17 empty] | octafluorocyclopentene | 21.943 | 93 | 626750 | 26.695 | 26.695 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 26.695 | [17 empty] | octafluorocyclopentene | 21.943 | 93 | 240403 | 10.136 | 10.136 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.136
trichlorofluoromethane | 21.239 | 101 | 6770 | 0.277 | 0.277 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 0.277 | [17 empty] | trichlorofluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 235289 | 9.971 | 9.971 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 9.971 |  |  | trichlorofluoromethane | 21.239 | 101 | 318m | 0.013 | 0.013 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 111m | 0.005 | 0.005 | ppb | 200 | 2.8438 | 2.8438 | 0.028098 | 0.079903 | BDL | [17 empty] | trichlorofluoromethane | 21.257 | 101 | 924m | 0.039 | 0.039 | ppb | 200 | 2.9041 | 2.9041 | 0.028098 | 0.081597 | 0.11326 | [17 empty] | trichlorofluoromethane | 21.257 | 101 | 756m | 0.033 | 0.033 | ppb | 200 | 2.9022 | 2.9022 | 0.028098 | 0.081546 | 0.095773 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 768m | 0.034 | 0.034 | ppb | 200 | 2.9022 | 2.9022 | 0.028098 | 0.081546 | 0.098676 | [17 empty] | trichlorofluoromethane | 21.24 | 101 | 26327 | 1.089 | 1.089 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 1.089 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 633705 | 26.999 | 26.999 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 26.999 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 240278 | 10.134 | 10.134 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.134
dodecafluoro-n-pentane | 21.012 | 69 | 270m | 0.005 | 0.005 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 549309 | 10.213 | 10.213 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.213 |  |  | dodecafluoro-n-pentane | 21.03 | 69 | 734 | 0.013 | 0.013 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 0.013 | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.010934 | 0.031093 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.010934 | 0.031752 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.010934 | 0.031732 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.010934 | 0.031732 | BDL | [17 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 63823m | 1.158 | 1.158 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 1.158 | [17 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 1472980 | 27.534 | 27.534 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 27.534 | [17 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 571942 | 10.583 | 10.583 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.583
1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 22.801 | 51 | 222716 | 10.484 | 10.484 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.484 |  |  | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.012238 | 0.034801 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.012238 | 0.03554 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.012238 | 0.035517 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.012238 | 0.035517 | BDL | [17 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 25427 | 1.168 | 1.168 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 1.168 | [17 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 597475 | 28.276 | 28.276 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 28.276 | [17 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 233362 | 10.933 | 10.933 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.933
tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 24.455 | 69 | 563397m | 9.847 | 9.847 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 9.847 |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.045372 | 0.12903 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.045372 | 0.13176 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.045372 | 0.13168 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.045372 | 0.13168 | BDL | [17 empty] | tetradecafluorohexane | 24.519 | 69 | 62919m | 1.073 | 1.073 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 1.073 | [17 empty] | tetradecafluorohexane | 24.495 | 69 | 1585568m | 27.861 | 27.861 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 27.861 | [17 empty] | tetradecafluorohexane | 24.456 | 69 | 580547m | 10.098 | 10.098 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.098
1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 26.105 | 69 | 186847 | 10.134 | 10.134 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.134 |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.02563 | 0.072884 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.02563 | 0.07443 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.02563 | 0.074383 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.02563 | 0.074383 | BDL | [17 empty] | 1H-Perfluoropentane | 26.134 | 69 | 21659m | 1.146 | 1.146 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 1.146 | [17 empty] | 1H-Perfluoropentane | 26.124 | 69 | 520346 | 28.372 | 28.372 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 28.372 | [17 empty] | 1H-Perfluoropentane | 26.105 | 69 | 196289 | 10.595 | 10.595 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.595
heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 26.173 | 101 | 466707m | 9.955 | 9.955 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 9.955 |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.0093935 | 0.026713 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.0093935 | 0.02728 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0093935 | 0.027262 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0093935 | 0.027262 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 26.202 | 101 | 52781m | 1.099 | 1.099 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 1.099 | [17 empty] | heptafluoropropyl 1,2,... | 26.193 | 101 | 1272926m | 27.298 | 27.298 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 27.298 | [17 empty] | heptafluoropropyl 1,2,... | 26.173 | 101 | 467901m | 9.933 | 9.933 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 9.933
hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 27.275 | 69 | 673979 | 10.006 | 10.006 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.006 |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.007102 | 0.020196 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.007102 | 0.020625 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.007102 | 0.020612 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.007102 | 0.020612 | BDL | [17 empty] | hexadecafluoroheptane | 27.314 | 69 | 78624 | 1.139 | 1.139 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 1.139 | [17 empty] | hexadecafluoroheptane | 27.294 | 69 | 1821445 | 27.187 | 27.187 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 27.187 | [17 empty] | hexadecafluoroheptane | 27.285 | 69 | 691987 | 10.225 | 10.225 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.225
1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 28.747 | 51 | 325895 | 10.136 | 10.136 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.136 |  |  | 1H-Perfluorohexane | 28.718 | 51 | 242m | 0.007 | 0.007 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.0083342 | 0.0237 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.0083342 | 0.024203 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0083342 | 0.024188 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.0083342 | 0.024188 | BDL | [17 empty] | 1H-Perfluorohexane | 28.776 | 51 | 35462 | 1.076 | 1.076 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 1.076 | [17 empty] | 1H-Perfluorohexane | 28.767 | 51 | 881997 | 27.579 | 27.579 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 27.579 | [17 empty] | 1H-Perfluorohexane | 28.757 | 51 | 337248 | 10.439 | 10.439 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.439
octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 29.537 | 119 | 409773 | 9.813 | 9.813 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.813 |  |  | octadecafluorooctane | 29.525 | 119 | 482 | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.027692 | 0.07875 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.027692 | 0.080421 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.027692 | 0.08037 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.027692 | 0.08037 | BDL | [17 empty] | octadecafluorooctane | 29.572 | 119 | 48975 | 1.145 | 1.145 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 1.145 | [17 empty] | octadecafluorooctane | 29.549 | 119 | 1069808 | 25.756 | 25.756 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 25.756 | [17 empty] | octadecafluorooctane | 29.548 | 119 | 407237 | 9.706 | 9.706 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.706
1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL | [17 empty] | 1H-Perfluoroheptane | 31.066 | 51 | 894988 | 9.994 | 9.994 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 9.994 |  |  | 1H-Perfluoroheptane | 31.078 | 51 | 1328 | 0.014 | 0.014 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 0.014 | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.010647 | 0.030278 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.010647 | 0.030921 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.010647 | 0.030901 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.010647 | 0.030901 | BDL | [17 empty] | 1H-Perfluoroheptane | 31.09 | 51 | 108458 | 1.182 | 1.182 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 1.182 | [17 empty] | 1H-Perfluoroheptane | 31.078 | 51 | 2402975 | 26.977 | 26.977 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 26.977 | [17 empty] | 1H-Perfluoroheptane | 31.09 | 51 | 923485 | 10.263 | 10.263 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.263
1H-Perfluorooctane | 33.505 | 51 | 2709m | 0.038 | 0.038 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.038 | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | BDL | [17 empty] | 1H-Perfluorooctane | 33.482 | 51 | 688094 | 10.024 | 10.024 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.024 |  |  | 1H-Perfluorooctane | 33.516 | 51 | 3901 | 0.054 | 0.054 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.054 | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.028576 | 0.081262 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.028576 | 0.082986 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.028576 | 0.082933 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.028576 | 0.082933 | BDL | [17 empty] | 1H-Perfluorooctane | 33.516 | 51 | 85607m | 1.217 | 1.217 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 1.217 | [17 empty] | 1H-Perfluorooctane | 33.505 | 51 | 1823581 | 26.708 | 26.708 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 26.708 | [17 empty] | 1H-Perfluorooctane | 33.505 | 51 | 704508 | 10.214 | 10.214 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.214
2H-peffluoro-5-methyl-... | 33.056 | 169 | 1597m | 0.008 | 0.008 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 33.033 | 169 | 1860293 | 9.797 | 9.797 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.797 |  |  | 2H-peffluoro-5-methyl-... | 33.056 | 169 | 2902 | 0.015 | 0.015 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8438 | 2.8438 | 0.037129 | 0.10559 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9041 | 2.9041 | 0.037129 | 0.10783 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.037129 | 0.10776 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9022 | 2.9022 | 0.037129 | 0.10776 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 33.056 | 169 | 217362 | 1.117 | 1.117 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 1.117 | [17 empty] | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 4903077 | 25.96 | 25.96 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 25.96 | [17 empty] | 2H-peffluoro-5-methyl-... | 33.056 | 169 | 1829586 | 9.589 | 9.589 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.589

## 112724
112724/112724_01.D/epatemp.txt (QT Reviewed): AirBlank_Vol 200 ml | [14 empty] | 112724/112724_02.D/epatemp.txt (QT Reviewed): Rainbow BatchBlank_CF4_Vol 20 ml | [14 empty] | 112724/112724_03.D/epatemp.txt (QT Reviewed): Rainbow BatchBlank_Rest_Vol 200 ml | [14 empty] | 112724/112724_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_initial_100ppbv_vol 20 ml | [14 empty] | 112724/112724_05.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | 112724/112724_06.D/epatemp.txt (QT Reviewed): Rainbow Post_BatchBlankCK_Rest_vol 200 ml | [14 empty] | 112724/112724_07.D/epatemp.txt (QT Reviewed): Rainbow Post Test N2_Blank_vol 20 mL | [14 empty] | 112724/112724_08.D/epatemp.txt (QT Reviewed): Rainbow Post Test N2_Blank_vol 200 mL | [14 empty] | 112724/112724_09.D/epatemp.txt (QT Reviewed): Rainbow Post Test N2_Blank_vol 20 mL | [14 empty] | 112724/112724_10.D/epatemp.txt (QT Reviewed): Rainbow Post Test N2_Blank_vol 200 mL | [14 empty] | 112724/112724_11.D/epatemp.txt (QT Reviewed): Utah Ambient air_20 mL | [14 empty] | 112724/112724_12.D/epatemp.txt (QT Reviewed): Utah Ambient air_200 mL | [14 empty] | 112724/112724_13.D/epatemp.txt (QT Reviewed): Rainbow Run2_Port6_PFBA inj_vol 20 mL | [14 empty] | 112724/112724_14.D/epatemp.txt (QT Reviewed): Rainbow Run2_Port6_PFBA inj_vol 200 mL | [14 empty] | 112724/112724_15.D/epatemp.txt (QT Reviewed): Rainbow Run3_Port8_PFBA inj_vol 20 mL | [14 empty] | 112724/112724_16.D/epatemp.txt (QT Reviewed): Rainbow Run3_Port8_PFBA inj_vol 200 mL | [14 empty] | 112724/112724_17.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [14 empty] | 112724/112724_18.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [14 empty] | 112724/112724_19.D/epatemp.txt (QT Reviewed): Post BatchBlankCK_CF4 20 mL | [14 empty] | CF4 Method |  | Lower Limit | Uppper Limit | Rest Method |  | Lower Limit | Uppper Limit
Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Average IS area | 1053174 | 631904.4 | 1474443.6 | Average IS area | 1032225 | 619335 | 1445115
1,4-difluorobenzene | 30.353 | 114 | 1092895 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1074385 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.342 | 114 | 1088855 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1063940 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.388 | 114 | 1049943 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.328 | 114 | 1097937 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 1086135 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.399 | 114 | 1077707 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1093978 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.399 | 114 | 1083116 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1091153 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.399 | 114 | 1090366 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1204173 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.4 | 114 | 1044161 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1212225 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.388 | 114 | 1026452 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1094387 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.388 | 114 | 1059081 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.492 | 114 | 1084513 | 50 |  | ppb | [6 empty] | Pass | Pass | Average RT | 30.503 | 30.4197 | 30.5863 | Average RT | 30.411 | 30.3277 | 30.4943
hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | BDL | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 6.994 | 69 | 82247 | 100.583 | 100.583 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 100.583 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 166556 | 9.816 | 9.816 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.816 | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.1777 | 2.1777 | 0.089699 | 0.19533 | BDL | Pass | Pass | hexafluoroethane | 8.105 | 119 | 76m | 0.004 | 0.004 | ppb | 200 | 2.1777 | 2.1777 | 0.0084458 | 0.018392 | BDL | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.8691 | 2.8691 | 0.089699 | 0.25736 | BDL | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.0084458 | 0.024232 | BDL | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | BDL | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.8986 | 2.8986 | 0.089699 | 0.26001 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 10304 | 0.611 | 0.611 | ppb | 200 | 2.8986 | 2.8986 | 0.0084458 | 0.024481 | 1.7711 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.8772 | 2.8772 | 0.089699 | 0.25808 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 7551 | 0.455 | 0.455 | ppb | 200 | 2.8772 | 2.8772 | 0.0084458 | 0.0243 | 1.3091 | Pass | Pass | tetrafluoromethane | 6.995 | 69 | 80568 | 95.788 | 95.788 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 95.788 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 168067 | 9.82 | 9.82 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.82 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | BDL | Pass | Pass
chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 8.515 | 85 | 44539 | 10.101 | 10.101 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.101 |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.029884 | 0.065078 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.029884 | 0.085743 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.029884 | 0.086624 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.029884 | 0.085982 | BDL | [17 empty] | chlorotrifluoromethane | 8.515 | 85 | 45132 | 10.147 | 10.147 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.147
trifluoromethane | 8.72 | 69 | 220m | 0.055 | 0.055 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 0.055 | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL | [17 empty] | trifluoromethane | 8.708 | 69 | 39839 | 10.434 | 10.434 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.434 |  |  | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL | [17 empty] | trifluoromethane | 8.708 | 69 | 107m | 0.027 | 0.027 | ppb | 200 | 2.1777 | 2.1777 | 0.0356 | 0.077525 | BDL | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.0356 | 0.10214 | BDL | [17 empty] | trifluoromethane | 8.696 | 69 | 95m | 0.024 | 0.024 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL | [17 empty] | trifluoromethane | 8.708 | 69 | 5662 | 1.491 | 1.491 | ppb | 200 | 2.8986 | 2.8986 | 0.0356 | 0.10319 | 4.3219 | [17 empty] | trifluoromethane | 8.709 | 69 | 83518 | 22.374 | 22.374 | ppb | 200 | 2.8772 | 2.8772 | 0.0356 | 0.10243 | 64.3736 | [17 empty] | trifluoromethane | 8.708 | 69 | 39903 | 10.36 | 10.36 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.36
octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 11.178 | 69 | 350883 | 10.273 | 10.273 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.273 |  |  | octafluoropropane | 11.178 | 69 | 410m | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 11.192 | 69 | 558 | 0.016 | 0.016 | ppb | 200 | 2.1777 | 2.1777 | 0.015822 | 0.034454 | 0.034843 | [17 empty] | octafluoropropane | 11.192 | 69 | 713m | 0.02 | 0.02 | ppb | 200 | 2.8691 | 2.8691 | 0.015822 | 0.045395 | 0.057383 | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.015822 | 0.045862 | BDL | [17 empty] | octafluoropropane | 11.192 | 69 | 698m | 0.021 | 0.021 | ppb | 200 | 2.8772 | 2.8772 | 0.015822 | 0.045522 | 0.06042 | [17 empty] | octafluoropropane | 11.178 | 69 | 355402 | 10.315 | 10.315 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.315
difluoromethane | 10.961 | 51 | 418 | 0.156 | 0.156 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 0.156 | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL | [17 empty] | difluoromethane | 10.975 | 51 | 26313 | 10.226 | 10.226 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.226 |  |  | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.007355 | 0.016017 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.007355 | 0.021103 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.007355 | 0.02132 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.007355 | 0.021162 | BDL | [17 empty] | difluoromethane | 10.961 | 51 | 26740 | 10.303 | 10.303 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.303
1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 1308m | 0.131 | 0.131 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.131 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | BDL | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 101499 | 10.594 | 10.594 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.594 |  |  | 1,1,1,2,2-pentafluoroe... | 12.827 | 51 | 156m | 0.016 | 0.016 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.016 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 254m | 0.026 | 0.026 | ppb | 200 | 2.1777 | 2.1777 | 0.013843 | 0.030145 | 0.056619 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.013843 | 0.039718 | BDL | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 269m | 0.027 | 0.027 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.027 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 257m | 0.027 | 0.027 | ppb | 200 | 2.8986 | 2.8986 | 0.013843 | 0.040126 | 0.078264 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 21229 | 2.267 | 2.267 | ppb | 200 | 2.8772 | 2.8772 | 0.013843 | 0.039828 | 6.5225 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 103038 | 10.662 | 10.662 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.662
1,1,2,2,3,3,4,4-octafl... | 15.953 | 100 | 554m | 0.016 | 0.016 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 0.016 | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 316 | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 339554 | 10.347 | 10.347 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.347 |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 565m | 0.016 | 0.016 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 0.016 | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 151m | 0.004 | 0.004 | ppb | 200 | 2.1777 | 2.1777 | 0.013563 | 0.029535 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 145m | 0.004 | 0.004 | ppb | 200 | 2.8691 | 2.8691 | 0.013563 | 0.038913 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.013563 | 0.039313 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.013563 | 0.039022 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 342381 | 10.343 | 10.343 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.343
fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 11.409 | 33 | 11952 | 10.707 | 10.707 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.707 |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.333 | 0.72516 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.333 | 0.95543 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.333 | 0.96525 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.333 | 0.95809 | BDL | [17 empty] | fluoromethane | 11.409 | 33 | 12430 | 11.039 | 11.039 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 11.039
tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 8.503 | 81 | 74938 | 10.162 | 10.162 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.162 |  |  | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.019104 | 0.041601 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.019104 | 0.054811 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.019104 | 0.055375 | BDL | [17 empty] | tetrafluoroethylene | 8.503 | 81 | 1397 | 0.194 | 0.194 | ppb | 200 | 2.8772 | 2.8772 | 0.019104 | 0.054964 | 0.55817 | [17 empty] | tetrafluoroethylene | 8.491 | 81 | 75945 | 10.21 | 10.21 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.21
hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 13.248 | 131 | 158162 | 9.682 | 9.682 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.682 |  |  | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.0086848 | 0.018913 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.0086848 | 0.024918 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.0086848 | 0.025174 | BDL | [17 empty] | hexafluoropropene | 13.248 | 131 | 175m | 0.011 | 0.011 | ppb | 200 | 2.8772 | 2.8772 | 0.0086848 | 0.024988 | 0.031649 | [17 empty] | hexafluoropropene | 13.248 | 131 | 157747 | 9.573 | 9.573 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.573
1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 10109 | 10.151 | 10.151 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 10.151 |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.034251 | 0.074588 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.034251 | 0.098272 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.034251 | 0.099282 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.034251 | 0.098546 | BDL | [17 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9975 | 9.93 | 9.93 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.93
hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 13.436 | 69 | 11453 | 10.893 | 10.893 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 10.893 |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.33918 | 0.73863 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.33918 | 0.97317 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.33918 | 0.98317 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.33918 | 0.97588 | BDL | [17 empty] | hexafluoropropene_oxide | 13.436 | 69 | 11018 | 10.389 | 10.389 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 10.389
chlorodifluoromethane | 13.739 | 51 | 2945m | 0.29 | 0.29 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 0.29 | [17 empty] | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | BDL | [17 empty] | chlorodifluoromethane | 13.74 | 51 | 102717 | 10.533 | 10.533 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.533 |  |  | chlorodifluoromethane | 13.739 | 51 | 297m | 0.029 | 0.029 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 0.029 | [17 empty] | chlorodifluoromethane | 13.74 | 51 | 945m | 0.094 | 0.094 | ppb | 200 | 2.1777 | 2.1777 | 0.024044 | 0.052359 | 0.2047 | [17 empty] | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.024044 | 0.068985 | BDL | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 1085 | 0.107 | 0.107 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 0.107 | [17 empty] | chlorodifluoromethane | 13.74 | 51 | 527m | 0.054 | 0.054 | ppb | 200 | 2.8986 | 2.8986 | 0.024044 | 0.069694 | 0.15653 | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 459m | 0.048 | 0.048 | ppb | 200 | 2.8772 | 2.8772 | 0.024044 | 0.069178 | 0.1381 | [17 empty] | chlorodifluoromethane | 13.74 | 51 | 104406 | 10.614 | 10.614 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.614
1,1,1,2-tetrafluoroethane | 15.939 | 33 | 1109 | 0.33 | 0.33 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 0.33 | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 35066 | 10.876 | 10.876 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.876 |  |  | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 251m | 0.076 | 0.076 | ppb | 200 | 2.1777 | 2.1777 | 0.02815 | 0.061302 | 0.1655 | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.02815 | 0.080768 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 229m | 0.068 | 0.068 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 0.068 | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.02815 | 0.081598 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.953 | 33 | 161m | 0.051 | 0.051 | ppb | 200 | 2.8772 | 2.8772 | 0.02815 | 0.080993 | 0.14674 | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 35875 | 11.031 | 11.031 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 11.031
decafluorobutane | 16.175 | 119 | 156m | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 16.161 | 119 | 131m | 0.007 | 0.007 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 16.175 | 119 | 176708 | 10.104 | 10.104 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.104 |  |  | decafluorobutane | 16.147 | 119 | 231m | 0.013 | 0.013 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.023141 | 0.050394 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.023141 | 0.066396 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.023141 | 0.067079 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.023141 | 0.066581 | BDL | [17 empty] | decafluorobutane | 16.161 | 119 | 178800 | 10.135 | 10.135 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.135
1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 150915 | 10.467 | 10.467 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.467 |  |  | 1H-Heptafluoropropane | 18.13 | 51 | 304m | 0.02 | 0.02 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 0.02 | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.016117 | 0.035097 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.016117 | 0.046242 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.016117 | 0.046718 | BDL | [17 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 2377 | 0.169 | 0.169 | ppb | 200 | 2.8772 | 2.8772 | 0.016117 | 0.046371 | 0.48624 | [17 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 157604 | 10.836 | 10.836 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.836
octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 21.943 | 93 | 244195 | 10.232 | 10.232 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.232 |  |  | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.0051927 | 0.011308 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.0051927 | 0.014899 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.0051927 | 0.015052 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.0051927 | 0.01494 | BDL | [17 empty] | octafluorocyclopentene | 21.943 | 93 | 249672 | 10.371 | 10.371 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.371
trichlorofluoromethane | 21.239 | 101 | 7038 | 0.283 | 0.283 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 0.283 | [17 empty] | trichlorofluoromethane | 21.24 | 101 | 268m | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 246979 | 10.352 | 10.352 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.352 |  |  | trichlorofluoromethane | 21.239 | 101 | 452m | 0.018 | 0.018 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 2105m | 0.086 | 0.086 | ppb | 200 | 2.1777 | 2.1777 | 0.028098 | 0.061187 | 0.18728 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 224m | 0.009 | 0.009 | ppb | 200 | 2.8691 | 2.8691 | 0.028098 | 0.080616 | BDL | [17 empty] | trichlorofluoromethane | 21.257 | 101 | 2703 | 0.109 | 0.109 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 0.109 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 886m | 0.037 | 0.037 | ppb | 200 | 2.8986 | 2.8986 | 0.028098 | 0.081445 | 0.10725 | [17 empty] | trichlorofluoromethane | 21.24 | 101 | 965 | 0.041 | 0.041 | ppb | 200 | 2.8772 | 2.8772 | 0.028098 | 0.080841 | 0.11796 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 248487 | 10.325 | 10.325 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.325
dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 581564 | 10.694 | 10.694 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.694 |  |  | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.010934 | 0.02381 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.010934 | 0.03137 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.010934 | 0.031693 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.010934 | 0.031458 | BDL | [17 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 587590m | 10.712 | 10.712 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.712
1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 231808 | 10.792 | 10.792 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.792 |  |  | 1H-Nonafluorobutane | 22.801 | 51 | 240m | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.012238 | 0.02665 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.012238 | 0.035112 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.012238 | 0.035473 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.012238 | 0.03521 | BDL | [17 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 241608 | 11.151 | 11.151 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.151
tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 24.456 | 69 | 596577m | 10.312 | 10.312 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.312 |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.045372 | 0.098806 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.045372 | 0.13018 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.045372 | 0.13152 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.045372 | 0.13054 | BDL | [17 empty] | tetradecafluorohexane | 24.416 | 69 | 601255m | 10.303 | 10.303 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.303
1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 26.095 | 69 | 190013m | 10.192 | 10.192 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.192 |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.02563 | 0.055813 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.02563 | 0.073535 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.02563 | 0.074291 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.02563 | 0.07374 | BDL | [17 empty] | 1H-Perfluoropentane | 26.076 | 69 | 204588m | 10.879 | 10.879 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.879
heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 26.173 | 101 | 489920m | 10.335 | 10.335 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.335 |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.0093935 | 0.020456 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.0093935 | 0.026951 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.0093935 | 0.027229 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.0093935 | 0.027027 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 26.154 | 101 | 489443m | 10.236 | 10.236 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.236
hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 27.275 | 69 | 711695 | 10.45 | 10.45 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.45 |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.007102 | 0.015466 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.007102 | 0.020377 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.007102 | 0.020586 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.007102 | 0.020434 | BDL | [17 empty] | hexadecafluoroheptane | 27.265 | 69 | 714489 | 10.4 | 10.4 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.4
1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 28.737 | 51 | 347360 | 10.685 | 10.685 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.685 |  |  | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.0083342 | 0.018149 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.0083342 | 0.023912 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.0083342 | 0.024158 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.0083342 | 0.023979 | BDL | [17 empty] | 1H-Perfluorohexane | 28.737 | 51 | 350739 | 10.696 | 10.696 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.696
octadecafluorooctane | 29.514 | 119 | 131m | 0.003 | 0.003 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 29.537 | 119 | 417927 | 9.898 | 9.898 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.898 |  |  | octadecafluorooctane | 29.5 | 119 | 430m | 0.01 | 0.01 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 29.56 | 119 | 105m | 0.002 | 0.002 | ppb | 200 | 2.1777 | 2.1777 | 0.027692 | 0.060305 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.027692 | 0.079454 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 29.572 | 119 | 128m | 0.003 | 0.003 | ppb | 200 | 2.8986 | 2.8986 | 0.027692 | 0.080271 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.027692 | 0.079676 | BDL | [17 empty] | octadecafluorooctane | 29.537 | 119 | 417977 | 9.814 | 9.814 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.814
1H-Perfluoroheptane | 31.078 | 51 | 517 | 0.005 | 0.005 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL | [17 empty] | 1H-Perfluoroheptane | 31.067 | 51 | 951905 | 10.513 | 10.513 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.513 |  |  | 1H-Perfluoroheptane | 31.064 | 51 | 1326 | 0.014 | 0.014 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 0.014 | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.010647 | 0.023186 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.010647 | 0.030549 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.010647 | 0.030863 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.010647 | 0.030634 | BDL | [17 empty] | 1H-Perfluoroheptane | 31.078 | 51 | 961481 | 10.527 | 10.527 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.527
1H-Perfluorooctane | 33.516 | 51 | 2321 | 0.032 | 0.032 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.032 | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | BDL | [17 empty] | 1H-Perfluorooctane | 33.482 | 51 | 722734 | 10.413 | 10.413 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.413 |  |  | 1H-Perfluorooctane | 33.503 | 51 | 5929 | 0.082 | 0.082 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.082 | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.028576 | 0.062228 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.028576 | 0.081988 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.028576 | 0.082831 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.028576 | 0.082217 | BDL | [17 empty] | 1H-Perfluorooctane | 33.505 | 51 | 726063 | 10.371 | 10.371 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.371
2H-peffluoro-5-methyl-... | 33.056 | 169 | 1328 | 0.007 | 0.007 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 1848463 | 9.628 | 9.628 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.628 |  |  | 2H-peffluoro-5-methyl-... | 33.043 | 169 | 3009 | 0.015 | 0.015 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.1777 | 2.1777 | 0.037129 | 0.080855 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8691 | 2.8691 | 0.037129 | 0.10653 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8986 | 2.8986 | 0.037129 | 0.10763 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8772 | 2.8772 | 0.037129 | 0.10683 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 1820840 | 9.402 | 9.402 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.402

## 112924
112924/112924_01.D/epatemp.txt (QT Reviewed): AirBlank_Vol 200 ml | [14 empty] | 112924/112924_02.D/epatemp.txt (QT Reviewed): Rainbow BatchBlank_CF4_Vol 20 ml | [14 empty] | 112924/112924_03.D/epatemp.txt (QT Reviewed): Rainbow BatchBlank_Rest_Vol 200 ml | [14 empty] | 112924/112924_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_initial_100ppbv_vol 20 ml | [14 empty] | 112924/112924_05.D/epatemp.txt (QT Reviewed): Rainbow Post_BatchBlankCK_CF4_vol 20 ml | [14 empty] | 112924/112924_06.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | 112924/112924_07.D/epatemp.txt (QT Reviewed): Rainbow Post_BatchBlankCK_Rest_vol 200 ml | [14 empty] | 112924/112924_08.D/epatemp.txt (QT Reviewed): Rainbow Run4_Port10_PFBA inj_vol 20 mL | [14 empty] | 112924/112924_09.D/epatemp.txt (QT Reviewed): Rainbow Run4_Port10_PFBA inj_vol 200 mL | [14 empty] | 112924/112924_10.D/epatemp.txt (QT Reviewed): Rainbow Run5_Port4_PFOA inj_vol 20 mL | [14 empty] | 112924/112924_11.D/epatemp.txt (QT Reviewed): Rainbow Run5_Port4_PFOA inj_vol 200 mL | [14 empty] | 112924/112924_12.D/epatemp.txt (QT Reviewed): Rainbow Run6_Port6_PFOA inj_vol 20 mL | [14 empty] | 112924/112924_13.D/epatemp.txt (QT Reviewed): Rainbow Run6_Port6_PFOA inj_vol 200 mL | [14 empty] | 112924/112924_14.D/epatemp.txt (QT Reviewed): Rainbow Run7_Port8_PFOA inj_vol 20 mL | [14 empty] | 112924/112924_15.D/epatemp.txt (QT Reviewed): Rainbow Run7_Port8_PFOA inj_vol 200 mL | [14 empty] | 112924/112924_16.D/epatemp.txt (QT Reviewed): Rainbow Run7_Port8_PFOA inj_DUP_vol 20 mL | [14 empty] | 112924/112924_17.D/epatemp.txt (QT Reviewed): Rainbow Run7_Port8_PFOA inj_DUP_vol 200 mL | [14 empty] | 112924/112924_18.D/epatemp.txt (QT Reviewed): Rainbow Run8_Port10_PFOA inj_vol 20 mL | [14 empty] | 112924/112924_19.D/epatemp.txt (QT Reviewed): Rainbow Run8_Port10_PFOA inj_vol 200 mL | [14 empty] | 112924/112924_20.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [14 empty] | 112924/112924_21.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [14 empty] | 112924/112924_22.D/epatemp.txt (QT Reviewed): Post BatchBlankCK_Rest 200 mL | [14 empty] | CF4 Method |  | Lower Limit | Uppper Limit | Rest Method |  | Lower Limit | Uppper Limit
Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Average IS area | 1053174 | 631904.4 | 1474443.6 | Average IS area | 1032225 | 619335 | 1445115
1,4-difluorobenzene | 30.319 | 114 | 1053423 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1040330 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.411 | 114 | 1040454 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.491 | 114 | 1035814 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 993155 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.365 | 114 | 1023085 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.4 | 114 | 1047614 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1161680 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.399 | 114 | 1013808 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1148309 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.4 | 114 | 1041079 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.468 | 114 | 1166314 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.4 | 114 | 998514 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.468 | 114 | 1171148 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.399 | 114 | 997251 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.468 | 114 | 1184505 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.388 | 114 | 986872 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.469 | 114 | 1194456 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.4 | 114 | 981411 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.457 | 114 | 1048353 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.354 | 114 | 1019948 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.398 | 114 | 1036914 | 5 |  | ppb | [6 empty] | Pass | Pass | Average RT | 30.503 | 30.4197 | 30.5863 | Average RT | 30.411 | 30.3277 | 30.4943
hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Fail | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | BDL | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 6.994 | 69 | 74377 | 93.428 | 93.428 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 93.428 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | #REF! | 1 | 0.089699 | 0.089699 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 163708 | 9.902 | 9.902 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.902 | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.9152 | 2.9152 | 0.089699 | 0.26149 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 3749 | 0.229 | 0.229 | ppb | 200 | 2.9152 | 2.9152 | 0.0084458 | 0.024621 | 0.66758 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.9013 | 2.9013 | 0.089699 | 0.26024 | BDL | Pass | Pass | hexafluoroethane | 8.117 | 119 | 5782 | 0.344 | 0.344 | ppb | 200 | 2.9013 | 2.9013 | 0.0084458 | 0.024504 | 0.99803 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 3 | 3 | 0.089699 | 0.2691 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 12117 | 0.751 | 0.751 | ppb | 200 | 3 | 3 | 0.0084458 | 0.025337 | 2.253 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.8791 | 2.8791 | 0.089699 | 0.25825 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 11245 | 0.698 | 0.698 | ppb | 200 | 2.8791 | 2.8791 | 0.0084458 | 0.024316 | 2.0096 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.8791 | 2.8791 | 0.089699 | 0.25825 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 11110 | 0.697 | 0.697 | ppb | 200 | 2.8791 | 2.8791 | 0.0084458 | 0.024316 | 2.0067 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 2.9096 | 2.9096 | 0.089699 | 0.26099 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 8020 | 0.506 | 0.506 | ppb | 200 | 2.9096 | 2.9096 | 0.0084458 | 0.024574 | 1.4723 | Pass | Pass | tetrafluoromethane | 6.995 | 69 | 72180 | 89.584 | 89.584 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 89.584 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 158786 | 9.634 | 9.634 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.634 | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass
chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [32 empty] | chlorotrifluoromethane | 8.515 | 85 | 44000 | 10.24 | 10.24 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.24 |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.029884 | 0.087119 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.029884 | 0.086702 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.029884 | 0.089653 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.029884 | 0.08604 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.029884 | 0.08604 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.029884 | 0.086952 | BDL | [17 empty] | chlorotrifluoromethane | 8.515 | 85 | 42686 | 9.965 | 9.965 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.965 |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL
trifluoromethane | 8.721 | 69 | 179m | 0.047 | 0.047 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 0.047 | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL | [32 empty] | trifluoromethane | 8.709 | 69 | 39204 | 10.537 | 10.537 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.537 |  |  | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL | [17 empty] | trifluoromethane | 8.708 | 69 | 215359 | 58.412 | 58.412 | ppb | 200 | 2.9152 | 2.9152 | 0.0356 | 0.10378 | 170.2829 | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.0356 | 0.10328 | BDL | [17 empty] | trifluoromethane | 8.709 | 69 | 2037 | 0.561 | 0.561 | ppb | 200 | 3 | 3 | 0.0356 | 0.1068 | 1.683 | [17 empty] | trifluoromethane | 8.708 | 69 | 42208 | 11.638 | 11.638 | ppb | 200 | 2.8791 | 2.8791 | 0.0356 | 0.1025 | 33.5067 | [17 empty] | trifluoromethane | 8.709 | 69 | 41411 | 11.539 | 11.539 | ppb | 200 | 2.8791 | 2.8791 | 0.0356 | 0.1025 | 33.2217 | [17 empty] | trifluoromethane | 8.708 | 69 | 146948 | 41.173 | 41.173 | ppb | 200 | 2.9096 | 2.9096 | 0.0356 | 0.10358 | 119.7978 | [17 empty] | trifluoromethane | 8.709 | 69 | 38963 | 10.504 | 10.504 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.504 |  |  | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL
octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [32 empty] | octafluoropropane | 11.178 | 69 | 344476 | 10.35 | 10.35 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.35 |  |  | octafluoropropane | 11.178 | 69 | 359m | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 11.192 | 69 | 683m | 0.021 | 0.021 | ppb | 200 | 2.9152 | 2.9152 | 0.015822 | 0.046123 | 0.061219 | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.015822 | 0.045903 | BDL | [17 empty] | octafluoropropane | 11.193 | 69 | 782m | 0.024 | 0.024 | ppb | 200 | 3 | 3 | 0.015822 | 0.047465 | 0.072 | [17 empty] | octafluoropropane | 11.178 | 69 | 3308m | 0.102 | 0.102 | ppb | 200 | 2.8791 | 2.8791 | 0.015822 | 0.045552 | 0.29367 | [17 empty] | octafluoropropane | 11.178 | 69 | 3294m | 0.103 | 0.103 | ppb | 200 | 2.8791 | 2.8791 | 0.015822 | 0.045552 | 0.29655 | [17 empty] | octafluoropropane | 11.178 | 69 | 4994 | 0.156 | 0.156 | ppb | 200 | 2.9096 | 2.9096 | 0.015822 | 0.046035 | 0.4539 | [17 empty] | octafluoropropane | 11.178 | 69 | 341069 | 10.279 | 10.279 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.279 |  |  | octafluoropropane | 11.178 | 69 | 302m | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL
difluoromethane | 10.975 | 51 | 102m | 0.04 | 0.04 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 0.04 | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL | [32 empty] | difluoromethane | 10.961 | 51 | 26298 | 10.489 | 10.489 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.489 |  |  | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL | [17 empty] | difluoromethane | 10.961 | 51 | 166m | 0.067 | 0.067 | ppb | 200 | 2.9152 | 2.9152 | 0.007355 | 0.021441 | 0.19532 | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.007355 | 0.021339 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.007355 | 0.022065 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.007355 | 0.021176 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.007355 | 0.021176 | BDL | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.007355 | 0.0214 | BDL | [17 empty] | difluoromethane | 10.961 | 51 | 25987 | 10.397 | 10.397 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.397 |  |  | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL
1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 651 | 0.068 | 0.068 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.068 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | BDL | [32 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 100169 | 10.73 | 10.73 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.73 |  |  | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 137m | 0.014 | 0.014 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.014 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 261470 | 28.265 | 28.265 | ppb | 200 | 2.9152 | 2.9152 | 0.013843 | 0.040355 | 82.3983 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 357 | 0.038 | 0.038 | ppb | 200 | 2.9013 | 2.9013 | 0.013843 | 0.040162 | 0.11025 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 175m | 0.019 | 0.019 | ppb | 200 | 3 | 3 | 0.013843 | 0.041529 | 0.057 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 537 | 0.059 | 0.059 | ppb | 200 | 2.8791 | 2.8791 | 0.013843 | 0.039855 | 0.16987 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 472m | 0.052 | 0.052 | ppb | 200 | 2.8791 | 2.8791 | 0.013843 | 0.039855 | 0.14971 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 6389 | 0.713 | 0.713 | ppb | 200 | 2.9096 | 2.9096 | 0.013843 | 0.040278 | 2.0746 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 97951 | 10.525 | 10.525 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.525 |  |  | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 119m | 0.013 | 0.013 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | BDL
1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [32 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 331990 | 10.382 | 10.382 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.382 |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.953 | 100 | 347 | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.013563 | 0.039538 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.013563 | 0.039348 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.013563 | 0.040688 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.013563 | 0.039048 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.013563 | 0.039048 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 982 | 0.032 | 0.032 | ppb | 200 | 2.9096 | 2.9096 | 0.013563 | 0.039462 | 0.093108 | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 326371 | 10.238 | 10.238 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.238 |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 291 | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL
fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [32 empty] | fluoromethane | 11.41 | 33 | 12348 | 11.352 | 11.352 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 11.352 |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.333 | 0.97076 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.333 | 0.96612 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.333 | 0.999 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.333 | 0.95873 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.333 | 0.95873 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.333 | 0.9689 | BDL | [17 empty] | fluoromethane | 11.41 | 33 | 11918 | 10.99 | 10.99 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.99 |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL
tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [32 empty] | tetrafluoroethylene | 8.491 | 81 | 73373 | 10.211 | 10.211 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.211 |  |  | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 8.503 | 81 | 16086 | 2.259 | 2.259 | ppb | 200 | 2.9152 | 2.9152 | 0.019104 | 0.055691 | 6.5854 | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.019104 | 0.055425 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.019104 | 0.057311 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.019104 | 0.055001 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.019104 | 0.055001 | BDL | [17 empty] | tetrafluoroethylene | 8.491 | 81 | 161m | 0.023 | 0.023 | ppb | 200 | 2.9096 | 2.9096 | 0.019104 | 0.055584 | 0.066921 | [17 empty] | tetrafluoroethylene | 8.491 | 81 | 72484 | 10.118 | 10.118 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.118 |  |  | tetrafluoroethylene | 8.491 | 81 | 65m | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL
hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [32 empty] | hexafluoropropene | 13.248 | 131 | 153597 | 9.649 | 9.649 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.649 |  |  | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 13.247 | 131 | 9708 | 0.615 | 0.615 | ppb | 200 | 2.9152 | 2.9152 | 0.0086848 | 0.025318 | 1.7929 | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.0086848 | 0.025197 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.0086848 | 0.026054 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.0086848 | 0.025004 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.0086848 | 0.025004 | BDL | [17 empty] | hexafluoropropene | 13.248 | 131 | 192m | 0.013 | 0.013 | ppb | 200 | 2.9096 | 2.9096 | 0.0086848 | 0.02527 | 0.037825 | [17 empty] | hexafluoropropene | 13.248 | 131 | 147789 | 9.313 | 9.313 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.313 |  |  | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL
1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [32 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9641 | 9.935 | 9.935 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.935 |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.034251 | 0.099849 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.034251 | 0.099371 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.034251 | 0.10275 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.034251 | 0.098612 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.034251 | 0.098612 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.034251 | 0.099658 | BDL | [17 empty] | 1,1,1-trifluoroethane | 13.219 | 64 | 9578 | 9.901 | 9.901 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.901 |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL
hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [32 empty] | hexafluoropropene_oxide | 13.436 | 69 | 11964 | 11.678 | 11.678 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 11.678 |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.33918 | 0.98879 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.33918 | 0.98406 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.33918 | 1.0175 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.33918 | 0.97653 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.33918 | 0.97653 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.33918 | 0.98689 | BDL | [17 empty] | hexafluoropropene_oxide | 13.421 | 69 | 9060m | 8.871 | 8.871 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 8.871 |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL
chlorodifluoromethane | 13.74 | 51 | 2610 | 0.267 | 0.267 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 0.267 | [17 empty] | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | BDL | [32 empty] | chlorodifluoromethane | 13.725 | 51 | 101004 | 10.629 | 10.629 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.629 |  |  | chlorodifluoromethane | 13.725 | 51 | 217m | 0.022 | 0.022 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | BDL | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 474 | 0.05 | 0.05 | ppb | 200 | 2.9152 | 2.9152 | 0.024044 | 0.070092 | 0.14576 | [17 empty] | chlorodifluoromethane | 13.74 | 51 | 499m | 0.052 | 0.052 | ppb | 200 | 2.9013 | 2.9013 | 0.024044 | 0.069757 | 0.15087 | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 419m | 0.045 | 0.045 | ppb | 200 | 3 | 3 | 0.024044 | 0.072131 | 0.135 | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 452 | 0.049 | 0.049 | ppb | 200 | 2.8791 | 2.8791 | 0.024044 | 0.069224 | 0.14107 | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 437m | 0.048 | 0.048 | ppb | 200 | 2.8791 | 2.8791 | 0.024044 | 0.069224 | 0.1382 | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 505 | 0.055 | 0.055 | ppb | 200 | 2.9096 | 2.9096 | 0.024044 | 0.069958 | 0.16003 | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 99617 | 10.515 | 10.515 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.515 |  |  | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | BDL
1,1,1,2-tetrafluoroethane | 15.939 | 33 | 613 | 0.189 | 0.189 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 0.189 | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL | [32 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 34201 | 10.886 | 10.886 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.886 |  |  | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 121m | 0.039 | 0.039 | ppb | 200 | 2.9152 | 2.9152 | 0.02815 | 0.082064 | 0.11369 | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.02815 | 0.081672 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 109m | 0.036 | 0.036 | ppb | 200 | 3 | 3 | 0.02815 | 0.084451 | 0.108 | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 141m | 0.046 | 0.046 | ppb | 200 | 2.8791 | 2.8791 | 0.02815 | 0.081047 | 0.13244 | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 105m | 0.035 | 0.035 | ppb | 200 | 2.8791 | 2.8791 | 0.02815 | 0.081047 | 0.10077 | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.02815 | 0.081907 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 33713 | 10.764 | 10.764 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.764 |  |  | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL
decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [32 empty] | decafluorobutane | 16.161 | 119 | 174318 | 10.229 | 10.229 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.229 |  |  | decafluorobutane | 16.161 | 119 | 94m | 0.005 | 0.005 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.023141 | 0.067462 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.023141 | 0.067139 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.023141 | 0.069424 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.023141 | 0.066626 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.023141 | 0.066626 | BDL | [17 empty] | decafluorobutane | 16.161 | 119 | 359 | 0.022 | 0.022 | ppb | 200 | 2.9096 | 2.9096 | 0.023141 | 0.067333 | BDL | [17 empty] | decafluorobutane | 16.161 | 119 | 170042 | 10.008 | 10.008 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.008 |  |  | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL
1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [32 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 150470 | 10.71 | 10.71 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.71 |  |  | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 18.112 | 51 | 109669 | 7.877 | 7.877 | ppb | 200 | 2.9152 | 2.9152 | 0.016117 | 0.046984 | 22.9631 | [17 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 549 | 0.038 | 0.038 | ppb | 200 | 2.9013 | 2.9013 | 0.016117 | 0.04676 | 0.11025 | [17 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 254 | 0.019 | 0.019 | ppb | 200 | 3 | 3 | 0.016117 | 0.048351 | 0.057 | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.016117 | 0.046402 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.016117 | 0.046402 | BDL | [17 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 1400 | 0.104 | 0.104 | ppb | 200 | 2.9096 | 2.9096 | 0.016117 | 0.046894 | 0.3026 | [17 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 151320 | 10.803 | 10.803 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.803 |  |  | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL
octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [32 empty] | octafluorocyclopentene | 21.943 | 93 | 245432 | 10.553 | 10.553 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.553 |  |  | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.0051927 | 0.015138 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.0051927 | 0.015065 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.0051927 | 0.015578 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.0051927 | 0.01495 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.0051927 | 0.01495 | BDL | [17 empty] | octafluorocyclopentene | 21.931 | 93 | 1476 | 0.066 | 0.066 | ppb | 200 | 2.9096 | 2.9096 | 0.0051927 | 0.015109 | 0.19203 | [17 empty] | octafluorocyclopentene | 21.931 | 93 | 236496 | 10.2 | 10.2 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.2 |  |  | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL
trichlorofluoromethane | 21.24 | 101 | 6649 | 0.278 | 0.278 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 0.278 | [17 empty] | trichlorofluoromethane | 21.24 | 101 | 82m | 0.003 | 0.003 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL | [32 empty] | trichlorofluoromethane | 21.24 | 101 | 245935 | 10.578 | 10.578 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.578 |  |  | trichlorofluoromethane | 21.24 | 101 | 243m | 0.01 | 0.01 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 881m | 0.038 | 0.038 | ppb | 200 | 2.9152 | 2.9152 | 0.028098 | 0.08191 | 0.11078 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 1046m | 0.044 | 0.044 | ppb | 200 | 2.9013 | 2.9013 | 0.028098 | 0.081518 | 0.12766 | [17 empty] | trichlorofluoromethane | 21.24 | 101 | 843m | 0.037 | 0.037 | ppb | 200 | 3 | 3 | 0.028098 | 0.084293 | 0.111 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 916m | 0.04 | 0.04 | ppb | 200 | 2.8791 | 2.8791 | 0.028098 | 0.080895 | 0.11516 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 888m | 0.04 | 0.04 | ppb | 200 | 2.8791 | 2.8791 | 0.028098 | 0.080895 | 0.11516 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 952 | 0.043 | 0.043 | ppb | 200 | 2.9096 | 2.9096 | 0.028098 | 0.081753 | 0.12511 | [17 empty] | trichlorofluoromethane | 21.24 | 101 | 237303 | 10.239 | 10.239 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.239 |  |  | trichlorofluoromethane | 21.239 | 101 | 241m | 0.01 | 0.01 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL
dodecafluoro-n-pentane | 21.012 | 69 | 250m | 0.005 | 0.005 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [32 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 591803 | 11.168 | 11.168 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 11.168 |  |  | dodecafluoro-n-pentane | 21.012 | 69 | 335m | 0.006 | 0.006 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.010934 | 0.031874 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.010934 | 0.031721 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.010934 | 0.032801 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.010934 | 0.031479 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.010934 | 0.031479 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.010934 | 0.031813 | BDL | [17 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 584950 | 11.073 | 11.073 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 11.073 |  |  | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL
1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [32 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 236240 | 11.287 | 11.287 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.287 |  |  | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.012238 | 0.035676 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.012238 | 0.035505 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.012238 | 0.036714 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.012238 | 0.035234 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.012238 | 0.035234 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.012238 | 0.035608 | BDL | [17 empty] | 1H-Nonafluorobutane | 22.79 | 51 | 229369 | 10.993 | 10.993 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.993 |  |  | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL
tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [32 empty] | tetradecafluorohexane | 24.408 | 69 | 609180m | 10.807 | 10.807 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.807 |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.045372 | 0.13227 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.045372 | 0.13164 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.045372 | 0.13612 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.045372 | 0.13063 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.045372 | 0.13063 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.045372 | 0.13202 | BDL | [17 empty] | tetradecafluorohexane | 24.408 | 69 | 595209m | 10.591 | 10.591 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.591 |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL
1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [32 empty] | 1H-Perfluoropentane | 26.066 | 69 | 208076m | 11.454 | 11.454 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 11.454 |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.02563 | 0.074715 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.02563 | 0.074358 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.02563 | 0.076889 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.02563 | 0.07379 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.02563 | 0.07379 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.02563 | 0.074572 | BDL | [17 empty] | 1H-Perfluoropentane | 26.066 | 69 | 194714m | 10.751 | 10.751 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.751 |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL
heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [32 empty] | heptafluoropropyl 1,2,... | 26.144 | 101 | 499494m | 10.814 | 10.814 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.814 |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.0093935 | 0.027384 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.0093935 | 0.027253 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.0093935 | 0.028181 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.0093935 | 0.027045 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.0093935 | 0.027045 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.0093935 | 0.027332 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 26.144 | 101 | 496995m | 10.793 | 10.793 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.793 |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL
hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [32 empty] | hexadecafluoroheptane | 27.255 | 69 | 704131 | 10.61 | 10.61 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.61 |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.007102 | 0.020704 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.007102 | 0.020605 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.007102 | 0.021306 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.007102 | 0.020447 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.007102 | 0.020447 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.007102 | 0.020664 | BDL | [17 empty] | hexadecafluoroheptane | 27.255 | 69 | 686926 | 10.383 | 10.383 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.383 |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL
1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [32 empty] | 1H-Perfluorohexane | 28.718 | 51 | 343655 | 10.848 | 10.848 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.848 |  |  | 1H-Perfluorohexane | 28.747 | 51 | 238m | 0.007 | 0.007 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.0083342 | 0.024296 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.0083342 | 0.02418 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.0083342 | 0.025003 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.0083342 | 0.023995 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.0083342 | 0.023995 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.0083342 | 0.024249 | BDL | [17 empty] | 1H-Perfluorohexane | 28.718 | 51 | 337809 | 10.696 | 10.696 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.696 |  |  | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL
octadecafluorooctane | 29.503 | 119 | 103m | 0.002 | 0.002 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [32 empty] | octadecafluorooctane | 29.514 | 119 | 408181 | 9.921 | 9.921 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.921 |  |  | octadecafluorooctane | 29.56 | 119 | 446 | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 29.548 | 119 | 106m | 0.003 | 0.003 | ppb | 200 | 2.9152 | 2.9152 | 0.027692 | 0.080729 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.027692 | 0.080343 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.027692 | 0.083077 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.027692 | 0.079729 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.027692 | 0.079729 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.027692 | 0.080574 | BDL | [17 empty] | octadecafluorooctane | 29.514 | 119 | 400089 | 9.754 | 9.754 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.754 |  |  | octadecafluorooctane | 29.569 | 119 | 419m | 0.01 | 0.01 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL
1H-Perfluoroheptane | 31.055 | 51 | 431 | 0.005 | 0.005 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL | [32 empty] | 1H-Perfluoroheptane | 31.055 | 51 | 933689 | 10.582 | 10.582 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.582 |  |  | 1H-Perfluoroheptane | 31.09 | 51 | 987 | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 0.011 | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.010647 | 0.031039 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.010647 | 0.030891 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.010647 | 0.031942 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.010647 | 0.030655 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.010647 | 0.030655 | BDL | [17 empty] | 1H-Perfluoroheptane | 31.09 | 51 | 2452 | 0.029 | 0.029 | ppb | 200 | 2.9096 | 2.9096 | 0.010647 | 0.03098 | 0.084379 | [17 empty] | 1H-Perfluoroheptane | 31.055 | 51 | 921534 | 10.477 | 10.477 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.477 |  |  | 1H-Perfluoroheptane | 31.065 | 51 | 950 | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 0.011
1H-Perfluorooctane | 33.505 | 51 | 2142m | 0.031 | 0.031 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.031 | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | BDL | [32 empty] | 1H-Perfluorooctane | 33.482 | 51 | 698436 | 10.327 | 10.327 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.327 |  |  | 1H-Perfluorooctane | 33.505 | 51 | 6654 | 0.096 | 0.096 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.096 | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.028576 | 0.083304 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.028576 | 0.082905 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.028576 | 0.085727 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.028576 | 0.082272 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.028576 | 0.082272 | BDL | [17 empty] | 1H-Perfluorooctane | 33.344 | 51 | 411 | 0.006 | 0.006 | ppb | 200 | 2.9096 | 2.9096 | 0.028576 | 0.083144 | BDL | [17 empty] | 1H-Perfluorooctane | 33.47 | 51 | 687014 | 10.189 | 10.189 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.189 |  |  | 1H-Perfluorooctane | 33.48 | 51 | 9619m | 0.14 | 0.14 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.14
2H-peffluoro-5-methyl-... | 33.056 | 169 | 1112 | 0.006 | 0.006 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [32 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 1765997 | 9.44 | 9.44 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.44 |  |  | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 2890 | 0.015 | 0.015 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9152 | 2.9152 | 0.037129 | 0.10824 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9013 | 2.9013 | 0.037129 | 0.10772 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 3 | 3 | 0.037129 | 0.11139 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.037129 | 0.1069 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.8791 | 2.8791 | 0.037129 | 0.1069 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 2.9096 | 2.9096 | 0.037129 | 0.10803 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 1681240 | 9.014 | 9.014 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.014 |  |  | 2H-peffluoro-5-methyl-... | 33.031 | 169 | 3224m | 0.017 | 0.017 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL

## 120324
120324/120324_01.D/epatemp.txt (QT Reviewed): AirBlank_Vol 200 ml | [14 empty] | 120324/120324_02.D/epatemp.txt (QT Reviewed): Rainbow BatchBlank_CF4_Vol 20 ml | [14 empty] | 120324/120324_03.D/epatemp.txt (QT Reviewed): Rainbow BatchBlank_Rest_Vol 200 ml | [14 empty] | 120324/120324_04.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_initial_100ppbv_vol 20 ml | [14 empty] | 120324/120324_05.D/epatemp.txt (Not Reviewed): Rainbow Post_BatchBlankCK_CF4_vol 20 ml | [14 empty] | 120324/120324_06.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_initial_10 ppbv_vol 200 ml | [14 empty] | 120324/120324_07.D/epatemp.txt (QT Reviewed): Rainbow Post_BatchBlankCK_Rest_vol 200 ml | [14 empty] | 120324/120324_08.D/epatemp.txt (QT Reviewed): Rainbow Run4_20x dilution_Port10_PFBA inj_vol 200 mL | [14 empty] | 120324/120324_09.D/epatemp.txt (QT Reviewed): Rainbow Run8__20xdilution_Port10_PFOA inj_vol 200 mL | [14 empty] | 120324/120324_10.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [14 empty] | 120324/120324_11.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [14 empty] | 120324/120324_12.D/epatemp.txt (QT Reviewed): Post BatchBlankCK_Rest 200 mL | [14 empty] | 120324/120324_13.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_100ppbv_vol 20 ml | [14 empty] | 120324/120324_14.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_10 ppbv_vol 200 ml | [14 empty] | 120324/120324_15.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_100ppbv_vol 20 ml | [14 empty] | 120324/120324_16.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_10 ppbv_vol 200 ml | [14 empty] | 120324/120324_17.D/epatemp.txt (QT Reviewed): AirBlank_Vol 200 ml | [14 empty] | CF4 Method |  | Lower Limit | Uppper Limit | Rest Method |  | Lower Limit | Uppper Limit
Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | ISTD (Pass/Fail) | RT (Pass/Fail) | Average IS area | 1053174 | 631904.4 | 1474443.6 | Average IS area | 1032225 | 619335 | 1445115
1,4-difluorobenzene | 30.308 | 114 | 923308 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.457 | 114 | 919831 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.4 | 114 | 930217 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.48 | 114 | 940577 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.457 | 114 | 966603 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.342 | 114 | 947736 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.377 | 114 | 958011 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.388 | 114 | 962127 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.4 | 114 | 969258 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.457 | 114 | 961607 | 50 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.319 | 114 | 957751 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.388 | 114 | 974921 | 5 |  | ppb | [6 empty] | Pass | Pass | 1,4-difluorobenzene | 30.457 | 114 | 1610785 | 50 |  | ppb | [6 empty] | Fail | Pass | 1,4-difluorobenzene | 30.331 | 114 | 1632766 | 5 |  | ppb | [6 empty] | Fail | Pass | 1,4-difluorobenzene | 30.457 | 114 | 1619642 | 50 |  | ppb | [6 empty] | Fail | Pass | 1,4-difluorobenzene | 30.4 | 114 | 1567572 | 5 |  | ppb | [6 empty] | Fail | Pass | 1,4-difluorobenzene | 30.308 | 114 | 1648653 | 5 |  | ppb | [6 empty] | Fail | Pass | Average RT | 30.503 | 30.4197 | 30.5863 | Average RT | 30.411 | 30.3277 | 30.4943
hexafluoroethane | 8.093 | 119 | 111m | 0.007 | 0.007 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Fail | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | BDL | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 6.982 | 69 | 61007 | 84.393 | 84.393 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 84.393 | Pass | Pass | tetrafluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 143478m | 9.368 | 9.368 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.368 | Pass | Pass | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | hexafluoroethane | 8.093 | 119 | 172m | 0.011 | 0.011 | ppb | 200 | 60.9505 | 60.9505 | 0.0084458 | 0.51478 | 0.67046 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 427m | 0.027 | 0.027 | ppb | 200 | 59.9016 | 59.9016 | 0.0084458 | 0.50592 | 1.6173 | Pass | Pass | tetrafluoromethane | 6.982 | 69 | 60393 | 81.716 | 81.716 | ppb | 20 | #REF! | #REF! | 0.089699 | #REF! | #REF! | Pass | Pass | hexafluoroethane | 8.093 | 119 | 151992 | 9.82 | 9.82 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.82 | Pass | Fail | hexafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Pass | tetrafluoromethane | 6.995 | 69 | 78411 | 63.337 | 63.337 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 63.337 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 202422 | 7.672 | 7.672 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 7.672 | Pass | Pass | tetrafluoromethane | 6.982 | 69 | 102555 | 82.387 | 82.387 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 82.387 | Pass | Pass | hexafluoroethane | 8.093 | 119 | 188373 | 7.436 | 7.436 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 7.436 | Pass | Pass | hexafluoroethane | 8.081 | 119 | 219m | 0.008 | 0.008 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | BDL | Pass | Fail
chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [32 empty] | chlorotrifluoromethane | 8.503 | 85 | 39853 | 10.013 | 10.013 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.013 |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.029884 | 1.8215 | BDL |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.029884 | 1.7901 | BDL | [17 empty] | chlorotrifluoromethane | 8.503 | 85 | 40944 | 10.179 | 10.179 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.179 |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL | [17 empty] | chlorotrifluoromethane | 8.503 | 85 | 54737 | 7.982 | 7.982 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 7.982 | [17 empty] | chlorotrifluoromethane | 8.503 | 85 | 50827 | 7.72 | 7.72 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 7.72 |  |  | chlorotrifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | BDL
trifluoromethane | 8.721 | 69 | 154m | 0.046 | 0.046 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 0.046 | [17 empty] | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL | [32 empty] | trifluoromethane | 8.708 | 69 | 35994 | 10.443 | 10.443 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.443 |  |  | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL |  |  | trifluoromethane | 8.696 | 69 | 9552 | 2.73 | 2.73 | ppb | 200 | 60.9505 | 60.9505 | 0.0356 | 2.1698 | 166.3949 |  |  | trifluoromethane | 8.697 | 69 | 6774 | 1.922 | 1.922 | ppb | 200 | 59.9016 | 59.9016 | 0.0356 | 2.1325 | 115.1308 | [17 empty] | trifluoromethane | 8.708 | 69 | 38590 | 11.079 | 11.079 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 11.079 |  |  | trifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | BDL | [17 empty] | trifluoromethane | 8.709 | 69 | 64802 | 10.913 | 10.913 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.913 | [17 empty] | trifluoromethane | 8.696 | 69 | 58034 | 10.18 | 10.18 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.18 |  |  | trifluoromethane | 8.697 | 69 | 387m | 0.065 | 0.065 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 0.065
octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [32 empty] | octafluoropropane | 11.178 | 69 | 310451 | 10.069 | 10.069 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.069 |  |  | octafluoropropane | 11.178 | 69 | 478 | 0.015 | 0.015 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL |  |  | octafluoropropane | 11.178 | 69 | 211m | 0.007 | 0.007 | ppb | 200 | 60.9505 | 60.9505 | 0.015822 | 0.96434 | BDL |  |  | octafluoropropane | 11.164 | 69 | 431 | 0.014 | 0.014 | ppb | 200 | 59.9016 | 59.9016 | 0.015822 | 0.94774 | BDL | [17 empty] | octafluoropropane | 11.178 | 69 | 324804 | 10.424 | 10.424 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.424 |  |  | octafluoropropane | 11.178 | 69 | 361 | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL | [17 empty] | octafluoropropane | 11.178 | 69 | 524511 | 9.875 | 9.875 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 9.875 | [17 empty] | octafluoropropane | 11.178 | 69 | 488509 | 9.579 | 9.579 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 9.579 |  |  | octafluoropropane | 11.164 | 69 | 816m | 0.015 | 0.015 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | BDL
difluoromethane | 10.975 | 51 | 143m | 0.063 | 0.063 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 0.063 | [17 empty] | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL | [32 empty] | difluoromethane | 10.961 | 51 | 23511 | 10.123 | 10.123 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.123 |  |  | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL |  |  | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.007355 | 0.44829 | BDL |  |  | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.007355 | 0.44058 | BDL | [17 empty] | difluoromethane | 10.961 | 51 | 24618 | 10.489 | 10.489 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.489 |  |  | difluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | BDL | [17 empty] | difluoromethane | 10.961 | 51 | 47550 | 11.883 | 11.883 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 11.883 | [17 empty] | difluoromethane | 10.946 | 51 | 41927 | 10.914 | 10.914 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.914 |  |  | difluoromethane | 10.961 | 51 | 1216 | 0.301 | 0.301 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 0.301
1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 527m | 0.063 | 0.063 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.063 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | BDL | [32 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 89435 | 10.342 | 10.342 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.342 |  |  | 1,1,1,2,2-pentafluoroe... | 12.813 | 51 | 105m | 0.012 | 0.012 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | BDL |  |  | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 11874 | 1.353 | 1.353 | ppb | 200 | 60.9505 | 60.9505 | 0.013843 | 0.84374 | 82.4661 |  |  | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 440m | 0.05 | 0.05 | ppb | 200 | 59.9016 | 59.9016 | 0.013843 | 0.82922 | 2.9951 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 93653 | 10.716 | 10.716 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.716 |  |  | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 97m | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | BDL | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 162839 | 10.93 | 10.93 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.93 | [17 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 149364 | 10.442 | 10.442 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.442 |  |  | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 2831 | 0.188 | 0.188 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 0.188
1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [32 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 301091 | 10.164 | 10.164 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.164 |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.911 | 100 | 274 | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL |  |  | 1,1,2,2,3,3,4,4-octafl... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.013563 | 0.82664 | BDL |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 123m | 0.004 | 0.004 | ppb | 200 | 59.9016 | 59.9016 | 0.013563 | 0.81242 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 309157 | 10.327 | 10.327 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.327 |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 273m | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | BDL | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 460072 | 9.015 | 9.015 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 9.015 | [17 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 433420 | 8.846 | 8.846 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 8.846 |  |  | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 841 | 0.016 | 0.016 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 0.016
fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [32 empty] | fluoromethane | 11.409 | 33 | 10542 | 10.462 | 10.462 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.462 |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.333 | 20.2965 | BDL |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.333 | 19.9472 | BDL | [17 empty] | fluoromethane | 11.409 | 33 | 11197 | 10.996 | 10.996 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.996 |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL | [17 empty] | fluoromethane | 11.41 | 33 | 23638 | 13.617 | 13.617 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 13.617 | [17 empty] | fluoromethane | 11.395 | 33 | 21599 | 12.959 | 12.959 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 12.959 |  |  | fluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | BDL
tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [32 empty] | tetrafluoroethylene | 8.491 | 81 | 65773 | 9.881 | 9.881 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 9.881 |  |  | tetrafluoroethylene | 8.491 | 81 | 71m | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL |  |  | tetrafluoroethylene | 8.491 | 81 | 749 | 0.111 | 0.111 | ppb | 200 | 60.9505 | 60.9505 | 0.019104 | 1.1644 | 6.7655 |  |  | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.019104 | 1.1443 | BDL | [17 empty] | tetrafluoroethylene | 8.491 | 81 | 67931 | 10.099 | 10.099 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.099 |  |  | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL | [17 empty] | tetrafluoroethylene | 8.491 | 81 | 107802 | 9.401 | 9.401 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 9.401 | [17 empty] | tetrafluoroethylene | 8.491 | 81 | 102340 | 9.295 | 9.295 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 9.295 |  |  | tetrafluoroethylene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | BDL
hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [32 empty] | hexafluoropropene | 13.247 | 131 | 140498 | 9.528 | 9.528 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.528 |  |  | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL |  |  | hexafluoropropene | 13.233 | 131 | 451 | 0.03 | 0.03 | ppb | 200 | 60.9505 | 60.9505 | 0.0086848 | 0.52934 | 1.8285 |  |  | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.0086848 | 0.52023 | BDL | [17 empty] | hexafluoropropene | 13.248 | 131 | 141708 | 9.51 | 9.51 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.51 |  |  | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL | [17 empty] | hexafluoropropene | 13.248 | 131 | 195129 | 7.681 | 7.681 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 7.681 | [17 empty] | hexafluoropropene | 13.248 | 131 | 181193 | 7.429 | 7.429 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 7.429 |  |  | hexafluoropropene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | BDL
1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [32 empty] | 1,1,1-trifluoroethane | 13.218 | 64 | 8909 | 9.911 | 9.911 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.911 |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.034251 | 2.0876 | BDL |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.034251 | 2.0517 | BDL | [17 empty] | 1,1,1-trifluoroethane | 13.219 | 64 | 9070 | 9.984 | 9.984 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.984 |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL | [17 empty] | 1,1,1-trifluoroethane | 13.219 | 64 | 13344 | 8.616 | 8.616 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 8.616 | [17 empty] | 1,1,1-trifluoroethane | 13.219 | 64 | 12652 | 8.509 | 8.509 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 8.509 |  |  | 1,1,1-trifluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | BDL
hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [32 empty] | hexafluoropropene_oxide | 13.421 | 69 | 9101m | 9.59 | 9.59 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.59 |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.33918 | 20.6734 | BDL |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.33918 | 20.3176 | BDL | [17 empty] | hexafluoropropene_oxide | 13.421 | 69 | 9160 | 9.551 | 9.551 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.551 |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL | [17 empty] | hexafluoropropene_oxide | 13.421 | 69 | 15674 | 9.587 | 9.587 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.587 | [17 empty] | hexafluoropropene_oxide | 13.421 | 69 | 11479 | 7.313 | 7.313 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 7.313 |  |  | hexafluoropropene_oxide | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | BDL
chlorodifluoromethane | 13.725 | 51 | 2234 | 0.26 | 0.26 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 0.26 | [17 empty] | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | BDL | [32 empty] | chlorodifluoromethane | 13.725 | 51 | 91033 | 10.341 | 10.341 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.341 |  |  | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | BDL |  |  | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.024044 | 1.4655 | BDL |  |  | chlorodifluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.024044 | 1.4403 | BDL | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 94482 | 10.621 | 10.621 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.621 |  |  | chlorodifluoromethane | 13.725 | 51 | 176m | 0.019 | 0.019 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | BDL | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 163017 | 10.749 | 10.749 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.749 | [17 empty] | chlorodifluoromethane | 13.725 | 51 | 152221 | 10.455 | 10.455 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.455 |  |  | chlorodifluoromethane | 13.725 | 51 | 11040 | 0.721 | 0.721 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 0.721
1,1,1,2-tetrafluoroethane | 15.925 | 33 | 446m | 0.157 | 0.157 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 0.157 | [17 empty] | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL | [32 empty] | 1,1,1,2-tetrafluoroethane | 15.911 | 33 | 30445 | 10.461 | 10.461 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.461 |  |  | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL |  |  | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.02815 | 1.7158 | BDL |  |  | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.02815 | 1.6863 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.911 | 33 | 32270 | 10.972 | 10.972 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.972 |  |  | 1,1,1,2-tetrafluoroethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | BDL | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.911 | 33 | 63208 | 12.607 | 12.607 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 12.607 | [17 empty] | 1,1,1,2-tetrafluoroethane | 15.911 | 33 | 58577 | 12.169 | 12.169 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 12.169 |  |  | 1,1,1,2-tetrafluoroethane | 15.911 | 33 | 1568m | 0.31 | 0.31 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 0.31
decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [32 empty] | decafluorobutane | 16.161 | 119 | 160648 | 10.176 | 10.176 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.176 |  |  | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL |  |  | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.023141 | 1.4105 | BDL |  |  | decafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.023141 | 1.3862 | BDL | [17 empty] | decafluorobutane | 16.161 | 119 | 164011 | 10.28 | 10.28 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.28 |  |  | decafluorobutane | 16.147 | 119 | 101m | 0.006 | 0.006 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL | [17 empty] | decafluorobutane | 16.161 | 119 | 245737 | 9.035 | 9.035 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 9.035 | [17 empty] | decafluorobutane | 16.161 | 119 | 225658 | 8.642 | 8.642 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 8.642 |  |  | decafluorobutane | 16.175 | 119 | 286m | 0.01 | 0.01 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | BDL
1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [32 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 135104 | 10.38 | 10.38 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.38 |  |  | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL |  |  | 1H-Heptafluoropropane | 18.095 | 51 | 5414 | 0.41 | 0.41 | ppb | 200 | 60.9505 | 60.9505 | 0.016117 | 0.98234 | 24.9897 |  |  | 1H-Heptafluoropropane | 18.095 | 51 | 397m | 0.03 | 0.03 | ppb | 200 | 59.9016 | 59.9016 | 0.016117 | 0.96543 | 1.797 | [17 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 142107 | 10.804 | 10.804 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.804 |  |  | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL | [17 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 250903 | 11.19 | 11.19 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 11.19 | [17 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 234237 | 10.881 | 10.881 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.881 |  |  | 1H-Heptafluoropropane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | BDL
octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [32 empty] | octafluorocyclopentene | 21.931 | 93 | 223215 | 10.361 | 10.361 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.361 |  |  | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL |  |  | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.0051927 | 0.31649 | BDL |  |  | octafluorocyclopentene | 21.931 | 93 | 89m | 0.004 | 0.004 | ppb | 200 | 59.9016 | 59.9016 | 0.0051927 | 0.31105 | BDL | [17 empty] | octafluorocyclopentene | 21.931 | 93 | 228756 | 10.507 | 10.507 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.507 |  |  | octafluorocyclopentene | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL | [17 empty] | octafluorocyclopentene | 21.931 | 93 | 348454 | 9.388 | 9.388 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 9.388 | [17 empty] | octafluorocyclopentene | 21.931 | 93 | 300875 | 8.444 | 8.444 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 8.444 |  |  | octafluorocyclopentene | 21.907 | 93 | 108m | 0.003 | 0.003 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | BDL
trichlorofluoromethane | 21.24 | 101 | 6012 | 0.287 | 0.287 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 0.287 | [17 empty] | trichlorofluoromethane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL | [32 empty] | trichlorofluoromethane | 21.239 | 101 | 227323 | 10.555 | 10.555 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.555 |  |  | trichlorofluoromethane | 21.239 | 101 | 223m | 0.01 | 0.01 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL |  |  | trichlorofluoromethane | 21.222 | 101 | 169m | 0.008 | 0.008 | ppb | 200 | 60.9505 | 60.9505 | 0.028098 | 1.7126 | BDL |  |  | trichlorofluoromethane | 21.24 | 101 | 133m | 0.006 | 0.006 | ppb | 200 | 59.9016 | 59.9016 | 0.028098 | 1.6831 | BDL | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 232544 | 10.685 | 10.685 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.685 |  |  | trichlorofluoromethane | 21.239 | 101 | 186m | 0.008 | 0.008 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | BDL | [17 empty] | trichlorofluoromethane | 21.24 | 101 | 323113 | 8.708 | 8.708 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 8.708 | [17 empty] | trichlorofluoromethane | 21.239 | 101 | 279880 | 7.857 | 7.857 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 7.857 |  |  | trichlorofluoromethane | 21.24 | 101 | 9146 | 0.244 | 0.244 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 0.244
dodecafluoro-n-pentane | 20.995 | 69 | 233m | 0.005 | 0.005 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [32 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 529884 | 10.794 | 10.794 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.794 |  |  | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL |  |  | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.010934 | 0.66641 | BDL |  |  | dodecafluoro-n-pentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.010934 | 0.65494 | BDL | [17 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 557251 | 11.233 | 11.233 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 11.233 |  |  | dodecafluoro-n-pentane | 21.012 | 69 | 310m | 0.006 | 0.006 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL | [17 empty] | dodecafluoro-n-pentane | 21.013 | 69 | 929977 | 10.997 | 10.997 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.997 | [17 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 785593 | 9.676 | 9.676 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 9.676 |  |  | dodecafluoro-n-pentane | 21.013 | 69 | 835m | 0.01 | 0.01 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | BDL
1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [32 empty] | 1H-Nonafluorobutane | 22.789 | 51 | 214533 | 11.065 | 11.065 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.065 |  |  | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL |  |  | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.012238 | 0.7459 | BDL |  |  | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.012238 | 0.73307 | BDL | [17 empty] | 1H-Nonafluorobutane | 22.79 | 51 | 220921 | 11.275 | 11.275 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.275 |  |  | 1H-Nonafluorobutane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | BDL | [17 empty] | 1H-Nonafluorobutane | 22.79 | 51 | 395634 | 11.844 | 11.844 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.844 | [17 empty] | 1H-Nonafluorobutane | 22.79 | 51 | 341136 | 10.638 | 10.638 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.638 |  |  | 1H-Nonafluorobutane | 22.79 | 51 | 485 | 0.014 | 0.014 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 0.014
tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [32 empty] | tetradecafluorohexane | 24.392 | 69 | 543626m | 10.41 | 10.41 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.41 |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.045372 | 2.7655 | BDL |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.045372 | 2.7179 | BDL | [17 empty] | tetradecafluorohexane | 24.353 | 69 | 556165m | 10.539 | 10.539 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.539 |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL | [17 empty] | tetradecafluorohexane | 24.321 | 69 | 901202m | 10.017 | 10.017 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.017 | [17 empty] | tetradecafluorohexane | 24.535 | 69 | 900513m | 10.426 | 10.426 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.426 |  |  | tetradecafluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | BDL
1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [32 empty] | 1H-Perfluoropentane | 26.056 | 69 | 180190m | 10.707 | 10.707 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.707 |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.02563 | 1.5621 | BDL |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.02563 | 1.5353 | BDL | [17 empty] | 1H-Perfluoropentane | 26.027 | 69 | 185143m | 10.887 | 10.887 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.887 |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL | [17 empty] | 1H-Perfluoropentane | 26.017 | 69 | 318063m | 10.971 | 10.971 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.971 | [17 empty] | 1H-Perfluoropentane | 26.124 | 69 | 274364 | 9.857 | 9.857 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 9.857 |  |  | 1H-Perfluoropentane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | BDL
heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [32 empty] | heptafluoropropyl 1,2,... | 26.124 | 101 | 455105m | 10.636 | 10.636 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.636 |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.0093935 | 0.57254 | BDL |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.0093935 | 0.56269 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 26.105 | 101 | 464944m | 10.753 | 10.753 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.753 |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL | [17 empty] | heptafluoropropyl 1,2,... | 26.095 | 101 | 676018m | 9.171 | 9.171 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 9.171 | [17 empty] | heptafluoropropyl 1,2,... | 26.193 | 101 | 648343m | 9.161 | 9.161 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 9.161 |  |  | heptafluoropropyl 1,2,... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | BDL
hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [32 empty] | hexadecafluoroheptane | 27.245 | 69 | 640340 | 10.416 | 10.416 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.416 |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.007102 | 0.43287 | BDL |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.007102 | 0.42542 | BDL | [17 empty] | hexadecafluoroheptane | 27.226 | 69 | 661489 | 10.648 | 10.648 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.648 |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL | [17 empty] | hexadecafluoroheptane | 27.226 | 69 | 1090906 | 10.3 | 10.3 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.3 | [17 empty] | hexadecafluoroheptane | 27.304 | 69 | 1031276 | 10.142 | 10.142 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.142 |  |  | hexadecafluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | BDL
1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [32 empty] | 1H-Perfluorohexane | 28.708 | 51 | 309545 | 10.548 | 10.548 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.548 |  |  | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL |  |  | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.0083342 | 0.50797 | BDL |  |  | 1H-Perfluorohexane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.0083342 | 0.49923 | BDL | [17 empty] | 1H-Perfluorohexane | 28.688 | 51 | 320709 | 10.814 | 10.814 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.814 |  |  | 1H-Perfluorohexane | 28.747 | 51 | 189m | 0.006 | 0.006 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | BDL | [17 empty] | 1H-Perfluorohexane | 28.689 | 51 | 582037 | 11.513 | 11.513 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 11.513 | [17 empty] | 1H-Perfluorohexane | 28.757 | 51 | 535801 | 11.039 | 11.039 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 11.039 |  |  | 1H-Perfluorohexane | 28.698 | 51 | 462 | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 0.009
octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [32 empty] | octadecafluorooctane | 29.514 | 119 | 372714 | 9.779 | 9.779 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.779 |  |  | octadecafluorooctane | 29.549 | 119 | 411 | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL |  |  | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 60.9505 | 60.9505 | 0.027692 | 1.6879 | BDL |  |  | octadecafluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.027692 | 1.6588 | BDL | [17 empty] | octadecafluorooctane | 29.502 | 119 | 377638 | 9.805 | 9.805 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.805 |  |  | octadecafluorooctane | 29.548 | 119 | 356m | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL | [17 empty] | octadecafluorooctane | 29.503 | 119 | 553885 | 8.435 | 8.435 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 8.435 | [17 empty] | octadecafluorooctane | 29.548 | 119 | 563341 | 8.936 | 8.936 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 8.936 |  |  | octadecafluorooctane | 29.491 | 119 | 765m | 0.012 | 0.012 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | BDL
1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL | [17 empty] | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | BDL | [32 empty] | 1H-Perfluoroheptane | 31.055 | 51 | 825974 | 10.106 | 10.106 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.106 |  |  | 1H-Perfluoroheptane | 31.055 | 51 | 911 | 0.011 | 0.011 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 0.011 |  |  | 1H-Perfluoroheptane | 31.078 | 51 | 419 | 0.005 | 0.005 | ppb | 200 | 60.9505 | 60.9505 | 0.010647 | 0.64896 | BDL |  |  | 1H-Perfluoroheptane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.010647 | 0.63779 | BDL | [17 empty] | 1H-Perfluoroheptane | 31.044 | 51 | 855374 | 10.356 | 10.356 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.356 |  |  | 1H-Perfluoroheptane | 31.078 | 51 | 1027 | 0.012 | 0.012 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 0.012 | [17 empty] | 1H-Perfluoroheptane | 31.055 | 51 | 1522640 | 10.813 | 10.813 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.813 | [17 empty] | 1H-Perfluoroheptane | 31.067 | 51 | 1519667 | 11.241 | 11.241 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 11.241 |  |  | 1H-Perfluoroheptane | 31.055 | 51 | 1974m | 0.014 | 0.014 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 0.014
1H-Perfluorooctane | 33.493 | 51 | 539m | 0.009 | 0.009 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | BDL | [17 empty] | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | BDL | [32 empty] | 1H-Perfluorooctane | 33.47 | 51 | 612025 | 9.769 | 9.769 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 9.769 |  |  | 1H-Perfluorooctane | 33.482 | 51 | 8022 | 0.127 | 0.127 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.127 |  |  | 1H-Perfluorooctane | 33.482 | 51 | 1131 | 0.018 | 0.018 | ppb | 200 | 60.9505 | 60.9505 | 0.028576 | 1.7417 | BDL |  |  | 1H-Perfluorooctane | 0 |  | 0 | N.D. | 0 | ppb | 200 | 59.9016 | 59.9016 | 0.028576 | 1.7117 | BDL | [17 empty] | 1H-Perfluorooctane | 33.47 | 51 | 627515 | 9.911 | 9.911 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 9.911 |  |  | 1H-Perfluorooctane | 33.505 | 51 | 10328 | 0.16 | 0.16 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.16 | [17 empty] | 1H-Perfluorooctane | 33.494 | 51 | 1087049 | 10.071 | 10.071 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.071 | [17 empty] | 1H-Perfluorooctane | 33.47 | 51 | 1222040 | 11.793 | 11.793 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 11.793 |  |  | 1H-Perfluorooctane | 33.471 | 51 | 16405m | 0.151 | 0.151 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 0.151
2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 0 |  | 0 | N.D. | 0 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [32 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 1589566 | 9.172 | 9.172 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.172 |  |  | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 3000 | 0.017 | 0.017 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL |  |  | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 659 | 0.004 | 0.004 | ppb | 200 | 60.9505 | 60.9505 | 0.037129 | 2.2631 | BDL |  |  | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 492m | 0.003 | 0.003 | ppb | 200 | 59.9016 | 59.9016 | 0.037129 | 2.2241 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 33.01 | 169 | 1561311 | 8.915 | 8.915 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 8.915 |  |  | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 3338 | 0.019 | 0.019 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL | [17 empty] | 2H-peffluoro-5-methyl-... | 33.033 | 169 | 2172108 | 7.275 | 7.275 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 7.275 | [17 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 2390649 | 8.34 | 8.34 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 8.34 |  |  | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 5218 | 0.017 | 0.017 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | BDL

## Cal Curve Levels
Updated with Recertified Riemer values-Cal Levels Used | Updated 10/24 | [9 empty] | Nominal | October 2024 Riemer
Description | Cylinder Conc ppbv | Level 10 CF4 |  | Level 50 CF4 | Level 100 CF4 | Level 150 CF4 | Level 200 CF4 | Level 300 CF4 | Level 2000 CF4 | Level 3000 CF4 | Cylinder Conc ppbv | Correction factor
|  | 1 | 2.5 | 5 | 10 | 15 | 20 | 30 | 200 | 300
| ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | ppbv | 10000 | 0.9838
Tetrafluoromethane | 9838 | 9.838 | 24.595 | 49.19 | 98.38 | 147.57 | 196.76 | 295.14 | 1967.6 | 2951.4 | 1000 | 0.979
Hexafluoroethane (FC-116) | 979 | 0.979 | 2.4475 | 4.895 | 9.79 | 14.685 | 19.58 | 29.37 | 195.8 | 293.7 | 1000 | 0.984
Chlorotrifluoromethane (CFC-13) | 984 | 0.984 | 2.46 | 4.92 | 9.84 | 14.76 | 19.68 | 29.52 | 196.8 | 295.2 | 1000 | 0.986
Trifluoromethane (Fluoroform) | 986 | 0.986 | 2.465 | 4.93 | 9.86 | 14.79 | 19.72 | 29.58 | 197.2 | 295.8 | 1000 | 0.979
Octafluoropropane | 979 | 0.979 | 2.4475 | 4.895 | 9.79 | 14.685 | 19.58 | 29.37 | 195.8 | 293.7 | 1000 | 0.972
Difluoromethane (HFC-32) | 972 | 0.972 | 2.43 | 4.86 | 9.72 | 14.58 | 19.44 | 29.16 | 194.4 | 291.6 | 1000 | 0.989
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | 989 | 0.989 | 2.4725 | 4.945 | 9.89 | 14.835 | 19.78 | 29.67 | 197.8 | 296.7 | 1000 | 0.972
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | 972 | 0.972 | 2.43 | 4.86 | 9.72 | 14.58 | 19.44 | 29.16 | 194.4 | 291.6 | 1000 | 0.989
Fluoromethane (HFC-41) | 989 | 0.989 | 2.4725 | 4.945 | 9.89 | 14.835 | 19.78 | 29.67 | 197.8 | 296.7 | 1000 | 0.984
Tetrafluoroethylene | 984 | 0.984 | 2.46 | 4.92 | 9.84 | 14.76 | 19.68 | 29.52 | 196.8 | 295.2 | 1000 | 0.977
Hexafluoropropene | 977 | 0.977 | 2.4425 | 4.885 | 9.77 | 14.655 | 19.54 | 29.31 | 195.4 | 293.1 | 1000 | 0.974
1,1,1-trifluoroethane | 974 | 0.974 | 2.435 | 4.87 | 9.74 | 14.61 | 19.48 | 29.22 | 194.8 | 292.2 | 1000 | 0.984
Hexafluoropropene Oxide | 984 | 0.984 | 2.46 | 4.92 | 9.84 | 14.76 | 19.68 | 29.52 | 196.8 | 295.2 | 1000 | 0.989
Chlorodifluoromethane (HCFC-22) | 989 | 0.989 | 2.4725 | 4.945 | 9.89 | 14.835 | 19.78 | 29.67 | 197.8 | 296.7 | 1000 | 0.989
1,1,1,2-Tetrafluoroethane | 989 | 0.989 | 2.4725 | 4.945 | 9.89 | 14.835 | 19.78 | 29.67 | 197.8 | 296.7 | 1000 | 0.974
Decafluorobutane (Perfluorobutane) | 974 | 0.974 | 2.435 | 4.87 | 9.74 | 14.61 | 19.48 | 29.22 | 194.8 | 292.2 | 1000 | 0.984
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | 984 | 0.984 | 2.46 | 4.92 | 9.84 | 14.76 | 19.68 | 29.52 | 196.8 | 295.2 | 1000 | 0.974
Octafluorocyclopentene (FC-C1418) | 974 | 0.974 | 2.435 | 4.87 | 9.74 | 14.61 | 19.48 | 29.22 | 194.8 | 292.2 | 1000 | 0.986
Trichlorofluoromethane (CFC-11) | 986 | 0.986 | 2.465 | 4.93 | 9.86 | 14.79 | 19.72 | 29.58 | 197.2 | 295.8 | 1000 | 0.986
Dodecafluoro-n-pentane | 986 | 0.986 | 2.465 | 4.93 | 9.86 | 14.79 | 19.72 | 29.58 | 197.2 | 295.8 | 1000 | 0.991
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | 991 | 0.991 | 2.4775 | 4.955 | 9.91 | 14.865 | 19.82 | 29.73 | 198.2 | 297.3 | 1000 | 0.983
Tetradecafluorohexane | 983 | 0.983 | 2.4575 | 4.915 | 9.83 | 14.745 | 19.66 | 29.49 | 196.6 | 294.9 | 1000 | 0.99
1,1,1,2,2,3,3,4,4,5,5-Undecafluoropentane (1H-Perfluoropentane) | 990 | 0.99 | 2.475 | 4.95 | 9.9 | 14.85 | 19.8 | 29.7 | 198 | 297 | 1000 | 0.983
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | 983 | 0.983 | 2.4575 | 4.915 | 9.83 | 14.745 | 19.66 | 29.49 | 196.6 | 294.9 | 1000 | 0.988
Hexadecafluoroheptane | 988 | 0.988 | 2.47 | 4.94 | 9.88 | 14.82 | 19.76 | 29.64 | 197.6 | 296.4 | 1000 | 0.986
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | 986 | 0.986 | 2.465 | 4.93 | 9.86 | 14.79 | 19.72 | 29.58 | 197.2 | 295.8 | 1000 | 0.981
Octadecafluorooctane (Perfluorooctane) | 981 | 0.981 | 2.4525 | 4.905 | 9.81 | 14.715 | 19.62 | 29.43 | 196.2 | 294.3 | 1000 | 0.988
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | 988 | 0.988 | 2.47 | 4.94 | 9.88 | 14.82 | 19.76 | 29.64 | 197.6 | 296.4 | 1000 | 0.981
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | 981 | 0.981 | 2.4525 | 4.905 | 9.81 | 14.715 | 19.62 | 29.43 | 196.2 | 294.3 | 1000 | 0.988
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | 988 | 0.988 | 2.47 | 4.94 | 9.88 | 14.82 | 19.76 | 29.64 | 197.6 | 296.4

## Sample log Rainbow
Rainbow Fall 2024 OTM-50 |  | Sample Date | Log in date | L | L | L | Orifice vs Lab Received | C of Custody (" Hg) (RTP) | C of Custody (psia) (RTP) | C of Custody ("Hg) | C of Custody (psia) | Entech 4700 (psia) | Entech 4700 (psia) | Using lab data |  |  | Canister Leaving lab | 2024-11-25
Silonite Canister ID | Sample Name | DATE | DATE | Field vol sampled | Field vol sampled | Entech 4700 Received vol sampled | Field vol vs Received vol | Lab Pressure before sampling | Lab Pressure before sampling | Field Vacuum after sampling | Field Pressure after sampling | Lab reception pressure | Lab fill final pressure | dilution factor | CO2 Amount/Barometric Pressure/Comments | Field pressure Hg" | Lab presure psia" | EPA atm pressure 3 pm "Hg
|  |  |  | via orifice flow/sample duration | via field pressure | Using lab reception pressures | Percent Diff (%) |  |  |  |  | (11/25/24) | (11/25/24) | alt adj not necessary-same day compare
36834 | N2 Blank | 2024-11-19 | 2024-11-25 | 4.3 | 4.3957 | 4.3102 | -0.237% | 29.92 | 0 | 8 | 10.7695 | 10.56 | 30.03 | 2.8438 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
40882 | Combustion Blank | 2024-11-19 | 2024-11-25 | 4.3 | 4.1952 | 4.2122 | 2.062% | 29.92 | 0 | 9 | 10.2782 | 10.32 | 29.97 | 2.9041 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
40880 | Run 1 Port 4 PFBA inj | 2024-11-20 | 2024-11-25 | 4.3 | 4.2955 | 4.2163 | 1.965% | 29.92 | 0 | 8.5 | 10.5239 | 10.33 | 29.98 | 2.9022 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
40864 | Run 2 Port 6 PFBA inj | 2024-11-20 | 2024-11-25 | 4.3 | 4.3957 | 4.2286 | 1.675% | 29.92 | 0 | 8 | 10.7695 | 10.36 | 30.03 | 2.8986 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
40860 | Run 3 Port 8 PFBA inj | 2024-11-19 | 2024-11-25 | 4.3 | 4.3957 | 4.2531 | 1.098% | 29.92 | 0 | 8 | 10.7695 | 10.42 | 29.98 | 2.8772 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
40879 | Run 4 Port 10 PFBA inj | 2024-11-19 | 2024-11-25 | 4.3 | 4.2955 | 4.1878 | 2.645% | 29.92 | 0 | 8.5 | 10.5239 | 10.26 | 29.91 | 2.9152 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
42891 | Run 5 Port 4 PFOA inj | 2024-11-19 | 2024-11-25 | 4.3 | 4.2955 | 4.2163 | 1.965% | 29.92 | 0 | 8.5 | 10.5239 | 10.33 | 29.97 | 2.9013 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
40877 | Run 6 Port 6 PFOA inj | 2024-11-19 | 2024-11-25 | 4.3 | 4.3556 | 4.2163 | 1.965% | 29.92 | 0 | 8.2 | 10.6713 | 10.33 | 29.96 | 3 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
40869 | Run 7 Port 8 PFOA inj | 2024-11-19 | 2024-11-25 | 4.3 | 4.3957 | 4.2531 | 1.098% | 29.92 | 0 | 8 | 10.7695 | 10.42 | 30 | 2.8791 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
42896 | Run 8 Port 10 PFOA inj | 2024-11-19 | 2024-11-25 | 4.3 | 4.1952 | 4.2 | 2.353% | 29.92 | 0 | 9 | 10.2782 | 10.29 | 29.94 | 2.9096 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
1750* | Post Test N2 Blank* | 2024-11-19 | 2024-11-25 | 4.3 | 4.2955 | 5.6286 | -26.76% | 29.92 | 0 | 8.5 | 10.5239 | 13.79 | 30.03 | 2.1777 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
40865 | Post Test N2 Blank | 2024-11-19 | 2024-11-25 | 4.3 | 4.3957 | 4.2735 | 0.6189% | 29.92 | 0 | 8 | 10.7695 | 10.47 | 30.04 | 2.8691 | 7 % CO2/29.92" Hg COC Field values | 29.92 | 0.04 | 29.651
42881 | 20X Dilution of 40879 | [10 empty] | 1.41 | 29.48 | 20.9078
42873 | 20X Dilution of 42896 | [10 empty] | 1.43 | 29.44 | 20.5874
* Suspected sample field leak
Failed OTM-50 RPD <5% criterion

## CO2 Bias Results
Rainbow CO2 Bias Checks |  |  | 112624/112624_15.D/epatemp.txt (QT Reviewed): 2.5% CO2_10ppbv_CF4_20 mL | [12 empty] | Compare to 10 ppbv(multiplied by 10) | 112624/112624_16.D/epatemp.txt (QT Reviewed): 2.5% CO2_1ppbv_Rest_200 mL | [12 empty] | Compare to 10 ppbv(multiplied by 10) | 112624/112624_17.D/epatemp.txt (QT Reviewed): 2.5% CO2_300ppbv_CF4_20 mL | [12 empty] | Compare to 10 ppbv( divide by 3) | 112624/112624_18.D/epatemp.txt (QT Reviewed): 2.5% CO2_30ppbv_CF4_200 mL | [12 empty] | Compare to 10 ppbv( divide by 3) | 112624/112624_19.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [12 empty] | 112624/112624_20.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [12 empty] | 112724/112724_17.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [12 empty] | 112724/112724_18.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [12 empty] | 112924/112924_20.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [12 empty] | 112924/112924_21.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml | [12 empty] | 120324/120324_10.D/epatemp.txt (QT Reviewed): Riemer PFAS CF4_Cal ck_closing_100ppbv_vol 20 ml | [12 empty] | 120324/120324_11.D/epatemp.txt (QT Reviewed): Riemer PFAS Rest_Calck_closing_10 ppbv_vol 200 ml
|  |  | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Adjusted ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Adjusted ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Adjusted ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Adjusted ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb | Compound | R.T. | QIon | Response | Conc | num Conc | Units | Volume (ml) | Air Dil. Factor | Net Dilution | MDL | Net MDL | Actual ppb
|  |  | 1,4-difluorobenzene | 30.48 | 114 | 1251881 | 50 |  | ppb | [7 empty] | 1,4-difluorobenzene | 30.423 | 114 | 1063954 | 5 |  | ppb | [7 empty] | 1,4-difluorobenzene | 30.48 | 114 | 1204137 | 50 |  | ppb | [7 empty] | 1,4-difluorobenzene | 30.423 | 114 | 1032864 | 5 |  | ppb | [7 empty] | 1,4-difluorobenzene | 30.48 | 114 | 1072880 | 50 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.4 | 114 | 1043367 | 5 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.469 | 114 | 1094387 | 50 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.388 | 114 | 1059081 | 5 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.457 | 114 | 1048353 | 50 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.354 | 114 | 1019948 | 5 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.457 | 114 | 961607 | 50 |  | ppb | [6 empty] | 1,4-difluorobenzene | 30.319 | 114 | 957751 | 5 |  | ppb
DATE | 2025-11-26 | 112624 | tetrafluoromethane | 6.958 | 69 | 8080 | 8.398 | 8.398 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 8.398 | 83.98 | hexafluoroethane | 8.093 | 119 | 15860 | 0.922 | 0.922 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 0.922 | 9.22 | tetrafluoromethane | 6.958 | 69 | 225778 | 243.964 | 243.964 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 243.964 | 81.3213 | hexafluoroethane | 8.093 | 119 | 433248 | 25.957 | 25.957 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 25.957 | 8.6523 | tetrafluoromethane | 6.995 | 69 | 81010 | 98.244 | 98.244 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 98.244 | hexafluoroethane | 8.093 | 119 | 162256 | 9.623 | 9.623 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.623 | tetrafluoromethane | 6.995 | 69 | 80568 | 95.788 | 95.788 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 95.788 | hexafluoroethane | 8.093 | 119 | 168067 | 9.82 | 9.82 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.82 | tetrafluoromethane | 6.995 | 69 | 72180 | 89.584 | 89.584 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 89.584 | hexafluoroethane | 8.093 | 119 | 158786 | 9.634 | 9.634 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.634 | tetrafluoromethane | 6.982 | 69 | 60393 | 81.716 | 81.716 | ppb | 20 | 1 | 1 | 0.089699 | 0.089699 | 81.716 | hexafluoroethane | 8.093 | 119 | 151992 | 9.82 | 9.82 | ppb | 200 | 1 | 1 | 0.0084458 | 0.0084458 | 9.82
Compound | 1ppbv % Recovery | 30 ppbv % Recovery | [14 empty] | chlorotrifluoromethane | 8.515 | 85 | 4531 | 1.014 | 1.014 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 1.014 | 10.14 | [14 empty] | chlorotrifluoromethane | 8.515 | 85 | 117921 | 27.184 | 27.184 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 27.184 | 9.0613 | [13 empty] | chlorotrifluoromethane | 8.515 | 85 | 43052 | 9.825 | 9.825 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.825 | [13 empty] | chlorotrifluoromethane | 8.515 | 85 | 45132 | 10.147 | 10.147 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.147 | [13 empty] | chlorotrifluoromethane | 8.515 | 85 | 42686 | 9.965 | 9.965 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 9.965 | [13 empty] | chlorotrifluoromethane | 8.503 | 85 | 40944 | 10.179 | 10.179 | ppb | 200 | 1 | 1 | 0.029884 | 0.029884 | 10.179
*Tetrafluoromethane | 85.48% | 82.77% | [14 empty] | trifluoromethane | 8.721 | 69 | 3888 | 1.005 | 1.005 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 1.005 | 10.05 | [14 empty] | trifluoromethane | 8.721 | 69 | 110549 | 29.431 | 29.431 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 29.431 | 9.8103 | [13 empty] | trifluoromethane | 8.708 | 69 | 38023 | 10.021 | 10.021 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.021 | [13 empty] | trifluoromethane | 8.708 | 69 | 39903 | 10.36 | 10.36 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.36 | [13 empty] | trifluoromethane | 8.709 | 69 | 38963 | 10.504 | 10.504 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 10.504 | [13 empty] | trifluoromethane | 8.708 | 69 | 38590 | 11.079 | 11.079 | ppb | 200 | 1 | 1 | 0.0356 | 0.0356 | 11.079
Hexafluoroethane (FC-116) | 95.81% | 89.91% | [14 empty] | octafluoropropane | 11.178 | 69 | 35880 | 1.037 | 1.037 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 1.037 | 10.37 | [14 empty] | octafluoropropane | 11.178 | 69 | 931465 | 27.721 | 27.721 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 27.721 | 9.2403 | [13 empty] | octafluoropropane | 11.178 | 69 | 339379 | 9.998 | 9.998 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 9.998 | [13 empty] | octafluoropropane | 11.178 | 69 | 355402 | 10.315 | 10.315 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.315 | [13 empty] | octafluoropropane | 11.178 | 69 | 341069 | 10.279 | 10.279 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.279 | [13 empty] | octafluoropropane | 11.178 | 69 | 324804 | 10.424 | 10.424 | ppb | 200 | 1 | 1 | 0.015822 | 0.015822 | 10.424
Chlorotrifluoromethane (CFC-13) | 103.2% | 92.23% | [14 empty] | difluoromethane | 10.975 | 51 | 2796 | 1.072 | 1.072 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 1.072 | 10.72 | [14 empty] | difluoromethane | 10.975 | 51 | 74339 | 29.369 | 29.369 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 29.369 | 9.7897 | [13 empty] | difluoromethane | 10.975 | 51 | 26160 | 10.231 | 10.231 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.231 | [13 empty] | difluoromethane | 10.961 | 51 | 26740 | 10.303 | 10.303 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.303 | [13 empty] | difluoromethane | 10.961 | 51 | 25987 | 10.397 | 10.397 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.397 | [13 empty] | difluoromethane | 10.961 | 51 | 24618 | 10.489 | 10.489 | ppb | 200 | 1 | 1 | 0.007355 | 0.007355 | 10.489
Trifluoromethane (Fluoroform) | 100.3% | 97.9% | [14 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 10047 | 1.035 | 1.035 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 1.035 | 10.35 | [14 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 269462 | 28.592 | 28.592 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 28.592 | 9.5307 | [13 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 97677 | 10.26 | 10.26 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.26 | [13 empty] | 1,1,1,2,2-pentafluoroe... | 12.842 | 51 | 103038 | 10.662 | 10.662 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.662 | [13 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 97951 | 10.525 | 10.525 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.525 | [13 empty] | 1,1,1,2,2-pentafluoroe... | 12.828 | 51 | 93653 | 10.716 | 10.716 | ppb | 200 | 1 | 1 | 0.013843 | 0.013843 | 10.716
Octafluoropropane | 103.7% | 92.42% | [14 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 35251 | 1.06 | 1.06 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 1.06 | 10.6 | [14 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 904368 | 28.013 | 28.013 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 28.013 | 9.3377 | [13 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.939 | 100 | 329640 | 10.108 | 10.108 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.108 | [13 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 342381 | 10.343 | 10.343 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.343 | [13 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 326371 | 10.238 | 10.238 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.238 | [13 empty] | 1,1,2,2,3,3,4,4-octafl... | 15.925 | 100 | 309157 | 10.327 | 10.327 | ppb | 200 | 1 | 1 | 0.013563 | 0.013563 | 10.327
Difluoromethane (HFC-32) | 104.8% | 95.69% | [14 empty] | fluoromethane | 11.424 | 33 | 902m | 0.797 | 0.797 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 0.797 | 7.97 | [14 empty] | fluoromethane | 11.424 | 33 | 23254m | 21.176 | 21.176 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 21.176 | 7.0587 | [13 empty] | fluoromethane | 11.409 | 33 | 11540m | 10.403 | 10.403 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.403 | [13 empty] | fluoromethane | 11.409 | 33 | 12430 | 11.039 | 11.039 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 11.039 | [13 empty] | fluoromethane | 11.41 | 33 | 11918 | 10.99 | 10.99 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.99 | [13 empty] | fluoromethane | 11.409 | 33 | 11197 | 10.996 | 10.996 | ppb | 200 | 1 | 1 | 0.333 | 0.333 | 10.996
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | 100.9% | 92.89% | [14 empty] | tetrafluoroethylene | 8.503 | 81 | 7407 | 0.991 | 0.991 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 0.991 | 9.91 | [14 empty] | tetrafluoroethylene | 8.503 | 81 | 196399 | 27.074 | 27.074 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 27.074 | 9.0247 | [13 empty] | tetrafluoroethylene | 8.491 | 81 | 72540 | 9.899 | 9.899 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 9.899 | [13 empty] | tetrafluoroethylene | 8.491 | 81 | 75945 | 10.21 | 10.21 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.21 | [13 empty] | tetrafluoroethylene | 8.491 | 81 | 72484 | 10.118 | 10.118 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.118 | [13 empty] | tetrafluoroethylene | 8.491 | 81 | 67931 | 10.099 | 10.099 | ppb | 200 | 1 | 1 | 0.019104 | 0.019104 | 10.099
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | 104.9% | 92.38% | [14 empty] | hexafluoropropene | 13.248 | 131 | 16284 | 0.984 | 0.984 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 0.984 | 9.84 | [14 empty] | hexafluoropropene | 13.248 | 131 | 423903m | 26.378 | 26.378 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 26.378 | 8.7927 | [13 empty] | hexafluoropropene | 13.248 | 131 | 154380 | 9.51 | 9.51 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.51 | [13 empty] | hexafluoropropene | 13.248 | 131 | 157747 | 9.573 | 9.573 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.573 | [13 empty] | hexafluoropropene | 13.248 | 131 | 147789 | 9.313 | 9.313 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.313 | [13 empty] | hexafluoropropene | 13.248 | 131 | 141708 | 9.51 | 9.51 | ppb | 200 | 1 | 1 | 0.0086848 | 0.0086848 | 9.51
Fluoromethane (HFC-41) | 76.61% | 67.85% | [14 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 1018 | 1.009 | 1.009 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 1.009 | 10.09 | [14 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 26780 | 27.336 | 27.336 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 27.336 | 9.112 | [13 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9732 | 9.834 | 9.834 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.834 | [13 empty] | 1,1,1-trifluoroethane | 13.233 | 64 | 9975 | 9.93 | 9.93 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.93 | [13 empty] | 1,1,1-trifluoroethane | 13.219 | 64 | 9578 | 9.901 | 9.901 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.901 | [13 empty] | 1,1,1-trifluoroethane | 13.219 | 64 | 9070 | 9.984 | 9.984 | ppb | 200 | 1 | 1 | 0.034251 | 0.034251 | 9.984
Tetrafluoroethylene | 100.1% | 91.17% | [14 empty] | hexafluoropropene_oxide | 13.45 | 69 | 1177m | 1.105 | 1.105 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 1.105 | 11.05 | [14 empty] | hexafluoropropene_oxide | 13.436 | 69 | 28219m | 27.284 | 27.284 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 27.284 | 9.0947 | [13 empty] | hexafluoropropene_oxide | 13.436 | 69 | 10233 | 9.794 | 9.794 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.794 | [13 empty] | hexafluoropropene_oxide | 13.436 | 69 | 11018 | 10.389 | 10.389 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 10.389 | [13 empty] | hexafluoropropene_oxide | 13.421 | 69 | 9060m | 8.871 | 8.871 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 8.871 | [13 empty] | hexafluoropropene_oxide | 13.421 | 69 | 9160 | 9.551 | 9.551 | ppb | 200 | 1 | 1 | 0.33918 | 0.33918 | 9.551
Hexafluoropropene | 103.5% | 92.46% | [14 empty] | chlorodifluoromethane | 13.74 | 51 | 10506 | 1.063 | 1.063 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 1.063 | 10.63 | [14 empty] | chlorodifluoromethane | 13.74 | 51 | 273630 | 28.523 | 28.523 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 28.523 | 9.5077 | [13 empty] | chlorodifluoromethane | 13.74 | 51 | 99702 | 10.288 | 10.288 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.288 | [13 empty] | chlorodifluoromethane | 13.74 | 51 | 104406 | 10.614 | 10.614 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.614 | [13 empty] | chlorodifluoromethane | 13.725 | 51 | 99617 | 10.515 | 10.515 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.515 | [13 empty] | chlorodifluoromethane | 13.725 | 51 | 94482 | 10.621 | 10.621 | ppb | 200 | 1 | 1 | 0.024044 | 0.024044 | 10.621
1,1,1-trifluoroethane | 102.6% | 92.66% | [14 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 3396 | 1.039 | 1.039 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 1.039 | 10.39 | [14 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 93306 | 29.418 | 29.418 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 29.418 | 9.806 | [13 empty] | 1,1,1,2-tetrafluoroethane | 15.939 | 33 | 33493 | 10.454 | 10.454 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.454 | [13 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 35875 | 11.031 | 11.031 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 11.031 | [13 empty] | 1,1,1,2-tetrafluoroethane | 15.925 | 33 | 33713 | 10.764 | 10.764 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.764 | [13 empty] | 1,1,1,2-tetrafluoroethane | 15.911 | 33 | 32270 | 10.972 | 10.972 | ppb | 200 | 1 | 1 | 0.02815 | 0.02815 | 10.972
Hexafluoropropene Oxide | 112.8% | 92.86% | [14 empty] | decafluorobutane | 16.161 | 119 | 19102 | 1.078 | 1.078 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 1.078 | 10.78 | [14 empty] | decafluorobutane | 16.161 | 119 | 458240m | 26.634 | 26.634 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 26.634 | 8.878 | [13 empty] | decafluorobutane | 16.175 | 119 | 173271 | 9.97 | 9.97 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 9.97 | [13 empty] | decafluorobutane | 16.161 | 119 | 178800 | 10.135 | 10.135 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.135 | [13 empty] | decafluorobutane | 16.161 | 119 | 170042 | 10.008 | 10.008 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.008 | [13 empty] | decafluorobutane | 16.161 | 119 | 164011 | 10.28 | 10.28 | ppb | 200 | 1 | 1 | 0.023141 | 0.023141 | 10.28
Chlorodifluoromethane (HCFC-22) | 103.3% | 92.42% | [14 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 15628 | 1.07 | 1.07 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 1.07 | 10.7 | [14 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 394456 | 27.809 | 27.809 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 27.809 | 9.2697 | [13 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 147074 | 10.264 | 10.264 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.264 | [13 empty] | 1H-Heptafluoropropane | 18.113 | 51 | 157604 | 10.836 | 10.836 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.836 | [13 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 151320 | 10.803 | 10.803 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.803 | [13 empty] | 1H-Heptafluoropropane | 18.095 | 51 | 142107 | 10.804 | 10.804 | ppb | 200 | 1 | 1 | 0.016117 | 0.016117 | 10.804
1,1,1,2-Tetrafluoroethane | 99.39% | 93.8% | [14 empty] | octafluorocyclopentene | 21.943 | 93 | 24941 | 1.031 | 1.031 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 1.031 | 10.31 | [14 empty] | octafluorocyclopentene | 21.943 | 93 | 626750 | 26.695 | 26.695 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 26.695 | 8.8983 | [13 empty] | octafluorocyclopentene | 21.943 | 93 | 240403 | 10.136 | 10.136 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.136 | [13 empty] | octafluorocyclopentene | 21.943 | 93 | 249672 | 10.371 | 10.371 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.371 | [13 empty] | octafluorocyclopentene | 21.931 | 93 | 236496 | 10.2 | 10.2 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.2 | [13 empty] | octafluorocyclopentene | 21.931 | 93 | 228756 | 10.507 | 10.507 | ppb | 200 | 1 | 1 | 0.0051927 | 0.0051927 | 10.507
Decafluorobutane (Perfluorobutane) | 108.1% | 89.05% | [14 empty] | trichlorofluoromethane | 21.24 | 101 | 26327 | 1.089 | 1.089 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 1.089 | 10.89 | [14 empty] | trichlorofluoromethane | 21.239 | 101 | 633705 | 26.999 | 26.999 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 26.999 | 8.9997 | [13 empty] | trichlorofluoromethane | 21.239 | 101 | 240278 | 10.134 | 10.134 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.134 | [13 empty] | trichlorofluoromethane | 21.239 | 101 | 248487 | 10.325 | 10.325 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.325 | [13 empty] | trichlorofluoromethane | 21.24 | 101 | 237303 | 10.239 | 10.239 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.239 | [13 empty] | trichlorofluoromethane | 21.239 | 101 | 232544 | 10.685 | 10.685 | ppb | 200 | 1 | 1 | 0.028098 | 0.028098 | 10.685
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | 104.2% | 90.31% | [14 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 63823m | 1.158 | 1.158 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 1.158 | 11.58 | [14 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 1472980 | 27.534 | 27.534 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 27.534 | 9.178 | [13 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 571942 | 10.583 | 10.583 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.583 | [13 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 587590m | 10.712 | 10.712 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 10.712 | [13 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 584950 | 11.073 | 11.073 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 11.073 | [13 empty] | dodecafluoro-n-pentane | 21.012 | 69 | 557251 | 11.233 | 11.233 | ppb | 200 | 1 | 1 | 0.010934 | 0.010934 | 11.233
Octafluorocyclopentene (FC-C1418) | 101.7% | 87.79% | [14 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 25427 | 1.168 | 1.168 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 1.168 | 11.68 | [14 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 597475 | 28.276 | 28.276 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 28.276 | 9.4253 | [13 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 233362 | 10.933 | 10.933 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.933 | [13 empty] | 1H-Nonafluorobutane | 22.802 | 51 | 241608 | 11.151 | 11.151 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.151 | [13 empty] | 1H-Nonafluorobutane | 22.79 | 51 | 229369 | 10.993 | 10.993 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 10.993 | [13 empty] | 1H-Nonafluorobutane | 22.79 | 51 | 220921 | 11.275 | 11.275 | ppb | 200 | 1 | 1 | 0.012238 | 0.012238 | 11.275
Trichlorofluoromethane (CFC-11) | 107.5% | 88.81% | [14 empty] | tetradecafluorohexane | 24.519 | 69 | 62919m | 1.073 | 1.073 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 1.073 | 10.73 | [14 empty] | tetradecafluorohexane | 24.495 | 69 | 1585568m | 27.861 | 27.861 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 27.861 | 9.287 | [13 empty] | tetradecafluorohexane | 24.456 | 69 | 580547m | 10.098 | 10.098 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.098 | [13 empty] | tetradecafluorohexane | 24.416 | 69 | 601255m | 10.303 | 10.303 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.303 | [13 empty] | tetradecafluorohexane | 24.408 | 69 | 595209m | 10.591 | 10.591 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.591 | [13 empty] | tetradecafluorohexane | 24.353 | 69 | 556165m | 10.539 | 10.539 | ppb | 200 | 1 | 1 | 0.045372 | 0.045372 | 10.539
Dodecafluoro-n-pentane | 109.4% | 86.72% | [14 empty] | 1H-Perfluoropentane | 26.134 | 69 | 21659m | 1.146 | 1.146 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 1.146 | 11.46 | [14 empty] | 1H-Perfluoropentane | 26.124 | 69 | 520346 | 28.372 | 28.372 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 28.372 | 9.4573 | [13 empty] | 1H-Perfluoropentane | 26.105 | 69 | 196289 | 10.595 | 10.595 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.595 | [13 empty] | 1H-Perfluoropentane | 26.076 | 69 | 204588m | 10.879 | 10.879 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.879 | [13 empty] | 1H-Perfluoropentane | 26.066 | 69 | 194714m | 10.751 | 10.751 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.751 | [13 empty] | 1H-Perfluoropentane | 26.027 | 69 | 185143m | 10.887 | 10.887 | ppb | 200 | 1 | 1 | 0.02563 | 0.02563 | 10.887
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | 106.8% | 86.21% | [14 empty] | heptafluoropropyl 1,2,... | 26.202 | 101 | 52781m | 1.099 | 1.099 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 1.099 | 10.99 | [14 empty] | heptafluoropropyl 1,2,... | 26.193 | 101 | 1272926m | 27.298 | 27.298 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 27.298 | 9.0993 | [13 empty] | heptafluoropropyl 1,2,... | 26.173 | 101 | 467901m | 9.933 | 9.933 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 9.933 | [13 empty] | heptafluoropropyl 1,2,... | 26.154 | 101 | 489443m | 10.236 | 10.236 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.236 | [13 empty] | heptafluoropropyl 1,2,... | 26.144 | 101 | 496995m | 10.793 | 10.793 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.793 | [13 empty] | heptafluoropropyl 1,2,... | 26.105 | 101 | 464944m | 10.753 | 10.753 | ppb | 200 | 1 | 1 | 0.0093935 | 0.0093935 | 10.753
Tetradecafluorohexane | 106.3% | 91.97% | [14 empty] | hexadecafluoroheptane | 27.314 | 69 | 78624 | 1.139 | 1.139 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 1.139 | 11.39 | [14 empty] | hexadecafluoroheptane | 27.294 | 69 | 1821445 | 27.187 | 27.187 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 27.187 | 9.0623 | [13 empty] | hexadecafluoroheptane | 27.285 | 69 | 691987 | 10.225 | 10.225 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.225 | [13 empty] | hexadecafluoroheptane | 27.265 | 69 | 714489 | 10.4 | 10.4 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.4 | [13 empty] | hexadecafluoroheptane | 27.255 | 69 | 686926 | 10.383 | 10.383 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.383 | [13 empty] | hexadecafluoroheptane | 27.226 | 69 | 661489 | 10.648 | 10.648 | ppb | 200 | 1 | 1 | 0.007102 | 0.007102 | 10.648
1,1,1,2,2,3,3,4,4,5,5-Undecafluoropentane (1H-Perfluoropentane) | 108.2% | 89.26% | [14 empty] | 1H-Perfluorohexane | 28.776 | 51 | 35462 | 1.076 | 1.076 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 1.076 | 10.76 | [14 empty] | 1H-Perfluorohexane | 28.767 | 51 | 881997 | 27.579 | 27.579 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 27.579 | 9.193 | [13 empty] | 1H-Perfluorohexane | 28.757 | 51 | 337248 | 10.439 | 10.439 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.439 | [13 empty] | 1H-Perfluorohexane | 28.737 | 51 | 350739 | 10.696 | 10.696 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.696 | [13 empty] | 1H-Perfluorohexane | 28.718 | 51 | 337809 | 10.696 | 10.696 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.696 | [13 empty] | 1H-Perfluorohexane | 28.688 | 51 | 320709 | 10.814 | 10.814 | ppb | 200 | 1 | 1 | 0.0083342 | 0.0083342 | 10.814
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | 110.6% | 91.61% | [14 empty] | octadecafluorooctane | 29.572 | 119 | 48975 | 1.145 | 1.145 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 1.145 | 11.45 | [14 empty] | octadecafluorooctane | 29.549 | 119 | 1069808 | 25.756 | 25.756 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 25.756 | 8.5853 | [13 empty] | octadecafluorooctane | 29.548 | 119 | 407237 | 9.706 | 9.706 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.706 | [13 empty] | octadecafluorooctane | 29.537 | 119 | 417977 | 9.814 | 9.814 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.814 | [13 empty] | octadecafluorooctane | 29.514 | 119 | 400089 | 9.754 | 9.754 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.754 | [13 empty] | octadecafluorooctane | 29.502 | 119 | 377638 | 9.805 | 9.805 | ppb | 200 | 1 | 1 | 0.027692 | 0.027692 | 9.805
Hexadecafluoroheptane | 111.4% | 88.63% | [14 empty] | 1H-Perfluoroheptane | 31.09 | 51 | 108458 | 1.182 | 1.182 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 1.182 | 11.82 | [14 empty] | 1H-Perfluoroheptane | 31.078 | 51 | 2402975 | 26.977 | 26.977 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 26.977 | 8.9923 | [13 empty] | 1H-Perfluoroheptane | 31.09 | 51 | 923485 | 10.263 | 10.263 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.263 | [13 empty] | 1H-Perfluoroheptane | 31.078 | 51 | 961481 | 10.527 | 10.527 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.527 | [13 empty] | 1H-Perfluoroheptane | 31.055 | 51 | 921534 | 10.477 | 10.477 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.477 | [13 empty] | 1H-Perfluoroheptane | 31.044 | 51 | 855374 | 10.356 | 10.356 | ppb | 200 | 1 | 1 | 0.010647 | 0.010647 | 10.356
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | 103.1% | 88.06% | [14 empty] | 1H-Perfluorooctane | 33.516 | 51 | 85607m | 1.217 | 1.217 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 1.217 | 12.17 | [14 empty] | 1H-Perfluorooctane | 33.505 | 51 | 1823581 | 26.708 | 26.708 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 26.708 | 8.9027 | [13 empty] | 1H-Perfluorooctane | 33.505 | 51 | 704508 | 10.214 | 10.214 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.214 | [13 empty] | 1H-Perfluorooctane | 33.505 | 51 | 726063 | 10.371 | 10.371 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.371 | [13 empty] | 1H-Perfluorooctane | 33.47 | 51 | 687014 | 10.189 | 10.189 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 10.189 | [13 empty] | 1H-Perfluorooctane | 33.47 | 51 | 627515 | 9.911 | 9.911 | ppb | 200 | 1 | 1 | 0.028576 | 0.028576 | 9.911
Octadecafluorooctane (Perfluorooctane) | 118% | 88.45% | [14 empty] | 2H-peffluoro-5-methyl-... | 33.056 | 169 | 217362 | 1.117 | 1.117 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 1.117 | 11.17 | [14 empty] | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 4903077 | 25.96 | 25.96 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 25.96 | 8.6533 | [13 empty] | 2H-peffluoro-5-methyl-... | 33.056 | 169 | 1829586 | 9.589 | 9.589 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.589 | [13 empty] | 2H-peffluoro-5-methyl-... | 33.045 | 169 | 1820840 | 9.402 | 9.402 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.402 | [13 empty] | 2H-peffluoro-5-methyl-... | 33.022 | 169 | 1681240 | 9.014 | 9.014 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 9.014 | [13 empty] | 2H-peffluoro-5-methyl-... | 33.01 | 169 | 1561311 | 8.915 | 8.915 | ppb | 200 | 1 | 1 | 0.037129 | 0.037129 | 8.915
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | 115.2% | 87.62%
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | 119.2% | 87.16%
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | 116.5% | 90.24%
*spiked values are 10x higher-10 ppbv and 300 ppbv
Values within the 30 % accuracy limits
Values NOT within the 30 % accuracy limits

## MDL
MDL Composite (ref TO15A section 17.3-17.7 for MDL spiking and blanking)
(E-288 Extractor Source MDL October 4th, 2023) |  | MDL | MDL
| Can ID | ppbv | ug/m3 | Comments | [9 empty] | Conversion
Description | CAS # | [11 empty] | molecular weight (ng/nmole) | ng/l=µg/m3 per 1 ppb
Tetrafluoromethane | 75-73-0 | 0.089699 | 0.32815 | [9 empty] | 88.0043 | 3.6583
Hexafluoroethane (FC-116) | 76-16-4 | 0.0084458 | 0.048454 | [9 empty] | 138.01 | 5.7371
Chlorotrifluoromethane (CFC-13) | 75-72-9 | 0.029884 | 0.12977 | [9 empty] | 104.46 | 4.3424
Trifluoromethane (Fluoroform) | 75-46-7 | 0.0356 | 0.10361 | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | [8 empty] | 70.01 | 2.9103
Octafluoropropane | 76-19-7 | 0.015822 | 0.12366 | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | [8 empty] | 188.02 | 7.816
Difluoromethane (HFC-32) | 75-10-5 | 0.007355 | 0.015906 | [9 empty] | 52.024 | 2.1626
1,1,1,2,2-Pentafluoroethane (Pentafluoroethane)(HFC-125) | 354-33-6 | 0.013843 | 0.069066 | [9 empty] | 120.02 | 4.9892
1,1,2,2,3,3,4,4-Octafluorocyclobutane (Octafluorocyclobutane) (FC-C318) | 115-25-3 | 0.013563 | 0.11278 | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | [8 empty] | 200.03 | 8.3152
Fluoromethane (HFC-41) | 593-53-3 | 0.333 | 0.12 |  | 0.33 | would be 200 ml MDL but won't actually be used with high CO2 data | [6 empty] | 34.03 | 1.4146
Tetrafluoroethylene | 116-14-3 | 0.019104 | 0.07943 | [9 empty] | 100.02 | 4.1578
Hexafluoropropene | 116-15-4 | 0.0086848 | 0.054161 | [9 empty] | 150.02 | 6.2363
1,1,1-trifluoroethane | 420-46-2 | 0.034251 | 0.11966 | [9 empty] | 84.04 | 3.4935
Hexafluoropropene Oxide | 428-59-1 | 0.33918 | 2.3408 | [9 empty] | 166.02 | 6.9014
Chlorodifluoromethane (HCFC-22) | 75-45-6 | 0.024044 | 0.086426 | [9 empty] | 86.47 | 3.5945
1,1,1,2-Tetrafluoroethane | 811-97-2 | 0.02815 | 0.1194 | [9 empty] | 102.03 | 4.2414
Decafluorobutane (Perfluorobutane) | 355-25-9 | 0.023141 | 0.22898 | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | [8 empty] | 238.03 | 9.8949
1,1,1,2,2,3,3-Heptafluoropropane (1H-Heptafluoropropane) | 2252-84-8 | 0.016117 | 0.11392 | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | [8 empty] | 170.03 | 7.0681
Octafluorocyclopentene (FC-C1418) | 559-40-0 | 0.0051927 | 0.045771 | [9 empty] | 212.042 | 8.8146
Trichlorofluoromethane (CFC-11) | 75-69-4 | 0.028098 | 0.16045 | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | [8 empty] | 137.37 | 5.7105
Dodecafluoro-n-pentane | 678-26-2 | 0.010934 | 0.13092 | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | [8 empty] | 288.036 | 11.9736
1,1,1,2,2,3,3,4,4-Nonafluorobutane (1H-Nonafluorobutane) | 375-17-7 | 0.012238 | 0.11194 | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | [8 empty] | 220.04 | 9.147
Tetradecafluorohexane | 355-42-0 | 0.045372 | 0.63758 | [9 empty] | 338.04 | 14.0523
1,1,1,2,2,3,3,4,4,5,5-Undecafluoropentane (1H-Perfluoropentane) | 375-61-1 | 0.02563 | 0.28771 | [9 empty] | 270.04 | 11.2255
Heptafluoropropyl 1,2,2,2-tetrafluoroethyl ether (E1) | 3330-15-2 | 0.0093935 | 0.1117 | [9 empty] | 286.04 | 11.8906
Hexadecafluoroheptane | 335-57-9 | 0.007102 | 0.11456 | [9 empty] | 388.051 | 16.1312
1,1,1,2,2,3,3,4,4,5,5,6,6-Tridecafluorohexane (1H-Perfluorohexane) | 355-37-3 | 0.0083342 | 0.11088 | [9 empty] | 320.05 | 13.3044
Octadecafluorooctane (Perfluorooctane) | 307-34-6 | 0.027692 | 0.50428 | [9 empty] | 438.06 | 18.2101
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7-Pentadecafluoroheptane (1H-Perfluoroheptane) | 375-83-7 | 0.010647 | 0.16379 | [9 empty] | 370.06 | 15.3833
1,1,1,2,2,3,3,4,4,5,5,6,6,7,7,8,8-Heptadecafluorooctane (1H-Perfluorooctane) | 335-65-9 | 0.028576 | 0.499 | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | [8 empty] | 420.07 | 17.4623
2H-Perfluoro-5-methyl-3,6-dioxanonane (E2) | 3330-14-1 | 0.037129 | 0.69775 | [9 empty] | 452.07 | 18.7925
Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) |  |  |  | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) |  |  | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) |  |  | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of all hits present and positive in each blank-see TO15A section 17.7)
Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) |  |  |  | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) |  |  | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) |  |  | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7) | Took the highest between MDL spiked and MDL blank (case of not all hits present and positive in each blank-see TO15A section 17.7)