## Tbl1-Tbl4_Lab2
IDC - Laboratory 2 (ODEQ)
Calibration and QCS
2021-03-11 | QCS Analysis, concentration 0.3 ug/l |  | Fortified | 0.3
|  |  |  | Ave Result |  | Criteria
| S21C060-SCV1 |  |  | 0.26195 | 87.32%
| S21C060-SCV2 |  |  | 0.25105 | 83.68%
| S21C060-SCV3 |  |  | 0.25585 | 85.28%
|  |  |  | Ave Recovery | 85.43% | ±20% of true value | PASS
|  |  |  | %RSD | 2.132% |  | PASS
Accuracy and Precision in Reagent Water
2021-03-11 | Reagent Water Fortified at 0.5 ug/l |  | Fortified | 0.5
|  |  |  | Ave Result | % Rec | Criteria/Method 546 Performance
| B21C148-BLK1 |  | Blank | 0.01255
| B21C148-BS1 |  |  | 0.5487 | 109.7%
| B21C148-BS2 |  |  | 0.512 | 102.4%
| B21C148-BS3 |  |  | 0.515 | 103%
| B21C148-BS4 |  |  | 0.47265 | 94.53%
| B21C148-BS5 |  |  | 0.4775 | 95.5%
| B21C148-BS6 |  |  | 0.44915 | 89.83%
| B21C148-BS7 |  |  | 0.4623 | 92.46%
|  |  |  | Ave Recovery | 98.21% | within 30% of true value | PASS
|  |  |  | %RSD | 7.171% | ≤ 15% | PASS
|  | Cfactor | 3.963 | HRPIR | 0.13956
|  |  |  | UpperPIR | 126.1% | </= 150% | PASS | calculated as a check
|  |  |  | LowerPIR | 70.3% | >/= 50% | PASS | calculated as a check
MRL Confirmation (from LCMRL data)
2021-03-11 | Reagent Water Fortified at 0.115 ug/l |  | Fortified | 0.115
|  |  |  | Ave Result | % Rec | Criteria/Method 546 Performance
| B21C148-BLK1 |  | Blank | 0.01255
| B21C148-MRLA |  |  | 0.1134 | 98.61%
| B21C148-MRLB |  |  | 0.1123 | 97.65%
| B21C148-MRLC |  |  | 0.1141 | 99.22%
| B21C148-MRLD |  |  | 0.10835 | 94.22%
| B21C148-MRLE |  |  | 0.10875 | 94.57%
| B21C148-MRLF |  |  | 0.0977 | 84.96%
| B21C148-MRLG |  |  | 0.10085 | 87.7%
|  |  |  | Ave Recovery | 93.84%
|  |  |  | %RSD | 5.893%
|  | Cfactor | 3.963 | HRPIR | 0.025206
|  |  |  | UpperPIR | 115.8% | </= 150% | PASS
|  |  |  | LowerPIR | 71.93% | >/= 50% | PASS
MRL Confirmation (from LCMRL data)
2021-03-11 | Reagent Water Fortified at 0.075 ug/l |  | Fortified | 0.075
|  |  |  | Ave Result | % Rec | Criteria/Method 546 Performance
| B21C148-BLK1 |  | Blank | 0.01255
| B21C148-MRL1 |  |  | 0.0809 | 107.9%
| B21C148-MRL2 |  |  | 0.06385 | 85.13%
| B21C148-MRL3 |  |  | 0.0746 | 99.47%
| B21C148-MRL4 |  |  | 0.06305 | 84.07%
| B21C148-MRL5 |  |  | 0.07035 | 93.8%
| B21C148-MRL6 |  |  | 0.0657 | 87.6%
| B21C148-MRL7 |  |  | 0.0512 | 68.27%
|  |  |  | Ave Recovery | 89.46%
|  |  |  | %RSD | 14.13%
|  | Cfactor | 3.963 | HRPIR | 0.037561
|  |  |  | UpperPIR | 139.5% | </= 150% | FAIL
|  |  |  | LowerPIR | 39.38% | >/= 50% | FAIL
Demonstration of Low System Background
2021-03-11 | Reagent Water |  |  | Result (ug/l) | Criterion
| B21C148-BLK1 |  |  | 0.01255 | < .15 | PASS
| B21C148-BLK2 |  |  | 0.0106 | < .15 | PASS
| B21C148-BLK3 |  |  | 0.00885 | < .15 | PASS
| B21C148-BLK4 |  |  | 0.00795 | < .15 | PASS
| B21C148-BLK5 |  |  | 0.00365 | < .15 | PASS
P&A In Finished Drinking Water
| Local Tap Fortified below MRL |  | Fortified | 0.115
2021-03-15 | Tap Water + Thiosulfate Fortified at 0.115 ug/l |  |  | Ave Result | % Rec | Criteria
| DW.115.1 |  |  | 0.1521 | 132.3%
| DW.115.2 |  |  | 0.1472 | 128%
| DW.115.3 |  |  | 0.1263 | 109.8%
| DW.115.4 |  |  | 0.127 | 110.4%
| DW.115.5 |  |  | 0.1259 | 109.5%
| DW.115.6 |  |  | 0.1468 | 127.7%
| DW.115.7 |  |  | 0.1298 | 112.9%
|  |  |  | Ave Recovery | 118.6%
|  |  |  | %RSD | 8.544%
|  | Cfactor | 3.963 | HRPIR | 0.046201
|  |  |  | UpperPIR | 158.8% | not required
|  |  |  | LowerPIR | 78.47% | not required
| Local Tap Fortified at MRL |  | Fortified | 0.3
2021-03-15 | Tap Water + Thiosulfate Fortified at 0.3 ug/l |  |  | Ave Result | % Rec | Criteria
| DW.300.1 |  |  | 0.38465 | 128.2%
| DW.300.2 |  |  | 0.33935 | 113.1%
| DW.300.3 |  |  | 0.33395 | 111.3%
| DW.300.4 |  |  | 0.3443 | 114.8%
| DW.300.5 |  |  | 0.38015 | 126.7%
| DW.300.6 |  |  | 0.3376 | 112.5%
| DW.300.7 |  |  | 0.33115 | 110.4%
|  |  |  | Ave Recovery | 116.7%
|  |  |  | %RSD | 6.409%
|  | Cfactor | 3.963 | HRPIR | 0.08894
|  |  |  | UpperPIR | 146.4% | not required
|  |  |  | LowerPIR | 87.07% | not required
| Local Tap Fortified Near Mid-Range |  | Fortified | 0.5
2021-03-15 | Tap Water + Thiosulfate Fortified at 0.5 ug/l |  |  | Ave Result | % Rec | Criteria
| DW.500.1 |  |  | 0.5764 | 115.3%
| DW.500.2 |  |  | 0.55 | 110%
| DW.500.3 |  |  | 0.5339 | 106.8%
| DW.500.4 |  |  | 0.5716 | 114.3%
| DW.500.5 |  |  | 0.50995 | 102%
| DW.500.6 |  |  | 0.5662 | 113.2%
| DW.500.7 |  |  | 0.51435 | 102.9%
|  |  |  | Ave Recovery | 109.2%
|  |  |  | %RSD | 4.988%
|  | Cfactor | 3.963 | HRPIR | 0.10794
|  |  |  | UpperPIR | 130.8% | not required
|  |  |  | LowerPIR | 87.62% | not required
Unfortified Local Tap Water
2021-03-15 | Tap Water + Thiosulfate
| DW.BLK.0.1 |  |  | 0.01305 | values neglected in above calculations
| DW.BLK.0.2 |  |  | 0.01285 | values neglected in above calculations
| DW.BLK.0.3 |  |  | 0.0101 | values neglected in above calculations

## Fig1_Output
≥MDL (0.064 µg/L) | May | June | July | August | September | October | November
2020 | 0 | 0 | 1 | 3 | 1 | 2 | 0 | 7
2021 | 0 | 0 | 1 | 1 | 1 | 4 | 0 | 7
2022 | 0 | 0 | 2 | 4 | 0 | 1 | 0 | 7
| 0 | 0 | 4 | 8 | 2 | 7 | 0
≥MRL (0.115 µg/L) | May | June | July | August | September | October | November
2020 | 0 | 0 | 2 | 3 | 0 | 0 | 0 | 5
2021 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 5
2022 | 0 | 0 | 4 | 4 | 0 | 3 | 3 | 14
| 0 | 0 | 7 | 8 | 2 | 4 | 3
≥HAL (0.3 µg/L) | May | June | July | August | September | October | November
2020 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 4
2021 | 1 | 0 | 0 | 4 | 2 | 0 | 0 | 7
2022 | 0 | 1 | 0 | 0 | 3 | 2 | 0 | 6
| 1 | 1 | 0 | 5 | 7 | 3 | 0
| May | June | July | August | September | October | November
≥MDL (0.064 µg/L) | 0 | 0 | 4 | 8 | 2 | 7 | 0
≥MRL (0.115 µg/L) | 0 | 0 | 7 | 8 | 2 | 4 | 3
≥HAL (0.3 µg/L) | 1 | 1 | 0 | 5 | 7 | 3 | 0

## Fig2_Output
| 2020 | 2021 | 2022 | Total
≥MDL (0.064 µg/L) | 2 | 2 | 3 | 7
≥MRL (0.115 µg/L) | 2 | 0 | 6 | 8
≥HAL (0.3 µg/L) | 4 | 5 | 3 | 12
| ≥MDL (0.064 µg/L) | ≥MRL (0.115 µg/L) | ≥HAL (0.3 µg/L) | Total Samples
2020 | 2 | 2 | 4
2021 | 2 | 0 | 5
2022 | 3 | 6 | 3
Total | 7 | 8 | 12

## Tbl5_SeqQC
Quality Control Samples (2022): Reporting Limit Standard (Low-CV)
Description: Calibration standard spiked at 0.1 µg/L and analyzed at least once per analytical sequence. Acceptable recovery ≥50% and ≤150% of the true value.
Results:
Count (n) | 51
Mean (µg/L) | 0.10237 | 102.4% | REC
StDev (µg/L) | 0.013195 | 12.89% | RSD |  | "DL": | 0.031713
HRPIR | 0.035679 | 34.85% | Range |  | Cfactor: | 2.7039
UpperPIR | 138% |  | </= 150% | PASS
LowerPIR | 66.69% |  | >/= 50% | PASS
All Replicates
ID | Analysis Date | Well 1 (µg/L) | Well 2 (µg/L) | Mean | Well 1 (ABS) | Well 2 (ABS) | ABS %CV | Spike (µg/L) | Status
S22F007-LCV1 | 2022-06-01 19:48 | 0.1019 | 0.1112 | 0.1065 | 1.366 | 1.343 | 0.84902 | 0.1 | PASS | 106.5%
S22F019-LCV1 | 2022-06-03 18:08 | 0.0991 | 0.105 | 0.102 | 1.475 | 1.461 | 0.47684 | 0.1 | PASS | 102%
S22F028-LCV1 | 2022-06-07 20:15 | 0.0958 | 0.1268 | 0.1113 | 1.511 | 1.439 | 2.4407 | 0.1 | PASS | 111.3%
S22F048-LCV1 | 2022-06-10 11:55 | 0.119 | 0.0894 | 0.1042 | 1.408 | 1.481 | 2.5268 | 0.1 | PASS | 104.2%
S22F059-LCV1 | 2022-06-14 20:12 | 0.0929 | 0.0898 | 0.0914 | 1.469 | 1.478 | 0.3054 | 0.1 | PASS | 91.4%
S22F075-LCV1 | 2022-06-16 12:49 | 0.1045 | 0.112 | 0.1083 | 1.411 | 1.394 | 0.60606 | 0.1 | PASS | 108.3%
S22F101-LCV1 | 2022-06-22 11:47 | 0.0693 | 0.0866 | 0.078 | 1.281 | 1.235 | 1.8283 | 0.1 | PASS | 78%
S22F112-LCV1 | 2022-06-23 12:41 | 0.1207 | 0.1139 | 0.1173 | 1.059 | 1.07 | 0.51667 | 0.1 | PASS | 117.3%
S22F120-LCV1 | 2022-06-28 12:38 | 0.1081 | 0.1111 | 0.1096 | 1.14 | 1.134 | 0.26385 | 0.1 | PASS | 109.6%
S22F136-LCV1 | 2022-06-29 10:58 | 0.1394 | 0.1052 | 0.1223 | 1.203 | 1.273 | 2.8271 | 0.1 | PASS | 122.3%
S22F144-LCV1 | 2022-06-29 20:45 | 0.1102 | 0.1031 | 0.1067 | 1.388 | 1.405 | 0.60866 | 0.1 | PASS | 106.7%
S22F149-LCV1 | 2022-06-30 12:21 | 0.1261 | 0.0945 | 0.1103 | 1.172 | 1.233 | 2.5364 | 0.1 | PASS | 110.3%
S22G002-LCV1 | 2022-07-01 12:02 | 0.0769 | 0.1057 | 0.0913 | 1.496 | 1.409 | 2.9948 | 0.1 | PASS | 91.3%
S22G020-LCV1 | 2022-07-07 13:27 | 0.0861 | 0.0893 | 0.0877 | 1.383 | 1.374 | 0.32644 | 0.1 | PASS | 87.7%
S22G028-LCV1 | 2022-07-08 11:23 | 0.0921 | 0.1014 | 0.0968 | 1.466 | 1.441 | 0.85999 | 0.1 | PASS | 96.8%
S22G041-LCV1 | 2022-07-12 20:04 | 0.0962 | 0.1009 | 0.0986 | 1.476 | 1.463 | 0.44233 | 0.1 | PASS | 98.6%
S22G048-LCV1 | 2022-07-14 11:05 | 0.1047 | 0.1087 | 0.1067 | 1.36 | 1.35 | 0.369 | 0.1 | PASS | 106.7%
S22G053-LCV1 | 2022-07-15 12:47 | 0.1527 | 0.1063 | 0.1295 | 0.86 | 0.95 | 4.9724 | 0.1 | PASS | 129.5%
S22G063-LCV1 | 2022-07-21 11:38 | 0.1157 | 0.1279 | 0.1218 | 1.232 | 1.203 | 1.191 | 0.1 | PASS | 121.8%
S22G065-LCV1 | 2022-07-19 18:52 | 0.1003 | 0.128 | 0.1142 | 1.335 | 1.271 | 2.4559 | 0.1 | PASS | 114.2%
S22G104-LCV1 | 2022-07-28 15:49 | 0.1191 | 0.1186 | 0.1189 | 1.094 | 1.095 | 0.045683 | 0.1 | PASS | 118.9%
S22G110-LCV1 | 2022-07-29 10:27 | 0.0976 | 0.1115 | 0.1046 | 1.35 | 1.312 | 1.4275 | 0.1 | PASS | 104.6%
S22H002-LCV1 | 2022-08-01 13:53 | 0.0909 | 0.1098 | 0.1004 | 1.168 | 1.12 | 2.0979 | 0.1 | PASS | 100.4%
S22H014-LCV1 | 2022-08-03 12:15 | 0.1195 | 0.1276 | 0.1236 | 1.189 | 1.17 | 0.80543 | 0.1 | PASS | 123.6%
S22H024-LCV1 | 2022-08-04 11:18 | 0.1006 | 0.1012 | 0.1009 | 0.809 | 0.808 | 0.061843 | 0.1 | PASS | 100.9%
S22H030-LCV1 | 2022-08-05 11:28 | 0.1203 | 0.11 | 0.1152 | 1.123 | 1.147 | 1.0573 | 0.1 | PASS | 115.2%
S22H039-LCV1 | 2022-08-10 12:57 | 0.1047 | 0.1163 | 0.1105 | 1.5 | 1.468 | 1.0782 | 0.1 | PASS | 110.5%
S22H046-LCV1 | 2022-08-11 18:42 | 0.0872 | 0.1072 | 0.0972 | 1.274 | 1.22 | 2.1652 | 0.1 | PASS | 97.2%
S22H054-LCV1 | 2022-08-12 14:52 | 0.0961 | 0.1066 | 0.1014 | 1.297 | 1.27 | 1.0518 | 0.1 | PASS | 101.4%
S22H067-LCV1 | 2022-08-16 16:50 | 0.0926 | 0.1168 | 0.1047 | 1.298 | 1.24 | 2.2853 | 0.1 | PASS | 104.7%
S22H073-LCV1 | 2022-08-18 11:15 | 0.0974 | 0.1128 | 0.1051 | 1.223 | 1.185 | 1.5781 | 0.1 | PASS | 105.1%
S22H078-LCV1 | 2022-08-18 17:51 | 0.0807 | 0.0844 | 0.0826 | 1.179 | 1.168 | 0.46868 | 0.1 | PASS | 82.6%
S22H087-LCV1 | 2022-08-23 11:33 | 0.087 | 0.1086 | 0.0978 | 1.166 | 1.107 | 2.5957 | 0.1 | PASS | 97.8%
S22H108-LCV1 | 2022-08-24 18:24 | 0.0885 | 0.0903 | 0.0894 | 1.238 | 1.233 | 0.20235 | 0.1 | PASS | 89.4%
S22H110-LCV1 | 2022-08-25 14:02 | 0.1266 | 0.1152 | 0.1209 | 1.152 | 1.173 | 0.90323 | 0.1 | PASS | 120.9%
S22H117-LCV1 | 2022-08-26 13:29 | 0.07 | 0.082 | 0.076 | 1.189 | 1.161 | 1.1915 | 0.1 | PASS | 76%
S22H140-LCV1 | 2022-08-31 11:23 | 0.0985 | 0.11 | 0.1043 | 1.024 | 0.996 | 1.3861 | 0.1 | PASS | 104.3%
S22H146-LCV1 | 2022-09-01 13:43 | 0.0995 | 0.0965 | 0.098 | 1.198 | 1.206 | 0.33278 | 0.1 | PASS | 98%
S22I020-LCV1 | 2022-09-06 12:40 | 0.1014 | 0.1024 | 0.1019 | 1.191 | 1.189 | 0.084034 | 0.1 | PASS | 101.9%
S22I032-LCV1 | 2022-09-08 11:38 | 0.093 | 0.0901 | 0.0916 | 1.194 | 1.201 | 0.29228 | 0.1 | PASS | 91.6%
S22I043-LCV1 | 2022-09-09 13:47 | 0.1242 | 0.083 | 0.1036 | 1.285 | 1.378 | 3.4923 | 0.1 | PASS | 103.6%
S22I058-LCV1 | 2022-09-13 20:26 | 0.1315 | 0.0958 | 0.1137 | 1.172 | 1.24 | 2.8192 | 0.1 | PASS | 113.7%
S22I066-LCV1 | 2022-09-15 10:32 | 0.0982 | 0.081 | 0.0896 | 1.176 | 1.216 | 1.6722 | 0.1 | PASS | 89.6%
S22I073-LCV1 | 2022-09-16 12:50 | 0.1039 | 0.0886 | 0.0963 | 1.306 | 1.343 | 1.3968 | 0.1 | PASS | 96.3%
S22I078-LCV1 | 2022-09-20 13:35 | 0.1131 | 0.1147 | 0.1139 | 1.197 | 1.194 | 0.12547 | 0.1 | PASS | 113.9%
S22I087-LCV1 | 2022-09-21 19:08 | 0.1023 | 0.112 | 0.1072 | 1.226 | 1.205 | 0.86384 | 0.1 | PASS | 107.2%
S22I098-LCV1 | 2022-09-22 14:00 | 0.1004 | 0.0904 | 0.0954 | 1.176 | 1.197 | 0.88496 | 0.1 | PASS | 95.4%
S22I106-LCV1 | 2022-09-22 18:57 | 0.0682 | 0.06 | 0.0641 | 1.249 | 1.27 | 0.83366 | 0.1 | PASS | 64.1%
S22I120-LCV1 | 2022-09-27 19:45 | 0.0763 | 0.0776 | 0.077 | 1.256 | 1.253 | 0.11957 | 0.1 | PASS | 77%
S22I126-LCV1 | 2022-09-28 18:01 | 0.0967 | 0.1116 | 0.1042 | 1.276 | 1.244 | 1.2698 | 0.1 | PASS | 104.2%
S22I140-LCV1 | 2022-09-29 19:39 | 0.0999 | 0.0927 | 0.0963 | 1.265 | 1.283 | 0.70644 | 0.1 | PASS | 96.3%
Quality Control Samples (2022): Secondary Source Standard (QCS)
Description: A QCS sample must be assayed with each new lot of calibration standards. The QCS must be obtained from a source independent from that of the calibration standards, however, the same vendor may be used for both sources, provided that the lot numbers differ. The percent recovery of the QCS spiked at or near the EC50 of the calibration must be ≥70% and ≤130% of the true value.
Results:
Count (n) | 7
Mean (µg/L) | 0.78999 | 105.3% | REC
StDev (µg/L) | 0.069236 | 8.764% | RSD
All Replicates
ID | Analysis Date | Well 1 (µg/L) | Well 2 (µg/L) | Mean | Well 1 (ABS) | Well 2 (ABS) | ABS %CV | Spike (µg/L) | Status
S22F028-SCV1 | 2022-06-07 20:15 | 0.8583 | 0.9327 | 0.8955 | 0.837 | 0.811 | 2.2 | 0.75 | PASS | 119.4%
S22G053-SCV1 | 2022-07-15 12:47 | 0.8344 | 0.8298 | 0.8321 | 0.474 | 0.475 | 0.1 | 0.75 | PASS | 110.9%
S22G065-SCV1 | 2022-07-19 18:52 | 0.8889 | 0.7346 | 0.8118 | 0.696 | 0.75 | 5.3 | 0.75 | PASS | 108.2%
S22G063-SCV1 | 2022-07-21 11:38 | 0.7178 | 0.7043 | 0.7111 | 0.691 | 0.696 | 0.5 | 0.75 | PASS | 94.81%
S22I020-SCV1 | 2022-09-06 12:40 | 0.6915 | 0.7279 | 0.7097 | 0.698 | 0.685 | 1.3 | 0.75 | PASS | 94.63%
S22I078-SCV1 | 2022-09-20 13:35 | 0.7758 | 0.8667 | 0.8213 | 0.718 | 0.69 | 2.8 | 0.75 | PASS | 109.5%
S22I087-SCV1 | 2022-09-21 19:08 | 0.8061 | 0.6907 | 0.7484 | 0.681 | 0.72 | 3.9 | 0.75 | PASS | 99.79%

## Tbl5_BatchQC
Type | Recovery | These data are in Table 5
Reagent | 133 |  | mean | u.sd | 0% | 25% | 50% | 75% | 100% | data:n
Reagent | 89.8 | Finished | 91.9544 | 15.9426 | 64.3 | 79.425 | 90.45 | 103.25 | 137 | 68 | oneway.test(Recovery ~ Type, data=Dataset) # Welch test
Reagent | 70.3 | Reagent | 88.9737 | 14.0338 | 60 | 79.35 | 86.8 | 96.85 | 135 | 315
Reagent | 83 | Source | 93.7525 | 15.9551 | 58.8 | 82.65 | 92.4 | 103 | 144 | 238 | One-way analysis of means (not assuming equal variances)
Reagent | 115 | Surface | 99.3635 | 18.9627 | 54.3 | 87 | 96.95 | 113.75 | 138 | 74
Reagent | 113 | [9 empty] | data: Recovery and Type
Reagent | 87.7 | summary.anova | [8 empty] | F = 9.1739, num df = 3.00, denom df = 185.33, p-value = 0.00001087
Reagent | 93.6 |  | mean | sd | p.value |  |  | summary(AnovaModel.1)
Reagent | 114 | Type=Finished | 91.9544 | 15.9426 | 7.83e-07 |  |  | Df Sum Sq Mean Sq F value Pr(>F)
Reagent | 115 | Type=Reagent | 88.9737 | 14.0338 |  |  |  | Type 3 7628 2542.8 10.62 0.000000783 ***
Reagent | 103 | Type=Source | 93.7525 | 15.9551 |  |  |  | Residuals 691 165452 239.4
Reagent | 109 | Type=Surface | 99.3635 | 18.9627 |  |  |  | ---
Reagent | 99.1 | [6 empty] | Signif. codes: 0 '***' 0.001 '**' 0.01 '*' 0.05 '.' 0.1 ' ' 1
Reagent | 106
Reagent | 88.2
Reagent | 135 | Figures below are not in manuscript
Reagent | 89.8
Reagent | 99.3
Reagent | 86.7
Reagent | 108
Reagent | 112
Reagent | 123
Reagent | 102
Reagent | 103
Reagent | 91.4
Reagent | 83
Reagent | 95.8
Reagent | 99.3
Reagent | 106
Reagent | 119
Reagent | 83.2
Reagent | 117
Reagent | 92.9
Reagent | 93.9
Reagent | 83.1
Reagent | 84.5
Reagent | 103
Reagent | 116
Reagent | 105
Reagent | 84
Reagent | 109
Reagent | 103
Reagent | 81
Reagent | 80.7
Reagent | 89.9
Reagent | 84
Reagent | 84
Reagent | 78.8
Reagent | 93.1
Reagent | 95.7
Reagent | 83.4
Reagent | 83.8
Reagent | 97.2
Reagent | 82.1
Reagent | 98.8
Reagent | 78.5
Reagent | 86.9
Reagent | 84.2
Reagent | 83.4
Reagent | 82.6
Reagent | 71.8
Reagent | 94.4
Reagent | 94.4
Reagent | 74.4
Reagent | 71
Reagent | 71.4
Reagent | 79.2
Reagent | 73.1
Reagent | 73.8
Reagent | 74.3
Reagent | 75.4
Reagent | 68.4
Reagent | 82
Reagent | 95.4
Reagent | 70.9
Reagent | 84.3
Reagent | 91.9
Reagent | 73.1
Reagent | 99.3
Reagent | 67.1
Reagent | 80.1
Reagent | 84.3
Reagent | 98.3
Reagent | 87.1
Reagent | 117
Reagent | 80.4
Reagent | 95.9
Reagent | 77.7
Reagent | 95.2
Reagent | 76
Reagent | 73.7
Reagent | 74.6
Reagent | 102
Reagent | 80.3
Reagent | 87.8
Reagent | 71.9
Reagent | 61.5
Reagent | 68.3
Reagent | 79.2
Reagent | 98.4
Reagent | 83.9
Reagent | 76
Reagent | 74.6
Reagent | 84
Reagent | 85.6
Reagent | 76.4
Reagent | 79.5
Reagent | 82.9
Reagent | 75.4
Reagent | 74.7
Reagent | 78.6
Reagent | 86.9
Reagent | 84.2
Reagent | 88.8
Reagent | 94.9
Reagent | 96.8
Reagent | 91.3
Reagent | 96.4
Reagent | 89.4
Reagent | 79.9
Reagent | 72.4
Reagent | 78.4
Reagent | 70
Reagent | 98.6
Reagent | 107
Reagent | 83.7
Reagent | 82
Reagent | 95.6
Reagent | 105
Reagent | 96.5
Reagent | 91.7
Reagent | 106
Reagent | 126
Reagent | 94
Reagent | 104
Reagent | 74.4
Reagent | 98.7
Reagent | 96.2
Reagent | 113
Reagent | 82.6
Reagent | 90.3
Reagent | 63
Reagent | 78.7
Reagent | 89.7
Reagent | 103
Reagent | 86
Reagent | 104
Reagent | 98.9
Reagent | 94.4
Reagent | 98
Reagent | 93
Reagent | 89.1
Reagent | 110
Reagent | 80.3
Reagent | 93.8
Reagent | 85.1
Reagent | 111
Reagent | 88
Reagent | 96.8
Reagent | 94.4
Reagent | 95.4
Reagent | 87.9
Reagent | 101
Reagent | 81.9
Reagent | 83.7
Reagent | 84.3
Reagent | 102
Reagent | 82.4
Reagent | 110
Reagent | 110
Reagent | 104
Reagent | 98.7
Reagent | 85.5
Reagent | 70.6
Reagent | 87
Reagent | 94.8
Reagent | 110
Reagent | 101
Reagent | 118
Reagent | 70.2
Reagent | 112
Reagent | 77.9
Reagent | 91.3
Reagent | 65.1
Reagent | 80.1
Reagent | 81.4
Reagent | 90.3
Reagent | 69.5
Reagent | 81.7
Reagent | 74.7
Reagent | 77.8
Reagent | 96.3
Reagent | 95.2
Reagent | 79.6
Reagent | 80.6
Reagent | 81
Reagent | 110
Reagent | 82.3
Reagent | 90.4
Reagent | 81.2
Reagent | 94
Reagent | 104
Reagent | 112
Reagent | 91.8
Reagent | 87.6
Reagent | 80.7
Reagent | 78.6
Reagent | 111
Reagent | 124
Reagent | 79.5
Reagent | 81.1
Reagent | 88.5
Reagent | 82.7
Reagent | 89.7
Reagent | 81
Reagent | 76.3
Reagent | 74.9
Reagent | 65.6
Reagent | 73.8
Reagent | 85.6
Reagent | 89.6
Reagent | 77.9
Reagent | 84
Reagent | 93.8
Reagent | 71.6
Reagent | 90
Reagent | 89.4
Reagent | 88.9
Reagent | 95.5
Reagent | 102
Reagent | 98.5
Reagent | 92
Reagent | 81.8
Reagent | 132
Reagent | 95.6
Reagent | 94.3
Reagent | 84.6
Reagent | 104
Reagent | 97.6
Reagent | 107
Reagent | 71.9
Reagent | 114
Reagent | 123
Reagent | 92.9
Reagent | 83.4
Reagent | 101
Reagent | 85.2
Reagent | 99.4
Reagent | 101
Reagent | 90.6
Reagent | 66.6
Reagent | 97
Reagent | 95.4
Reagent | 132
Reagent | 111
Reagent | 87.1
Reagent | 69.3
Reagent | 86.1
Reagent | 85.4
Reagent | 75.8
Reagent | 78.9
Reagent | 68.9
Reagent | 61.7
Reagent | 83.1
Reagent | 65.1
Reagent | 91.6
Reagent | 81.4
Reagent | 78.3
Reagent | 84.6
Reagent | 96.5
Reagent | 95.6
Reagent | 77.2
Reagent | 86.8
Reagent | 86.8
Reagent | 79.5
Reagent | 76.7
Reagent | 76.8
Reagent | 86
Reagent | 96.9
Reagent | 87
Reagent | 85
Reagent | 77.5
Reagent | 64.9
Reagent | 74.7
Reagent | 67.9
Reagent | 88.8
Reagent | 60
Reagent | 80.1
Reagent | 86.3
Reagent | 77.9
Reagent | 86.5
Reagent | 96.1
Reagent | 96.1
Reagent | 77.6
Reagent | 66.7
Reagent | 78.3
Reagent | 68.3
Reagent | 76.1
Reagent | 70.9
Reagent | 76.6
Reagent | 71.5
Reagent | 98.1
Reagent | 107
Reagent | 88.9
Reagent | 85.4
Reagent | 85.3
Reagent | 60.3
Reagent | 93.5
Reagent | 72.3
Reagent | 81.3
Reagent | 110
Reagent | 84.1
Reagent | 83.2
Reagent | 75.7
Reagent | 82.5
Source | 105
Source | 95.3
Source | 88
Source | 88.6
Source | 95.9
Source | 99.8
Source | 125
Source | 92.7
Source | 91.3
Source | 74.7
Source | 81.4
Source | 81.8
Source | 108
Source | 124
Source | 85.4
Source | 88.3
Source | 88.2
Source | 94.2
Source | 85.7
Source | 75.9
Source | 79.8
Source | 68.8
Source | 119
Source | 97.1
Source | 109
Source | 96.3
Source | 81.3
Source | 81
Source | 87.9
Source | 89.8
Source | 78
Source | 93.5
Source | 86.4
Source | 87.9
Source | 93.8
Source | 99.8
Source | 94.9
Source | 122
Source | 81.5
Source | 70.6
Source | 83.3
Source | 85.3
Source | 102
Source | 98.7
Source | 108
Source | 133
Source | 88.2
Source | 71
Source | 79.3
Source | 87.9
Source | 63.6
Source | 69.6
Source | 88.9
Source | 97.5
Source | 94.4
Source | 103
Source | 83.8
Source | 94.7
Source | 86.7
Source | 79.8
Source | 94
Source | 95.8
Source | 143
Source | 135
Source | 105
Source | 105
Source | 90.1
Source | 104
Source | 63.3
Source | 82.6
Source | 104
Source | 112
Source | 115
Source | 116
Source | 107
Source | 118
Source | 88.2
Source | 79.8
Source | 99.9
Source | 108
Source | 142
Source | 106
Source | 82.6
Source | 101
Source | 92.5
Source | 95.7
Source | 76.3
Source | 80.8
Source | 128
Source | 99.3
Source | 84.5
Source | 92.8
Source | 107
Source | 99.7
Source | 128
Source | 127
Source | 104
Source | 83.6
Source | 107
Source | 99.2
Source | 83.5
Source | 82.2
Source | 81.5
Source | 95.4
Source | 83.3
Source | 83.7
Source | 108
Source | 76.9
Source | 105
Source | 124
Source | 73.4
Source | 68.4
Source | 83.2
Source | 86.4
Source | 100
Source | 90.6
Source | 108
Source | 77.9
Source | 75.7
Source | 83.8
Source | 112
Source | 125
Source | 90.1
Source | 81.2
Source | 96.6
Source | 99.7
Source | 84.9
Source | 74.2
Source | 101
Source | 77.1
Source | 97.4
Source | 91.4
Source | 69.1
Source | 80.9
Source | 76.6
Source | 77.8
Source | 58.8
Source | 78.9
Source | 94.8
Source | 86.7
Source | 83
Source | 91.4
Source | 64.6
Source | 73.8
Source | 87.3
Source | 77.1
Source | 99.3
Source | 105
Source | 103
Source | 108
Source | 89.9
Source | 91.5
Source | 86.8
Source | 94.6
Source | 108
Source | 88
Source | 106
Source | 138
Source | 78.6
Source | 82.2
Source | 70.6
Source | 71.8
Source | 97.5
Source | 106
Source | 113
Source | 108
Source | 99.5
Source | 103
Source | 84.1
Source | 66.6
Source | 101
Source | 107
Source | 81.1
Source | 65.3
Source | 101
Source | 92.4
Source | 126
Source | 107
Source | 127
Source | 132
Source | 87
Source | 98.5
Source | 86.4
Source | 79.2
Source | 92.3
Source | 88.2
Source | 81.1
Source | 89.9
Source | 84.1
Source | 87.5
Source | 89.2
Source | 92.4
Source | 93.5
Source | 91.6
Source | 88.2
Source | 73.7
Source | 87.9
Source | 82.8
Source | 66.9
Source | 72
Source | 93.3
Source | 105
Source | 99.6
Source | 113
Source | 74.1
Source | 92.1
Source | 105
Source | 86.2
Source | 78.1
Source | 85.1
Source | 78.2
Source | 99.5
Source | 96.9
Source | 100
Source | 109
Source | 112
Source | 72.3
Source | 70.6
Source | 90.4
Source | 109
Source | 106
Source | 93.7
Source | 81.1
Source | 98.3
Source | 95.3
Source | 76.6
Source | 96.3
Source | 85.4
Source | 100
Source | 101
Source | 102
Source | 98.3
Source | 117
Source | 93.7
Source | 97.6
Source | 97.9
Source | 117
Source | 144
Finished | 86
Finished | 89.5
Finished | 79.2
Finished | 74.3
Finished | 72.2
Finished | 96.7
Finished | 98.1
Finished | 109
Finished | 96.3
Finished | 105
Finished | 81.2
Finished | 81.3
Finished | 104
Finished | 122
Finished | 83.9
Finished | 105
Finished | 103
Finished | 69.7
Finished | 77.6
Finished | 93.8
Finished | 76.7
Finished | 78.1
Finished | 65.5
Finished | 101
Finished | 87.7
Finished | 80.3
Finished | 71.9
Finished | 112
Finished | 84.4
Finished | 89.2
Finished | 115
Finished | 111
Finished | 116
Finished | 64.3
Finished | 99.2
Finished | 124
Finished | 79.5
Finished | 73.2
Finished | 93.3
Finished | 103
Finished | 137
Finished | 82.6
Finished | 65.7
Finished | 107
Finished | 101
Finished | 81.5
Finished | 75.5
Finished | 78.5
Finished | 70.5
Finished | 89.8
Finished | 99.1
Finished | 77.6
Finished | 92.7
Finished | 99.1
Finished | 79.5
Finished | 85.3
Finished | 93.4
Finished | 104
Finished | 113
Finished | 90.3
Finished | 104
Finished | 102
Finished | 92
Finished | 90.6
Finished | 115
Finished | 116
Finished | 86
Finished | 71.1
Surface | 138
Surface | 137
Surface | 135
Surface | 132
Surface | 129
Surface | 126
Surface | 125
Surface | 125
Surface | 124
Surface | 123
Surface | 122
Surface | 122
Surface | 120
Surface | 120
Surface | 119
Surface | 117
Surface | 115
Surface | 114
Surface | 114
Surface | 113
Surface | 112
Surface | 109
Surface | 109
Surface | 108
Surface | 106
Surface | 105
Surface | 105
Surface | 105
Surface | 105
Surface | 105
Surface | 105
Surface | 103
Surface | 103
Surface | 100
Surface | 99
Surface | 97.8
Surface | 97.3
Surface | 96.6
Surface | 96.4
Surface | 95.6
Surface | 95.6
Surface | 94
Surface | 92.9
Surface | 92.8
Surface | 92.5
Surface | 92.3
Surface | 92.2
Surface | 91.4
Surface | 91.1
Surface | 90.2
Surface | 89.6
Surface | 89.1
Surface | 88.8
Surface | 88.3
Surface | 87.3
Surface | 86.9
Surface | 85
Surface | 84.5
Surface | 84.2
Surface | 84.1
Surface | 83.5
Surface | 83.1
Surface | 82.8
Surface | 82.2
Surface | 82
Surface | 81
Surface | 77.8
Surface | 70.3
Surface | 70
Surface | 68.7
Surface | 67.9
Surface | 63.5
Surface | 60.3
Surface | 54.3

## Tbl1-Tbl4_Lab2_LIMS_Data
Request Time | Name/ID | Test | REPLICATE | ABSORBANCE// | ABS_MEAN | ABS_CV | CONCENTRATION | CONC_MEAN | Absorbance | Concentration | Interpretation | Position | Reading
2021-03-11 18:52 | MCT SAES Std 0 | Microcystins-ADDA-SAES | 1 | 2.532 | 2.5025 | 1.7 | 0 | 0.8 | 2.532 Abs | 0.0 ng/L | R^2=0.99757 | RK1:23->A01@2 | T=230,A=+2.5320
2021-03-11 18:52 | MCT SAES Std 0 | Microcystins-ADDA-SAES | 2 | 2.473 | 2.5025 | 1.7 | 1.6 | 0.8 | 2.473 Abs [2.5025] {1.7 CV} | 1.6 ng/L [0.8] {141.4 CV} | R^2=0.99757 | RK1:23->B01@2 | T=229,A=+2.4732
2021-03-11 18:52 | MCT SAES Std 1 | Microcystins-ADDA-SAES | 1 | 2.119 | 2.1215 | 0.2 | 42.3 | 41.8 | 2.119 Abs | 42.3 ng/L | R^2=0.99757 | RK1:24->C01@2 | T=228,A=+2.1195
2021-03-11 18:52 | MCT SAES Std 1 | Microcystins-ADDA-SAES | 2 | 2.124 | 2.1215 | 0.2 | 41.4 | 41.8 | 2.124 Abs [2.1215] {0.2 CV} | 41.4 ng/L [41.8] {1.5 CV} | R^2=0.99757 | RK1:24->D01@2 | T=227,A=+2.1237
2021-03-11 18:52 | MCT SAES Std 2 | Microcystins-ADDA-SAES | 1 | 1.724 | 1.6795 | 3.7 | 148.5 | 167.5 | 1.724 Abs | 148.5 ng/L | R^2=0.99757 | RK1:25->E01@2 | T=226,A=+1.7242
2021-03-11 18:52 | MCT SAES Std 2 | Microcystins-ADDA-SAES | 2 | 1.635 | 1.6795 | 3.7 | 186.5 | 167.5 | 1.635 Abs [1.6795] {3.7 CV} | 186.5 ng/L [167.5] {16.0 CV} | R^2=0.99757 | RK1:25->F01@3 | T=225,A=+1.6346
2021-03-11 18:52 | MCT SAES Std 3 | Microcystins-ADDA-SAES | 1 | 1.354 | 1.302 | 5.6 | 367.6 | 419.4 | 1.354 Abs | 367.6 ng/L | R^2=0.99757 | RK1:26->G01@3 | T=224,A=+1.3540
2021-03-11 18:52 | MCT SAES Std 3 | Microcystins-ADDA-SAES | 2 | 1.25 | 1.302 | 5.6 | 471.2 | 419.4 | 1.250 Abs [1.3020] {5.6 CV} | 471.2 ng/L [419.4] {17.5 CV} | R^2=0.99757 | RK1:26->H01@3 | T=223,A=+1.2502
2021-03-11 18:52 | MCT SAES Std 4 | Microcystins-ADDA-SAES | 1 | 0.914 | 0.863 | 8.4 | 1117.1 | 1315.6 | 0.914 Abs | 1117.1 ng/L | R^2=0.99757 | RK1:27->A02@2 | T=226,A=+0.9137
2021-03-11 18:52 | MCT SAES Std 4 | Microcystins-ADDA-SAES | 2 | 0.812 | 0.863 | 8.4 | 1514.1 | 1315.6 | 0.812 Abs [0.8630] {8.4 CV} | 1514.1 ng/L [1315.6] {21.3 CV} | R^2=0.99757 | RK1:27->B02@2 | T=225,A=+0.8122
2021-03-11 18:52 | MCT SAES Std 5 | Microcystins-ADDA-SAES | 1 | 0.493 | 0.491 | 0.6 | > | > | 0.493 Abs | > 5000.0 ng/L |  | RK1:28->C02@2 | T=224,A=+0.4925
2021-03-11 18:52 | MCT SAES Std 5 | Microcystins-ADDA-SAES | 2 | 0.489 | 0.491 | 0.6 | > | > | 0.489 Abs [0.4910] {0.6 CV} | > 5000.0 ng/L |  | RK1:28->D02@2 | T=223,A=+0.4890
2021-03-11 18:54 | B21C148-BLK1 | Microcystins-ADDA-SAES | 1 | 2.377 | 2.337 | 2.4 | 8.5 | 12.6 | 2.377 Abs | 8.5 ng/L |  | RK2:01->E02@2 | T=222,A=+2.3767
2021-03-11 18:54 | B21C148-BLK1 | Microcystins-ADDA-SAES | 2 | 2.297 | 2.337 | 2.4 | 16.6 | 12.6 | 2.297 Abs [2.3370] {2.4 CV} | 16.6 ng/L [12.6] {45.6 CV} |  | RK2:01->F02@3 | T=221,A=+2.2970
2021-03-11 18:54 | S21C060-LCV1 | Microcystins-ADDA-SAES | 1 | 1.891 | 1.8955 | 0.3 | 93.3 | 92.1 | 1.891 Abs | 93.3 ng/L |  | RK2:02->G02@3 | T=220,A=+1.8912
2021-03-11 18:54 | S21C060-LCV1 | Microcystins-ADDA-SAES | 2 | 1.9 | 1.8955 | 0.3 | 90.8 | 92.1 | 1.900 Abs [1.8955] {0.3 CV} | 90.8 ng/L [92.1] {1.9 CV} |  | RK2:02->H02@3 | T=219,A=+1.8996
2021-03-11 18:54 | S21C060-SCV1 | Microcystins-ADDA-SAES | 1 | 1.554 | 1.4995 | 5.1 | 227.8 | 262 | 1.554 Abs | 227.8 ng/L |  | RK2:03->A03@2 | T=217,A=+1.5542
2021-03-11 18:54 | S21C060-SCV1 | Microcystins-ADDA-SAES | 2 | 1.445 | 1.4995 | 5.1 | 296.1 | 262 | 1.445 Abs [1.4995] {5.1 CV} | 296.1 ng/L [262.0] {18.4 CV} |  | RK2:03->B03@2 | T=216,A=+1.4452
2021-03-11 18:54 | B21C148-MRL1 | Microcystins-ADDA-SAES | 1 | 1.926 | 1.9375 | 0.8 | 83.8 | 80.9 | 1.926 Abs | 83.8 ng/L |  | RK2:04->C03@2 | T=215,A=+1.9261
2021-03-11 18:54 | B21C148-MRL1 | Microcystins-ADDA-SAES | 2 | 1.949 | 1.9375 | 0.8 | 78 | 80.9 | 1.949 Abs [1.9375] {0.8 CV} | 78.0 ng/L [80.9] {5.1 CV} |  | RK2:04->D03@2 | T=214,A=+1.9489
2021-03-11 18:54 | B21C148-MRLA | Microcystins-ADDA-SAES | 1 | 1.846 | 1.824 | 1.7 | 106.4 | 113.4 | 1.846 Abs | 106.4 ng/L |  | RK2:05->E03@2 | T=213,A=+1.8456
2021-03-11 18:54 | B21C148-MRLA | Microcystins-ADDA-SAES | 2 | 1.802 | 1.824 | 1.7 | 120.4 | 113.4 | 1.802 Abs [1.8240] {1.7 CV} | 120.4 ng/L [113.4] {8.7 CV} |  | RK2:05->F03@3 | T=212,A=+1.8024
2021-03-11 18:54 | B21C148-BS1 | Microcystins-ADDA-SAES | 1 | 1.204 | 1.1875 | 2 | 526.6 | 548.7 | 1.204 Abs | 526.6 ng/L |  | RK2:06->G03@3 | T=211,A=+1.2044
2021-03-11 18:54 | B21C148-BS1 | Microcystins-ADDA-SAES | 2 | 1.171 | 1.1875 | 2 | 570.8 | 548.7 | 1.171 Abs [1.1875] {2.0 CV} | 570.8 ng/L [548.7] {5.7 CV} |  | RK2:06->H03@3 | T=210,A=+1.1712
2021-03-11 18:54 | B21C148-BLK2 | Microcystins-ADDA-SAES | 1 | 2.343 | 2.3545 | 0.7 | 11.7 | 10.6 | 2.343 Abs | 11.7 ng/L |  | RK2:07->A04@1 | T=210,A=+2.3431
2021-03-11 18:54 | B21C148-BLK2 | Microcystins-ADDA-SAES | 2 | 2.366 | 2.3545 | 0.7 | 9.5 | 10.6 | 2.366 Abs [2.3545] {0.7 CV} | 9.5 ng/L [10.6] {14.7 CV} |  | RK2:07->B04@1 | T=209,A=+2.3661
2021-03-11 18:54 | B21C148-MRL2 | Microcystins-ADDA-SAES | 1 | 2.008 | 2.0095 | 0.1 | 64.2 | 63.8 | 2.008 Abs | 64.2 ng/L |  | RK2:08->C04@1 | T=208,A=+2.0082
2021-03-11 18:54 | B21C148-MRL2 | Microcystins-ADDA-SAES | 2 | 2.011 | 2.0095 | 0.1 | 63.5 | 63.8 | 2.011 Abs [2.0095] {0.1 CV} | 63.5 ng/L [63.8] {0.8 CV} |  | RK2:08->D04@1 | T=207,A=+2.0114
2021-03-11 18:54 | B21C148-MRLB | Microcystins-ADDA-SAES | 1 | 1.86 | 1.828 | 2.5 | 102.2 | 112.3 | 1.860 Abs | 102.2 ng/L |  | RK2:09->E04@1 | T=206,A=+1.8595
2021-03-11 18:54 | B21C148-MRLB | Microcystins-ADDA-SAES | 2 | 1.796 | 1.828 | 2.5 | 122.4 | 112.3 | 1.796 Abs [1.8280] {2.5 CV} | 122.4 ng/L [112.3] {12.7 CV} |  | RK2:09->F04@4 | T=205,A=+1.7960
2021-03-11 18:54 | B21C148-BS2 | Microcystins-ADDA-SAES | 1 | 1.234 | 1.216 | 2.1 | 489.7 | 512 | 1.234 Abs | 489.7 ng/L |  | RK2:10->G04@4 | T=204,A=+1.2341
2021-03-11 18:54 | B21C148-BS2 | Microcystins-ADDA-SAES | 2 | 1.198 | 1.216 | 2.1 | 534.3 | 512 | 1.198 Abs [1.2160] {2.1 CV} | 534.3 ng/L [512.0] {6.2 CV} |  | RK2:10->H04@4 | T=203,A=+1.1979
2021-03-11 18:54 | S21C060-SCV2 | Microcystins-ADDA-SAES | 1 | 1.547 | 1.515 | 3 | 231.7 | 251 | 1.547 Abs | 231.7 ng/L |  | RK2:11->A05@1 | T=202,A=+1.5470
2021-03-11 18:54 | S21C060-SCV2 | Microcystins-ADDA-SAES | 2 | 1.483 | 1.515 | 3 | 270.4 | 251 | 1.483 Abs [1.5150] {3.0 CV} | 270.4 ng/L [251.0] {10.9 CV} |  | RK2:11->B05@1 | T=201,A=+1.4832
2021-03-11 18:54 | B21C148-MRL3 | Microcystins-ADDA-SAES | 1 | 1.938 | 1.9635 | 1.8 | 80.8 | 74.6 | 1.938 Abs | 80.8 ng/L |  | RK2:12->C05@1 | T=200,A=+1.9376
2021-03-11 18:54 | B21C148-MRL3 | Microcystins-ADDA-SAES | 2 | 1.989 | 1.9635 | 1.8 | 68.4 | 74.6 | 1.989 Abs [1.9635] {1.8 CV} | 68.4 ng/L [74.6] {11.8 CV} |  | RK2:12->D05@1 | T=199,A=+1.9891
2021-03-11 18:54 | B21C148-MRLC | Microcystins-ADDA-SAES | 1 | 1.856 | 1.8225 | 2.6 | 103.4 | 114.1 | 1.856 Abs | 103.4 ng/L |  | RK2:13->E05@1 | T=198,A=+1.8564
2021-03-11 18:54 | B21C148-MRLC | Microcystins-ADDA-SAES | 2 | 1.789 | 1.8225 | 2.6 | 124.8 | 114.1 | 1.789 Abs [1.8225] {2.6 CV} | 124.8 ng/L [114.1] {13.3 CV} |  | RK2:13->F05@4 | T=197,A=+1.7891
2021-03-11 18:54 | B21C148-BS3 | Microcystins-ADDA-SAES | 1 | 1.252 | 1.215 | 4.3 | 468.9 | 515 | 1.252 Abs | 468.9 ng/L |  | RK2:14->G05@4 | T=196,A=+1.2519
2021-03-11 18:54 | B21C148-BS3 | Microcystins-ADDA-SAES | 2 | 1.178 | 1.215 | 4.3 | 561.1 | 515 | 1.178 Abs [1.2150] {4.3 CV} | 561.1 ng/L [515.0] {12.7 CV} |  | RK2:14->H05@4 | T=195,A=+1.1781
2021-03-11 18:54 | B21C148-BLK3 | Microcystins-ADDA-SAES | 1 | 2.394 | 2.3735 | 1.2 | 7 | 8.97 | 2.394 Abs | 7.0 ng/L |  | RK2:15->A06@1 | T=196,A=+2.3941
2021-03-11 18:54 | B21C148-BLK3 | Microcystins-ADDA-SAES | 2 | 2.353 | 2.3735 | 1.2 | 10.7 | 8.97 | 2.353 Abs [2.3735] {1.2 CV} | 10.7 ng/L [8.9] {29.6 CV} |  | RK2:15->B06@1 | T=195,A=+2.3530
2021-03-11 18:54 | B21C148-MRL4 | Microcystins-ADDA-SAES | 1 | 2.031 | 2.0135 | 1.2 | 59.2 | 63.1 | 2.031 Abs | 59.2 ng/L |  | RK2:16->C06@1 | T=194,A=+2.0307
2021-03-11 18:54 | B21C148-MRL4 | Microcystins-ADDA-SAES | 2 | 1.996 | 2.0135 | 1.2 | 66.9 | 63.1 | 1.996 Abs [2.0135] {1.2 CV} | 66.9 ng/L [63.1] {8.6 CV} |  | RK2:16->D06@1 | T=193,A=+1.9956
2021-03-11 18:54 | B21C148-MRLD | Microcystins-ADDA-SAES | 1 | 1.897 | 1.8425 | 4.2 | 91.6 | 108.3 | 1.897 Abs | 91.6 ng/L |  | RK2:17->E06@1 | T=192,A=+1.8974
2021-03-11 18:54 | B21C148-MRLD | Microcystins-ADDA-SAES | 2 | 1.788 | 1.8425 | 4.2 | 125.1 | 108.3 | 1.788 Abs [1.8425] {4.2 CV} | 125.1 ng/L [108.3] {21.9 CV} |  | RK2:17->F06@4 | T=191,A=+1.7875
2021-03-11 18:54 | B21C148-BS4 | Microcystins-ADDA-SAES | 1 | 1.264 | 1.249 | 1.7 | 455.6 | 472.7 | 1.264 Abs | 455.6 ng/L |  | RK2:18->G06@4 | T=190,A=+1.2642
2021-03-11 18:54 | B21C148-BS4 | Microcystins-ADDA-SAES | 2 | 1.234 | 1.249 | 1.7 | 489.7 | 472.7 | 1.234 Abs [1.2490] {1.7 CV} | 489.7 ng/L [472.7] {5.1 CV} |  | RK2:18->H06@4 | T=189,A=+1.2339
2021-03-11 18:54 | B21C148-BLK4 | Microcystins-ADDA-SAES | 1 | 2.384 | 2.383 | 0.1 | 7.9 | 7.9 | 2.384 Abs | 7.9 ng/L |  | RK2:19->A07@2 | T=186,A=+2.3839
2021-03-11 18:54 | B21C148-BLK4 | Microcystins-ADDA-SAES | 2 | 2.382 | 2.383 | 0.1 | 8 | 7.9 | 2.382 Abs [2.3830] {0.1 CV} | 8.0 ng/L [7.9] {0.9 CV} |  | RK2:19->B07@2 | T=185,A=+2.3817
2021-03-11 18:54 | B21C148-MRL5 | Microcystins-ADDA-SAES | 1 | 1.992 | 1.981 | 0.8 | 67.8 | 70.4 | 1.992 Abs | 67.8 ng/L |  | RK2:20->C07@2 | T=184,A=+1.9924
2021-03-11 18:54 | B21C148-MRL5 | Microcystins-ADDA-SAES | 2 | 1.97 | 1.981 | 0.8 | 72.9 | 70.4 | 1.970 Abs [1.9810] {0.8 CV} | 72.9 ng/L [70.4] {5.1 CV} |  | RK2:20->D07@2 | T=183,A=+1.9699
2021-03-11 18:54 | B21C148-MRLE | Microcystins-ADDA-SAES | 1 | 1.873 | 1.8395 | 2.6 | 98.4 | 108.8 | 1.873 Abs | 98.4 ng/L |  | RK2:21->E07@2 | T=182,A=+1.8735
2021-03-11 18:54 | B21C148-MRLE | Microcystins-ADDA-SAES | 2 | 1.806 | 1.8395 | 2.6 | 119.1 | 108.8 | 1.806 Abs [1.8395] {2.6 CV} | 119.1 ng/L [108.8] {13.5 CV} |  | RK2:21->F07@3 | T=181,A=+1.8055
2021-03-11 18:54 | B21C148-BS5 | Microcystins-ADDA-SAES | 1 | 1.28 | 1.246 | 3.9 | 438.5 | 477.5 | 1.280 Abs | 438.5 ng/L |  | RK2:22->G07@3 | T=180,A=+1.2797
2021-03-11 18:54 | B21C148-BS5 | Microcystins-ADDA-SAES | 2 | 1.212 | 1.246 | 3.9 | 516.5 | 477.5 | 1.212 Abs [1.2460] {3.9 CV} | 516.5 ng/L [477.5] {11.6 CV} |  | RK2:22->H07@3 | T=179,A=+1.2119
2021-03-11 18:54 | S21C060-SCV3 | Microcystins-ADDA-SAES | 1 | 1.484 | 1.5065 | 2.1 | 269.7 | 255.9 | 1.484 Abs | 269.7 ng/L |  | RK2:23->A08@2 | T=177,A=+1.4839
2021-03-11 18:54 | S21C060-SCV3 | Microcystins-ADDA-SAES | 2 | 1.529 | 1.5065 | 2.1 | 242 | 255.9 | 1.529 Abs [1.5065] {2.1 CV} | 242.0 ng/L [255.9] {7.7 CV} |  | RK2:23->B08@2 | T=176,A=+1.5289
2021-03-11 18:54 | B21C148-MRL6 | Microcystins-ADDA-SAES | 1 | 2.019 | 2.0015 | 1.2 | 61.8 | 65.7 | 2.019 Abs | 61.8 ng/L |  | RK2:24->C08@2 | T=175,A=+2.0185
2021-03-11 18:54 | B21C148-MRL6 | Microcystins-ADDA-SAES | 2 | 1.984 | 2.0015 | 1.2 | 69.6 | 65.7 | 1.984 Abs [2.0015] {1.2 CV} | 69.6 ng/L [65.7] {8.4 CV} |  | RK2:24->D08@2 | T=174,A=+1.9840
2021-03-11 18:54 | B21C148-MRLF | Microcystins-ADDA-SAES | 1 | 1.9 | 1.876 | 1.8 | 90.8 | 97.76 | 1.900 Abs | 90.8 ng/L |  | RK2:25->E08@2 | T=173,A=+1.8997
2021-03-11 18:54 | B21C148-MRLF | Microcystins-ADDA-SAES | 2 | 1.852 | 1.876 | 1.8 | 104.6 | 97.76 | 1.852 Abs [1.8760] {1.8 CV} | 104.6 ng/L [97.7] {10.0 CV} |  | RK2:25->F08@3 | T=172,A=+1.8516
2021-03-11 18:54 | B21C148-BS6 | Microcystins-ADDA-SAES | 1 | 1.291 | 1.2705 | 2.3 | 427.1 | 449.2 | 1.291 Abs | 427.1 ng/L |  | RK2:26->G08@3 | T=171,A=+1.2906
2021-03-11 18:54 | B21C148-BS6 | Microcystins-ADDA-SAES | 2 | 1.25 | 1.2705 | 2.3 | 471.2 | 449.2 | 1.250 Abs [1.2705] {2.3 CV} | 471.2 ng/L [449.2] {6.9 CV} |  | RK2:26->H08@3 | T=170,A=+1.2501
2021-03-11 18:54 | B21C148-BLK5 | Microcystins-ADDA-SAES | 1 | 2.456 | 2.439 | 1 | 2.5 | 3.7 | 2.456 Abs | 2.5 ng/L |  | RK2:27->A09@2 | T=167,A=+2.4559
2021-03-11 18:54 | B21C148-BLK5 | Microcystins-ADDA-SAES | 2 | 2.422 | 2.439 | 1 | 4.8 | 3.7 | 2.422 Abs [2.4390] {1.0 CV} | 4.8 ng/L [3.7] {44.6 CV} |  | RK2:27->B09@2 | T=166,A=+2.4224
2021-03-11 18:54 | B21C148-MRL7 | Microcystins-ADDA-SAES | 1 | 2.029 | 2.0725 | 3 | 59.6 | 51.2 | 2.029 Abs | 59.6 ng/L |  | RK2:28->C09@2 | T=165,A=+2.0294
2021-03-11 18:54 | B21C148-MRL7 | Microcystins-ADDA-SAES | 2 | 2.116 | 2.0725 | 3 | 42.8 | 51.2 | 2.116 Abs [2.0725] {3.0 CV} | 42.8 ng/L [51.2] {23.2 CV} |  | RK2:28->D09@2 | T=164,A=+2.1161
2021-03-11 18:54 | B21C148-MRLG | Microcystins-ADDA-SAES | 1 | 1.898 | 1.8655 | 2.5 | 91.3 | 100.9 | 1.898 Abs | 91.3 ng/L |  | RK2:29->E09@2 | T=163,A=+1.8977
2021-03-11 18:54 | B21C148-MRLG | Microcystins-ADDA-SAES | 2 | 1.833 | 1.8655 | 2.5 | 110.4 | 100.9 | 1.833 Abs [1.8655] {2.5 CV} | 110.4 ng/L [100.9] {13.4 CV} |  | RK2:29->F09@3 | T=162,A=+1.8331
2021-03-11 18:54 | B21C148-BS7 | Microcystins-ADDA-SAES | 1 | 1.225 | 1.2595 | 3.9 | 500.5 | 462.3 | 1.225 Abs | 500.5 ng/L |  | RK2:30->G09@3 | T=161,A=+1.2246
2021-03-11 18:54 | B21C148-BS7 | Microcystins-ADDA-SAES | 2 | 1.294 | 1.2595 | 3.9 | 424.1 | 462.3 | 1.294 Abs [1.2595] {3.9 CV} | 424.1 ng/L [462.3] {11.7 CV} |  | RK2:30->H09@3 | T=160,A=+1.2941
2021-03-15 16:16 | MCT SAES Std 0 | Microcystins-ADDA-SAES | 1 | 2.249 | 2.311 | 3.8 | 4.6 | 2.3 | 2.249 Abs | 4.6 ng/L | R^2=0.99541 | RK1:23->A01@2 | T=257,A=+2.2489
2021-03-15 16:16 | MCT SAES Std 0 | Microcystins-ADDA-SAES | 2 | 2.373 | 2.311 | 3.8 | 0 | 2.3 | 2.373 Abs [2.3110] {3.8 CV} | 0.0 ng/L [2.3] {141.4 CV} | R^2=0.99541 | RK1:23->B01@2 | T=256,A=+2.3730
2021-03-15 16:16 | MCT SAES Std 1 | Microcystins-ADDA-SAES | 1 | 2.004 | 2 | 0.3 | 36.9 | 37.6 | 2.004 Abs | 36.9 ng/L | R^2=0.99541 | RK1:24->C01@2 | T=255,A=+2.0039
2021-03-15 16:16 | MCT SAES Std 1 | Microcystins-ADDA-SAES | 2 | 1.996 | 2 | 0.3 | 38.3 | 37.6 | 1.996 Abs [2.0000] {0.3 CV} | 38.3 ng/L [37.6] {2.6 CV} | R^2=0.99541 | RK1:24->D01@2 | T=254,A=+1.9962
2021-03-15 16:16 | MCT SAES Std 2 | Microcystins-ADDA-SAES | 1 | 1.555 | 1.529 | 2.4 | 172.1 | 184.9 | 1.555 Abs | 172.1 ng/L | R^2=0.99541 | RK1:25->E01@2 | T=253,A=+1.5549
2021-03-15 16:16 | MCT SAES Std 2 | Microcystins-ADDA-SAES | 2 | 1.503 | 1.529 | 2.4 | 197.7 | 184.9 | 1.503 Abs [1.5290] {2.4 CV} | 197.7 ng/L [184.9] {9.8 CV} | R^2=0.99541 | RK1:25->F01@3 | T=252,A=+1.5035
2021-03-15 16:16 | MCT SAES Std 3 | Microcystins-ADDA-SAES | 1 | 1.235 | 1.2385 | 0.4 | 389.7 | 386.4 | 1.235 Abs | 389.7 ng/L | R^2=0.99541 | RK1:26->G01@3 | T=251,A=+1.2352
2021-03-15 16:16 | MCT SAES Std 3 | Microcystins-ADDA-SAES | 2 | 1.242 | 1.2385 | 0.4 | 383 | 386.4 | 1.242 Abs [1.2385] {0.4 CV} | 383.0 ng/L [386.4] {1.2 CV} | R^2=0.99541 | RK1:26->H01@3 | T=250,A=+1.2416
2021-03-15 16:16 | MCT SAES Std 4 | Microcystins-ADDA-SAES | 1 | 0.802 | 0.7755 | 4.8 | 1240.3 | 1352.6 | 0.802 Abs | 1240.3 ng/L | R^2=0.99541 | RK1:27->A02@2 | T=253,A=+0.8021
2021-03-15 16:16 | MCT SAES Std 4 | Microcystins-ADDA-SAES | 2 | 0.749 | 0.7755 | 4.8 | 1464.9 | 1352.6 | 0.749 Abs [0.7755] {4.8 CV} | 1464.9 ng/L [1352.6] {11.7 CV} | R^2=0.99541 | RK1:27->B02@2 | T=252,A=+0.7487
2021-03-15 16:16 | MCT SAES Std 5 | Microcystins-ADDA-SAES | 1 | 0.454 | 0.4355 | 6 | > | > | 0.454 Abs | > 5000.0 ng/L |  | RK1:28->C02@2 | T=251,A=+0.4542
2021-03-15 16:16 | MCT SAES Std 5 | Microcystins-ADDA-SAES | 2 | 0.417 | 0.4355 | 6 | > | > | 0.417 Abs [0.4355] {6.0 CV} | > 5000.0 ng/L |  | RK1:28->D02@2 | T=250,A=+0.4175
2021-03-15 16:18 | RW.BLK.0.1 | Microcystins-ADDA-SAES | 1 | 2.241 | 2.191 | 3.2 | 5.3 | 10.6 | 2.241 Abs | 5.3 ng/L |  | RK2:01->E02@2 | T=249,A=+2.2406
2021-03-15 16:18 | RW.BLK.0.1 | Microcystins-ADDA-SAES | 2 | 2.141 | 2.191 | 3.2 | 16 | 10.6 | 2.141 Abs [2.1910] {3.2 CV} | 16.0 ng/L [10.6] {71.0 CV} |  | RK2:01->F02@3 | T=248,A=+2.1407
2021-03-15 16:18 | LowCV.75.1 | Microcystins-ADDA-SAES | 1 | 1.686 | 1.744 | 4.7 | 118.9 | 100.7 | 1.686 Abs | 118.9 ng/L |  | RK2:02->G02@3 | T=247,A=+1.6856
2021-03-15 16:18 | LowCV.75.1 | Microcystins-ADDA-SAES | 2 | 1.802 | 1.744 | 4.7 | 82.5 | 100.7 | 1.802 Abs [1.7440] {4.7 CV} | 82.5 ng/L [100.7] {25.6 CV} |  | RK2:02->H02@3 | T=246,A=+1.8021
2021-03-15 16:18 | DW.115.1 | Microcystins-ADDA-SAES | 1 | 1.619 | 1.6005 | 1.6 | 144.3 | 152.1 | 1.619 Abs | 144.3 ng/L |  | RK2:03->A03@2 | T=248,A=+1.6192
2021-03-15 16:18 | DW.115.1 | Microcystins-ADDA-SAES | 2 | 1.582 | 1.6005 | 1.6 | 159.9 | 152.1 | 1.582 Abs [1.6005] {1.6 CV} | 159.9 ng/L [152.1] {7.3 CV} |  | RK2:03->B03@2 | T=247,A=+1.5817
2021-03-15 16:18 | DW.300.1 | Microcystins-ADDA-SAES | 1 | 1.225 | 1.2405 | 1.8 | 399.5 | 384.6 | 1.225 Abs | 399.5 ng/L |  | RK2:04->C03@2 | T=246,A=+1.2254
2021-03-15 16:18 | DW.300.1 | Microcystins-ADDA-SAES | 2 | 1.256 | 1.2405 | 1.8 | 369.8 | 384.6 | 1.256 Abs [1.2405] {1.8 CV} | 369.8 ng/L [384.6] {5.5 CV} |  | RK2:04->D03@2 | T=245,A=+1.2562
2021-03-15 16:18 | DW.BLK.0.1 | Microcystins-ADDA-SAES | 1 | 2.168 | 2.165 | 0.2 | 12.7 | 13 | 2.168 Abs | 12.7 ng/L |  | RK2:05->E03@2 | T=244,A=+2.1679
2021-03-15 16:18 | DW.BLK.0.1 | Microcystins-ADDA-SAES | 2 | 2.162 | 2.165 | 0.2 | 13.4 | 13 | 2.162 Abs [2.1650] {0.2 CV} | 13.4 ng/L [13.0] {3.8 CV} |  | RK2:05->F03@3 | T=243,A=+2.1625
2021-03-15 16:18 | DW.500.1 | Microcystins-ADDA-SAES | 1 | 1.064 | 1.08 | 2.1 | 600 | 576.4 | 1.064 Abs | 600.0 ng/L |  | RK2:06->G03@3 | T=242,A=+1.0641
2021-03-15 16:18 | DW.500.1 | Microcystins-ADDA-SAES | 2 | 1.096 | 1.08 | 2.1 | 552.8 | 576.4 | 1.096 Abs [1.0800] {2.1 CV} | 552.8 ng/L [576.4] {5.8 CV} |  | RK2:06->H03@3 | T=241,A=+1.0956
2021-03-15 16:18 | RW.115.1 | Microcystins-ADDA-SAES | 1 | 1.555 | 1.5825 | 2.5 | 172.1 | 160.1 | 1.555 Abs | 172.1 ng/L |  | RK2:07->A04@1 | T=239,A=+1.5547
2021-03-15 16:18 | RW.115.1 | Microcystins-ADDA-SAES | 2 | 1.61 | 1.5825 | 2.5 | 148 | 160.1 | 1.610 Abs [1.5825] {2.5 CV} | 148.0 ng/L [160.1] {10.6 CV} |  | RK2:07->B04@1 | T=238,A=+1.6096
2021-03-15 16:18 | DW.300.2 | Microcystins-ADDA-SAES | 1 | 1.292 | 1.2905 | 0.2 | 338.1 | 339.4 | 1.292 Abs | 338.1 ng/L |  | RK2:08->C04@1 | T=237,A=+1.2921
2021-03-15 16:18 | DW.300.2 | Microcystins-ADDA-SAES | 2 | 1.289 | 1.2905 | 0.2 | 340.6 | 339.4 | 1.289 Abs [1.2905] {0.2 CV} | 340.6 ng/L [339.4] {0.5 CV} |  | RK2:08->D04@1 | T=236,A=+1.2886
2021-03-15 16:18 | RW.BLK.0.2 | Microcystins-ADDA-SAES | 1 | 2.209 | 2.196 | 0.8 | 8.3 | 9.6 | 2.209 Abs | 8.3 ng/L |  | RK2:09->E04@1 | T=235,A=+2.2087
2021-03-15 16:18 | RW.BLK.0.2 | Microcystins-ADDA-SAES | 2 | 2.183 | 2.196 | 0.8 | 11 | 9.6 | 2.183 Abs [2.1960] {0.8 CV} | 11.0 ng/L [9.6] {19.8 CV} |  | RK2:09->F04@4 | T=234,A=+2.1832
2021-03-15 16:18 | RW.500.1 | Microcystins-ADDA-SAES | 1 | 1.072 | 1.0845 | 1.6 | 587.8 | 569.6 | 1.072 Abs | 587.8 ng/L |  | RK2:10->G04@4 | T=233,A=+1.0722
2021-03-15 16:18 | RW.500.1 | Microcystins-ADDA-SAES | 2 | 1.097 | 1.0845 | 1.6 | 551.4 | 569.6 | 1.097 Abs [1.0845] {1.6 CV} | 551.4 ng/L [569.6] {4.5 CV} |  | RK2:10->H04@4 | T=232,A=+1.0970
2021-03-15 16:18 | DW.115.2 | Microcystins-ADDA-SAES | 1 | 1.605 | 1.612 | 0.6 | 150.1 | 147.2 | 1.605 Abs | 150.1 ng/L |  | RK2:11->A05@1 | T=229,A=+1.6045
2021-03-15 16:18 | DW.115.2 | Microcystins-ADDA-SAES | 2 | 1.619 | 1.612 | 0.6 | 144.3 | 147.2 | 1.619 Abs [1.6120] {0.6 CV} | 144.3 ng/L [147.2] {2.8 CV} |  | RK2:11->B05@1 | T=228,A=+1.6189
2021-03-15 16:18 | DW.300.3 | Microcystins-ADDA-SAES | 1 | 1.304 | 1.297 | 0.8 | 328.1 | 334 | 1.304 Abs | 328.1 ng/L |  | RK2:12->C05@1 | T=227,A=+1.3045
2021-03-15 16:18 | DW.300.3 | Microcystins-ADDA-SAES | 2 | 1.29 | 1.297 | 0.8 | 339.8 | 334 | 1.290 Abs [1.2970] {0.8 CV} | 339.8 ng/L [334.0] {2.5 CV} |  | RK2:12->D05@1 | T=226,A=+1.2904
2021-03-15 16:18 | DW.BLK.0.2 | Microcystins-ADDA-SAES | 1 | 2.201 | 2.1685 | 2.1 | 9.1 | 12.9 | 2.201 Abs | 9.1 ng/L |  | RK2:13->E05@1 | T=225,A=+2.2014
2021-03-15 16:18 | DW.BLK.0.2 | Microcystins-ADDA-SAES | 2 | 2.136 | 2.1685 | 2.1 | 16.6 | 12.9 | 2.136 Abs [2.1685] {2.1 CV} | 16.6 ng/L [12.9] {41.3 CV} |  | RK2:13->F05@4 | T=224,A=+2.1362
2021-03-15 16:18 | DW.500.2 | Microcystins-ADDA-SAES | 1 | 1.097 | 1.098 | 0.1 | 551.4 | 550 | 1.097 Abs | 551.4 ng/L |  | RK2:14->G05@4 | T=223,A=+1.0970
2021-03-15 16:18 | DW.500.2 | Microcystins-ADDA-SAES | 2 | 1.099 | 1.098 | 0.1 | 548.6 | 550 | 1.099 Abs [1.0980] {0.1 CV} | 548.6 ng/L [550.0] {0.4 CV} |  | RK2:14->H05@4 | T=222,A=+1.0985
2021-03-15 16:18 | RW.115.2 | Microcystins-ADDA-SAES | 1 | 1.611 | 1.5955 | 1.4 | 147.6 | 154.2 | 1.611 Abs | 147.6 ng/L |  | RK2:15->A06@1 | T=222,A=+1.6106
2021-03-15 16:18 | RW.115.2 | Microcystins-ADDA-SAES | 2 | 1.58 | 1.5955 | 1.4 | 160.8 | 154.2 | 1.580 Abs [1.5955] {1.4 CV} | 160.8 ng/L [154.2] {6.1 CV} |  | RK2:15->B06@1 | T=221,A=+1.5796
2021-03-15 16:18 | DW.300.4 | Microcystins-ADDA-SAES | 1 | 1.325 | 1.2865 | 4.2 | 311.3 | 344.3 | 1.325 Abs | 311.3 ng/L |  | RK2:16->C06@1 | T=220,A=+1.3254
2021-03-15 16:18 | DW.300.4 | Microcystins-ADDA-SAES | 2 | 1.248 | 1.2865 | 4.2 | 377.3 | 344.3 | 1.248 Abs [1.2865] {4.2 CV} | 377.3 ng/L [344.3] {13.6 CV} |  | RK2:16->D06@1 | T=219,A=+1.2479
2021-03-15 16:18 | RW.BLK.0.3 | Microcystins-ADDA-SAES | 1 | 2.248 | 2.238 | 0.6 | 4.7 | 5.6 | 2.248 Abs | 4.7 ng/L |  | RK2:17->E06@1 | T=218,A=+2.2484
2021-03-15 16:18 | RW.BLK.0.3 | Microcystins-ADDA-SAES | 2 | 2.228 | 2.238 | 0.6 | 6.5 | 5.6 | 2.228 Abs [2.2380] {0.6 CV} | 6.5 ng/L [5.6] {22.7 CV} |  | RK2:17->F06@4 | T=217,A=+2.2285
2021-03-15 16:18 | RW.500.2 | Microcystins-ADDA-SAES | 1 | 1.171 | 1.135 | 4.5 | 457.3 | 502.9 | 1.171 Abs | 457.3 ng/L |  | RK2:18->G06@4 | T=216,A=+1.1714
2021-03-15 16:18 | RW.500.2 | Microcystins-ADDA-SAES | 2 | 1.099 | 1.135 | 4.5 | 548.6 | 502.9 | 1.099 Abs [1.1350] {4.5 CV} | 548.6 ng/L [502.9] {12.8 CV} |  | RK2:18->H06@4 | T=215,A=+1.0991
2021-03-15 16:18 | DW.115.3 | Microcystins-ADDA-SAES | 1 | 1.665 | 1.6655 | 0 | 126.5 | 126.3 | 1.665 Abs | 126.5 ng/L |  | RK2:19->A07@2 | T=212,A=+1.6647
2021-03-15 16:18 | DW.115.3 | Microcystins-ADDA-SAES | 2 | 1.666 | 1.6655 | 0 | 126.1 | 126.3 | 1.666 Abs [1.6655] {0.0 CV} | 126.1 ng/L [126.3] {0.2 CV} |  | RK2:19->B07@2 | T=211,A=+1.6664
2021-03-15 16:18 | DW.300.5 | Microcystins-ADDA-SAES | 1 | 1.248 | 1.245 | 0.3 | 377.3 | 380.1 | 1.248 Abs | 377.3 ng/L |  | RK2:20->C07@2 | T=210,A=+1.2485
2021-03-15 16:18 | DW.300.5 | Microcystins-ADDA-SAES | 2 | 1.242 | 1.245 | 0.3 | 383 | 380.1 | 1.242 Abs [1.2450] {0.3 CV} | 383.0 ng/L [380.1] {1.1 CV} |  | RK2:20->D07@2 | T=209,A=+1.2421
2021-03-15 16:18 | DW.BLK.0.3 | Microcystins-ADDA-SAES | 1 | 2.258 | 2.1985 | 3.8 | 3.9 | 10.1 | 2.258 Abs | 3.9 ng/L |  | RK2:21->E07@2 | T=208,A=+2.2579
2021-03-15 16:18 | DW.BLK.0.3 | Microcystins-ADDA-SAES | 2 | 2.139 | 2.1985 | 3.8 | 16.3 | 10.1 | 2.139 Abs [2.1985] {3.8 CV} | 16.3 ng/L [10.1] {86.8 CV} |  | RK2:21->F07@3 | T=207,A=+2.1394
2021-03-15 16:18 | DW.500.3 | Microcystins-ADDA-SAES | 1 | 1.153 | 1.112 | 5.2 | 478.5 | 533.9 | 1.153 Abs | 478.5 ng/L |  | RK2:22->G07@3 | T=206,A=+1.1530
2021-03-15 16:18 | DW.500.3 | Microcystins-ADDA-SAES | 2 | 1.071 | 1.112 | 5.2 | 589.3 | 533.9 | 1.071 Abs [1.1120] {5.2 CV} | 589.3 ng/L [533.9] {14.7 CV} |  | RK2:22->H07@3 | T=205,A=+1.0711
2021-03-15 16:18 | RW.115.3 | Microcystins-ADDA-SAES | 1 | 1.598 | 1.6465 | 4.2 | 153 | 134.4 | 1.598 Abs | 153.0 ng/L |  | RK2:23->A08@2 | T=202,A=+1.5981
2021-03-15 16:18 | RW.115.3 | Microcystins-ADDA-SAES | 2 | 1.695 | 1.6465 | 4.2 | 115.8 | 134.4 | 1.695 Abs [1.6465] {4.2 CV} | 115.8 ng/L [134.4] {19.6 CV} |  | RK2:23->B08@2 | T=201,A=+1.6950
2021-03-15 16:18 | DW.300.6 | Microcystins-ADDA-SAES | 1 | 1.321 | 1.2935 | 3 | 314.5 | 337.6 | 1.321 Abs | 314.5 ng/L |  | RK2:24->C08@2 | T=200,A=+1.3215
2021-03-15 16:18 | DW.300.6 | Microcystins-ADDA-SAES | 2 | 1.266 | 1.2935 | 3 | 360.7 | 337.6 | 1.266 Abs [1.2935] {3.0 CV} | 360.7 ng/L [337.6] {9.7 CV} |  | RK2:24->D08@2 | T=199,A=+1.2657
2021-03-15 16:18 | RW.500.3 | Microcystins-ADDA-SAES | 1 | 1.104 | 1.094 | 1.3 | 541.6 | 555.8 | 1.104 Abs | 541.6 ng/L |  | RK2:25->E08@2 | T=198,A=+1.1042
2021-03-15 16:18 | RW.500.3 | Microcystins-ADDA-SAES | 2 | 1.084 | 1.094 | 1.3 | 570 | 555.8 | 1.084 Abs [1.0940] {1.3 CV} | 570.0 ng/L [555.8] {3.6 CV} |  | RK2:25->F08@3 | T=197,A=+1.0835
2021-03-15 16:18 | DW.500.4 | Microcystins-ADDA-SAES | 1 | 1.091 | 1.083 | 1 | 559.9 | 571.6 | 1.091 Abs | 559.9 ng/L |  | RK2:26->G08@3 | T=196,A=+1.0914
2021-03-15 16:18 | DW.500.4 | Microcystins-ADDA-SAES | 2 | 1.075 | 1.083 | 1 | 583.3 | 571.6 | 1.075 Abs [1.0830] {1.0 CV} | 583.3 ng/L [571.6] {2.9 CV} |  | RK2:26->H08@3 | T=195,A=+1.0752
2021-03-15 16:18 | DW.115.4 | Microcystins-ADDA-SAES | 1 | 1.694 | 1.6645 | 2.5 | 116.1 | 127 | 1.694 Abs | 116.1 ng/L |  | RK2:27->A09@2 | T=191,A=+1.6936
2021-03-15 16:18 | DW.115.4 | Microcystins-ADDA-SAES | 2 | 1.635 | 1.6645 | 2.5 | 137.9 | 127 | 1.635 Abs [1.6645] {2.5 CV} | 137.9 ng/L [127.0] {12.1 CV} |  | RK2:27->B09@2 | T=190,A=+1.6347
2021-03-15 16:18 | DW.300.7 | Microcystins-ADDA-SAES | 1 | 1.314 | 1.3005 | 1.5 | 320 | 331.1 | 1.314 Abs | 320.0 ng/L |  | RK2:28->C09@2 | T=189,A=+1.3139
2021-03-15 16:18 | DW.300.7 | Microcystins-ADDA-SAES | 2 | 1.287 | 1.3005 | 1.5 | 342.3 | 331.1 | 1.287 Abs [1.3005] {1.5 CV} | 342.3 ng/L [331.1] {4.8 CV} |  | RK2:28->D09@2 | T=188,A=+1.2871
2021-03-15 16:18 | DW.500.5 | Microcystins-ADDA-SAES | 1 | 1.142 | 1.128 | 1.8 | 491.9 | 510 | 1.142 Abs | 491.9 ng/L |  | RK2:29->E09@2 | T=187,A=+1.1419
2021-03-15 16:18 | DW.500.5 | Microcystins-ADDA-SAES | 2 | 1.114 | 1.128 | 1.8 | 528 | 510 | 1.114 Abs [1.1280] {1.8 CV} | 528.0 ng/L [510.0] {5.0 CV} |  | RK2:29->F09@3 | T=186,A=+1.1137
2021-03-15 16:18 | DW.500.6 | Microcystins-ADDA-SAES | 1 | 1.07 | 1.087 | 2.2 | 590.8 | 566.2 | 1.070 Abs | 590.8 ng/L |  | RK2:30->G09@3 | T=185,A=+1.0699
2021-03-15 16:18 | DW.500.6 | Microcystins-ADDA-SAES | 2 | 1.104 | 1.087 | 2.2 | 541.6 | 566.2 | 1.104 Abs [1.0870] {2.2 CV} | 541.6 ng/L [566.2] {6.1 CV} |  | RK2:30->H09@3 | T=184,A=+1.1044
2021-03-15 16:18 | DW.115.5 | Microcystins-ADDA-SAES | 1 | 1.632 | 1.668 | 3.1 | 139.1 | 125.9 | 1.632 Abs | 139.1 ng/L |  | RK2:31->A10@1 | T=184,A=+1.6315
2021-03-15 16:18 | DW.115.5 | Microcystins-ADDA-SAES | 2 | 1.704 | 1.668 | 3.1 | 112.7 | 125.9 | 1.704 Abs [1.6680] {3.1 CV} | 112.7 ng/L [125.9] {14.8 CV} |  | RK2:31->B10@1 | T=183,A=+1.7035
2021-03-15 16:18 | DW.115.6 | Microcystins-ADDA-SAES | 1 | 1.621 | 1.613 | 0.7 | 143.5 | 146.8 | 1.621 Abs | 143.5 ng/L |  | RK2:32->C10@1 | T=182,A=+1.6207
2021-03-15 16:18 | DW.115.6 | Microcystins-ADDA-SAES | 2 | 1.605 | 1.613 | 0.7 | 150.1 | 146.8 | 1.605 Abs [1.6130] {0.7 CV} | 150.1 ng/L [146.8] {3.2 CV} |  | RK2:32->D10@1 | T=181,A=+1.6045
2021-03-15 16:18 | DW.500.7 | Microcystins-ADDA-SAES | 1 | 1.114 | 1.1245 | 1.3 | 528 | 514.3 | 1.114 Abs | 528.0 ng/L |  | RK2:33->E10@1 | T=180,A=+1.1137
2021-03-15 16:18 | DW.500.7 | Microcystins-ADDA-SAES | 2 | 1.135 | 1.1245 | 1.3 | 500.7 | 514.3 | 1.135 Abs [1.1245] {1.3 CV} | 500.7 ng/L [514.3] {3.8 CV} |  | RK2:33->F10@4 | T=179,A=+1.1345
2021-03-15 16:18 | DW.115.7 | Microcystins-ADDA-SAES | 1 | 1.705 | 1.6585 | 4 | 112.4 | 129.8 | 1.705 Abs | 112.4 ng/L |  | RK2:34->G10@4 | T=178,A=+1.7046
2021-03-15 16:18 | DW.115.7 | Microcystins-ADDA-SAES | 2 | 1.612 | 1.6585 | 4 | 147.2 | 129.8 | 1.612 Abs [1.6585] {4.0 CV} | 147.2 ng/L [129.8] {19.0 CV} |  | RK2:34->H10@4 | T=177,A=+1.6116

## Fig1_Fig2_LIMS_Data
ANALYTE | Result | DL | RL | TL | HAL | oDL | oRL | oTL | oHAL | UNITS | SAMPDATE | CLIENT | Month | SAMPLENAME | LABSAMPID | RPTMATRIX | DOY | WeekOY | SAMPDATE | PREPDATE | ANADATE | BATCH | ANALYST
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 05:15 | Hiland WC - Shady Cove | May | PWS01520:SRC-AA | 2005029-01 | Source Water | 123 | 17 | 2020-05-04 05:15 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 05:30 | Anglers Cove/SCHWC | May | PWS01483:SRC-AB | 2005017-01 | Source Water | 123 | 17 | 2020-05-04 05:30 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 06:20 | City of Albany | May | PWS00012:SRC-BA | 2005020-01 | Source Water | 123 | 17 | 2020-05-04 06:20 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 08:00 | City of Pendleton | May | PWS00613:SRC-AM | 2005024-01 | Source Water | 123 | 17 | 2020-05-04 08:00 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 08:00 | Glide Water Association | May | PWS00326:SRC-AA | 2005028-01 | Source Water | 123 | 17 | 2020-05-04 08:00 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 08:03 | Roseburg Forest Products - Dillard | May | PWS94300:SRC-AA | 2005034-01 | Source Water | 123 | 17 | 2020-05-04 08:03 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 09:00 | Winston-Dillard Water District | May | PWS00957:SRC-AA | 2005039-01 | Source Water | 123 | 17 | 2020-05-04 09:00 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 09:05 | Camp Baker BSA | May | PWS91786:SRC-AA | 2005041-01 | Source Water | 123 | 17 | 2020-05-04 09:05 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 09:25 | Jackson Co Pks Emigrant Lake | May | PWS90730:CH-A | 2005030-01 | Source Water | 123 | 17 | 2020-05-04 09:25 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 09:30 | City of Powers | May | PWS00672:SRC-AA | 2005025-01 | Source Water | 123 | 17 | 2020-05-04 09:30 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 09:40 | Eugene Water & Electric Board | May | PWS00287:SRC-AA | 2005027-01 | Source Water | 123 | 17 | 2020-05-04 09:40 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 09:53 | USFS Tiller Ranger Station | May | PWS01092:SRC-AA | 2005038-01 | Source Water | 123 | 17 | 2020-05-04 09:53 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 10:10 | Springfield Utility Board | May | PWS00837:SRC-BB | 2005036-01 | Source Water | 123 | 17 | 2020-05-04 10:10 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 10:21 | Salem Public Works | May | PWS00731:SRC-AA | 2005035-01 | Source Water | 123 | 17 | 2020-05-04 10:21 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 10:23 | Lake Oswego - Tigard Water Supply | May | PWS00457:SRC-AA | 2005031-01 | Source Water | 123 | 17 | 2020-05-04 10:23 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 10:30 | Clarks Branch Water Association | May | PWS00548:SRC-AA | 2005040-01 | Source Water | 123 | 17 | 2020-05-04 10:30 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 10:45 | Springfield Utility Board | May | PWS00837:CH-A | 2005036-02 | Source Water | 123 | 17 | 2020-05-04 10:45 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 11:17 | North Clackamas County Water Commission | May | PWS00580:SRC-BA | 2005032-01 | Source Water | 123 | 17 | 2020-05-04 11:17 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 12:15 | City of Cottage Grove | May | PWS00236:SRC-BA | 2005021-01 | Source Water | 123 | 17 | 2020-05-04 12:15 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 12:15 | City of Wilsonville | May | PWS00954:SRC-HA | 2005026-01 | Source Water | 123 | 17 | 2020-05-04 12:15 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 12:59 | PP&L-Toketee Village | May | PWS01012:SRC-AA | 2005033-01 | Source Water | 123 | 17 | 2020-05-04 12:59 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 13:30 | Buell-Red Prairie Water Association | May | PWS01174:SRC-AA | 2005018-01 | Source Water | 123 | 17 | 2020-05-04 13:30 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 13:30 | City of Gates | May | PWS00317:SRC-AA | 2005022-01 | Source Water | 123 | 17 | 2020-05-04 13:30 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 13:40 | Canby Utility/Veolia Water | May | PWS00157:CH-A | 2005019-01 | Source Water | 123 | 17 | 2020-05-04 13:40 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 14:05 | USFS Steamboat Work Center | May | PWS01091:SRC-AA | 2005037-01 | Source Water | 123 | 17 | 2020-05-04 14:05 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-04 15:50 | City of Grants Pass | May | PWS00342:SRC-AA | 2005023-01 | Source Water | 123 | 17 | 2020-05-04 15:50 | 2020-05-05 11:35 | 2020-05-06 13:30 | B20E022 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 08:00 | City of Creswell | May | PWS00246:SRC-BK | 2005066-01 | Source Water | 130 | 18 | 2020-05-11 08:00 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 08:35 | City of Jefferson | May | PWS00408:SRC-AA | 2005070-01 | Source Water | 130 | 18 | 2020-05-11 08:35 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 08:46 | City of Roseburg | May | PWS00720:SRC-AA | 2005075-01 | Source Water | 130 | 18 | 2020-05-11 08:46 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 09:00 | Shangri La Water District | May | PWS00835:SRC-AA | 2005084-01 | Source Water | 130 | 18 | 2020-05-11 09:00 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 09:04 | City of Gold Hill | May | PWS00333:SRC-AA | 2005069-01 | Source Water | 130 | 18 | 2020-05-11 09:04 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 09:30 | Josephine Co Pks Lake Selmac 1 | May | PWS90186:SRC-AA | 2005077-01 | Source Water | 130 | 18 | 2020-05-11 09:30 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 09:30 | Umpqua Ranch Co-op | May | PWS00714:SRC-AH | 2005089-01 | Source Water | 130 | 18 | 2020-05-11 09:30 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 09:50 | Lyons Mehama Water District | May | PWS00493:SRC-AA | 2005079-01 | Source Water | 130 | 18 | 2020-05-11 09:50 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 09:55 | Umpqua Basin Water Association | May | PWS00719:SRC-AA | 2005088-01 | Source Water | 130 | 18 | 2020-05-11 09:55 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 10:00 | Josephine Co Pks Lake Selmac 2 | May | PWS94645:SRC-AA | 2005078-01 | Source Water | 130 | 18 | 2020-05-11 10:00 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 10:01 | Medford Water Commission | May | PWS00513:SRC-BA | 2005080-01 | Source Water | 130 | 18 | 2020-05-11 10:01 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 10:10 | City of Estacada | May | PWS00279:SRC-AA | 2005068-01 | Source Water | 130 | 18 | 2020-05-11 10:10 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 10:20 | South Fork Water Board - Oregon City | May | PWS00591:SRC-AA | 2005085-01 | Source Water | 130 | 18 | 2020-05-11 10:20 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 10:30 | Roberts Creek Water District | May | PWS00717:SRC-AA | 2005082-01 | Source Water | 130 | 18 | 2020-05-11 10:30 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 10:40 | City of Rogue River | May | PWS00712:SRC-BA | 2005074-01 | Source Water | 130 | 18 | 2020-05-11 10:40 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 10:47 | City of Lowell | May | PWS00492:SRC-BA | 2005071-01 | Source Water | 130 | 18 | 2020-05-11 10:47 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 10:50 | Rainbow Water District | May | PWS00839:SRC-BB | 2005081-01 | Source Water | 130 | 18 | 2020-05-11 10:50 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 11:00 | City of Newport | May | PWS00566:CH-A | 2005073-01 | Source Water | 130 | 18 | 2020-05-11 11:00 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 11:20 | City of Myrtle Creek | May | PWS00550:SRC-AA | 2005072-01 | Source Water | 130 | 18 | 2020-05-11 11:20 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 11:30 | Stayton Water Supply | May | PWS00843:SRC-AA | 2005086-01 | Source Water | 130 | 18 | 2020-05-11 11:30 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 11:55 | Ashland Water Department | May | PWS00047:SRC-AA | 2005092-01 | Source Water | 130 | 18 | 2020-05-11 11:55 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 12:10 | Tri-City JW & SA | May | PWS00549:SRC-AA | 2005087-01 | Source Water | 130 | 18 | 2020-05-11 12:10 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 12:15 | Joint Water Commission | May | PWS00379:SRC-AA | 2005076-01 | Source Water | 130 | 18 | 2020-05-11 12:15 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 14:15 | City of Drain Shops | May | PWS00260:SRC-AA | 2005067-01 | Source Water | 130 | 18 | 2020-05-11 14:15 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 15:30 | Seaside Water Department | May | PWS00799:CH-A | 2005083-01 | Source Water | 130 | 18 | 2020-05-11 15:30 | 2020-05-12 12:15 | 2020-05-13 16:48 | B20E065 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-11 19:00 | City of Monroe | May | PWS00540:SRC-BA | 2005103-01 | Source Water | 130 | 18 | 2020-05-11 19:00 | 2020-05-13 12:10 | 2020-05-14 17:02 | B20E081 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 | North Clackamas County Water Commission | May | PWS00580:SRC-BA | 2005135-01 | Source Water | 137 | 19 | 2020-05-18 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 05:15 | Hiland WC - Shady Cove | May | PWS01520:SRC-AA | 2005132-01 | Source Water | 137 | 19 | 2020-05-18 05:15 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 05:30 | Anglers Cove/SCHWC | May | PWS01483:SRC-AB | 2005118-01 | Source Water | 137 | 19 | 2020-05-18 05:30 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 06:20 | USFS Tiller Ranger Station | May | PWS01092:SRC-AA | 2005141-01 | Source Water | 137 | 19 | 2020-05-18 06:20 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 07:10 | City of Albany | May | PWS00012:SRC-BA | 2005122-01 | Source Water | 137 | 19 | 2020-05-18 07:10 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 07:20 | Camp Baker BSA | May | PWS91786:SRC-AA | 2005120-01 | Source Water | 137 | 19 | 2020-05-18 07:20 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 07:22 | City of Wilsonville | May | PWS00954:SRC-HA | 2005128-01 | Source Water | 137 | 19 | 2020-05-18 07:22 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 07:45 | Glide Water Association | May | PWS00326:SRC-AA | 2005131-01 | Source Water | 137 | 19 | 2020-05-18 07:45 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 07:58 | Jackson Co Pks Emigrant Lake | May | PWS90730:CH-A | 2005133-01 | Source Water | 137 | 19 | 2020-05-18 07:58 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 08:09 | City of Pendleton | May | PWS00613:SRC-AM | 2005126-01 | Source Water | 137 | 19 | 2020-05-18 08:09 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 08:10 | Roseburg Forest Products - Dillard | May | PWS94300:SRC-AA | 2005137-01 | Source Water | 137 | 19 | 2020-05-18 08:10 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 08:20 | Clarks Branch Water Association | May | PWS00548:SRC-AA | 2005129-01 | Source Water | 137 | 19 | 2020-05-18 08:20 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 08:45 | Winston-Dillard Water District | May | PWS00957:SRC-AA | 2005142-01 | Source Water | 137 | 19 | 2020-05-18 08:45 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 09:25 | Lake Oswego - Tigard Water Supply | May | PWS00457:SRC-AA | 2005134-01 | Source Water | 137 | 19 | 2020-05-18 09:25 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 09:50 | City of Powers | May | PWS00672:SRC-AA | 2005127-01 | Source Water | 137 | 19 | 2020-05-18 09:50 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 09:50 | Springfield Utility Board | May | PWS00837:CH-A | 2005139-02 | Source Water | 137 | 19 | 2020-05-18 09:50 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 10:00 | Eugene Water & Electric Board | May | PWS00287:SRC-AA | 2005130-01 | Source Water | 137 | 19 | 2020-05-18 10:00 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 10:10 | Salem Public Works | May | PWS00731:SRC-AA | 2005138-01 | Source Water | 137 | 19 | 2020-05-18 10:10 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 10:30 | Buell-Red Prairie Water Association | May | PWS01174:SRC-AA | 2005119-01 | Source Water | 137 | 19 | 2020-05-18 10:30 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 10:30 | Springfield Utility Board | May | PWS00837:SRC-BB | 2005139-01 | Source Water | 137 | 19 | 2020-05-18 10:30 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 10:31 | USFS Steamboat Work Center | May | PWS01091:SRC-AA | 2005140-01 | Source Water | 137 | 19 | 2020-05-18 10:31 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 10:55 | PP&L-Toketee Village | May | PWS01012:SRC-AA | 2005136-01 | Source Water | 137 | 19 | 2020-05-18 10:55 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 12:35 | City of Cottage Grove | May | PWS00236:SRC-BA | 2005123-01 | Source Water | 137 | 19 | 2020-05-18 12:35 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 13:30 | City of Gates | May | PWS00317:SRC-AA | 2005124-01 | Source Water | 137 | 19 | 2020-05-18 13:30 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 13:40 | Canby Utility/Veolia Water | May | PWS00157:CH-A | 2005121-01 | Source Water | 137 | 19 | 2020-05-18 13:40 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-18 14:40 | City of Grants Pass | May | PWS00342:SRC-AA | 2005125-01 | Source Water | 137 | 19 | 2020-05-18 14:40 | 2020-05-19 11:45 | 2020-05-20 15:51 | B20E105 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 08:30 | City of Jefferson | May | PWS00408:SRC-AA | 2005183-01 | Source Water | 145 | 20 | 2020-05-26 08:30 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 08:32 | Joint Water Commission | May | PWS00379:SRC-AA | 2005189-01 | Source Water | 145 | 20 | 2020-05-26 08:32 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 09:00 | Shangri La Water District | May | PWS00835:SRC-AA | 2005197-01 | Source Water | 145 | 20 | 2020-05-26 09:00 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 09:00 | Umpqua Ranch Co-op | May | PWS00714:SRC-AH | 2005202-01 | Source Water | 145 | 20 | 2020-05-26 09:00 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 10:00 | Josephine Co Pks Lake Selmac 2 | May | PWS94645:SRC-AA | 2005191-01 | Source Water | 145 | 20 | 2020-05-26 10:00 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 10:10 | South Fork Water Board - Oregon City | May | PWS00591:SRC-AA | 2005198-01 | Source Water | 145 | 20 | 2020-05-26 10:10 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 10:20 | City of Estacada | May | PWS00279:SRC-AA | 2005181-01 | Source Water | 145 | 20 | 2020-05-26 10:20 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 10:30 | Josephine Co Pks Lake Selmac 1 | May | PWS90186:SRC-AA | 2005190-01 | Source Water | 145 | 20 | 2020-05-26 10:30 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 10:30 | Lyons Mehama Water District | May | PWS00493:SRC-AA | 2005192-01 | Source Water | 145 | 20 | 2020-05-26 10:30 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 10:40 | City of Myrtle Creek | May | PWS00550:SRC-AA | 2005185-01RE1 | Source Water | 145 | 20 | 2020-05-26 10:40 | 2020-05-27 12:30 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 10:40 | Umpqua Basin Water Association | May | PWS00719:SRC-AA | 2005201-01 | Source Water | 145 | 20 | 2020-05-26 10:40 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 11:09 | City of Newport | May | PWS00566:CH-A | 2005186-01 | Source Water | 145 | 20 | 2020-05-26 11:09 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 11:15 | City of Lowell | May | PWS00492:SRC-BA | 2005184-01 | Source Water | 145 | 20 | 2020-05-26 11:15 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 11:15 | City of Rogue River | May | PWS00712:SRC-BA | 2005187-01 | Source Water | 145 | 20 | 2020-05-26 11:15 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 11:30 | Roberts Creek Water District | May | PWS00717:SRC-AA | 2005195-01 | Source Water | 145 | 20 | 2020-05-26 11:30 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 11:30 | Stayton Water Supply | May | PWS00843:SRC-AA | 2005199-01 | Source Water | 145 | 20 | 2020-05-26 11:30 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 11:36 | City of Roseburg | May | PWS00720:SRC-AA | 2005188-01 | Source Water | 145 | 20 | 2020-05-26 11:36 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 12:15 | Ashland Water Department | May | PWS00047:SRC-AA | 2005203-01 | Source Water | 145 | 20 | 2020-05-26 12:15 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 13:10 | Tri-City JW & SA | May | PWS00549:SRC-AA | 2005200-01 | Source Water | 145 | 20 | 2020-05-26 13:10 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 13:30 | Seaside Water Department | May | PWS00799:CH-A | 2005196-01 | Source Water | 145 | 20 | 2020-05-26 13:30 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 14:10 | City of Drain Shops | May | PWS00260:SRC-AA | 2005204-01 | Source Water | 145 | 20 | 2020-05-26 14:10 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-26 15:47 | Rainbow Water District | May | PWS00839:SRC-BB | 2005194-01 | Source Water | 145 | 20 | 2020-05-26 15:47 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-27 10:17 | City of Gold Hill | May | PWS00333:SRC-AA | 2005182-01 | Source Water | 146 | 20 | 2020-05-27 10:17 | 2020-05-28 12:00 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-27 10:55 | Medford Water Commission | May | PWS00513:SRC-BA | 2005193-01 | Source Water | 146 | 20 | 2020-05-27 10:55 | 2020-05-27 12:30 | 2020-05-28 13:20 | B20E168 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-28 09:45 | City of Creswell | May | PWS00246:SRC-BK | 2006053-01 | Source Water | 147 | 21 | 2020-05-28 09:45 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-05-29 19:00 | City of Monroe | May | PWS00540:SRC-BA | 2006054-01 | Source Water | 148 | 21 | 2020-05-29 19:00 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 05:15 | Hiland WC - Shady Cove | June | PWS01520:SRC-AA | 2006030-01 | Source Water | 150 | 21 | 2020-06-01 05:15 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 05:30 | Anglers Cove/SCHWC | June | PWS01483:SRC-AB | 2006001-01 | Source Water | 150 | 21 | 2020-06-01 05:30 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 06:30 | City of Albany | June | PWS00012:SRC-BA | 2006009-01 | Source Water | 150 | 21 | 2020-06-01 06:30 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 07:15 | City of Pendleton | June | PWS00613:SRC-AM | 2006017-01 | Source Water | 150 | 21 | 2020-06-01 07:15 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 07:30 | Camp Baker BSA | June | PWS91786:SRC-AA | 2006005-01 | Source Water | 150 | 21 | 2020-06-01 07:30 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 08:00 | Glide Water Association | June | PWS00326:SRC-AA | 2006028-01 | Source Water | 150 | 21 | 2020-06-01 08:00 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 08:15 | Roseburg Forest Products - Dillard | June | PWS94300:SRC-AA | 2006040-01 | Source Water | 150 | 21 | 2020-06-01 08:15 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 08:16 | City of Wilsonville | June | PWS00954:SRC-HA | 2006021-01 | Source Water | 150 | 21 | 2020-06-01 08:16 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 08:23 | USFS Tiller Ranger Station | June | PWS01092:SRC-AA | 2006048-01 | Source Water | 150 | 21 | 2020-06-01 08:23 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 08:30 | Jackson Co Pks Emigrant Lake | June | PWS90730:CH-A | 2006032-01 | Source Water | 150 | 21 | 2020-06-01 08:30 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 09:20 | Salem Public Works | June | PWS00731:SRC-AA | 2006042-01 | Source Water | 150 | 21 | 2020-06-01 09:20 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 09:20 | Springfield Utility Board | June | PWS00837:CH-A | 2006044-02 | Source Water | 150 | 21 | 2020-06-01 09:20 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 09:43 | Clarks Branch Water Association | June | PWS00548:SRC-AA | 2006024-01 | Source Water | 150 | 21 | 2020-06-01 09:43 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 09:45 | Winston-Dillard Water District | June | PWS00957:SRC-AA | 2006050-01 | Source Water | 150 | 21 | 2020-06-01 09:45 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 10:00 | City of Powers | June | PWS00672:SRC-AA | 2006019-01 | Source Water | 150 | 21 | 2020-06-01 10:00 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 10:45 | North Clackamas County Water Commission | June | PWS00580:SRC-BA | 2006036-01 | Source Water | 150 | 21 | 2020-06-01 10:45 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 10:50 | Lake Oswego - Tigard Water Supply | June | PWS00457:SRC-AA | 2006034-01 | Source Water | 150 | 21 | 2020-06-01 10:50 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 11:00 | Eugene Water & Electric Board | June | PWS00287:SRC-AA | 2006026-01 | Source Water | 150 | 21 | 2020-06-01 11:00 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 11:11 | PP&L-Toketee Village | June | PWS01012:SRC-AA | 2006038-01 | Source Water | 150 | 21 | 2020-06-01 11:11 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 11:50 | Springfield Utility Board | June | PWS00837:SRC-BB | 2006044-01 | Source Water | 150 | 21 | 2020-06-01 11:50 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 12:22 | USFS Steamboat Work Center | June | PWS01091:SRC-AA | 2006046-01 | Source Water | 150 | 21 | 2020-06-01 12:22 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 12:30 | Buell-Red Prairie Water Association | June | PWS01174:SRC-AA | 2006003-01 | Source Water | 150 | 21 | 2020-06-01 12:30 | 2020-06-03 11:00 | 2020-06-08 16:30 | B20F053 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 13:30 | City of Gates | June | PWS00317:SRC-AA | 2006013-01 | Source Water | 150 | 21 | 2020-06-01 13:30 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 13:45 | Canby Utility/Veolia Water | June | PWS00157:CH-A | 2006007-01 | Source Water | 150 | 21 | 2020-06-01 13:45 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 14:10 | City of Cottage Grove | June | PWS00236:SRC-BA | 2006011-01 | Source Water | 150 | 21 | 2020-06-01 14:10 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-01 15:45 | City of Grants Pass | June | PWS00342:SRC-AA | 2006015-01 | Source Water | 150 | 21 | 2020-06-01 15:45 | 2020-06-02 13:20 | 2020-06-04 10:47 | B20E173 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 06:20 | Joint Water Commission | June | PWS00379:SRC-AA | 2006104-01 | Source Water | 157 | 22 | 2020-06-08 06:20 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 07:30 | City of Jefferson | June | PWS00408:SRC-AA | 2006092-01 | Source Water | 157 | 22 | 2020-06-08 07:30 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 08:20 | City of Creswell | June | PWS00246:SRC-BK | 2006086-01 | Source Water | 157 | 22 | 2020-06-08 08:20 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 09:15 | Susan Creek Mobile Home Park | June | PWS94508:SRC-AA | 2006142-01 | Source Water | 157 | 22 | 2020-06-08 09:15 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 09:30 | Josephine Co Pks Lake Selmac 1 | June | PWS90186:SRC-AA | 2006106-01 | Source Water | 157 | 22 | 2020-06-08 09:30 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 09:30 | Shangri La Water District | June | PWS00835:SRC-AA | 2006120-01 | Source Water | 157 | 22 | 2020-06-08 09:30 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 09:30 | Tri-City JW & SA | June | PWS00549:SRC-AA | 2006126-01 | Source Water | 157 | 22 | 2020-06-08 09:30 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 09:40 | Josephine Co Pks Lake Selmac 2 | June | PWS94645:SRC-AA | 2006108-01 | Source Water | 157 | 22 | 2020-06-08 09:40 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 09:45 | Umpqua Basin Water Association | June | PWS00719:SRC-AA | 2006128-01 | Source Water | 157 | 22 | 2020-06-08 09:45 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 09:55 | South Fork Water Board - Oregon City | June | PWS00591:SRC-AA | 2006122-01 | Source Water | 157 | 22 | 2020-06-08 09:55 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 10:04 | City of Roseburg | June | PWS00720:SRC-AA | 2006102-01 | Source Water | 157 | 22 | 2020-06-08 10:04 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 10:15 | Medford Water Commission | June | PWS00513:SRC-BA | 2006112-01 | Source Water | 157 | 22 | 2020-06-08 10:15 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 10:45 | City of Estacada | June | PWS00279:SRC-AA | 2006088-01 | Source Water | 157 | 22 | 2020-06-08 10:45 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 10:59 | Rainbow Water District | June | PWS00839:SRC-BB | 2006114-01 | Source Water | 157 | 22 | 2020-06-08 10:59 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 11:00 | City of Rogue River | June | PWS00712:SRC-BA | 2006100-01 | Source Water | 157 | 22 | 2020-06-08 11:00 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 11:00 | Stayton Water Supply | June | PWS00843:SRC-AA | 2006124-01 | Source Water | 157 | 22 | 2020-06-08 11:00 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 11:15 | City of Myrtle Creek | June | PWS00550:SRC-AA | 2006096-01 | Source Water | 157 | 22 | 2020-06-08 11:15 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 11:20 | City of Lowell | June | PWS00492:SRC-BA | 2006094-01 | Source Water | 157 | 22 | 2020-06-08 11:20 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 11:45 | Lyons Mehama Water District | June | PWS00493:SRC-AA | 2006110-01 | Source Water | 157 | 22 | 2020-06-08 11:45 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 11:45 | Roberts Creek Water District | June | PWS00717:SRC-AA | 2006116-01 | Source Water | 157 | 22 | 2020-06-08 11:45 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 12:00 | Ashland Water Department | June | PWS00047:SRC-AA | 2006132-01 | Source Water | 157 | 22 | 2020-06-08 12:00 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 13:00 | Umpqua Ranch Co-op | June | PWS00714:SRC-AH | 2006130-01 | Source Water | 157 | 22 | 2020-06-08 13:00 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 16:31 | City of Gold Hill | June | PWS00333:SRC-AA | 2006090-01 | Source Water | 157 | 22 | 2020-06-08 16:31 | 2020-06-09 13:15 | 2020-06-10 16:21 | B20F075 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-08 20:30 | City of Monroe | June | PWS00540:SRC-BA | 2006135-01 | Source Water | 157 | 22 | 2020-06-08 20:30 | 2020-06-10 12:45 | 2020-06-15 12:57 | B20F078 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-09 11:30 | City of Drain Shops | June | PWS00260:SRC-AA | 2006133-01 | Source Water | 158 | 22 | 2020-06-09 11:30 | 2020-06-10 12:45 | 2020-06-15 12:57 | B20F078 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-09 11:57 | Seaside Water Department | June | PWS00799:CH-A | 2006118-01RE1 | Source Water | 158 | 22 | 2020-06-09 11:57 | 2020-06-10 12:45 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-09 13:25 | City of Newport | June | PWS00566:CH-A | 2006098-01 | Source Water | 158 | 22 | 2020-06-09 13:25 | 2020-06-10 12:45 | 2020-06-15 12:57 | B20F078 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 05:15 | Hiland WC - Shady Cove | June | PWS01520:SRC-AA | 2006203-01 | Source Water | 164 | 23 | 2020-06-15 05:15 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 05:30 | Anglers Cove/SCHWC | June | PWS01483:SRC-AB | 2006174-01 | Source Water | 164 | 23 | 2020-06-15 05:30 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 06:39 | City of Albany | June | PWS00012:SRC-BA | 2006182-01 | Source Water | 164 | 23 | 2020-06-15 06:39 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 07:07 | City of Wilsonville | June | PWS00954:SRC-HA | 2006194-01 | Source Water | 164 | 23 | 2020-06-15 07:07 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 08:15 | City of Pendleton | June | PWS00613:SRC-AM | 2006190-01 | Source Water | 164 | 23 | 2020-06-15 08:15 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 08:15 | Glide Water Association | June | PWS00326:SRC-AA | 2006201-01 | Source Water | 164 | 23 | 2020-06-15 08:15 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 08:20 | Roseburg Forest Products - Dillard | June | PWS94300:SRC-AA | 2006213-01 | Source Water | 164 | 23 | 2020-06-15 08:20 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 08:35 | Camp Baker BSA | June | PWS91786:SRC-AA | 2006178-01 | Source Water | 164 | 23 | 2020-06-15 08:35 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 08:35 | Salem Public Works | June | PWS00731:SRC-AA | 2006215-01 | Source Water | 164 | 23 | 2020-06-15 08:35 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 09:29 | USFS Tiller Ranger Station | June | PWS01092:SRC-AA | 2006221-01 | Source Water | 164 | 23 | 2020-06-15 09:29 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 09:30 | Lake Oswego - Tigard Water Supply | June | PWS00457:SRC-AA | 2006207-01 | Source Water | 164 | 23 | 2020-06-15 09:30 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 10:00 | Springfield Utility Board | June | PWS00837:CH-A | 2006217-02 | Source Water | 164 | 23 | 2020-06-15 10:00 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 10:15 | North Clackamas County Water Commission | June | PWS00580:SRC-BA | 2006209-01 | Source Water | 164 | 23 | 2020-06-15 10:15 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 10:20 | City of Powers | June | PWS00672:SRC-AA | 2006192-01 | Source Water | 164 | 23 | 2020-06-15 10:20 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 10:20 | Winston-Dillard Water District | June | PWS00957:SRC-AA | 2006223-01 | Source Water | 164 | 23 | 2020-06-15 10:20 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 10:30 | Springfield Utility Board | June | PWS00837:SRC-BB | 2006217-01 | Source Water | 164 | 23 | 2020-06-15 10:30 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 10:35 | Clarks Branch Water Association | June | PWS00548:SRC-AA | 2006197-01 | Source Water | 164 | 23 | 2020-06-15 10:35 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 10:59 | Eugene Water & Electric Board | June | PWS00287:SRC-AA | 2006199-01 | Source Water | 164 | 23 | 2020-06-15 10:59 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 11:30 | Buell-Red Prairie Water Association | June | PWS01174:SRC-AA | 2006176-01 | Source Water | 164 | 23 | 2020-06-15 11:30 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 12:05 | USFS Steamboat Work Center | June | PWS01091:SRC-AA | 2006219-01 | Source Water | 164 | 23 | 2020-06-15 12:05 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 12:51 | PP&L-Toketee Village | June | PWS01012:SRC-AA | 2006211-01 | Source Water | 164 | 23 | 2020-06-15 12:51 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 13:30 | City of Gates | June | PWS00317:SRC-AA | 2006186-01 | Source Water | 164 | 23 | 2020-06-15 13:30 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 14:45 | City of Cottage Grove | June | PWS00236:SRC-BA | 2006184-01 | Source Water | 164 | 23 | 2020-06-15 14:45 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-15 16:10 | City of Grants Pass | June | PWS00342:SRC-AA | 2006188-01 | Source Water | 164 | 23 | 2020-06-15 16:10 | 2020-06-16 12:30 | 2020-06-17 16:17 | B20F142 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-16 13:05 | Canby Utility/Veolia Water | June | PWS00157:CH-A | 2006180-01 | Source Water | 165 | 23 | 2020-06-16 13:05 | 2020-06-17 15:00 | 2020-06-22 15:17 | B20F169 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-17 12:06 | Jackson Co Pks Emigrant Lake | June | PWS90730:CH-A | 2006205-01 | Source Water | 166 | 23 | 2020-06-17 12:06 | 2020-06-18 12:45 | 2020-06-22 15:17 | B20F169 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 07:55 | City of Jefferson | June | PWS00408:SRC-AA | 2006262-01 | Source Water | 171 | 24 | 2020-06-22 07:55 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 08:24 | Joint Water Commission | June | PWS00379:SRC-AA | 2006274-01 | Source Water | 171 | 24 | 2020-06-22 08:24 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 08:45 | City of Newport | June | PWS00566:CH-A | 2006268-01 | Source Water | 171 | 24 | 2020-06-22 08:45 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 09:00 | Susan Creek Mobile Home Park | June | PWS94508:SRC-AA | 2006320-01 | Source Water | 171 | 24 | 2020-06-22 09:00 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 09:30 | Medford Water Commission | June | PWS00513:SRC-BA | 2006282-01 | Source Water | 171 | 24 | 2020-06-22 09:30 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 09:30 | Rainbow Water District | June | PWS00839:SRC-BB | 2006284-01 | Source Water | 171 | 24 | 2020-06-22 09:30 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 09:30 | Shangri La Water District | June | PWS00835:SRC-AA | 2006290-01 | Source Water | 171 | 24 | 2020-06-22 09:30 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 09:30 | Tri-City JW & SA | June | PWS00549:SRC-AA | 2006296-01 | Source Water | 171 | 24 | 2020-06-22 09:30 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 09:45 | City of Creswell | June | PWS00246:SRC-BK | 2006256-01 | Source Water | 171 | 24 | 2020-06-22 09:45 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 09:45 | City of Estacada | June | PWS00279:SRC-AA | 2006258-01 | Source Water | 171 | 24 | 2020-06-22 09:45 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 09:50 | South Fork Water Board - Oregon City | June | PWS00591:SRC-AA | 2006292-01 | Source Water | 171 | 24 | 2020-06-22 09:50 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 09:55 | Umpqua Basin Water Association | June | PWS00719:SRC-AA | 2006298-01 | Source Water | 171 | 24 | 2020-06-22 09:55 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 10:12 | City of Roseburg | June | PWS00720:SRC-AA | 2006272-01 | Source Water | 171 | 24 | 2020-06-22 10:12 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 10:40 | City of Myrtle Creek | June | PWS00550:SRC-AA | 2006266-01 | Source Water | 171 | 24 | 2020-06-22 10:40 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 10:45 | Josephine Co Pks Lake Selmac 1 | June | PWS90186:SRC-AA | 2006276-01 | Source Water | 171 | 24 | 2020-06-22 10:45 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 11:00 | Roberts Creek Water District | June | PWS00717:SRC-AA | 2006286-01 | Source Water | 171 | 24 | 2020-06-22 11:00 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 11:00 | Stayton Water Supply | June | PWS00843:SRC-AA | 2006294-01 | Source Water | 171 | 24 | 2020-06-22 11:00 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 11:00 | Umpqua Ranch Co-op | June | PWS00714:SRC-AH | 2006300-01 | Source Water | 171 | 24 | 2020-06-22 11:00 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 11:15 | City of Rogue River | June | PWS00712:SRC-BA | 2006270-01 | Source Water | 171 | 24 | 2020-06-22 11:15 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 11:30 | City of Drain Shops | June | PWS00260:SRC-AA | 2006303-01 | Source Water | 171 | 24 | 2020-06-22 11:30 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 11:45 | City of Lowell | June | PWS00492:SRC-BA | 2006264-01 | Source Water | 171 | 24 | 2020-06-22 11:45 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 12:45 | Ashland Water Department | June | PWS00047:SRC-AA | 2006302-01 | Source Water | 171 | 24 | 2020-06-22 12:45 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 14:00 | Seaside Water Department | June | PWS00799:CH-A | 2006288-01 | Source Water | 171 | 24 | 2020-06-22 14:00 | 2020-06-23 13:15 | 2020-06-24 16:35 | B20F217 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-22 21:05 | City of Monroe | June | PWS00540:SRC-BA | 2006305-01 | Source Water | 171 | 24 | 2020-06-22 21:05 | 2020-06-24 14:30 | 2020-06-29 16:17 | B20F221 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-23 09:30 | Lyons Mehama Water District | June | PWS00493:SRC-AA | 2006280-01 | Source Water | 172 | 24 | 2020-06-23 09:30 | 2020-06-24 14:30 | 2020-06-29 16:17 | B20F221 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-24 09:45 | Josephine Co Pks Lake Selmac 2 | June | PWS94645:SRC-AA | 2006278-01 | Source Water | 173 | 24 | 2020-06-24 09:45 | 2020-06-25 13:45 | 2020-06-29 16:17 | B20F221 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-24 10:55 | City of Gold Hill | June | PWS00333:SRC-AA | 2006260-01 | Source Water | 173 | 24 | 2020-06-24 10:55 | 2020-06-25 13:45 | 2020-06-29 16:17 | B20F221 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 05:15 | Hiland WC - Shady Cove | June | PWS01520:SRC-AA | 2006368-01 | Source Water | 178 | 25 | 2020-06-29 05:15 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 05:30 | Anglers Cove/SCHWC | June | PWS01483:SRC-AB | 2006339-01 | Source Water | 178 | 25 | 2020-06-29 05:30 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 06:30 | USFS Tiller Ranger Station | June | PWS01092:SRC-AA | 2006386-01RE2 | Source Water | 178 | 25 | 2020-06-29 06:30 | 2020-06-30 12:30 | 2020-07-07 17:56 | B20G023 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 07:20 | Camp Baker BSA | June | PWS91786:SRC-AA | 2006343-01 | Source Water | 178 | 25 | 2020-06-29 07:20 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 07:22 | Clarks Branch Water Association | June | PWS00548:SRC-AA | 2006362-01 | Source Water | 178 | 25 | 2020-06-29 07:22 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 08:10 | City of Pendleton | June | PWS00613:SRC-AM | 2006355-01 | Source Water | 178 | 25 | 2020-06-29 08:10 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 08:12 | Roseburg Forest Products - Dillard | June | PWS94300:SRC-AA | 2006378-01 | Source Water | 178 | 25 | 2020-06-29 08:12 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 08:24 | USFS Steamboat Work Center | June | PWS01091:SRC-AA | 2006384-01 | Source Water | 178 | 25 | 2020-06-29 08:24 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 08:27 | City of Wilsonville | June | PWS00954:SRC-HA | 2006359-01 | Source Water | 178 | 25 | 2020-06-29 08:27 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 08:30 | Glide Water Association | June | PWS00326:SRC-AA | 2006366-01 | Source Water | 178 | 25 | 2020-06-29 08:30 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 08:45 | Salem Public Works | June | PWS00731:SRC-AA | 2006380-01 | Source Water | 178 | 25 | 2020-06-29 08:45 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 09:05 | Winston-Dillard Water District | June | PWS00957:SRC-AA | 2006388-01 | Source Water | 178 | 25 | 2020-06-29 09:05 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 09:20 | PP&L-Toketee Village | June | PWS01012:SRC-AA | 2006376-01 | Source Water | 178 | 25 | 2020-06-29 09:20 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 09:30 | Lake Oswego - Tigard Water Supply | June | PWS00457:SRC-AA | 2006372-01 | Source Water | 178 | 25 | 2020-06-29 09:30 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 10:00 | Eugene Water & Electric Board | June | PWS00287:SRC-AA | 2006364-01 | Source Water | 178 | 25 | 2020-06-29 10:00 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 10:12 | City of Grants Pass | June | PWS00342:SRC-AA | 2006353-01 | Source Water | 178 | 25 | 2020-06-29 10:12 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 10:15 | City of Powers | June | PWS00672:SRC-AA | 2006357-01 | Source Water | 178 | 25 | 2020-06-29 10:15 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 10:30 | Springfield Utility Board | June | PWS00837:CH-A | 2006382-02 | Source Water | 178 | 25 | 2020-06-29 10:30 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 10:45 | North Clackamas County Water Commission | June | PWS00580:SRC-BA | 2006374-01 | Source Water | 178 | 25 | 2020-06-29 10:45 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 11:15 | Springfield Utility Board | June | PWS00837:SRC-BB | 2006382-01 | Source Water | 178 | 25 | 2020-06-29 11:15 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 12:00 | Buell-Red Prairie Water Association | June | PWS01174:SRC-AA | 2006341-01 | Source Water | 178 | 25 | 2020-06-29 12:00 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 12:17 | Jackson Co Pks Emigrant Lake | June | PWS90730:CH-A | 2006370-01 | Source Water | 178 | 25 | 2020-06-29 12:17 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 12:35 | City of Cottage Grove | June | PWS00236:SRC-BA | 2006349-01 | Source Water | 178 | 25 | 2020-06-29 12:35 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 13:25 | Canby Utility/Veolia Water | June | PWS00157:CH-A | 2006345-01 | Source Water | 178 | 25 | 2020-06-29 13:25 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-29 13:30 | City of Gates | June | PWS00317:SRC-AA | 2006351-01 | Source Water | 178 | 25 | 2020-06-29 13:30 | 2020-06-30 12:30 | 2020-07-01 15:09 | B20F287 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-06-30 08:30 | City of Albany | June | PWS00012:SRC-BA | 2006347-01 | Source Water | 179 | 25 | 2020-06-30 08:30 | 2020-07-01 11:30 | 2020-07-02 13:03 | B20G012 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 07:05 | City of Myrtle Creek | July | PWS00550:SRC-AA | 2007038-01 | Source Water | 185 | 26 | 2020-07-06 07:05 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 08:24 | Joint Water Commission | July | PWS00379:SRC-AA | 2007046-01 | Source Water | 185 | 26 | 2020-07-06 08:24 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 08:30 | City of Jefferson | July | PWS00408:SRC-AA | 2007034-01 | Source Water | 185 | 26 | 2020-07-06 08:30 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 08:30 | City of Newport | July | PWS00566:CH-A | 2007040-01 | Source Water | 185 | 26 | 2020-07-06 08:30 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 08:30 | Shangri La Water District | July | PWS00835:SRC-AA | 2007062-01 | Source Water | 185 | 26 | 2020-07-06 08:30 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 09:00 | City of Creswell | July | PWS00246:SRC-BK | 2007028-01 | Source Water | 185 | 26 | 2020-07-06 09:00 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 09:30 | Stayton Water Supply | July | PWS00843:SRC-AA | 2007066-01 | Source Water | 185 | 26 | 2020-07-06 09:30 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 09:45 | Umpqua Ranch Co-op | July | PWS00714:SRC-AH | 2007072-01 | Source Water | 185 | 26 | 2020-07-06 09:45 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 09:52 | City of Roseburg | July | PWS00720:SRC-AA | 2007044-01 | Source Water | 185 | 26 | 2020-07-06 09:52 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 10:00 | City of Estacada | July | PWS00279:SRC-AA | 2007030-01 | Source Water | 185 | 26 | 2020-07-06 10:00 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 10:00 | Lyons Mehama Water District | July | PWS00493:SRC-AA | 2007052-01 | Source Water | 185 | 26 | 2020-07-06 10:00 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 10:00 | Susan Creek Mobile Home Park | July | PWS94508:SRC-AA | 2007079-01 | Source Water | 185 | 26 | 2020-07-06 10:00 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 10:15 | Josephine Co Pks Lake Selmac 1 | July | PWS90186:SRC-AA | 2007048-01 | Source Water | 185 | 26 | 2020-07-06 10:15 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 10:25 | South Fork Water Board - Oregon City | July | PWS00591:SRC-AA | 2007064-01 | Source Water | 185 | 26 | 2020-07-06 10:25 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 10:30 | Medford Water Commission | July | PWS00513:SRC-BA | 2007054-01 | Source Water | 185 | 26 | 2020-07-06 10:30 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 10:35 | Josephine Co Pks Lake Selmac 2 | July | PWS94645:SRC-AA | 2007050-01 | Source Water | 185 | 26 | 2020-07-06 10:35 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 11:00 | Tri-City JW & SA | July | PWS00549:SRC-AA | 2007068-01 | Source Water | 185 | 26 | 2020-07-06 11:00 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 11:05 | City of Rogue River | July | PWS00712:SRC-BA | 2007042-01 | Source Water | 185 | 26 | 2020-07-06 11:05 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 11:10 | Umpqua Basin Water Association | July | PWS00719:SRC-AA | 2007070-01 | Source Water | 185 | 26 | 2020-07-06 11:10 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 11:20 | City of Lowell | July | PWS00492:SRC-BA | 2007036-01 | Source Water | 185 | 26 | 2020-07-06 11:20 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 11:30 | Roberts Creek Water District | July | PWS00717:SRC-AA | 2007058-01 | Source Water | 185 | 26 | 2020-07-06 11:30 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 12:30 | Ashland Water Department | July | PWS00047:SRC-AA | 2007074-01 | Source Water | 185 | 26 | 2020-07-06 12:30 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 12:40 | Rainbow Water District | July | PWS00839:SRC-BB | 2007056-01 | Source Water | 185 | 26 | 2020-07-06 12:40 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 14:30 | Seaside Water Department | July | PWS00799:CH-A | 2007060-01 | Source Water | 185 | 26 | 2020-07-06 14:30 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-06 15:00 | City of Gold Hill | July | PWS00333:SRC-AA | 2007032-01 | Source Water | 185 | 26 | 2020-07-06 15:00 | 2020-07-07 13:00 | 2020-07-08 13:28 | B20G039 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-07 11:30 | City of Drain Shops | July | PWS00260:SRC-AA | 2007075-01 | Source Water | 186 | 26 | 2020-07-07 11:30 | 2020-07-08 12:00 | 2020-07-10 12:33 | B20G050 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-07 12:55 | City of Monroe | July | PWS00540:SRC-BA | 2007077-01 | Source Water | 186 | 26 | 2020-07-07 12:55 | 2020-07-08 12:00 | 2020-07-10 12:33 | B20G050 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 05:15 | Hiland WC - Shady Cove | July | PWS01520:SRC-AA | 2007134-01 | Source Water | 192 | 27 | 2020-07-13 05:15 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 05:30 | Anglers Cove/SCHWC | July | PWS01483:SRC-AB | 2007105-01 | Source Water | 192 | 27 | 2020-07-13 05:30 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 06:25 | City of Albany | July | PWS00012:SRC-BA | 2007113-01 | Source Water | 192 | 27 | 2020-07-13 06:25 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 06:45 | Camp Baker BSA | July | PWS91786:SRC-AA | 2007109-01 | Source Water | 192 | 27 | 2020-07-13 06:45 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 07:30 | Glide Water Association | July | PWS00326:SRC-AA | 2007132-01 | Source Water | 192 | 27 | 2020-07-13 07:30 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 08:10 | Roseburg Forest Products - Dillard | July | PWS94300:SRC-AA | 2007144-01 | Source Water | 192 | 27 | 2020-07-13 08:10 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.07 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-07-13 08:45 | Salem Public Works | July | PWS00731:SRC-AA | 2007146-01 | Source Water | 192 | 27 | 2020-07-13 08:45 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 09:05 | USFS Tiller Ranger Station | July | PWS01092:SRC-AA | 2007152-01 | Source Water | 192 | 27 | 2020-07-13 09:05 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 09:10 | City of Wilsonville | July | PWS00954:SRC-HA | 2007125-01 | Source Water | 192 | 27 | 2020-07-13 09:10 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 09:14 | City of Powers | July | PWS00672:SRC-AA | 2007123-01 | Source Water | 192 | 27 | 2020-07-13 09:14 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 09:20 | Springfield Utility Board | July | PWS00837:CH-A | 2007148-02 | Source Water | 192 | 27 | 2020-07-13 09:20 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 09:20 | Winston-Dillard Water District | July | PWS00957:SRC-AA | 2007154-01 | Source Water | 192 | 27 | 2020-07-13 09:20 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 09:58 | North Clackamas County Water Commission | July | PWS00580:SRC-BA | 2007140-01 | Source Water | 192 | 27 | 2020-07-13 09:58 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 10:00 | Jackson Co Pks Emigrant Lake | July | PWS90730:CH-A | 2007136-01 | Source Water | 192 | 27 | 2020-07-13 10:00 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 10:15 | Lake Oswego - Tigard Water Supply | July | PWS00457:SRC-AA | 2007138-01 | Source Water | 192 | 27 | 2020-07-13 10:15 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 10:19 | Clarks Branch Water Association | July | PWS00548:SRC-AA | 2007128-01 | Source Water | 192 | 27 | 2020-07-13 10:19 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 10:30 | Buell-Red Prairie Water Association | July | PWS01174:SRC-AA | 2007107-01 | Source Water | 192 | 27 | 2020-07-13 10:30 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 11:26 | USFS Steamboat Work Center | July | PWS01091:SRC-AA | 2007150-01 | Source Water | 192 | 27 | 2020-07-13 11:26 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 12:00 | PP&L-Toketee Village | July | PWS01012:SRC-AA | 2007142-01 | Source Water | 192 | 27 | 2020-07-13 12:00 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 12:45 | Springfield Utility Board | July | PWS00837:SRC-BB | 2007148-01 | Source Water | 192 | 27 | 2020-07-13 12:45 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 13:00 | City of Cottage Grove | July | PWS00236:SRC-BA | 2007115-01 | Source Water | 192 | 27 | 2020-07-13 13:00 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 13:00 | City of Pendleton | July | PWS00613:SRC-AM | 2007121-01 | Source Water | 192 | 27 | 2020-07-13 13:00 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 13:02 | Eugene Water & Electric Board | July | PWS00287:SRC-AA | 2007130-01 | Source Water | 192 | 27 | 2020-07-13 13:02 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.12 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2020-07-13 13:30 | City of Gates | July | PWS00317:SRC-AA | 2007117-01 | Source Water | 192 | 27 | 2020-07-13 13:30 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-13 14:40 | Canby Utility/Veolia Water | July | PWS00157:CH-A | 2007111-01 | Source Water | 192 | 27 | 2020-07-13 14:40 | 2020-07-14 14:30 | 2020-07-15 16:45 | B20G094 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-14 10:02 | City of Grants Pass | July | PWS00342:SRC-AA | 2007119-01 | Source Water | 193 | 27 | 2020-07-14 10:02 | 2020-07-15 12:00 | 2020-07-17 13:26 | B20G107 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 08:00 | City of Jefferson | July | PWS00408:SRC-AA | 2007202-01 | Source Water | 199 | 28 | 2020-07-20 08:00 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 08:30 | Joint Water Commission | July | PWS00379:SRC-AA | 2007214-01 | Source Water | 199 | 28 | 2020-07-20 08:30 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 09:05 | Josephine Co Pks Lake Selmac 1 | July | PWS90186:SRC-AA | 2007216-01 | Source Water | 199 | 28 | 2020-07-20 09:05 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 09:15 | City of Creswell | July | PWS00246:SRC-BK | 2007196-01 | Source Water | 199 | 28 | 2020-07-20 09:15 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 09:27 | Rainbow Water District | July | PWS00839:SRC-BB | 2007224-01 | Source Water | 199 | 28 | 2020-07-20 09:27 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 09:30 | Josephine Co Pks Lake Selmac 2 | July | PWS94645:SRC-AA | 2007218-01 | Source Water | 199 | 28 | 2020-07-20 09:30 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 09:30 | Shangri La Water District | July | PWS00835:SRC-AA | 2007230-01 | Source Water | 199 | 28 | 2020-07-20 09:30 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 09:30 | Susan Creek Mobile Home Park | July | PWS94508:SRC-AA | 2007250-01 | Source Water | 199 | 28 | 2020-07-20 09:30 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 09:30 | Tri-City JW & SA | July | PWS00549:SRC-AA | 2007236-01 | Source Water | 199 | 28 | 2020-07-20 09:30 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 09:30 | Umpqua Ranch Co-op | July | PWS00714:SRC-AH | 2007240-01 | Source Water | 199 | 28 | 2020-07-20 09:30 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 09:40 | City of Newport | July | PWS00566:CH-A | 2007208-01 | Source Water | 199 | 28 | 2020-07-20 09:40 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 10:20 | City of Myrtle Creek | July | PWS00550:SRC-AA | 2007206-01 | Source Water | 199 | 28 | 2020-07-20 10:20 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 10:20 | South Fork Water Board - Oregon City | July | PWS00591:SRC-AA | 2007232-01 | Source Water | 199 | 28 | 2020-07-20 10:20 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 10:30 | City of Estacada | July | PWS00279:SRC-AA | 2007198-01 | Source Water | 199 | 28 | 2020-07-20 10:30 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 10:30 | Douglas County Parks - Whistler's Bend | July | PWS93944:SRC-AB | 2007249-01 | Source Water | 199 | 28 | 2020-07-20 10:30 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 10:35 | Medford Water Commission | July | PWS00513:SRC-BA | 2007222-01 | Source Water | 199 | 28 | 2020-07-20 10:35 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 11:00 | City of Rogue River | July | PWS00712:SRC-BA | 2007210-01 | Source Water | 199 | 28 | 2020-07-20 11:00 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 11:00 | Roberts Creek Water District | July | PWS00717:SRC-AA | 2007226-01 | Source Water | 199 | 28 | 2020-07-20 11:00 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.18 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2020-07-20 11:00 | Stayton Water Supply | July | PWS00843:SRC-AA | 2007234-01 | Source Water | 199 | 28 | 2020-07-20 11:00 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 11:40 | City of Roseburg | July | PWS00720:SRC-AA | 2007212-01 | Source Water | 199 | 28 | 2020-07-20 11:40 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 11:45 | Umpqua Basin Water Association | July | PWS00719:SRC-AA | 2007238-01 | Source Water | 199 | 28 | 2020-07-20 11:45 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 11:55 | Ashland Water Department | July | PWS00047:SRC-AA | 2007242-01 | Source Water | 199 | 28 | 2020-07-20 11:55 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 12:00 | City of Lowell | July | PWS00492:SRC-BA | 2007204-01 | Source Water | 199 | 28 | 2020-07-20 12:00 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 13:00 | Seaside Water Department | July | PWS00799:CH-A | 2007228-01 | Source Water | 199 | 28 | 2020-07-20 13:00 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 14:24 | City of Gold Hill | July | PWS00333:SRC-AA | 2007200-01 | Source Water | 199 | 28 | 2020-07-20 14:24 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-20 14:30 | Lyons Mehama Water District | July | PWS00493:SRC-AA | 2007220-01 | Source Water | 199 | 28 | 2020-07-20 14:30 | 2020-07-21 13:15 | 2020-07-22 16:03 | B20G157 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-21 09:00 | City of Monroe | July | PWS00540:SRC-BA | 2007245-01 | Source Water | 200 | 28 | 2020-07-21 09:00 | 2020-07-22 11:40 | 2020-07-23 15:21 | B20G171 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-21 11:30 | City of Drain Shops | July | PWS00260:SRC-AA | 2007243-01 | Source Water | 200 | 28 | 2020-07-21 11:30 | 2020-07-22 11:40 | 2020-07-23 15:21 | B20G171 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 05:45 | City of Albany | July | PWS00012:SRC-BA | 2007303-01 | Source Water | 206 | 29 | 2020-07-27 05:45 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.1 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-07-27 06:20 | Salem Public Works | July | PWS00731:SRC-AA | 2007336-01 | Source Water | 206 | 29 | 2020-07-27 06:20 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 06:50 | Camp Baker BSA | July | PWS91786:SRC-AA | 2007299-01 | Source Water | 206 | 29 | 2020-07-27 06:50 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 07:08 | City of Pendleton | July | PWS00613:SRC-AM | 2007311-01 | Source Water | 206 | 29 | 2020-07-27 07:08 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 07:45 | USFS Tiller Ranger Station | July | PWS01092:SRC-AA | 2007342-01 | Source Water | 206 | 29 | 2020-07-27 07:45 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 08:00 | Glide Water Association | July | PWS00326:SRC-AA | 2007322-01 | Source Water | 206 | 29 | 2020-07-27 08:00 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 08:42 | City of Wilsonville | July | PWS00954:SRC-HA | 2007315-01 | Source Water | 206 | 29 | 2020-07-27 08:42 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 09:00 | Jackson Co Pks Emigrant Lake | July | PWS90730:CH-A | 2007326-01 | Source Water | 206 | 29 | 2020-07-27 09:00 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 09:00 | Winston-Dillard Water District | July | PWS00957:SRC-AA | 2007344-01 | Source Water | 206 | 29 | 2020-07-27 09:00 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 09:15 | City of Powers | July | PWS00672:SRC-AA | 2007313-01 | Source Water | 206 | 29 | 2020-07-27 09:15 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 09:15 | North Clackamas County Water Commission | July | PWS00580:SRC-BA | 2007330-01 | Source Water | 206 | 29 | 2020-07-27 09:15 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 09:50 | Springfield Utility Board | July | PWS00837:CH-A | 2007338-02 | Source Water | 206 | 29 | 2020-07-27 09:50 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 10:17 | USFS Steamboat Work Center | July | PWS01091:SRC-AA | 2007340-01 | Source Water | 206 | 29 | 2020-07-27 10:17 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 10:18 | Lake Oswego - Tigard Water Supply | July | PWS00457:SRC-AA | 2007328-01 | Source Water | 206 | 29 | 2020-07-27 10:18 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 10:44 | PP&L-Toketee Village | July | PWS01012:SRC-AA | 2007332-01 | Source Water | 206 | 29 | 2020-07-27 10:44 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 10:50 | Eugene Water & Electric Board | July | PWS00287:SRC-AA | 2007320-01 | Source Water | 206 | 29 | 2020-07-27 10:50 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 11:30 | Buell-Red Prairie Water Association | July | PWS01174:SRC-AA | 2007297-01 | Source Water | 206 | 29 | 2020-07-27 11:30 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 11:30 | Springfield Utility Board | July | PWS00837:SRC-BB | 2007338-01 | Source Water | 206 | 29 | 2020-07-27 11:30 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 12:20 | Clarks Branch Water Association | July | PWS00548:SRC-AA | 2007318-01 | Source Water | 206 | 29 | 2020-07-27 12:20 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 13:10 | City of Grants Pass | July | PWS00342:SRC-AA | 2007309-01 | Source Water | 206 | 29 | 2020-07-27 13:10 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-27 13:30 | Canby Utility/Veolia Water | July | PWS00157:CH-A | 2007301-01 | Source Water | 206 | 29 | 2020-07-27 13:30 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.16 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2020-07-27 13:30 | City of Gates | July | PWS00317:SRC-AA | 2007307-01 | Source Water | 206 | 29 | 2020-07-27 13:30 | 2020-07-28 12:00 | 2020-07-29 11:53 | B20G221 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-29 13:00 | City of Cottage Grove | July | PWS00236:SRC-BA | 2007366-01 | Source Water | 208 | 29 | 2020-07-29 13:00 | 2020-07-30 11:30 | 2020-07-31 15:45 | B20G236 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-30 14:00 | Hiland WC - Shady Cove | July | PWS01520:SRC-AA | 2008072-01 | Source Water | 209 | 29 | 2020-07-30 14:00 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-30 14:15 | Anglers Cove/SCHWC | July | PWS01483:SRC-AB | 2008068-01 | Source Water | 209 | 29 | 2020-07-30 14:15 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-07-31 12:11 | Roseburg Forest Products - Dillard | July | PWS94300:SRC-AA | 2008067-01 | Source Water | 210 | 30 | 2020-07-31 12:11 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 08:20 | City of Creswell | August | PWS00246:SRC-BK | 2008014-01 | Source Water | 212 | 30 | 2020-08-03 08:20 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 08:34 | Joint Water Commission | August | PWS00379:SRC-AA | 2008032-01 | Source Water | 212 | 30 | 2020-08-03 08:34 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 09:30 | City of Jefferson | August | PWS00408:SRC-AA | 2008020-01 | Source Water | 212 | 30 | 2020-08-03 09:30 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 09:30 | Shangri La Water District | August | PWS00835:SRC-AA | 2008048-01 | Source Water | 212 | 30 | 2020-08-03 09:30 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 09:30 | Tri-City JW & SA | August | PWS00549:SRC-AA | 2008054-01 | Source Water | 212 | 30 | 2020-08-03 09:30 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 09:42 | City of Newport | August | PWS00566:CH-A | 2008026-01 | Source Water | 212 | 30 | 2020-08-03 09:42 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 09:50 | South Fork Water Board - Oregon City | August | PWS00591:SRC-AA | 2008050-01 | Source Water | 212 | 30 | 2020-08-03 09:50 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 10:00 | Roberts Creek Water District | August | PWS00717:SRC-AA | 2008044-01 | Source Water | 212 | 30 | 2020-08-03 10:00 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 10:00 | Susan Creek Mobile Home Park | August | PWS94508:SRC-AA | 2008070-01 | Source Water | 212 | 30 | 2020-08-03 10:00 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 10:00 | Umpqua Ranch Co-op | August | PWS00714:SRC-AH | 2008058-01 | Source Water | 212 | 30 | 2020-08-03 10:00 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 10:30 | City of Estacada | August | PWS00279:SRC-AA | 2008016-01 | Source Water | 212 | 30 | 2020-08-03 10:30 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 10:35 | Josephine Co Pks Lake Selmac 1 | August | PWS90186:SRC-AA | 2008034-01 | Source Water | 212 | 30 | 2020-08-03 10:35 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 10:50 | City of Myrtle Creek | August | PWS00550:SRC-AA | 2008024-01 | Source Water | 212 | 30 | 2020-08-03 10:50 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 11:11 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2008036-01 | Source Water | 212 | 30 | 2020-08-03 11:11 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 11:12 | Medford Water Commission | August | PWS00513:SRC-BA | 2008040-01 | Source Water | 212 | 30 | 2020-08-03 11:12 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 11:25 | Rainbow Water District | August | PWS00839:SRC-BB | 2008042-01 | Source Water | 212 | 30 | 2020-08-03 11:25 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 11:30 | Douglas County Parks - Whistler's Bend | August | PWS93944:SRC-AB | 2008074-01 | Source Water | 212 | 30 | 2020-08-03 11:30 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 11:45 | City of Rogue River | August | PWS00712:SRC-BA | 2008028-01 | Source Water | 212 | 30 | 2020-08-03 11:45 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 12:15 | Ashland Water Department | August | PWS00047:CH-A | 2008060-01 | Source Water | 212 | 30 | 2020-08-03 12:15 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 12:36 | City of Roseburg | August | PWS00720:SRC-AA | 2008030-01 | Source Water | 212 | 30 | 2020-08-03 12:36 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 12:50 | Umpqua Basin Water Association | August | PWS00719:SRC-AA | 2008056-01 | Source Water | 212 | 30 | 2020-08-03 12:50 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 13:00 | City of Drain Shops | August | PWS00260:SRC-AA | 2008061-01 | Source Water | 212 | 30 | 2020-08-03 13:00 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 13:45 | Seaside Water Department | August | PWS00799:CH-A | 2008046-01 | Source Water | 212 | 30 | 2020-08-03 13:45 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 13:50 | City of Lowell | August | PWS00492:SRC-BA | 2008022-01 | Source Water | 212 | 30 | 2020-08-03 13:50 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-03 15:40 | City of Gold Hill | August | PWS00333:SRC-AA | 2008018-01 | Source Water | 212 | 30 | 2020-08-03 15:40 | 2020-08-04 13:45 | 2020-08-05 13:51 | B20H032 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-04 10:35 | City of Monroe | August | PWS00540:SRC-BA | 2008063-01 | Source Water | 213 | 30 | 2020-08-04 10:35 | 2020-08-05 11:15 | 2020-08-10 12:45 | B20H047 | KMC
Microcystins ADDA, Total | 0.15 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2020-08-04 14:30 | Stayton Water Supply | August | PWS00843:SRC-AA | 2008082-01 | Source Water | 213 | 30 | 2020-08-04 14:30 | 2020-08-05 11:15 | 2020-08-10 12:45 | B20H047 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-05 10:14 | Lyons Mehama Water District | August | PWS00493:SRC-AA | 2008087-01 | Source Water | 214 | 30 | 2020-08-05 10:14 | 2020-08-06 10:40 | 2020-08-10 12:45 | B20H047 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 05:15 | Hiland WC - Shady Cove | August | PWS01520:SRC-AA | 2008133-01 | Source Water | 219 | 31 | 2020-08-10 05:15 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 05:30 | Anglers Cove/SCHWC | August | PWS01483:SRC-AB | 2008104-01 | Source Water | 219 | 31 | 2020-08-10 05:30 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.09 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-08-10 06:24 | Salem Public Works | August | PWS00731:SRC-AA | 2008145-01 | Source Water | 219 | 31 | 2020-08-10 06:24 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 07:20 | Camp Baker BSA | August | PWS91786:SRC-AA | 2008108-01 | Source Water | 219 | 31 | 2020-08-10 07:20 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 07:35 | City of Albany | August | PWS00012:SRC-BA | 2008112-01 | Source Water | 219 | 31 | 2020-08-10 07:35 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 07:40 | City of Pendleton | August | PWS00613:SRC-AM | 2008120-01 | Source Water | 219 | 31 | 2020-08-10 07:40 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 08:00 | Glide Water Association | August | PWS00326:SRC-AA | 2008131-01 | Source Water | 219 | 31 | 2020-08-10 08:00 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 08:10 | Roseburg Forest Products - Dillard | August | PWS94300:SRC-AA | 2008143-01 | Source Water | 219 | 31 | 2020-08-10 08:10 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 08:12 | City of Wilsonville | August | PWS00954:SRC-HA | 2008124-01 | Source Water | 219 | 31 | 2020-08-10 08:12 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 09:00 | Eugene Water & Electric Board | August | PWS00287:SRC-AA | 2008129-01 | Source Water | 219 | 31 | 2020-08-10 09:00 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 09:00 | Winston-Dillard Water District | August | PWS00957:SRC-AA | 2008153-01 | Source Water | 219 | 31 | 2020-08-10 09:00 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 09:16 | Jackson Co Pks Emigrant Lake | August | PWS90730:CH-A | 2008135-01 | Source Water | 219 | 31 | 2020-08-10 09:16 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 09:40 | Springfield Utility Board | August | PWS00837:CH-A | 2008147-02 | Source Water | 219 | 31 | 2020-08-10 09:40 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 09:45 | Lake Oswego - Tigard Water Supply | August | PWS00457:SRC-AA | 2008137-01 | Source Water | 219 | 31 | 2020-08-10 09:45 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 10:15 | North Clackamas County Water Commission | August | PWS00580:SRC-BA | 2008139-01 | Source Water | 219 | 31 | 2020-08-10 10:15 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 11:11 | USFS Tiller Ranger Station | August | PWS01092:SRC-AA | 2008151-01 | Source Water | 219 | 31 | 2020-08-10 11:11 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 12:00 | Springfield Utility Board | August | PWS00837:SRC-BB | 2008147-01 | Source Water | 219 | 31 | 2020-08-10 12:00 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 12:22 | Clarks Branch Water Association | August | PWS00548:SRC-AA | 2008127-01 | Source Water | 219 | 31 | 2020-08-10 12:22 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.07 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-08-10 12:30 | Buell-Red Prairie Water Association | August | PWS01174:SRC-AA | 2008106-01 | Source Water | 219 | 31 | 2020-08-10 12:30 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 13:27 | USFS Steamboat Work Center | August | PWS01091:SRC-AA | 2008149-01 | Source Water | 219 | 31 | 2020-08-10 13:27 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 13:30 | City of Cottage Grove | August | PWS00236:SRC-BA | 2008114-01 | Source Water | 219 | 31 | 2020-08-10 13:30 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 13:35 | Canby Utility/Veolia Water | August | PWS00157:CH-A | 2008110-01 | Source Water | 219 | 31 | 2020-08-10 13:35 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.15 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2020-08-10 13:35 | City of Gates | August | PWS00317:SRC-AA | 2008116-01 | Source Water | 219 | 31 | 2020-08-10 13:35 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-10 14:10 | PP&L-Toketee Village | August | PWS01012:SRC-AA | 2008141-01 | Source Water | 219 | 31 | 2020-08-10 14:10 | 2020-08-11 12:00 | 2020-08-12 12:14 | B20H117 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-11 09:00 | City of Powers | August | PWS00672:SRC-AA | 2008122-01 | Source Water | 220 | 31 | 2020-08-11 09:00 | 2020-08-12 16:00 | 2020-08-13 17:20 | B20H144 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-11 12:05 | City of Grants Pass | August | PWS00342:SRC-AA | 2008165-01 | Source Water | 220 | 31 | 2020-08-11 12:05 | 2020-08-12 16:00 | 2020-08-13 17:20 | B20H144 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 | City of Creswell | August | PWS00246:SRC-BK | 2008182-01 | Source Water | 226 | 32 | 2020-08-17 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 | Lyons Mehama Water District | August | PWS00493:SRC-AA | 2008207-01 | Source Water | 226 | 32 | 2020-08-17 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 07:00 | Susan Creek Mobile Home Park | August | PWS94508:SRC-AA | 2008223-01 | Source Water | 226 | 32 | 2020-08-17 07:00 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 08:47 | Medford Water Commission | August | PWS00513:SRC-BA | 2008209-01 | Source Water | 226 | 32 | 2020-08-17 08:47 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 09:00 | City of Jefferson | August | PWS00408:SRC-AA | 2008188-01 | Source Water | 226 | 32 | 2020-08-17 09:00 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 09:00 | Joint Water Commission | August | PWS00379:SRC-AA | 2008201-01 | Source Water | 226 | 32 | 2020-08-17 09:00 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 09:00 | Shangri La Water District | August | PWS00835:SRC-AA | 2008217-01 | Source Water | 226 | 32 | 2020-08-17 09:00 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 09:30 | Tri-City JW & SA | August | PWS00549:SRC-AA | 2008224-01 | Source Water | 226 | 32 | 2020-08-17 09:30 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 09:35 | City of Gold Hill | August | PWS00333:SRC-AA | 2008186-01 | Source Water | 226 | 32 | 2020-08-17 09:35 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 09:40 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2008205-01 | Source Water | 226 | 32 | 2020-08-17 09:40 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 09:40 | South Fork Water Board - Oregon City | August | PWS00591:SRC-AA | 2008219-01 | Source Water | 226 | 32 | 2020-08-17 09:40 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 09:45 | Umpqua Ranch Co-op | August | PWS00714:SRC-AH | 2008228-01 | Source Water | 226 | 32 | 2020-08-17 09:45 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 09:50 | City of Newport | August | PWS00566:CH-A | 2008194-01 | Source Water | 226 | 32 | 2020-08-17 09:50 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 10:15 | Josephine Co Pks Lake Selmac 1 | August | PWS90186:SRC-AA | 2008203-01 | Source Water | 226 | 32 | 2020-08-17 10:15 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 10:35 | City of Myrtle Creek | August | PWS00550:SRC-AA | 2008192-01 | Source Water | 226 | 32 | 2020-08-17 10:35 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 10:45 | City of Estacada | August | PWS00279:SRC-AA | 2008184-01 | Source Water | 226 | 32 | 2020-08-17 10:45 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 11:00 | Stayton Water Supply | August | PWS00843:SRC-AA | 2008221-01 | Source Water | 226 | 32 | 2020-08-17 11:00 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 11:10 | City of Rogue River | August | PWS00712:SRC-BA | 2008196-01 | Source Water | 226 | 32 | 2020-08-17 11:10 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 11:20 | Umpqua Basin Water Association | August | PWS00719:SRC-AA | 2008226-01 | Source Water | 226 | 32 | 2020-08-17 11:20 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 11:30 | Roberts Creek Water District | August | PWS00717:SRC-AA | 2008213-01 | Source Water | 226 | 32 | 2020-08-17 11:30 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 12:00 | City of Drain Shops | August | PWS00260:SRC-AA | 2008231-01 | Source Water | 226 | 32 | 2020-08-17 12:00 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 12:30 | Ashland Water Department | August | PWS00047:CH-A | 2008230-01 | Source Water | 226 | 32 | 2020-08-17 12:30 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 12:53 | City of Roseburg | August | PWS00720:SRC-AA | 2008198-01 | Source Water | 226 | 32 | 2020-08-17 12:53 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 13:00 | Seaside Water Department | August | PWS00799:CH-A | 2008215-01 | Source Water | 226 | 32 | 2020-08-17 13:00 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 13:30 | Douglas County Parks - Whistler's Bend | August | PWS93944:SRC-AB | 2008200-01 | Source Water | 226 | 32 | 2020-08-17 13:30 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 13:46 | Rainbow Water District | August | PWS00839:SRC-BB | 2008211-01 | Source Water | 226 | 32 | 2020-08-17 13:46 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-17 14:20 | City of Lowell | August | PWS00492:SRC-BA | 2008190-01 | Source Water | 226 | 32 | 2020-08-17 14:20 | 2020-08-18 12:45 | 2020-08-20 10:35 | B20H176 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-18 12:50 | City of Monroe | August | PWS00540:SRC-BA | 2008233-01 | Source Water | 227 | 32 | 2020-08-18 12:50 | 2020-08-19 17:12 | 2020-08-20 15:03 | B20H193 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 05:15 | Hiland WC - Shady Cove | August | PWS01520:SRC-AA | 2008294-01 | Source Water | 233 | 33 | 2020-08-24 05:15 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 05:30 | Anglers Cove/SCHWC | August | PWS01483:SRC-AB | 2008272-01 | Source Water | 233 | 33 | 2020-08-24 05:30 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 06:30 | City of Albany | August | PWS00012:SRC-BA | 2008278-01 | Source Water | 233 | 33 | 2020-08-24 06:30 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 06:30 | Salem Public Works | August | PWS00731:SRC-AA | 2008303-01 | Source Water | 233 | 33 | 2020-08-24 06:30 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 07:15 | Camp Baker BSA | August | PWS91786:SRC-AA | 2008276-01 | Source Water | 233 | 33 | 2020-08-24 07:15 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 07:26 | City of Pendleton | August | PWS00613:SRC-AM | 2008285-01 | Source Water | 233 | 33 | 2020-08-24 07:26 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 08:10 | Glide Water Association | August | PWS00326:SRC-AA | 2008292-01 | Source Water | 233 | 33 | 2020-08-24 08:10 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 08:10 | Roseburg Forest Products - Dillard | August | PWS94300:SRC-AA | 2008302-01 | Source Water | 233 | 33 | 2020-08-24 08:10 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 08:45 | North Clackamas County Water Commission | August | PWS00580:SRC-BA | 2008299-01 | Source Water | 233 | 33 | 2020-08-24 08:45 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 09:00 | City of Powers | August | PWS00672:SRC-AA | 2008286-01 | Source Water | 233 | 33 | 2020-08-24 09:00 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 09:09 | City of Wilsonville | August | PWS00954:SRC-HA | 2008287-01 | Source Water | 233 | 33 | 2020-08-24 09:09 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 09:11 | Jackson Co Pks Emigrant Lake | August | PWS90730:CH-A | 2008296-01 | Source Water | 233 | 33 | 2020-08-24 09:11 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 09:44 | USFS Tiller Ranger Station | August | PWS01092:SRC-AA | 2008309-01 | Source Water | 233 | 33 | 2020-08-24 09:44 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 09:45 | Eugene Water & Electric Board | August | PWS00287:SRC-AA | 2008290-01 | Source Water | 233 | 33 | 2020-08-24 09:45 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 09:45 | Winston-Dillard Water District | August | PWS00957:SRC-AA | 2008310-01 | Source Water | 233 | 33 | 2020-08-24 09:45 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 10:00 | Lake Oswego - Tigard Water Supply | August | PWS00457:SRC-AA | 2008298-01 | Source Water | 233 | 33 | 2020-08-24 10:00 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 10:15 | Springfield Utility Board | August | PWS00837:CH-A | 2008305-02 | Source Water | 233 | 33 | 2020-08-24 10:15 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 10:47 | Clarks Branch Water Association | August | PWS00548:SRC-AA | 2008289-01 | Source Water | 233 | 33 | 2020-08-24 10:47 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 11:36 | USFS Steamboat Work Center | August | PWS01091:SRC-AA | 2008307-01 | Source Water | 233 | 33 | 2020-08-24 11:36 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 12:00 | Springfield Utility Board | August | PWS00837:SRC-BB | 2008305-01 | Source Water | 233 | 33 | 2020-08-24 12:00 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.06 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 12:30 | Buell-Red Prairie Water Association | August | PWS01174:SRC-AA | 2008274-01 | Source Water | 233 | 33 | 2020-08-24 12:30 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 12:52 | PP&L-Toketee Village | August | PWS01012:SRC-AA | 2008300-01 | Source Water | 233 | 33 | 2020-08-24 12:52 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 13:30 | Canby Utility/Veolia Water | August | PWS00157:CH-A | 2008277-01 | Source Water | 233 | 33 | 2020-08-24 13:30 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 14:00 | City of Gates | August | PWS00317:SRC-AA | 2008281-01 | Source Water | 233 | 33 | 2020-08-24 14:00 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 14:15 | City of Cottage Grove | August | PWS00236:SRC-BA | 2008280-01 | Source Water | 233 | 33 | 2020-08-24 14:15 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-24 16:10 | City of Grants Pass | August | PWS00342:SRC-AA | 2008283-01 | Source Water | 233 | 33 | 2020-08-24 16:10 | 2020-08-25 12:40 | 2020-08-26 11:48 | B20H252 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 08:10 | City of Jefferson | August | PWS00408:SRC-AA | 2009006-01 | Source Water | 240 | 34 | 2020-08-31 08:10 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 08:24 | Joint Water Commission | August | PWS00379:SRC-AA | 2009018-01 | Source Water | 240 | 34 | 2020-08-31 08:24 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 09:00 | City of Creswell | August | PWS00246:SRC-BK | 2009002-01 | Source Water | 240 | 34 | 2020-08-31 09:00 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.07 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-08-31 09:15 | Josephine Co Pks Lake Selmac 1 | August | PWS90186:SRC-AA | 2009019-01 | Source Water | 240 | 34 | 2020-08-31 09:15 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 09:30 | Shangri La Water District | August | PWS00835:SRC-AA | 2009031-01 | Source Water | 240 | 34 | 2020-08-31 09:30 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 09:30 | Tri-City JW & SA | August | PWS00549:SRC-AA | 2009038-01 | Source Water | 240 | 34 | 2020-08-31 09:30 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 09:35 | City of Estacada | August | PWS00279:SRC-AA | 2009003-01 | Source Water | 240 | 34 | 2020-08-31 09:35 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 09:40 | South Fork Water Board - Oregon City | August | PWS00591:SRC-AA | 2009033-01 | Source Water | 240 | 34 | 2020-08-31 09:40 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 09:41 | Rainbow Water District | August | PWS00839:SRC-BB | 2009027-01 | Source Water | 240 | 34 | 2020-08-31 09:41 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 09:45 | City of Gold Hill | August | PWS00333:SRC-AA | 2009004-01 | Source Water | 240 | 34 | 2020-08-31 09:45 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 09:50 | City of Newport | August | PWS00566:CH-A | 2009011-01 | Source Water | 240 | 34 | 2020-08-31 09:50 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.14 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2020-08-31 10:00 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2009021-01 | Source Water | 240 | 34 | 2020-08-31 10:00 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 10:00 | Susan Creek Mobile Home Park | August | PWS94508:SRC-AA | 2009036-01 | Source Water | 240 | 34 | 2020-08-31 10:00 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 10:30 | City of Myrtle Creek | August | PWS00550:SRC-AA | 2009010-01 | Source Water | 240 | 34 | 2020-08-31 10:30 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 10:37 | Medford Water Commission | August | PWS00513:SRC-BA | 2009025-01 | Source Water | 240 | 34 | 2020-08-31 10:37 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 11:00 | City of Rogue River | August | PWS00712:SRC-BA | 2009012-01 | Source Water | 240 | 34 | 2020-08-31 11:00 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 11:00 | Stayton Water Supply | August | PWS00843:SRC-AA | 2009034-01 | Source Water | 240 | 34 | 2020-08-31 11:00 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 11:00 | Umpqua Ranch Co-op | August | PWS00714:SRC-AH | 2009041-01 | Source Water | 240 | 34 | 2020-08-31 11:00 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 11:02 | Umpqua Basin Water Association | August | PWS00719:SRC-AA | 2009039-01 | Source Water | 240 | 34 | 2020-08-31 11:02 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 11:30 | City of Lowell | August | PWS00492:SRC-BA | 2009008-01 | Source Water | 240 | 34 | 2020-08-31 11:30 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 11:30 | Roberts Creek Water District | August | PWS00717:SRC-AA | 2009029-01 | Source Water | 240 | 34 | 2020-08-31 11:30 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 12:17 | Lyons Mehama Water District | August | PWS00493:SRC-AA | 2009023-01 | Source Water | 240 | 34 | 2020-08-31 12:17 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 1.68 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2020-08-31 12:30 | Ashland Water Department | August | PWS00047:CH-A | 2009043-01 | Source Water | 240 | 34 | 2020-08-31 12:30 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 12:57 | City of Roseburg | August | PWS00720:SRC-AA | 2009014-01 | Source Water | 240 | 34 | 2020-08-31 12:57 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 13:00 | Seaside Water Department | August | PWS00799:CH-A | 2009030-01 | Source Water | 240 | 34 | 2020-08-31 13:00 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-08-31 13:45 | Douglas County Parks - Whistler's Bend | August | PWS93944:SRC-AB | 2009016-01 | Source Water | 240 | 34 | 2020-08-31 13:45 | 2020-09-01 12:45 | 2020-09-02 18:46 | B20I014 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-01 12:30 | City of Drain Shops | September | PWS00260:SRC-AA | 2009044-01 | Source Water | 240 | 34 | 2020-09-01 12:30 | 2020-09-02 11:00 | 2020-09-04 15:50 | B20I026 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-01 12:45 | City of Monroe | September | PWS00540:SRC-BA | 2009045-01 | Source Water | 240 | 34 | 2020-09-01 12:45 | 2020-09-02 11:00 | 2020-09-04 15:50 | B20I026 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 05:40 | Camp Baker BSA | September | PWS91786:SRC-AA | 2009097-01RE1 | Source Water | 247 | 35 | 2020-09-08 05:40 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 08:10 | USFS Tiller Ranger Station | September | PWS01092:SRC-AA | 2009092-01RE1 | Source Water | 247 | 35 | 2020-09-08 08:10 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 08:15 | City of Pendleton | September | PWS00613:SRC-AM | 2009078-01RE1 | Source Water | 247 | 35 | 2020-09-08 08:15 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 08:15 | Roseburg Forest Products - Dillard | September | PWS94300:SRC-AA | 2009088-01RE1 | Source Water | 247 | 35 | 2020-09-08 08:15 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 08:25 | Glide Water Association | September | PWS00326:SRC-AA | 2009082-01RE1 | Source Water | 247 | 35 | 2020-09-08 08:25 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 08:30 | City of Powers | September | PWS00672:SRC-AA | 2009079-01RE1 | Source Water | 247 | 35 | 2020-09-08 08:30 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 08:30 | Winston-Dillard Water District | September | PWS00957:SRC-AA | 2009093-01RE1 | Source Water | 247 | 35 | 2020-09-08 08:30 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 09:00 | Clarks Branch Water Association | September | PWS00548:SRC-AA | 2009102-01RE1 | Source Water | 247 | 35 | 2020-09-08 09:00 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 09:10 | Springfield Utility Board | September | PWS00837:CH-A | 2009090-02RE1 | Source Water | 247 | 35 | 2020-09-08 09:10 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 09:16 | North Clackamas County Water Commission | September | PWS00580:SRC-BA | 2009086-01RE1 | Source Water | 247 | 35 | 2020-09-08 09:16 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 09:30 | Buell-Red Prairie Water Association | September | PWS01174:SRC-AA | 2009072-01RE1 | Source Water | 247 | 35 | 2020-09-08 09:30 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 09:45 | Salem Public Works | September | PWS00731:SRC-AA | 2009089-01RE1 | Source Water | 247 | 35 | 2020-09-08 09:45 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 09:47 | City of Wilsonville | September | PWS00954:SRC-HA | 2009080-01RE1 | Source Water | 247 | 35 | 2020-09-08 09:47 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 10:00 | Eugene Water & Electric Board | September | PWS00287:SRC-AA | 2009081-01RE1 | Source Water | 247 | 35 | 2020-09-08 10:00 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 10:30 | Lake Oswego - Tigard Water Supply | September | PWS00457:SRC-AA | 2009085-01RE1 | Source Water | 247 | 35 | 2020-09-08 10:30 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 10:35 | Springfield Utility Board | September | PWS00837:SRC-BB | 2009090-01RE1 | Source Water | 247 | 35 | 2020-09-08 10:35 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 11:18 | PP&L-Toketee Village | September | PWS01012:SRC-AA | 2009087-01RE1 | Source Water | 247 | 35 | 2020-09-08 11:18 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 11:56 | USFS Steamboat Work Center | September | PWS01091:SRC-AA | 2009091-01RE1 | Source Water | 247 | 35 | 2020-09-08 11:56 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 13:00 | City of Cottage Grove | September | PWS00236:SRC-BA | 2009075-01RE1 | Source Water | 247 | 35 | 2020-09-08 13:00 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 13:55 | Canby Utility/Veolia Water | September | PWS00157:CH-A | 2009073-01RE1 | Source Water | 247 | 35 | 2020-09-08 13:55 | 2020-09-09 12:00 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-08 15:15 | City of Grants Pass | September | PWS00342:SRC-AA | 2009077-01 | Source Water | 247 | 35 | 2020-09-08 15:15 | 2020-09-10 10:30 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-09 06:45 | City of Albany | September | PWS00012:SRC-BA | 2009074-01 | Source Water | 248 | 35 | 2020-09-09 06:45 | 2020-09-10 10:30 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0.07 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-09-09 08:01 | Jackson Co Pks Emigrant Lake | September | PWS90730:CH-A | 2009084-01 | Source Water | 248 | 35 | 2020-09-09 08:01 | 2020-09-10 10:30 | 2020-09-11 14:32 | B20I093 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-10 11:00 | Ashland Water Department | September | PWS00047:SRC-AA | 2009179-01 | Source Water | 249 | 35 | 2020-09-10 11:00 | 2020-09-16 12:00 | 2020-09-18 16:22 | B20I133 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 08:30 | City of Jefferson | September | PWS00408:SRC-AA | 2009139-01 | Source Water | 253 | 36 | 2020-09-14 08:30 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 08:40 | City of Creswell | September | PWS00246:SRC-BK | 2009135-01 | Source Water | 253 | 36 | 2020-09-14 08:40 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 09:25 | Joint Water Commission | September | PWS00379:SRC-AA | 2009151-01 | Source Water | 253 | 36 | 2020-09-14 09:25 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.4 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2020-09-14 09:30 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2009152-01 | Source Water | 253 | 36 | 2020-09-14 09:30 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 09:30 | Tri-City JW & SA | September | PWS00549:SRC-AA | 2009171-01 | Source Water | 253 | 36 | 2020-09-14 09:30 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.28 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2020-09-14 09:50 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2009154-01 | Source Water | 253 | 36 | 2020-09-14 09:50 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 09:50 | South Fork Water Board - Oregon City | September | PWS00591:SRC-AA | 2009166-01 | Source Water | 253 | 36 | 2020-09-14 09:50 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 10:25 | City of Newport | September | PWS00566:CH-A | 2009144-01 | Source Water | 253 | 36 | 2020-09-14 10:25 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 10:30 | Umpqua Basin Water Association | September | PWS00719:SRC-AA | 2009172-01 | Source Water | 253 | 36 | 2020-09-14 10:30 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 11:00 | Stayton Water Supply | September | PWS00843:SRC-AA | 2009167-01 | Source Water | 253 | 36 | 2020-09-14 11:00 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 11:30 | City of Myrtle Creek | September | PWS00550:SRC-AA | 2009143-01 | Source Water | 253 | 36 | 2020-09-14 11:30 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 11:30 | Roberts Creek Water District | September | PWS00717:SRC-AA | 2009162-01 | Source Water | 253 | 36 | 2020-09-14 11:30 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 12:00 | City of Drain Shops | September | PWS00260:SRC-AA | 2009177-01 | Source Water | 253 | 36 | 2020-09-14 12:00 | 2020-09-16 12:00 | 2020-09-18 16:22 | B20I133 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 12:00 | City of Rogue River | September | PWS00712:SRC-BA | 2009145-01 | Source Water | 253 | 36 | 2020-09-14 12:00 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 12:15 | Rainbow Water District | September | PWS00839:SRC-BB | 2009160-01 | Source Water | 253 | 36 | 2020-09-14 12:15 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 12:45 | City of Lowell | September | PWS00492:SRC-BA | 2009141-01 | Source Water | 253 | 36 | 2020-09-14 12:45 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 12:47 | City of Roseburg | September | PWS00720:SRC-AA | 2009147-01 | Source Water | 253 | 36 | 2020-09-14 12:47 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 13:00 | Seaside Water Department | September | PWS00799:CH-A | 2009163-01 | Source Water | 253 | 36 | 2020-09-14 13:00 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 13:30 | Douglas County Parks - Whistler's Bend | September | PWS93944:SRC-AB | 2009149-01 | Source Water | 253 | 36 | 2020-09-14 13:30 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-14 15:20 | City of Gold Hill | September | PWS00333:SRC-AA | 2009137-01 | Source Water | 253 | 36 | 2020-09-14 15:20 | 2020-09-15 15:40 | 2020-09-16 15:59 | B20I111 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-16 09:00 | Lyons Mehama Water District | September | PWS00493:SRC-AA | 2009200-01 | Source Water | 255 | 36 | 2020-09-16 09:00 | 2020-09-17 11:00 | 2020-09-18 16:22 | B20I133 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-17 12:25 | City of Monroe | September | PWS00540:SRC-BA | 2009210-01 | Source Water | 256 | 36 | 2020-09-17 12:25 | 2020-09-18 11:15 | 2020-09-18 16:22 | B20I133 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-18 14:20 | Ashland Water Department | September | PWS00047:SRC-AA | 2009257-01 | Source Water | 257 | 36 | 2020-09-18 14:20 | 2020-09-22 12:45 | 2020-09-22 20:04 | B20I158 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-18 16:00 | Medford Water Commission | September | PWS00513:SRC-BA | 2009266-01 | Source Water | 257 | 36 | 2020-09-18 16:00 | 2020-09-22 12:45 | 2020-09-22 20:04 | B20I158 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 05:15 | Hiland WC - Shady Cove | September | PWS01520:SRC-AA | 2009226-01 | Source Water | 260 | 37 | 2020-09-21 05:15 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 05:30 | Anglers Cove/SCHWC | September | PWS01483:SRC-AB | 2009237-01 | Source Water | 260 | 37 | 2020-09-21 05:30 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 06:50 | Camp Baker BSA | September | PWS91786:SRC-AA | 2009240-01 | Source Water | 260 | 37 | 2020-09-21 06:50 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 07:26 | City of Pendleton | September | PWS00613:SRC-AM | 2009221-01 | Source Water | 260 | 37 | 2020-09-21 07:26 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 07:40 | Glide Water Association | September | PWS00326:SRC-AA | 2009225-01 | Source Water | 260 | 37 | 2020-09-21 07:40 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 08:03 | Roseburg Forest Products - Dillard | September | PWS94300:SRC-AA | 2009231-01 | Source Water | 260 | 37 | 2020-09-21 08:03 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 08:20 | City of Powers | September | PWS00672:SRC-AA | 2009222-01 | Source Water | 260 | 37 | 2020-09-21 08:20 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 08:45 | City of Wilsonville | September | PWS00954:SRC-HA | 2009223-01 | Source Water | 260 | 37 | 2020-09-21 08:45 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 09:00 | USFS Tiller Ranger Station | September | PWS01092:SRC-AA | 2009235-01 | Source Water | 260 | 37 | 2020-09-21 09:00 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 09:10 | North Clackamas County Water Commission | September | PWS00580:SRC-BA | 2009229-01 | Source Water | 260 | 37 | 2020-09-21 09:10 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 09:15 | Winston-Dillard Water District | September | PWS00957:SRC-AA | 2009236-01 | Source Water | 260 | 37 | 2020-09-21 09:15 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.63 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2020-09-21 09:30 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2009260-01 | Source Water | 260 | 37 | 2020-09-21 09:30 | 2020-09-22 12:45 | 2020-09-22 20:04 | B20I158 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 09:40 | Springfield Utility Board | September | PWS00837:CH-A | 2009233-02 | Source Water | 260 | 37 | 2020-09-21 09:40 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.11 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-09-21 09:43 | Jackson Co Pks Emigrant Lake | September | PWS90730:CH-A | 2009227-01 | Source Water | 260 | 37 | 2020-09-21 09:43 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.4 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2020-09-21 10:00 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2009264-02 | Source Water | 260 | 37 | 2020-09-21 10:00 | 2020-09-22 12:45 | 2020-09-22 20:04 | B20I158 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 10:15 | Salem Public Works | September | PWS00731:SRC-AA | 2009232-01 | Source Water | 260 | 37 | 2020-09-21 10:15 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 10:20 | Eugene Water & Electric Board | September | PWS00287:SRC-AA | 2009224-01 | Source Water | 260 | 37 | 2020-09-21 10:20 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 10:29 | Clarks Branch Water Association | September | PWS00548:SRC-AA | 2009245-01 | Source Water | 260 | 37 | 2020-09-21 10:29 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 10:30 | Buell-Red Prairie Water Association | September | PWS01174:SRC-AA | 2009215-01 | Source Water | 260 | 37 | 2020-09-21 10:30 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 10:45 | Lake Oswego - Tigard Water Supply | September | PWS00457:SRC-AA | 2009228-01 | Source Water | 260 | 37 | 2020-09-21 10:45 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 12:05 | Ashland Water Department | September | PWS00047:SRC-AA | 2009255-01 | Source Water | 260 | 37 | 2020-09-21 12:05 | 2020-09-22 12:45 | 2020-09-22 20:04 | B20I158 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 12:30 | Springfield Utility Board | September | PWS00837:SRC-BB | 2009233-01 | Source Water | 260 | 37 | 2020-09-21 12:30 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 13:00 | City of Cottage Grove | September | PWS00236:SRC-BA | 2009218-01 | Source Water | 260 | 37 | 2020-09-21 13:00 | 2020-09-23 10:45 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 13:00 | Milo Academy | September | PWS00250:SRC-AB | 2009262-01 | Source Water | 260 | 37 | 2020-09-21 13:00 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 13:20 | PP&L-Toketee Village | September | PWS01012:SRC-AA | 2009230-01 | Source Water | 260 | 37 | 2020-09-21 13:20 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 13:55 | City of Gates | September | PWS00317:SRC-AA | 2009219-01 | Source Water | 260 | 37 | 2020-09-21 13:55 | 2020-09-23 10:45 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 14:20 | Canby Utility/Veolia Water | September | PWS00157:CH-A | 2009216-01 | Source Water | 260 | 37 | 2020-09-21 14:20 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-21 15:30 | City of Grants Pass | September | PWS00342:SRC-AA | 2009220-01 | Source Water | 260 | 37 | 2020-09-21 15:30 | 2020-09-22 13:20 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-22 09:00 | City of Albany | September | PWS00012:SRC-BA | 2009217-01 | Source Water | 261 | 37 | 2020-09-22 09:00 | 2020-09-23 10:45 | 2020-09-23 16:23 | B20I172 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 07:38 | Medford Water Commission | September | PWS00513:SRC-BA | 2009315-01 | Source Water | 267 | 38 | 2020-09-28 07:38 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 08:20 | City of Jefferson | September | PWS00408:SRC-AA | 2009296-01 | Source Water | 267 | 38 | 2020-09-28 08:20 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 08:33 | Joint Water Commission | September | PWS00379:SRC-AA | 2009308-01 | Source Water | 267 | 38 | 2020-09-28 08:33 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 08:48 | Lyons Mehama Water District | September | PWS00493:SRC-AA | 2009313-01 | Source Water | 267 | 38 | 2020-09-28 08:48 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 09:00 | Shangri La Water District | September | PWS00835:SRC-AA | 2009321-01 | Source Water | 267 | 38 | 2020-09-28 09:00 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 09:30 | Tri-City JW & SA | September | PWS00549:SRC-AA | 2009328-01 | Source Water | 267 | 38 | 2020-09-28 09:30 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.14 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2020-09-28 09:45 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2009309-01 | Source Water | 267 | 38 | 2020-09-28 09:45 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 10:00 | City of Newport | September | PWS00566:CH-A | 2009301-01 | Source Water | 267 | 38 | 2020-09-28 10:00 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 10:00 | South Fork Water Board - Oregon City | September | PWS00591:SRC-AA | 2009323-01 | Source Water | 267 | 38 | 2020-09-28 10:00 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.19 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2020-09-28 10:15 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2009311-01 | Source Water | 267 | 38 | 2020-09-28 10:15 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 10:30 | City of Myrtle Creek | September | PWS00550:SRC-AA | 2009300-01 | Source Water | 267 | 38 | 2020-09-28 10:30 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 10:30 | Roberts Creek Water District | September | PWS00717:SRC-AA | 2009319-01 | Source Water | 267 | 38 | 2020-09-28 10:30 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 10:30 | Umpqua Basin Water Association | September | PWS00719:SRC-AA | 2009329-01 | Source Water | 267 | 38 | 2020-09-28 10:30 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 10:35 | City of Estacada | September | PWS00279:SRC-AA | 2009293-01 | Source Water | 267 | 38 | 2020-09-28 10:35 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 10:45 | City of Lowell | September | PWS00492:SRC-BA | 2009298-01 | Source Water | 267 | 38 | 2020-09-28 10:45 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 11:00 | Stayton Water Supply | September | PWS00843:SRC-AA | 2009324-01 | Source Water | 267 | 38 | 2020-09-28 11:00 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 11:03 | Rainbow Water District | September | PWS00839:SRC-BB | 2009317-01 | Source Water | 267 | 38 | 2020-09-28 11:03 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 11:30 | Seaside Water Department | September | PWS00799:CH-A | 2009320-01 | Source Water | 267 | 38 | 2020-09-28 11:30 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 11:40 | City of Rogue River | September | PWS00712:SRC-BA | 2009302-01 | Source Water | 267 | 38 | 2020-09-28 11:40 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 12:00 | City of Drain Shops | September | PWS00260:SRC-AA | 2009334-01 | Source Water | 267 | 38 | 2020-09-28 12:00 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 12:27 | City of Roseburg | September | PWS00720:SRC-AA | 2009304-01 | Source Water | 267 | 38 | 2020-09-28 12:27 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 12:30 | Ashland Water Department | September | PWS00047:SRC-AA | 2009333-01 | Source Water | 267 | 38 | 2020-09-28 12:30 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-28 14:00 | Douglas County Parks - Whistler's Bend | September | PWS93944:SRC-AB | 2009306-01 | Source Water | 267 | 38 | 2020-09-28 14:00 | 2020-09-29 12:20 | 2020-09-30 14:16 | B20I207 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-29 07:30 | City of Gold Hill | September | PWS00333:SRC-AA | 2009294-01 | Source Water | 268 | 38 | 2020-09-29 07:30 | 2020-09-30 12:00 | 2020-10-02 16:02 | B20I216 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-09-29 12:45 | City of Monroe | September | PWS00540:SRC-BA | 2009335-01 | Source Water | 268 | 38 | 2020-09-29 12:45 | 2020-09-30 12:00 | 2020-10-02 16:02 | B20I216 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-01 13:15 | City of Creswell | October | PWS00246:SRC-BK | 2010021-01 | Source Water | 270 | 38 | 2020-10-01 13:15 | 2020-10-02 11:30 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 05:15 | Anglers Cove/SCHWC | October | PWS01483:SRC-AB | 2010049-01 | Source Water | 274 | 39 | 2020-10-05 05:15 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 05:30 | Hiland WC - Shady Cove | October | PWS01520:SRC-AA | 2010038-01 | Source Water | 274 | 39 | 2020-10-05 05:30 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 07:00 | Camp Baker BSA | October | PWS91786:SRC-AA | 2010052-01 | Source Water | 274 | 39 | 2020-10-05 07:00 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 08:05 | Roseburg Forest Products - Dillard | October | PWS94300:SRC-AA | 2010043-01 | Source Water | 274 | 39 | 2020-10-05 08:05 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 08:45 | Glide Water Association | October | PWS00326:SRC-AA | 2010037-01 | Source Water | 274 | 39 | 2020-10-05 08:45 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 08:45 | Springfield Utility Board | October | PWS00837:CH-A | 2010045-02 | Source Water | 274 | 39 | 2020-10-05 08:45 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 08:50 | City of Wilsonville | October | PWS00954:SRC-HA | 2010035-01 | Source Water | 274 | 39 | 2020-10-05 08:50 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 09:00 | City of Powers | October | PWS00672:SRC-AA | 2010034-01 | Source Water | 274 | 39 | 2020-10-05 09:00 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 09:15 | City of Pendleton | October | PWS00613:SRC-AM | 2010033-01 | Source Water | 274 | 39 | 2020-10-05 09:15 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 09:30 | Buell-Red Prairie Water Association | October | PWS01174:SRC-AA | 2010027-01 | Source Water | 274 | 39 | 2020-10-05 09:30 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 09:30 | Milo Academy | October | PWS00250:SRC-AB | 2010067-01 | Source Water | 274 | 39 | 2020-10-05 09:30 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.67 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2020-10-05 09:50 | Jackson Co Pks Emigrant Lake | October | PWS90730:CH-A | 2010039-01 | Source Water | 274 | 39 | 2020-10-05 09:50 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 10:08 | USFS Tiller Ranger Station | October | PWS01092:SRC-AA | 2010047-01 | Source Water | 274 | 39 | 2020-10-05 10:08 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.11 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-10-05 10:10 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2010068-01 | Source Water | 274 | 39 | 2020-10-05 10:10 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 10:15 | Salem Public Works | October | PWS00731:SRC-AA | 2010044-01 | Source Water | 274 | 39 | 2020-10-05 10:15 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 10:15 | Winston-Dillard Water District | October | PWS00957:SRC-AA | 2010048-01 | Source Water | 274 | 39 | 2020-10-05 10:15 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 10:25 | Eugene Water & Electric Board | October | PWS00287:SRC-AA | 2010036-01 | Source Water | 274 | 39 | 2020-10-05 10:25 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 10:35 | Lake Oswego - Tigard Water Supply | October | PWS00457:SRC-AA | 2010040-01 | Source Water | 274 | 39 | 2020-10-05 10:35 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.1 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-10-05 10:40 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2010070-01 | Source Water | 274 | 39 | 2020-10-05 10:40 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 11:20 | Clarks Branch Water Association | October | PWS00548:SRC-AA | 2010057-01 | Source Water | 274 | 39 | 2020-10-05 11:20 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 12:36 | USFS Steamboat Work Center | October | PWS01091:SRC-AA | 2010046-01 | Source Water | 274 | 39 | 2020-10-05 12:36 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 13:00 | City of Cottage Grove | October | PWS00236:SRC-BA | 2010030-01 | Source Water | 274 | 39 | 2020-10-05 13:00 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 13:10 | PP&L-Toketee Village | October | PWS01012:SRC-AA | 2010042-01 | Source Water | 274 | 39 | 2020-10-05 13:10 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 13:20 | Canby Utility/Veolia Water | October | PWS00157:CH-A | 2010028-01 | Source Water | 274 | 39 | 2020-10-05 13:20 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 13:40 | City of Gates | October | PWS00317:SRC-AA | 2010031-01 | Source Water | 274 | 39 | 2020-10-05 13:40 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-05 14:45 | City of Grants Pass | October | PWS00342:SRC-AA | 2010032-01 | Source Water | 274 | 39 | 2020-10-05 14:45 | 2020-10-06 12:40 | 2020-10-07 13:38 | B20J064 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-06 07:45 | City of Albany | October | PWS00012:SRC-BA | 2010029-01 | Source Water | 275 | 39 | 2020-10-06 07:45 | 2020-10-07 11:30 | 2020-10-12 13:36 | B20J104 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 08:17 | Lyons Mehama Water District | October | PWS00493:SRC-AA | 2010138-01 | Source Water | 281 | 40 | 2020-10-12 08:17 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 08:30 | City of Jefferson | October | PWS00408:SRC-AA | 2010121-01 | Source Water | 281 | 40 | 2020-10-12 08:30 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 08:30 | Joint Water Commission | October | PWS00379:SRC-AA | 2010133-01 | Source Water | 281 | 40 | 2020-10-12 08:30 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 08:30 | North Clackamas County Water Commission | October | PWS00580:SRC-BA | 2010163-01 | Source Water | 281 | 40 | 2020-10-12 08:30 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 08:40 | City of Creswell | October | PWS00246:SRC-BK | 2010117-01 | Source Water | 281 | 40 | 2020-10-12 08:40 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.06 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 09:30 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2010134-01 | Source Water | 281 | 40 | 2020-10-12 09:30 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 09:30 | Tri-City JW & SA | October | PWS00549:SRC-AA | 2010153-01 | Source Water | 281 | 40 | 2020-10-12 09:30 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 09:35 | Umpqua Basin Water Association | October | PWS00719:SRC-AA | 2010154-01 | Source Water | 281 | 40 | 2020-10-12 09:35 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.06 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 09:45 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2010136-01 | Source Water | 281 | 40 | 2020-10-12 09:45 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 09:55 | South Fork Water Board - Oregon City | October | PWS00591:SRC-AA | 2010148-01 | Source Water | 281 | 40 | 2020-10-12 09:55 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 10:00 | City of Estacada | October | PWS00279:SRC-AA | 2010118-01 | Source Water | 281 | 40 | 2020-10-12 10:00 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 10:03 | Medford Water Commission | October | PWS00513:SRC-BA | 2010140-01 | Source Water | 281 | 40 | 2020-10-12 10:03 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 10:15 | Springfield Utility Board | October | PWS00837:SRC-BB | 2010162-01 | Source Water | 281 | 40 | 2020-10-12 10:15 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 10:30 | Umpqua Ranch Co-op | October | PWS00714:SRC-AH | 2010156-01 | Source Water | 281 | 40 | 2020-10-12 10:30 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 11:00 | City of Myrtle Creek | October | PWS00550:SRC-AA | 2010125-01 | Source Water | 281 | 40 | 2020-10-12 11:00 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 11:00 | Roberts Creek Water District | October | PWS00717:SRC-AA | 2010144-01 | Source Water | 281 | 40 | 2020-10-12 11:00 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 11:00 | Seaside Water Department | October | PWS00799:CH-A | 2010145-01 | Source Water | 281 | 40 | 2020-10-12 11:00 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 11:00 | Stayton Water Supply | October | PWS00843:SRC-AA | 2010149-01 | Source Water | 281 | 40 | 2020-10-12 11:00 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 12:00 | City of Lowell | October | PWS00492:SRC-BA | 2010123-01 | Source Water | 281 | 40 | 2020-10-12 12:00 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 12:10 | City of Newport | October | PWS00566:CH-A | 2010126-01 | Source Water | 281 | 40 | 2020-10-12 12:10 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 12:15 | City of Gold Hill | October | PWS00333:SRC-AA | 2010119-01 | Source Water | 281 | 40 | 2020-10-12 12:15 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.09 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-10-12 12:16 | Jackson Co Pks Emigrant Lake | October | PWS90730:CH-A | 2010164-01 | Source Water | 281 | 40 | 2020-10-12 12:16 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 12:30 | Ashland Water Department | October | PWS00047:SRC-AA | 2010158-01 | Source Water | 281 | 40 | 2020-10-12 12:30 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-12 12:39 | City of Roseburg | October | PWS00720:SRC-AA | 2010129-01 | Source Water | 281 | 40 | 2020-10-12 12:39 | 2020-10-13 12:40 | 2020-10-14 12:02 | B20J156 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-13 11:10 | City of Rogue River | October | PWS00712:SRC-BA | 2010127-01 | Source Water | 282 | 40 | 2020-10-13 11:10 | 2020-10-14 11:10 | 2020-10-16 15:35 | B20J188 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-13 13:15 | City of Drain Shops | October | PWS00260:SRC-AA | 2010159-01 | Source Water | 282 | 40 | 2020-10-13 13:15 | 2020-10-14 11:10 | 2020-10-16 15:35 | B20J188 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-14 09:00 | Shangri La Water District | October | PWS00835:SRC-AA | 2010146-01 | Source Water | 283 | 40 | 2020-10-14 09:00 | 2020-10-15 13:45 | 2020-10-16 15:35 | B20J188 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-14 10:49 | Rainbow Water District | October | PWS00839:SRC-BB | 2010142-01 | Source Water | 283 | 40 | 2020-10-14 10:49 | 2020-10-15 13:45 | 2020-10-16 15:35 | B20J188 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-14 12:45 | City of Monroe | October | PWS00540:SRC-BA | 2010160-01 | Source Water | 283 | 40 | 2020-10-14 12:45 | 2020-10-15 13:45 | 2020-10-16 15:35 | B20J188 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 05:15 | Anglers Cove/SCHWC | October | PWS01483:SRC-AB | 2010221-01 | Source Water | 288 | 41 | 2020-10-19 05:15 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 05:30 | Hiland WC - Shady Cove | October | PWS01520:SRC-AA | 2010210-01 | Source Water | 288 | 41 | 2020-10-19 05:30 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.08 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2020-10-19 06:30 | Jackson Co Pks Emigrant Lake | October | PWS90730:CH-A | 2010211-01 | Source Water | 288 | 41 | 2020-10-19 06:30 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 07:08 | City of Pendleton | October | PWS00613:SRC-AM | 2010205-01 | Source Water | 288 | 41 | 2020-10-19 07:08 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 07:20 | Camp Baker BSA | October | PWS91786:SRC-AA | 2010224-01 | Source Water | 288 | 41 | 2020-10-19 07:20 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 08:00 | Roseburg Forest Products - Dillard | October | PWS94300:SRC-AA | 2010215-01 | Source Water | 288 | 41 | 2020-10-19 08:00 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 08:40 | North Clackamas County Water Commission | October | PWS00580:SRC-BA | 2010213-01 | Source Water | 288 | 41 | 2020-10-19 08:40 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 08:40 | Winston-Dillard Water District | October | PWS00957:SRC-AA | 2010220-01 | Source Water | 288 | 41 | 2020-10-19 08:40 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 08:45 | City of Wilsonville | October | PWS00954:SRC-HA | 2010207-01 | Source Water | 288 | 41 | 2020-10-19 08:45 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 09:10 | Springfield Utility Board | October | PWS00837:CH-A | 2010217-02 | Source Water | 288 | 41 | 2020-10-19 09:10 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 09:15 | Glide Water Association | October | PWS00326:SRC-AA | 2010209-01 | Source Water | 288 | 41 | 2020-10-19 09:15 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 09:30 | Milo Academy | October | PWS00250:SRC-AB | 2010239-01 | Source Water | 288 | 41 | 2020-10-19 09:30 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 10:00 | Eugene Water & Electric Board | October | PWS00287:SRC-AA | 2010208-01 | Source Water | 288 | 41 | 2020-10-19 10:00 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 10:05 | USFS Tiller Ranger Station | October | PWS01092:SRC-AA | 2010219-01 | Source Water | 288 | 41 | 2020-10-19 10:05 | 2020-10-20 16:00 | 2020-10-22 18:53 | B20J281 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 10:25 | Salem Public Works | October | PWS00731:SRC-AA | 2010216-01 | Source Water | 288 | 41 | 2020-10-19 10:25 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 11:00 | Lake Oswego - Tigard Water Supply | October | PWS00457:SRC-AA | 2010212-01 | Source Water | 288 | 41 | 2020-10-19 11:00 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 11:05 | Clarks Branch Water Association | October | PWS00548:SRC-AA | 2010229-01 | Source Water | 288 | 41 | 2020-10-19 11:05 | 2020-10-20 16:00 | 2020-10-22 18:53 | B20J281 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 11:30 | Buell-Red Prairie Water Association | October | PWS01174:SRC-AA | 2010199-01 | Source Water | 288 | 41 | 2020-10-19 11:30 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 12:00 | City of Gates | October | PWS00317:SRC-AA | 2010203-01 | Source Water | 288 | 41 | 2020-10-19 12:00 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 12:20 | USFS Steamboat Work Center | October | PWS01091:SRC-AA | 2010218-01 | Source Water | 288 | 41 | 2020-10-19 12:20 | 2020-10-20 16:00 | 2020-10-22 18:53 | B20J281 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 12:40 | Springfield Utility Board | October | PWS00837:SRC-BB | 2010217-01 | Source Water | 288 | 41 | 2020-10-19 12:40 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 13:10 | PP&L-Toketee Village | October | PWS01012:SRC-AA | 2010214-01 | Source Water | 288 | 41 | 2020-10-19 13:10 | 2020-10-20 16:00 | 2020-10-22 18:53 | B20J281 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 13:25 | Canby Utility/Veolia Water | October | PWS00157:CH-A | 2010200-01 | Source Water | 288 | 41 | 2020-10-19 13:25 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 14:00 | City of Cottage Grove | October | PWS00236:SRC-BA | 2010202-01 | Source Water | 288 | 41 | 2020-10-19 14:00 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-19 14:50 | City of Grants Pass | October | PWS00342:SRC-AA | 2010204-01 | Source Water | 288 | 41 | 2020-10-19 14:50 | 2020-10-20 12:00 | 2020-10-21 11:24 | B20J260 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-23 12:30 | City of Powers | October | PWS00672:SRC-AA | 2010317-01 | Source Water | 292 | 41 | 2020-10-23 12:30 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 08:17 | Lyons Mehama Water District | October | PWS00493:SRC-AA | 2010291-01 | Source Water | 295 | 42 | 2020-10-26 08:17 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 08:22 | Joint Water Commission | October | PWS00379:SRC-AA | 2010286-01 | Source Water | 295 | 42 | 2020-10-26 08:22 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 08:30 | City of Creswell | October | PWS00246:SRC-BK | 2010271-01 | Source Water | 295 | 42 | 2020-10-26 08:30 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 09:00 | City of Jefferson | October | PWS00408:SRC-AA | 2010275-01 | Source Water | 295 | 42 | 2020-10-26 09:00 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 09:00 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2010287-01 | Source Water | 295 | 42 | 2020-10-26 09:00 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 09:20 | Umpqua Basin Water Association | October | PWS00719:SRC-AA | 2010306-01 | Source Water | 295 | 42 | 2020-10-26 09:20 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 09:30 | Tri-City JW & SA | October | PWS00549:SRC-AA | 2010305-01 | Source Water | 295 | 42 | 2020-10-26 09:30 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 09:45 | South Fork Water Board - Oregon City | October | PWS00591:SRC-AA | 2010301-01 | Source Water | 295 | 42 | 2020-10-26 09:45 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 09:50 | City of Estacada | October | PWS00279:SRC-AA | 2010272-01 | Source Water | 295 | 42 | 2020-10-26 09:50 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 10:00 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2010289-01 | Source Water | 295 | 42 | 2020-10-26 10:00 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 10:31 | Medford Water Commission | October | PWS00513:SRC-BA | 2010293-01 | Source Water | 295 | 42 | 2020-10-26 10:31 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 11:00 | City of Newport | October | PWS00566:CH-A | 2010280-01 | Source Water | 295 | 42 | 2020-10-26 11:00 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 11:00 | Roberts Creek Water District | October | PWS00717:SRC-AA | 2010297-01 | Source Water | 295 | 42 | 2020-10-26 11:00 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 11:00 | Stayton Water Supply | October | PWS00843:SRC-AA | 2010302-01 | Source Water | 295 | 42 | 2020-10-26 11:00 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 11:10 | City of Myrtle Creek | October | PWS00550:SRC-AA | 2010279-01 | Source Water | 295 | 42 | 2020-10-26 11:10 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 11:30 | Seaside Water Department | October | PWS00799:CH-A | 2010298-01 | Source Water | 295 | 42 | 2020-10-26 11:30 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 12:00 | City of Rogue River | October | PWS00712:SRC-BA | 2010281-01 | Source Water | 295 | 42 | 2020-10-26 12:00 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 12:30 | Ashland Water Department | October | PWS00047:SRC-AA | 2010310-03 | Source Water | 295 | 42 | 2020-10-26 12:30 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 12:41 | City of Roseburg | October | PWS00720:SRC-AA | 2010283-01 | Source Water | 295 | 42 | 2020-10-26 12:41 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 12:50 | Rainbow Water District | October | PWS00839:SRC-BB | 2010295-01 | Source Water | 295 | 42 | 2020-10-26 12:50 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-26 13:00 | City of Drain Shops | October | PWS00260:SRC-AA | 2010311-01 | Source Water | 295 | 42 | 2020-10-26 13:00 | 2020-10-27 12:30 | 2020-10-28 13:51 | B20J355 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-27 12:10 | City of Lowell | October | PWS00492:SRC-BA | 2010277-01 | Source Water | 296 | 42 | 2020-10-27 12:10 | 2020-10-28 12:30 | 2020-10-30 13:55 | B20J375 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-27 15:30 | City of Gold Hill | October | PWS00333:SRC-AA | 2010273-01 | Source Water | 296 | 42 | 2020-10-27 15:30 | 2020-10-28 12:30 | 2020-10-30 13:55 | B20J375 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-28 12:15 | City of Monroe | October | PWS00540:SRC-BA | 2010312-01 | Source Water | 297 | 42 | 2020-10-28 12:15 | 2020-10-29 11:30 | 2020-10-30 13:55 | B20J375 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2020-10-29 14:00 | Shangri La Water District | October | PWS00835:SRC-AA | 2010336-01 | Source Water | 298 | 42 | 2020-10-29 14:00 | 2020-10-30 12:00 | 2020-11-02 12:06 | B20J398 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 04:31 | City of Wilsonville | May | PWS00954:SRC-HA | 2105028-01 | Source Water | 122 | 17 | 2021-05-03 04:31 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 05:15 | Anglers Cove/SCHWC | May | PWS01483:SRC-AB | 2105018-01 | Source Water | 122 | 17 | 2021-05-03 05:15 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 05:30 | Hiland WC - Shady Cove | May | PWS01520:SRC-AA | 2105032-01 | Source Water | 122 | 17 | 2021-05-03 05:30 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 07:18 | USFS Tiller Ranger Station | May | PWS01092:SRC-AA | 2105043-01 | Source Water | 122 | 17 | 2021-05-03 07:18 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 07:30 | City of Albany | May | PWS00012:SRC-BA | 2105022-01 | Source Water | 122 | 17 | 2021-05-03 07:30 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 07:59 | Clarks Branch Water Association | May | PWS00548:SRC-AA | 2105029-01 | Source Water | 122 | 17 | 2021-05-03 07:59 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 08:00 | Camp Baker BSA | May | PWS91786:SRC-AA | 2105020-01 | Source Water | 122 | 17 | 2021-05-03 08:00 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 08:12 | City of Pendleton | May | PWS00613:SRC-AM | 2105026-01 | Source Water | 122 | 17 | 2021-05-03 08:12 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 08:20 | Winston-Dillard Water District | May | PWS00957:SRC-AA | 2105047-01 | Source Water | 122 | 17 | 2021-05-03 08:20 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 08:33 | Joint Water Commission | May | PWS00379:SRC-AA | 2105034-01 | Source Water | 122 | 17 | 2021-05-03 08:33 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 08:50 | Springfield Utility Board | May | PWS00837:CH-A | 2105040-02 | Source Water | 122 | 17 | 2021-05-03 08:50 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 09:20 | City of Powers | May | PWS00672:SRC-AA | 2105027-01 | Source Water | 122 | 17 | 2021-05-03 09:20 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 09:20 | USFS Steamboat Work Center | May | PWS01091:SRC-AA | 2105042-01 | Source Water | 122 | 17 | 2021-05-03 09:20 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 09:30 | Glide Water Association | May | PWS00326:SRC-AA | 2105031-01 | Source Water | 122 | 17 | 2021-05-03 09:30 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 09:50 | North Clackamas County Water Commission | May | PWS00580:SRC-BA | 2105036-01 | Source Water | 122 | 17 | 2021-05-03 09:50 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 10:16 | USFS Horseshoe Bend CG | May | PWS94179:SRC-AA | 2105041-01 | Source Water | 122 | 17 | 2021-05-03 10:16 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 10:30 | Roseburg Forest Products - Dillard | May | PWS94300:SRC-AA | 2105038-01 | Source Water | 122 | 17 | 2021-05-03 10:30 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 11:06 | PP&L-Toketee Village | May | PWS01012:SRC-AA | 2105037-01 | Source Water | 122 | 17 | 2021-05-03 11:06 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 11:15 | Eugene Water & Electric Board | May | PWS00287:SRC-AA | 2105030-01 | Source Water | 122 | 17 | 2021-05-03 11:15 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 11:15 | Salem Public Works | May | PWS00731:SRC-AA | 2105039-01 | Source Water | 122 | 17 | 2021-05-03 11:15 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.67 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-05-03 11:30 | Buell-Red Prairie Water Association | May | PWS01174:SRC-AA | 2105019-01 | Source Water | 122 | 17 | 2021-05-03 11:30 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 11:30 | Springfield Utility Board | May | PWS00837:SRC-BB | 2105040-01 | Source Water | 122 | 17 | 2021-05-03 11:30 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 13:00 | Lake Oswego - Tigard Water Supply | May | PWS00457:SRC-AA | 2105035-01 | Source Water | 122 | 17 | 2021-05-03 13:00 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 13:30 | City of Gates | May | PWS00317:SRC-AA | 2105024-01 | Source Water | 122 | 17 | 2021-05-03 13:30 | 2021-05-05 10:30 | 2021-05-07 18:03 | B21E089 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 13:45 | Canby Utility/Veolia Water | May | PWS00157:CH-A | 2105021-01 | Source Water | 122 | 17 | 2021-05-03 13:45 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 15:30 | City of Cottage Grove | May | PWS00236:SRC-BA | 2105023-01 | Source Water | 122 | 17 | 2021-05-03 15:30 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-03 16:15 | City of Grants Pass | May | PWS00342:SRC-AA | 2105025-01 | Source Water | 122 | 17 | 2021-05-03 16:15 | 2021-05-04 12:30 | 2021-05-05 13:43 | B21E026 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 06:00 | City of Newport | May | PWS00566:CH-A | 2105077-01 | Source Water | 129 | 18 | 2021-05-10 06:00 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 08:11 | Jackson Co Pks Emigrant Lake | May | PWS90730:CH-A | 2105102-01 | Source Water | 129 | 18 | 2021-05-10 08:11 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 09:00 | Shangri La Water District | May | PWS00835:SRC-AA | 2105087-01 | Source Water | 129 | 18 | 2021-05-10 09:00 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 09:15 | City of Creswell | May | PWS00246:SRC-BK | 2105069-01 | Source Water | 129 | 18 | 2021-05-10 09:15 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 09:30 | City of Estacada | May | PWS00279:SRC-AA | 2105071-01 | Source Water | 129 | 18 | 2021-05-10 09:30 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 09:30 | Umpqua Ranch Co-op | May | PWS00714:SRC-AH | 2105092-01 | Source Water | 129 | 18 | 2021-05-10 09:30 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 09:45 | City of Jefferson | May | PWS00408:SRC-AA | 2105073-01 | Source Water | 129 | 18 | 2021-05-10 09:45 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 10:00 | South Fork Water Board - Oregon City | May | PWS00591:SRC-AA | 2105088-01 | Source Water | 129 | 18 | 2021-05-10 10:00 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 10:05 | Umpqua Basin Water Association | May | PWS00719:SRC-AA | 2105091-01 | Source Water | 129 | 18 | 2021-05-10 10:05 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 10:30 | Josephine Co Pks Lake Selmac 2 | May | PWS94645:SRC-AA | 2105081-01 | Source Water | 129 | 18 | 2021-05-10 10:30 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 10:47 | City of Myrtle Creek | May | PWS00550:SRC-AA | 2105076-01 | Source Water | 129 | 18 | 2021-05-10 10:47 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 11:00 | Josephine Co Pks Lake Selmac 1 | May | PWS90186:SRC-AA | 2105080-01 | Source Water | 129 | 18 | 2021-05-10 11:00 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 11:00 | Stayton Water Supply | May | PWS00843:SRC-AA | 2105089-01 | Source Water | 129 | 18 | 2021-05-10 11:00 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 11:30 | Buell-Red Prairie Water Association | May | PWS01174:SRC-AA | 2105097-01 | Source Water | 129 | 18 | 2021-05-10 11:30 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 11:30 | Roberts Creek Water District | May | PWS00717:SRC-AA | 2105085-01 | Source Water | 129 | 18 | 2021-05-10 11:30 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 12:00 | Ashland Water Department | May | PWS00047:SRC-AA | 2105068-01 | Source Water | 129 | 18 | 2021-05-10 12:00 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 12:19 | City of Roseburg | May | PWS00720:SRC-AA | 2105079-01 | Source Water | 129 | 18 | 2021-05-10 12:19 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 12:45 | Tri-City JW & SA | May | PWS00549:SRC-AA | 2105090-01 | Source Water | 129 | 18 | 2021-05-10 12:45 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 12:53 | Douglas County Parks - Whistler's Bend | May | PWS93944:SRC-AB | 2105099-01 | Source Water | 129 | 18 | 2021-05-10 12:53 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 13:00 | City of Drain Shops | May | PWS00260:SRC-AA | 2105070-01 | Source Water | 129 | 18 | 2021-05-10 13:00 | 2021-05-12 10:30 | 2021-05-17 19:07 | B21E135 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 13:00 | City of Rogue River | May | PWS00712:SRC-BA | 2105078-01 | Source Water | 129 | 18 | 2021-05-10 13:00 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 13:13 | Rainbow Water District | May | PWS00839:SRC-BB | 2105084-01 | Source Water | 129 | 18 | 2021-05-10 13:13 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 13:15 | Medford Water Commission | May | PWS00513:SRC-BA | 2105083-01 | Source Water | 129 | 18 | 2021-05-10 13:15 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 13:30 | Lyons Mehama Water District | May | PWS00493:SRC-AA | 2105082-01 | Source Water | 129 | 18 | 2021-05-10 13:30 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 14:00 | City of Lowell | May | PWS00492:SRC-BA | 2105074-01 | Source Water | 129 | 18 | 2021-05-10 14:00 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 14:00 | Seaside Water Department | May | PWS00799:CH-A | 2105086-01 | Source Water | 129 | 18 | 2021-05-10 14:00 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-10 14:40 | City of Gold Hill | May | PWS00333:SRC-AA | 2105072-01 | Source Water | 129 | 18 | 2021-05-10 14:40 | 2021-05-11 12:00 | 2021-05-12 13:52 | B21E112 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-11 11:45 | City of Monroe | May | PWS00540:SRC-BA | 2105075-01 | Source Water | 130 | 18 | 2021-05-11 11:45 | 2021-05-12 10:30 | 2021-05-17 19:07 | B21E135 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 05:15 | Anglers Cove/SCHWC | May | PWS01483:SRC-AB | 2105152-01 | Source Water | 136 | 19 | 2021-05-17 05:15 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 05:30 | Hiland WC - Shady Cove | May | PWS01520:SRC-AA | 2105173-01 | Source Water | 136 | 19 | 2021-05-17 05:30 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 06:50 | City of Albany | May | PWS00012:SRC-BA | 2105158-01 | Source Water | 136 | 19 | 2021-05-17 06:50 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 08:00 | City of Pendleton | May | PWS00613:SRC-AM | 2105165-01 | Source Water | 136 | 19 | 2021-05-17 08:00 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 08:00 | Glide Water Association | May | PWS00326:SRC-AA | 2105171-01 | Source Water | 136 | 19 | 2021-05-17 08:00 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 08:00 | Roseburg Forest Products - Dillard | May | PWS94300:SRC-AA | 2105182-01 | Source Water | 136 | 19 | 2021-05-17 08:00 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 08:13 | Joint Water Commission | May | PWS00379:SRC-AA | 2105175-01 | Source Water | 136 | 19 | 2021-05-17 08:13 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 08:50 | Lake Oswego - Tigard Water Supply | May | PWS00457:SRC-AA | 2105176-01 | Source Water | 136 | 19 | 2021-05-17 08:50 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 09:00 | Camp Baker BSA | May | PWS91786:SRC-AA | 2105155-01 | Source Water | 136 | 19 | 2021-05-17 09:00 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 09:00 | City of Powers | May | PWS00672:SRC-AA | 2105166-01 | Source Water | 136 | 19 | 2021-05-17 09:00 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 09:00 | Winston-Dillard Water District | May | PWS00957:SRC-AA | 2105190-01 | Source Water | 136 | 19 | 2021-05-17 09:00 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 09:05 | Jackson Co Pks Emigrant Lake | May | PWS90730:CH-A | 2105198-01 | Source Water | 136 | 19 | 2021-05-17 09:05 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 09:17 | USFS Tiller Ranger Station | May | PWS01092:SRC-AA | 2105189-01 | Source Water | 136 | 19 | 2021-05-17 09:17 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 10:09 | Clarks Branch Water Association | May | PWS00548:SRC-AA | 2105168-01 | Source Water | 136 | 19 | 2021-05-17 10:09 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 10:15 | North Clackamas County Water Commission | May | PWS00580:SRC-BA | 2105178-01 | Source Water | 136 | 19 | 2021-05-17 10:15 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 10:30 | Buell-Red Prairie Water Association | May | PWS01174:SRC-AA | 2105154-01 | Source Water | 136 | 19 | 2021-05-17 10:30 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 11:25 | Salem Public Works | May | PWS00731:SRC-AA | 2105183-01 | Source Water | 136 | 19 | 2021-05-17 11:25 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 12:15 | City of Wilsonville | May | PWS00954:SRC-HA | 2105167-01 | Source Water | 136 | 19 | 2021-05-17 12:15 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 12:15 | Eugene Water & Electric Board | May | PWS00287:SRC-AA | 2105169-01 | Source Water | 136 | 19 | 2021-05-17 12:15 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 12:15 | USFS Steamboat Work Center | May | PWS01091:SRC-AA | 2105187-01 | Source Water | 136 | 19 | 2021-05-17 12:15 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 13:36 | PP&L-Toketee Village | May | PWS01012:SRC-AA | 2105180-01 | Source Water | 136 | 19 | 2021-05-17 13:36 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 13:40 | Springfield Utility Board | May | PWS00837:SRC-BB | 2105185-01 | Source Water | 136 | 19 | 2021-05-17 13:40 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 13:58 | USFS Horseshoe Bend CG | May | PWS94179:SRC-AA | 2105186-01 | Source Water | 136 | 19 | 2021-05-17 13:58 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 14:00 | City of Gates | May | PWS00317:SRC-AA | 2105161-01 | Source Water | 136 | 19 | 2021-05-17 14:00 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 14:20 | Canby Utility/Veolia Water | May | PWS00157:CH-A | 2105156-01 | Source Water | 136 | 19 | 2021-05-17 14:20 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 14:20 | Springfield Utility Board | May | PWS00837:CH-A | 2105185-02 | Source Water | 136 | 19 | 2021-05-17 14:20 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 15:00 | City of Cottage Grove | May | PWS00236:SRC-BA | 2105160-01 | Source Water | 136 | 19 | 2021-05-17 15:00 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-17 15:55 | City of Grants Pass | May | PWS00342:SRC-AA | 2105163-01 | Source Water | 136 | 19 | 2021-05-17 15:55 | 2021-05-18 12:45 | 2021-05-19 17:02 | B21E209 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 08:05 | Tri-City JW & SA | May | PWS00549:SRC-AA | 2105267-01 | Source Water | 143 | 20 | 2021-05-24 08:05 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 08:33 | South Fork Water Board - Oregon City | May | PWS00591:SRC-AA | 2105263-01 | Source Water | 143 | 20 | 2021-05-24 08:33 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 09:00 | City of Creswell | May | PWS00246:SRC-BK | 2105235-01 | Source Water | 143 | 20 | 2021-05-24 09:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 09:00 | Umpqua Ranch Co-op | May | PWS00714:SRC-AH | 2105270-01 | Source Water | 143 | 20 | 2021-05-24 09:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 09:30 | Shangri La Water District | May | PWS00835:SRC-AA | 2105262-01 | Source Water | 143 | 20 | 2021-05-24 09:30 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 10:00 | City of Estacada | May | PWS00279:SRC-AA | 2105237-01 | Source Water | 143 | 20 | 2021-05-24 10:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 10:00 | Roberts Creek Water District | May | PWS00717:SRC-AA | 2105260-01 | Source Water | 143 | 20 | 2021-05-24 10:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 10:09 | City of Roseburg | May | PWS00720:SRC-AA | 2105249-01 | Source Water | 143 | 20 | 2021-05-24 10:09 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 10:10 | City of Jefferson | May | PWS00408:SRC-AA | 2105241-01 | Source Water | 143 | 20 | 2021-05-24 10:10 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 10:30 | City of Myrtle Creek | May | PWS00550:SRC-AA | 2105245-01 | Source Water | 143 | 20 | 2021-05-24 10:30 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 10:35 | Umpqua Basin Water Association | May | PWS00719:SRC-AA | 2105268-01 | Source Water | 143 | 20 | 2021-05-24 10:35 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 10:45 | City of Newport | May | PWS00566:CH-A | 2105246-01 | Source Water | 143 | 20 | 2021-05-24 10:45 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 11:00 | Ashland Water Department | May | PWS00047:SRC-AA | 2105233-01 | Source Water | 143 | 20 | 2021-05-24 11:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 11:00 | City of Rogue River | May | PWS00712:SRC-BA | 2105247-01 | Source Water | 143 | 20 | 2021-05-24 11:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 12:10 | Lyons Mehama Water District | May | PWS00493:SRC-AA | 2105255-01 | Source Water | 143 | 20 | 2021-05-24 12:10 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 12:20 | City of Lowell | May | PWS00492:SRC-BA | 2105243-01 | Source Water | 143 | 20 | 2021-05-24 12:20 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 13:00 | City of Drain Shops | May | PWS00260:SRC-AA | 2105236-01 | Source Water | 143 | 20 | 2021-05-24 13:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 1e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 13:00 | Stayton Water Supply | May | PWS00843:SRC-AA | 2105265-01 | Source Water | 143 | 20 | 2021-05-24 13:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 13:43 | Douglas County Parks - Whistler's Bend | May | PWS93944:SRC-AB | 2105273-01 | Source Water | 143 | 20 | 2021-05-24 13:43 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 14:00 | Medford Water Commission | May | PWS00513:SRC-BA | 2105257-01 | Source Water | 143 | 20 | 2021-05-24 14:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 9e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 14:00 | Seaside Water Department | May | PWS00799:CH-A | 2105261-01 | Source Water | 143 | 20 | 2021-05-24 14:00 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 14:20 | Josephine Co Pks Lake Selmac 2 | May | PWS94645:SRC-AA | 2105253-01 | Source Water | 143 | 20 | 2021-05-24 14:20 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 14:27 | Rainbow Water District | May | PWS00839:SRC-BB | 2105259-01 | Source Water | 143 | 20 | 2021-05-24 14:27 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-24 14:40 | Josephine Co Pks Lake Selmac 1 | May | PWS90186:SRC-AA | 2105251-01 | Source Water | 143 | 20 | 2021-05-24 14:40 | 2021-05-25 12:15 | 2021-05-26 14:55 | B21E298 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-25 11:50 | City of Monroe | May | PWS00540:SRC-BA | 2105244-01 | Source Water | 144 | 20 | 2021-05-25 11:50 | 2021-05-26 12:30 | 2021-05-27 18:51 | B21E324 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-05-25 15:00 | City of Gold Hill | May | PWS00333:SRC-AA | 2105239-01 | Source Water | 144 | 20 | 2021-05-25 15:00 | 2021-05-26 12:30 | 2021-05-27 18:51 | B21E324 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 05:15 | Anglers Cove/SCHWC | June | PWS01483:SRC-AB | 2106006-01 | Source Water | 150 | 21 | 2021-06-01 05:15 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 1e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 05:30 | Hiland WC - Shady Cove | June | PWS01520:SRC-AA | 2106020-01 | Source Water | 150 | 21 | 2021-06-01 05:30 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 07:00 | Camp Baker BSA | June | PWS91786:SRC-AA | 2106008-01 | Source Water | 150 | 21 | 2021-06-01 07:00 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 07:30 | Glide Water Association | June | PWS00326:SRC-AA | 2106019-01 | Source Water | 150 | 21 | 2021-06-01 07:30 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 08:00 | City of Pendleton | June | PWS00613:SRC-AM | 2106014-01 | Source Water | 150 | 21 | 2021-06-01 08:00 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 08:00 | Jackson Co Pks Emigrant Lake | June | PWS90730:CH-A | 2106021-01 | Source Water | 150 | 21 | 2021-06-01 08:00 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 08:02 | Roseburg Forest Products - Dillard | June | PWS94300:SRC-AA | 2106026-01 | Source Water | 150 | 21 | 2021-06-01 08:02 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 08:30 | Joint Water Commission | June | PWS00379:SRC-AA | 2106022-01 | Source Water | 150 | 21 | 2021-06-01 08:30 | 2021-06-02 13:30 | 2021-06-02 18:18 | B21E361 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 08:40 | Winston-Dillard Water District | June | PWS00957:SRC-AA | 2106032-01 | Source Water | 150 | 21 | 2021-06-01 08:40 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 08:45 | City of Powers | June | PWS00672:SRC-AA | 2106015-01 | Source Water | 150 | 21 | 2021-06-01 08:45 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 09:30 | USFS Tiller Ranger Station | June | PWS01092:SRC-AA | 2106031-01 | Source Water | 150 | 21 | 2021-06-01 09:30 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 09:35 | North Clackamas County Water Commission | June | PWS00580:SRC-BA | 2106024-01 | Source Water | 150 | 21 | 2021-06-01 09:35 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 09:40 | Springfield Utility Board | June | PWS00837:CH-A | 2106028-02 | Source Water | 150 | 21 | 2021-06-01 09:40 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 10:15 | Salem Public Works | June | PWS00731:SRC-AA | 2106027-01 | Source Water | 150 | 21 | 2021-06-01 10:15 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 10:30 | Buell-Red Prairie Water Association | June | PWS01174:SRC-AA | 2106007-01 | Source Water | 150 | 21 | 2021-06-01 10:30 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 10:30 | Clarks Branch Water Association | June | PWS00548:SRC-AA | 2106017-01 | Source Water | 150 | 21 | 2021-06-01 10:30 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 11:10 | Eugene Water & Electric Board | June | PWS00287:SRC-AA | 2106018-01 | Source Water | 150 | 21 | 2021-06-01 11:10 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 11:40 | USFS Steamboat Work Center | June | PWS01091:SRC-AA | 2106030-01 | Source Water | 150 | 21 | 2021-06-01 11:40 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 12:25 | City of Wilsonville | June | PWS00954:SRC-HA | 2106016-01 | Source Water | 150 | 21 | 2021-06-01 12:25 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 12:45 | USFS Horseshoe Bend CG | June | PWS94179:SRC-AA | 2106029-01 | Source Water | 150 | 21 | 2021-06-01 12:45 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 12:55 | Canby Utility/Veolia Water | June | PWS00157:CH-A | 2106009-01 | Source Water | 150 | 21 | 2021-06-01 12:55 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 13:00 | Springfield Utility Board | June | PWS00837:SRC-BB | 2106028-01 | Source Water | 150 | 21 | 2021-06-01 13:00 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 1e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 13:05 | Lake Oswego - Tigard Water Supply | June | PWS01542:SRC-AA | 2106023-01 | Source Water | 150 | 21 | 2021-06-01 13:05 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 13:18 | PP&L-Toketee Village | June | PWS01012:SRC-AA | 2106025-01 | Source Water | 150 | 21 | 2021-06-01 13:18 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 14:00 | City of Gates | June | PWS00317:SRC-AA | 2106012-01 | Source Water | 150 | 21 | 2021-06-01 14:00 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 15:30 | City of Cottage Grove | June | PWS00236:SRC-BA | 2106011-01 | Source Water | 150 | 21 | 2021-06-01 15:30 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-01 16:50 | City of Grants Pass | June | PWS00342:SRC-AA | 2106013-01 | Source Water | 150 | 21 | 2021-06-01 16:50 | 2021-06-02 13:30 | 2021-06-03 14:44 | B21F029 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-03 07:05 | City of Albany | June | PWS00012:SRC-BA | 2106010-01RE1 | Source Water | 152 | 21 | 2021-06-03 07:05 | 2021-06-03 10:15 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 08:15 | City of Gold Hill | June | PWS00333:SRC-AA | 2106065-01 | Source Water | 156 | 22 | 2021-06-07 08:15 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 08:35 | City of Creswell | June | PWS00246:SRC-BK | 2106062-01 | Source Water | 156 | 22 | 2021-06-07 08:35 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 08:38 | Lyons Mehama Water District | June | PWS00493:SRC-AA | 2106075-01 | Source Water | 156 | 22 | 2021-06-07 08:38 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 09:10 | City of Jefferson | June | PWS00408:SRC-AA | 2106066-01 | Source Water | 156 | 22 | 2021-06-07 09:10 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 09:30 | Shangri La Water District | June | PWS00835:SRC-AA | 2106080-01 | Source Water | 156 | 22 | 2021-06-07 09:30 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 10:10 | City of Drain Shops | June | PWS00260:SRC-AA | 2106063-01 | Source Water | 156 | 22 | 2021-06-07 10:10 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 10:20 | City of Lowell | June | PWS00492:SRC-BA | 2106067-01 | Source Water | 156 | 22 | 2021-06-07 10:20 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 10:30 | City of Myrtle Creek | June | PWS00550:SRC-AA | 2106069-01 | Source Water | 156 | 22 | 2021-06-07 10:30 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 10:45 | City of Newport | June | PWS00566:CH-A | 2106070-01 | Source Water | 156 | 22 | 2021-06-07 10:45 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 10:50 | South Fork Water Board - Oregon City | June | PWS00591:SRC-AA | 2106081-01 | Source Water | 156 | 22 | 2021-06-07 10:50 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 11:00 | City of Rogue River | June | PWS00712:SRC-BA | 2106071-01 | Source Water | 156 | 22 | 2021-06-07 11:00 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 11:10 | City of Estacada | June | PWS00279:SRC-AA | 2106064-01 | Source Water | 156 | 22 | 2021-06-07 11:10 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 11:30 | Roberts Creek Water District | June | PWS00717:SRC-AA | 2106078-01 | Source Water | 156 | 22 | 2021-06-07 11:30 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 11:30 | Tri-City JW & SA | June | PWS00549:SRC-AA | 2106083-01 | Source Water | 156 | 22 | 2021-06-07 11:30 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 11:40 | Rainbow Water District | June | PWS00839:SRC-BB | 2106077-01 | Source Water | 156 | 22 | 2021-06-07 11:40 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 11:45 | Umpqua Basin Water Association | June | PWS00719:SRC-AA | 2106084-01 | Source Water | 156 | 22 | 2021-06-07 11:45 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 12:11 | City of Roseburg | June | PWS00720:SRC-AA | 2106072-01 | Source Water | 156 | 22 | 2021-06-07 12:11 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 12:15 | Ashland Water Department | June | PWS00047:SRC-AA | 2106061-01 | Source Water | 156 | 22 | 2021-06-07 12:15 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 12:30 | City of Monroe | June | PWS00540:SRC-BA | 2106068-01 | Source Water | 156 | 22 | 2021-06-07 12:30 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 13:00 | Stayton Water Supply | June | PWS00843:SRC-AA | 2106082-01 | Source Water | 156 | 22 | 2021-06-07 13:00 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 13:40 | Medford Water Commission | June | PWS00513:SRC-BA | 2106076-01 | Source Water | 156 | 22 | 2021-06-07 13:40 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-07 14:00 | Seaside Water Department | June | PWS00799:CH-A | 2106079-01 | Source Water | 156 | 22 | 2021-06-07 14:00 | 2021-06-08 13:00 | 2021-06-08 22:06 | B21F113 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-08 09:30 | Umpqua Ranch Co-op | June | PWS00714:SRC-AH | 2106085-01 | Source Water | 157 | 22 | 2021-06-08 09:30 | 2021-06-09 12:30 | 2021-06-14 15:33 | B21F132 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-09 14:30 | Josephine Co Pks Lake Selmac 2 | June | PWS94645:SRC-AA | 2106074-01 | Source Water | 158 | 22 | 2021-06-09 14:30 | 2021-06-10 11:30 | 2021-06-14 15:33 | B21F132 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-09 15:00 | Josephine Co Pks Lake Selmac 1 | June | PWS90186:SRC-AA | 2106073-01 | Source Water | 158 | 22 | 2021-06-09 15:00 | 2021-06-10 11:30 | 2021-06-14 15:33 | B21F132 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 05:15 | Anglers Cove/SCHWC | June | PWS01483:SRC-AB | 2106147-01 | Source Water | 163 | 23 | 2021-06-14 05:15 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 05:30 | Hiland WC - Shady Cove | June | PWS01520:SRC-AA | 2106168-01 | Source Water | 163 | 23 | 2021-06-14 05:30 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 06:15 | City of Albany | June | PWS00012:SRC-BA | 2106153-01 | Source Water | 163 | 23 | 2021-06-14 06:15 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 08:00 | Glide Water Association | June | PWS00326:SRC-AA | 2106166-01 | Source Water | 163 | 23 | 2021-06-14 08:00 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 08:00 | USFS Tiller Ranger Station | June | PWS01092:SRC-AA | 2106185-01 | Source Water | 163 | 23 | 2021-06-14 08:00 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 08:05 | Roseburg Forest Products - Dillard | June | PWS94300:SRC-AA | 2106178-01 | Source Water | 163 | 23 | 2021-06-14 08:05 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 08:35 | Joint Water Commission | June | PWS00379:SRC-AA | 2106171-01 | Source Water | 163 | 23 | 2021-06-14 08:35 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 09:00 | Salem Public Works | June | PWS00731:SRC-AA | 2106180-01 | Source Water | 163 | 23 | 2021-06-14 09:00 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 09:05 | Clarks Branch Water Association | June | PWS00548:SRC-AA | 2106162-01 | Source Water | 163 | 23 | 2021-06-14 09:05 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 09:10 | Camp Baker BSA | June | PWS91786:SRC-AA | 2106149-01 | Source Water | 163 | 23 | 2021-06-14 09:10 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 09:10 | Springfield Utility Board | June | PWS00837:CH-A | 2106181-02 | Source Water | 163 | 23 | 2021-06-14 09:10 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 09:30 | City of Powers | June | PWS00672:SRC-AA | 2106160-01 | Source Water | 163 | 23 | 2021-06-14 09:30 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 09:30 | Jackson Co Pks Emigrant Lake | June | PWS90730:CH-A | 2106169-01 | Source Water | 163 | 23 | 2021-06-14 09:30 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 10:30 | Buell-Red Prairie Water Association | June | PWS01174:SRC-AA | 2106148-01 | Source Water | 163 | 23 | 2021-06-14 10:30 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 10:30 | North Clackamas County Water Commission | June | PWS00580:SRC-BA | 2106175-01 | Source Water | 163 | 23 | 2021-06-14 10:30 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 10:35 | Springfield Utility Board | June | PWS00837:SRC-BB | 2106181-01 | Source Water | 163 | 23 | 2021-06-14 10:35 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 10:37 | Douglas County Parks - Whistler's Bend | June | PWS93944:SRC-AB | 2106189-01 | Source Water | 163 | 23 | 2021-06-14 10:37 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 10:44 | PP&L-Toketee Village | June | PWS01012:SRC-AA | 2106177-01 | Source Water | 163 | 23 | 2021-06-14 10:44 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 11:21 | USFS Horseshoe Bend CG | June | PWS94179:SRC-AA | 2106182-01 | Source Water | 163 | 23 | 2021-06-14 11:21 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 11:30 | City of Pendleton | June | PWS00613:SRC-AM | 2106159-01 | Source Water | 163 | 23 | 2021-06-14 11:30 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 11:30 | Lake Oswego - Tigard Water Supply | June | PWS01542:SRC-AA | 2106173-01 | Source Water | 163 | 23 | 2021-06-14 11:30 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 12:00 | USFS Steamboat Work Center | June | PWS01091:SRC-AA | 2106184-01 | Source Water | 163 | 23 | 2021-06-14 12:00 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 12:40 | Eugene Water & Electric Board | June | PWS00287:SRC-AA | 2106164-01 | Source Water | 163 | 23 | 2021-06-14 12:40 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 13:10 | City of Wilsonville | June | PWS00954:SRC-HA | 2106161-01 | Source Water | 163 | 23 | 2021-06-14 13:10 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 13:55 | Canby Utility/Veolia Water | June | PWS00157:CH-A | 2106151-01 | Source Water | 163 | 23 | 2021-06-14 13:55 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 14:00 | City of Gates | June | PWS00317:SRC-AA | 2106156-01 | Source Water | 163 | 23 | 2021-06-14 14:00 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 14:00 | Winston-Dillard Water District | June | PWS00957:SRC-AA | 2106186-01 | Source Water | 163 | 23 | 2021-06-14 14:00 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 15:30 | City of Cottage Grove | June | PWS00236:SRC-BA | 2106154-01 | Source Water | 163 | 23 | 2021-06-14 15:30 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-14 16:50 | City of Grants Pass | June | PWS00342:SRC-AA | 2106158-01 | Source Water | 163 | 23 | 2021-06-14 16:50 | 2021-06-15 15:00 | 2021-06-16 16:41 | B21F208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 08:30 | City of Creswell | June | PWS00246:SRC-BK | 2106228-01 | Source Water | 170 | 24 | 2021-06-21 08:30 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 08:30 | Tri-City JW & SA | June | PWS00549:SRC-AA | 2106260-01 | Source Water | 170 | 24 | 2021-06-21 08:30 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 08:42 | South Fork Water Board - Oregon City | June | PWS00591:SRC-AA | 2106256-01 | Source Water | 170 | 24 | 2021-06-21 08:42 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 09:00 | City of Monroe | June | PWS00540:SRC-BA | 2106237-01 | Source Water | 170 | 24 | 2021-06-21 09:00 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 09:00 | Shangri La Water District | June | PWS00835:SRC-AA | 2106255-01 | Source Water | 170 | 24 | 2021-06-21 09:00 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 09:53 | City of Jefferson | June | PWS00408:SRC-AA | 2106234-01 | Source Water | 170 | 24 | 2021-06-21 09:53 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 09:55 | Douglas County Parks - Whistler's Bend | June | PWS93944:SRC-AB | 2106269-01 | Source Water | 170 | 24 | 2021-06-21 09:55 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 10:15 | City of Estacada | June | PWS00279:SRC-AA | 2106230-01 | Source Water | 170 | 24 | 2021-06-21 10:15 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 10:30 | City of Myrtle Creek | June | PWS00550:SRC-AA | 2106238-01 | Source Water | 170 | 24 | 2021-06-21 10:30 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 10:45 | Umpqua Basin Water Association | June | PWS00719:SRC-AA | 2106261-01 | Source Water | 170 | 24 | 2021-06-21 10:45 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 10:50 | City of Lowell | June | PWS00492:SRC-BA | 2106236-01 | Source Water | 170 | 24 | 2021-06-21 10:50 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 10:55 | City of Newport | June | PWS00566:CH-A | 2106239-01 | Source Water | 170 | 24 | 2021-06-21 10:55 | 2021-06-23 10:45 | 2021-06-28 13:54 | B21F309 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 11:00 | Roberts Creek Water District | June | PWS00717:SRC-AA | 2106253-01 | Source Water | 170 | 24 | 2021-06-21 11:00 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 11:40 | City of Rogue River | June | PWS00712:SRC-BA | 2106240-01 | Source Water | 170 | 24 | 2021-06-21 11:40 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 12:05 | Lyons Mehama Water District | June | PWS00493:SRC-AA | 2106248-01 | Source Water | 170 | 24 | 2021-06-21 12:05 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 12:06 | City of Roseburg | June | PWS00720:SRC-AA | 2106242-01 | Source Water | 170 | 24 | 2021-06-21 12:06 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 12:15 | Ashland Water Department | June | PWS00047:SRC-AA | 2106226-01 | Source Water | 170 | 24 | 2021-06-21 12:15 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 13:00 | Stayton Water Supply | June | PWS00843:SRC-AA | 2106258-01 | Source Water | 170 | 24 | 2021-06-21 13:00 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 13:00 | Umpqua Ranch Co-op | June | PWS00714:SRC-AH | 2106263-01 | Source Water | 170 | 24 | 2021-06-21 13:00 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 14:25 | Medford Water Commission | June | PWS00513:SRC-BA | 2106250-01 | Source Water | 170 | 24 | 2021-06-21 14:25 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-21 15:25 | Rainbow Water District | June | PWS00839:SRC-BB | 2106252-01 | Source Water | 170 | 24 | 2021-06-21 15:25 | 2021-06-22 12:45 | 2021-06-23 13:57 | B21F293 | KMC
Microcystins ADDA, Total | 1e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-22 12:00 | City of Drain Shops | June | PWS00260:SRC-AA | 2106229-01 | Source Water | 171 | 24 | 2021-06-22 12:00 | 2021-06-23 10:45 | 2021-06-28 13:54 | B21F309 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-22 14:00 | Seaside Water Department | June | PWS00799:CH-A | 2106254-01 | Source Water | 171 | 24 | 2021-06-22 14:00 | 2021-06-23 10:45 | 2021-06-28 13:54 | B21F309 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-23 11:45 | City of Gold Hill | June | PWS00333:SRC-AA | 2106287-01 | Source Water | 172 | 24 | 2021-06-23 11:45 | 2021-06-24 14:45 | 2021-06-28 13:54 | B21F309 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-24 09:30 | Josephine Co Pks Lake Selmac 2 | June | PWS94645:SRC-AA | 2106299-01 | Source Water | 173 | 24 | 2021-06-24 09:30 | 2021-06-25 11:00 | 2021-06-28 13:54 | B21F309 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-24 09:40 | Josephine Co Pks Lake Selmac 1 | June | PWS90186:SRC-AA | 2106297-01 | Source Water | 173 | 24 | 2021-06-24 09:40 | 2021-06-25 11:00 | 2021-06-28 13:54 | B21F309 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 05:15 | Anglers Cove/SCHWC | June | PWS01483:SRC-AB | 2106304-01 | Source Water | 177 | 25 | 2021-06-28 05:15 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 05:30 | Hiland WC - Shady Cove | June | PWS01520:SRC-AA | 2106318-01 | Source Water | 177 | 25 | 2021-06-28 05:30 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 07:05 | Joint Water Commission | June | PWS00379:SRC-AA | 2106319-01 | Source Water | 177 | 25 | 2021-06-28 07:05 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 07:49 | Jackson Co Pks Emigrant Lake | June | PWS90730:CH-A | 2106331-01 | Source Water | 177 | 25 | 2021-06-28 07:49 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 08:10 | Roseburg Forest Products - Dillard | June | PWS94300:SRC-AA | 2106323-01 | Source Water | 177 | 25 | 2021-06-28 08:10 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 1e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 09:00 | City of Powers | June | PWS00672:SRC-AA | 2106313-01 | Source Water | 177 | 25 | 2021-06-28 09:00 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 09:00 | Springfield Utility Board | June | PWS00837:CH-A | 2106325-02 | Source Water | 177 | 25 | 2021-06-28 09:00 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 09:15 | Winston-Dillard Water District | June | PWS00957:SRC-AA | 2106329-01 | Source Water | 177 | 25 | 2021-06-28 09:15 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 09:20 | Springfield Utility Board | June | PWS00837:SRC-BB | 2106325-01 | Source Water | 177 | 25 | 2021-06-28 09:20 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 10:00 | Salem Public Works | June | PWS00731:SRC-AA | 2106324-01 | Source Water | 177 | 25 | 2021-06-28 10:00 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 10:40 | Camp Baker BSA | June | PWS91786:SRC-AA | 2106306-01 | Source Water | 177 | 25 | 2021-06-28 10:40 | 2021-06-30 10:37 | 2021-07-06 13:45 | B21G010 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 11:30 | Lake Oswego - Tigard Water Supply | June | PWS01542:SRC-AA | 2106320-01 | Source Water | 177 | 25 | 2021-06-28 11:30 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 11:50 | North Clackamas County Water Commission | June | PWS00580:SRC-BA | 2106321-01 | Source Water | 177 | 25 | 2021-06-28 11:50 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 3e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 12:00 | Eugene Water & Electric Board | June | PWS00287:SRC-AA | 2106316-01 | Source Water | 177 | 25 | 2021-06-28 12:00 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 3e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-28 15:30 | City of Cottage Grove | June | PWS00236:SRC-BA | 2106309-01 | Source Water | 177 | 25 | 2021-06-28 15:30 | 2021-06-29 12:00 | 2021-06-30 12:07 | B21F389 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-06-30 14:30 | City of Albany | June | PWS00012:SRC-BA | 2107064-01 | Source Water | 179 | 25 | 2021-06-30 14:30 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-01 08:30 | Clarks Branch Water Association | July | PWS00548:SRC-AA | 2107070-01 | Source Water | 180 | 25 | 2021-07-01 08:30 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-01 11:30 | USFS Steamboat Work Center | July | PWS01091:SRC-AA | 2107072-01 | Source Water | 180 | 25 | 2021-07-01 11:30 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 5e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-01 12:04 | PP&L-Toketee Village | July | PWS01012:SRC-AA | 2107071-01 | Source Water | 180 | 25 | 2021-07-01 12:04 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-01 13:00 | USFS Horseshoe Bend CG | July | PWS94179:SRC-AA | 2107074-01 | Source Water | 180 | 25 | 2021-07-01 13:00 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-01 14:15 | Canby Utility/Veolia Water | July | PWS00157:CH-A | 2107066-01 | Source Water | 180 | 25 | 2021-07-01 14:15 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-01 14:30 | USFS Tiller Ranger Station | July | PWS01092:SRC-AA | 2107073-01 | Source Water | 180 | 25 | 2021-07-01 14:30 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-02 11:00 | City of Gates | July | PWS00317:SRC-AA | 2107068-01 | Source Water | 181 | 25 | 2021-07-02 11:00 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-02 11:55 | Buell-Red Prairie Water Association | July | PWS01174:SRC-AA | 2107065-01 | Source Water | 181 | 25 | 2021-07-02 11:55 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-04 15:40 | City of Grants Pass | July | PWS00342:SRC-AA | 2107067-01 | Source Water | 183 | 26 | 2021-07-04 15:40 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 06:55 | City of Wilsonville | July | PWS00954:SRC-HA | 2107069-01 | Source Water | 185 | 26 | 2021-07-06 06:55 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 07:10 | Medford Water Commission | July | PWS00513:SRC-BA | 2107048-01 | Source Water | 185 | 26 | 2021-07-06 07:10 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 08:42 | Lyons Mehama Water District | July | PWS00493:SRC-AA | 2107047-01 | Source Water | 185 | 26 | 2021-07-06 08:42 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 08:50 | City of Creswell | July | PWS00246:SRC-BK | 2107034-01 | Source Water | 185 | 26 | 2021-07-06 08:50 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 09:00 | Umpqua Basin Water Association | July | PWS00719:SRC-AA | 2107056-01 | Source Water | 185 | 26 | 2021-07-06 09:00 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 09:13 | City of Monroe | July | PWS00540:SRC-BA | 2107040-01 | Source Water | 185 | 26 | 2021-07-06 09:13 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 09:20 | Tri-City JW & SA | July | PWS00549:SRC-AA | 2107055-01 | Source Water | 185 | 26 | 2021-07-06 09:20 | 2021-07-07 15:00 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 09:30 | City of Estacada | July | PWS00279:SRC-AA | 2107036-01 | Source Water | 185 | 26 | 2021-07-06 09:30 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 10:03 | City of Jefferson | July | PWS00408:SRC-AA | 2107038-01 | Source Water | 185 | 26 | 2021-07-06 10:03 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 10:04 | South Fork Water Board - Oregon City | July | PWS00591:SRC-AA | 2107053-01 | Source Water | 185 | 26 | 2021-07-06 10:04 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 10:28 | Douglas County Parks - Whistler's Bend | July | PWS93944:SRC-AB | 2107062-01 | Source Water | 185 | 26 | 2021-07-06 10:28 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 10:30 | City of Myrtle Creek | July | PWS00550:SRC-AA | 2107041-01 | Source Water | 185 | 26 | 2021-07-06 10:30 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 10:50 | City of Newport | July | PWS00566:CH-A | 2107042-01 | Source Water | 185 | 26 | 2021-07-06 10:50 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 11:00 | Umpqua Ranch Co-op | July | PWS00714:SRC-AH | 2107057-01 | Source Water | 185 | 26 | 2021-07-06 11:00 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 11:30 | Roberts Creek Water District | July | PWS00717:SRC-AA | 2107050-01 | Source Water | 185 | 26 | 2021-07-06 11:30 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 11:50 | City of Lowell | July | PWS00492:SRC-BA | 2107039-01 | Source Water | 185 | 26 | 2021-07-06 11:50 | 2021-07-08 11:20 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 12:10 | City of Rogue River | July | PWS00712:SRC-BA | 2107043-01 | Source Water | 185 | 26 | 2021-07-06 12:10 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 12:19 | City of Roseburg | July | PWS00720:SRC-AA | 2107044-01 | Source Water | 185 | 26 | 2021-07-06 12:19 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 12:30 | Ashland Water Department | July | PWS00047:SRC-AA | 2107033-01 | Source Water | 185 | 26 | 2021-07-06 12:30 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 12:35 | Rainbow Water District | July | PWS00839:SRC-BB | 2107049-01 | Source Water | 185 | 26 | 2021-07-06 12:35 | 2021-07-08 11:20 | 2021-07-08 21:21 | B21G062 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 12:50 | City of Drain Shops | July | PWS00260:SRC-AA | 2107035-01 | Source Water | 185 | 26 | 2021-07-06 12:50 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 13:00 | Stayton Water Supply | July | PWS00843:SRC-AA | 2107054-01 | Source Water | 185 | 26 | 2021-07-06 13:00 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 3e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 14:00 | Seaside Water Department | July | PWS00799:CH-A | 2107051-01 | Source Water | 185 | 26 | 2021-07-06 14:00 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 14:15 | Josephine Co Pks Lake Selmac 1 | July | PWS90186:SRC-AA | 2107045-01 | Source Water | 185 | 26 | 2021-07-06 14:15 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-06 14:30 | Josephine Co Pks Lake Selmac 2 | July | PWS94645:SRC-AA | 2107046-01 | Source Water | 185 | 26 | 2021-07-06 14:30 | 2021-07-07 12:30 | 2021-07-08 13:21 | B21G060 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-08 13:00 | Shangri La Water District | July | PWS00835:SRC-AA | 2107090-01 | Source Water | 187 | 26 | 2021-07-08 13:00 | 2021-07-09 11:45 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-09 15:25 | Glide Water Association | July | PWS00326:SRC-AA | 2107147-01 | Source Water | 188 | 26 | 2021-07-09 15:25 | 2021-07-14 11:45 | 2021-07-19 14:35 | B21G123 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 05:15 | Anglers Cove/SCHWC | July | PWS01483:SRC-AB | 2107100-01 | Source Water | 191 | 27 | 2021-07-12 05:15 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 05:30 | Hiland WC - Shady Cove | July | PWS01520:SRC-AA | 2107121-01 | Source Water | 191 | 27 | 2021-07-12 05:30 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 07:25 | Springfield Utility Board | July | PWS00837:CH-A | 2107134-02 | Source Water | 191 | 27 | 2021-07-12 07:25 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 07:36 | City of Pendleton | July | PWS00613:SRC-AM | 2107112-01 | Source Water | 191 | 27 | 2021-07-12 07:36 | 2021-07-14 11:45 | 2021-07-19 14:35 | B21G123 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 07:47 | Glide Water Association | July | PWS00326:SRC-AA | 2107119-01 | Source Water | 191 | 27 | 2021-07-12 07:47 | 2021-07-14 11:45 | 2021-07-23 16:45 | B21G238 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 07:55 | Springfield Utility Board | July | PWS00837:SRC-BB | 2107134-01 | Source Water | 191 | 27 | 2021-07-12 07:55 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 08:01 | Roseburg Forest Products - Dillard | July | PWS94300:SRC-AA | 2107131-01 | Source Water | 191 | 27 | 2021-07-12 08:01 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 08:10 | City of Gold Hill | July | PWS00333:SRC-AA | 2107142-01 | Source Water | 191 | 27 | 2021-07-12 08:10 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 08:10 | USFS Tiller Ranger Station | July | PWS01092:SRC-AA | 2107138-01 | Source Water | 191 | 27 | 2021-07-12 08:10 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 08:28 | Joint Water Commission | July | PWS00379:SRC-AA | 2107124-01 | Source Water | 191 | 27 | 2021-07-12 08:28 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 08:55 | Camp Baker BSA | July | PWS91786:SRC-AA | 2107102-01 | Source Water | 191 | 27 | 2021-07-12 08:55 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 09:06 | Clarks Branch Water Association | July | PWS00548:SRC-AA | 2107115-01 | Source Water | 191 | 27 | 2021-07-12 09:06 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 09:20 | City of Albany | July | PWS00012:SRC-BA | 2107106-01 | Source Water | 191 | 27 | 2021-07-12 09:20 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 09:21 | City of Wilsonville | July | PWS00954:SRC-HA | 2107114-01 | Source Water | 191 | 27 | 2021-07-12 09:21 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 09:30 | North Clackamas County Water Commission | July | PWS00580:SRC-BA | 2107128-01 | Source Water | 191 | 27 | 2021-07-12 09:30 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 09:30 | Winston-Dillard Water District | July | PWS00957:SRC-AA | 2107139-01 | Source Water | 191 | 27 | 2021-07-12 09:30 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 10:00 | Salem Public Works | July | PWS00731:SRC-AA | 2107133-01 | Source Water | 191 | 27 | 2021-07-12 10:00 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 11:00 | Buell-Red Prairie Water Association | July | PWS01174:SRC-AA | 2107101-01 | Source Water | 191 | 27 | 2021-07-12 11:00 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 11:30 | USFS Steamboat Work Center | July | PWS01091:SRC-AA | 2107137-01 | Source Water | 191 | 27 | 2021-07-12 11:30 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 11:50 | City of Powers | July | PWS00672:SRC-AA | 2107113-01 | Source Water | 191 | 27 | 2021-07-12 11:50 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 11:50 | Lake Oswego - Tigard Water Supply | July | PWS01542:SRC-AA | 2107126-01 | Source Water | 191 | 27 | 2021-07-12 11:50 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 11:50 | PP&L-Toketee Village | July | PWS01012:SRC-AA | 2107130-01 | Source Water | 191 | 27 | 2021-07-12 11:50 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 12:02 | Eugene Water & Electric Board | July | PWS00287:SRC-AA | 2107117-01 | Source Water | 191 | 27 | 2021-07-12 12:02 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 13:30 | Canby Utility/Veolia Water | July | PWS00157:CH-A | 2107104-01 | Source Water | 191 | 27 | 2021-07-12 13:30 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 13:30 | City of Grants Pass | July | PWS00342:SRC-AA | 2107111-01 | Source Water | 191 | 27 | 2021-07-12 13:30 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-12 13:45 | City of Gates | July | PWS00317:SRC-AA | 2107109-01 | Source Water | 191 | 27 | 2021-07-12 13:45 | 2021-07-13 12:00 | 2021-07-14 17:07 | B21G112 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-13 12:30 | City of Cottage Grove | July | PWS00236:SRC-BA | 2107107-01 | Source Water | 192 | 27 | 2021-07-13 12:30 | 2021-07-14 11:45 | 2021-07-19 14:35 | B21G123 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 07:10 | Medford Water Commission | July | PWS00513:SRC-BA | 2107205-01 | Source Water | 198 | 28 | 2021-07-19 07:10 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 08:50 | South Fork Water Board - Oregon City | July | PWS00591:SRC-AA | 2107211-01 | Source Water | 198 | 28 | 2021-07-19 08:50 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 09:00 | City of Creswell | July | PWS00246:SRC-BK | 2107182-01 | Source Water | 198 | 28 | 2021-07-19 09:00 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 09:00 | Shangri La Water District | July | PWS00835:SRC-AA | 2107209-01 | Source Water | 198 | 28 | 2021-07-19 09:00 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 09:00 | Tri-City JW & SA | July | PWS00549:SRC-AA | 2107214-01 | Source Water | 198 | 28 | 2021-07-19 09:00 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 09:30 | City of Gold Hill | July | PWS00333:SRC-AA | 2107187-01 | Source Water | 198 | 28 | 2021-07-19 09:30 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 09:30 | Umpqua Ranch Co-op | July | PWS00714:SRC-AH | 2107218-01 | Source Water | 198 | 28 | 2021-07-19 09:30 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 09:45 | City of Monroe | July | PWS00540:SRC-BA | 2107191-01 | Source Water | 198 | 28 | 2021-07-19 09:45 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 09:55 | City of Jefferson | July | PWS00408:SRC-AA | 2107189-01 | Source Water | 198 | 28 | 2021-07-19 09:55 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 10:10 | City of Estacada | July | PWS00279:SRC-AA | 2107185-01 | Source Water | 198 | 28 | 2021-07-19 10:10 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 10:30 | City of Myrtle Creek | July | PWS00550:SRC-AA | 2107192-01 | Source Water | 198 | 28 | 2021-07-19 10:30 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.06 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 10:30 | City of Newport | July | PWS00566:CH-A | 2107193-01 | Source Water | 198 | 28 | 2021-07-19 10:30 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 10:45 | Umpqua Basin Water Association | July | PWS00719:SRC-AA | 2107216-01 | Source Water | 198 | 28 | 2021-07-19 10:45 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 11:00 | City of Lowell | July | PWS00492:SRC-BA | 2107190-01 | Source Water | 198 | 28 | 2021-07-19 11:00 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 11:00 | Roberts Creek Water District | July | PWS00717:SRC-AA | 2107207-01 | Source Water | 198 | 28 | 2021-07-19 11:00 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 11:05 | City of Drain Shops | July | PWS00260:SRC-AA | 2107183-01 | Source Water | 198 | 28 | 2021-07-19 11:05 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.07 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-07-19 12:00 | Josephine Co Pks Lake Selmac 1 | July | PWS90186:SRC-AA | 2107199-01 | Source Water | 198 | 28 | 2021-07-19 12:00 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 12:16 | City of Roseburg | July | PWS00720:SRC-AA | 2107197-01 | Source Water | 198 | 28 | 2021-07-19 12:16 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 12:30 | Ashland Water Department | July | PWS00047:SRC-AA | 2107181-01 | Source Water | 198 | 28 | 2021-07-19 12:30 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 12:30 | City of Rogue River | July | PWS00712:SRC-BA | 2107195-01 | Source Water | 198 | 28 | 2021-07-19 12:30 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 12:47 | Rainbow Water District | July | PWS00839:SRC-BB | 2107206-01 | Source Water | 198 | 28 | 2021-07-19 12:47 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.12 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2021-07-19 13:00 | Josephine Co Pks Lake Selmac 2 | July | PWS94645:SRC-AA | 2107201-01 | Source Water | 198 | 28 | 2021-07-19 13:00 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 13:00 | Stayton Water Supply | July | PWS00843:SRC-AA | 2107213-01 | Source Water | 198 | 28 | 2021-07-19 13:00 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-19 14:00 | Seaside Water Department | July | PWS00799:CH-A | 2107208-01 | Source Water | 198 | 28 | 2021-07-19 14:00 | 2021-07-20 12:15 | 2021-07-21 15:08 | B21G208 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-20 08:23 | Lyons Mehama Water District | July | PWS00493:SRC-AA | 2107203-01 | Source Water | 199 | 28 | 2021-07-20 08:23 | 2021-07-21 16:30 | 2021-07-23 16:45 | B21G238 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-21 10:12 | Douglas County Parks - Whistler's Bend | July | PWS93944:SRC-AB | 2107229-01 | Source Water | 200 | 28 | 2021-07-21 10:12 | 2021-07-22 12:00 | 2021-07-23 16:45 | B21G238 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 | North Clackamas County Water Commission | July | PWS00580:SRC-BA | 2107263-01 | Source Water | 205 | 29 | 2021-07-26 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 05:15 | Anglers Cove/SCHWC | July | PWS01483:SRC-AB | 2107246-01 | Source Water | 205 | 29 | 2021-07-26 05:15 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 05:30 | Hiland WC - Shady Cove | July | PWS01520:SRC-AA | 2107260-01 | Source Water | 205 | 29 | 2021-07-26 05:30 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 08:10 | Roseburg Forest Products - Dillard | July | PWS94300:SRC-AA | 2107265-01 | Source Water | 205 | 29 | 2021-07-26 08:10 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 08:18 | Joint Water Commission | July | PWS00379:SRC-AA | 2107261-01 | Source Water | 205 | 29 | 2021-07-26 08:18 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 08:20 | Camp Baker BSA | July | PWS91786:SRC-AA | 2107248-01 | Source Water | 205 | 29 | 2021-07-26 08:20 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 08:20 | Springfield Utility Board | July | PWS00837:CH-A | 2107267-02 | Source Water | 205 | 29 | 2021-07-26 08:20 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 08:22 | Salem Public Works | July | PWS00731:SRC-AA | 2107266-01 | Source Water | 205 | 29 | 2021-07-26 08:22 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 08:55 | Springfield Utility Board | July | PWS00837:SRC-BB | 2107267-01 | Source Water | 205 | 29 | 2021-07-26 08:55 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 09:00 | City of Albany | July | PWS00012:SRC-BA | 2107250-01 | Source Water | 205 | 29 | 2021-07-26 09:00 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 09:30 | Winston-Dillard Water District | July | PWS00957:SRC-AA | 2107271-01 | Source Water | 205 | 29 | 2021-07-26 09:30 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 09:35 | City of Wilsonville | July | PWS00954:SRC-HA | 2107256-01 | Source Water | 205 | 29 | 2021-07-26 09:35 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 09:40 | USFS Tiller Ranger Station | July | PWS01092:SRC-AA | 2107270-01 | Source Water | 205 | 29 | 2021-07-26 09:40 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 09:45 | Milo Academy | July | PWS00250:SRC-AB | 2107274-01 | Source Water | 205 | 29 | 2021-07-26 09:45 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 10:30 | Eugene Water & Electric Board | July | PWS00287:SRC-AA | 2107258-01 | Source Water | 205 | 29 | 2021-07-26 10:30 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 10:35 | Glide Water Association | July | PWS00326:SRC-AA | 2107259-01 | Source Water | 205 | 29 | 2021-07-26 10:35 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 10:45 | Lake Oswego - Tigard Water Supply | July | PWS01542:SRC-AA | 2107262-01 | Source Water | 205 | 29 | 2021-07-26 10:45 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.06 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 11:00 | Buell-Red Prairie Water Association | July | PWS01174:SRC-AA | 2107247-01 | Source Water | 205 | 29 | 2021-07-26 11:00 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 11:12 | Clarks Branch Water Association | July | PWS00548:SRC-AA | 2107257-01 | Source Water | 205 | 29 | 2021-07-26 11:12 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 11:36 | City of Pendleton | July | PWS00613:SRC-AM | 2107254-01 | Source Water | 205 | 29 | 2021-07-26 11:36 | 2021-07-28 13:00 | 2021-08-02 14:28 | B21G302 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 12:15 | City of Cottage Grove | July | PWS00236:SRC-BA | 2107251-01 | Source Water | 205 | 29 | 2021-07-26 12:15 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 13:00 | Canby Utility/Veolia Water | July | PWS00157:CH-A | 2107249-01 | Source Water | 205 | 29 | 2021-07-26 13:00 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 13:08 | PP&L-Toketee Village | July | PWS01012:SRC-AA | 2107264-01 | Source Water | 205 | 29 | 2021-07-26 13:08 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 13:30 | City of Grants Pass | July | PWS00342:SRC-AA | 2107253-01 | Source Water | 205 | 29 | 2021-07-26 13:30 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 13:44 | USFS Steamboat Work Center | July | PWS01091:SRC-AA | 2107269-01 | Source Water | 205 | 29 | 2021-07-26 13:44 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-07-26 14:00 | City of Gates | July | PWS00317:SRC-AA | 2107252-01 | Source Water | 205 | 29 | 2021-07-26 14:00 | 2021-07-27 12:30 | 2021-07-28 17:50 | B21G284 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 08:30 | City of Creswell | August | PWS00246:SRC-BK | 2108006-01 | Source Water | 211 | 30 | 2021-08-02 08:30 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 09:15 | City of Myrtle Creek | August | PWS00550:SRC-AA | 2108013-01 | Source Water | 211 | 30 | 2021-08-02 09:15 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 09:30 | Shangri La Water District | August | PWS00835:SRC-AA | 2108024-01 | Source Water | 211 | 30 | 2021-08-02 09:30 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 09:38 | City of Monroe | August | PWS00540:SRC-BA | 2108012-01 | Source Water | 211 | 30 | 2021-08-02 09:38 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 09:48 | Lyons Mehama Water District | August | PWS00493:SRC-AA | 2108019-01 | Source Water | 211 | 30 | 2021-08-02 09:48 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 10:30 | City of Estacada | August | PWS00279:SRC-AA | 2108008-01 | Source Water | 211 | 30 | 2021-08-02 10:30 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 10:30 | City of Lowell | August | PWS00492:SRC-BA | 2108011-01 | Source Water | 211 | 30 | 2021-08-02 10:30 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.09 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-08-02 10:30 | City of Newport | August | PWS00566:CH-A | 2108014-01 | Source Water | 211 | 30 | 2021-08-02 10:30 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 10:32 | City of Jefferson | August | PWS00408:SRC-AA | 2108010-01 | Source Water | 211 | 30 | 2021-08-02 10:32 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 11:00 | City of Rogue River | August | PWS00712:SRC-BA | 2108015-01 | Source Water | 211 | 30 | 2021-08-02 11:00 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 11:15 | Umpqua Basin Water Association | August | PWS00719:SRC-AA | 2108028-01 | Source Water | 211 | 30 | 2021-08-02 11:15 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 11:30 | Roberts Creek Water District | August | PWS00717:SRC-AA | 2108022-01 | Source Water | 211 | 30 | 2021-08-02 11:30 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 11:30 | Tri-City JW & SA | August | PWS00549:SRC-AA | 2108027-01 | Source Water | 211 | 30 | 2021-08-02 11:30 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 11:50 | Medford Water Commission | August | PWS00513:SRC-BA | 2108020-01 | Source Water | 211 | 30 | 2021-08-02 11:50 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 12:22 | City of Roseburg | August | PWS00720:SRC-AA | 2108016-01 | Source Water | 211 | 30 | 2021-08-02 12:22 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 12:30 | Ashland Water Department | August | PWS00047:SRC-AA | 2108005-01 | Source Water | 211 | 30 | 2021-08-02 12:30 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.21 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2021-08-02 12:30 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2108018-01 | Source Water | 211 | 30 | 2021-08-02 12:30 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.11 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-08-02 13:00 | Josephine Co Pks Lake Selmac 1 | August | PWS90186:SRC-AA | 2108017-01 | Source Water | 211 | 30 | 2021-08-02 13:00 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 14:00 | Seaside Water Department | August | PWS00799:CH-A | 2108023-01 | Source Water | 211 | 30 | 2021-08-02 14:00 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-02 18:00 | Stayton Water Supply | August | PWS00843:SRC-AA | 2108026-01 | Source Water | 211 | 30 | 2021-08-02 18:00 | 2021-08-03 11:30 | 2021-08-04 17:08 | B21H034 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-03 12:10 | South Fork Water Board - Oregon City | August | PWS00591:SRC-AA | 2108040-01 | Source Water | 212 | 30 | 2021-08-03 12:10 | 2021-08-04 11:45 | 2021-08-05 20:33 | B21H051 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-03 12:45 | City of Drain Shops | August | PWS00260:SRC-AA | 2108007-01 | Source Water | 212 | 30 | 2021-08-03 12:45 | 2021-08-04 11:45 | 2021-08-05 20:33 | B21H051 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-03 13:10 | Rainbow Water District | August | PWS00839:SRC-BB | 2108041-01 | Source Water | 212 | 30 | 2021-08-03 13:10 | 2021-08-04 11:45 | 2021-08-05 20:33 | B21H051 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-03 16:55 | City of Gold Hill | August | PWS00333:SRC-AA | 2108009-01 | Source Water | 212 | 30 | 2021-08-03 16:55 | 2021-08-04 11:45 | 2021-08-05 20:33 | B21H051 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 05:15 | Anglers Cove/SCHWC | August | PWS01483:SRC-AB | 2108088-01 | Source Water | 218 | 31 | 2021-08-09 05:15 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 05:30 | Hiland WC - Shady Cove | August | PWS01520:SRC-AA | 2108109-01 | Source Water | 218 | 31 | 2021-08-09 05:30 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 08:25 | Springfield Utility Board | August | PWS00837:CH-A | 2108122-02 | Source Water | 218 | 31 | 2021-08-09 08:25 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 08:27 | Joint Water Commission | August | PWS00379:SRC-AA | 2108112-01 | Source Water | 218 | 31 | 2021-08-09 08:27 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 08:27 | Salem Public Works | August | PWS00731:SRC-AA | 2108121-01 | Source Water | 218 | 31 | 2021-08-09 08:27 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 08:45 | Roseburg Forest Products - Dillard | August | PWS94300:SRC-AA | 2108119-01 | Source Water | 218 | 31 | 2021-08-09 08:45 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 08:50 | Winston-Dillard Water District | August | PWS00957:SRC-AA | 2108127-01 | Source Water | 218 | 31 | 2021-08-09 08:50 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 09:00 | Glide Water Association | August | PWS00326:SRC-AA | 2108107-01 | Source Water | 218 | 31 | 2021-08-09 09:00 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 09:40 | Camp Baker BSA | August | PWS91786:SRC-AA | 2108090-01 | Source Water | 218 | 31 | 2021-08-09 09:40 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 09:43 | USFS Tiller Ranger Station | August | PWS01092:SRC-AA | 2108126-01 | Source Water | 218 | 31 | 2021-08-09 09:43 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 09:55 | Springfield Utility Board | August | PWS00837:SRC-BB | 2108122-01 | Source Water | 218 | 31 | 2021-08-09 09:55 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 10:00 | City of Powers | August | PWS00672:SRC-AA | 2108101-01 | Source Water | 218 | 31 | 2021-08-09 10:00 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 10:10 | City of Wilsonville | August | PWS00954:SRC-HA | 2108102-01 | Source Water | 218 | 31 | 2021-08-09 10:10 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 10:10 | Lake Oswego - Tigard Water Supply | August | PWS01542:SRC-AA | 2108114-01 | Source Water | 218 | 31 | 2021-08-09 10:10 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 10:20 | North Clackamas County Water Commission | August | PWS00580:SRC-BA | 2108116-01 | Source Water | 218 | 31 | 2021-08-09 10:20 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 10:39 | Clarks Branch Water Association | August | PWS00548:SRC-AA | 2108103-01 | Source Water | 218 | 31 | 2021-08-09 10:39 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.34 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-08-09 11:00 | Buell-Red Prairie Water Association | August | PWS01174:SRC-AA | 2108089-01 | Source Water | 218 | 31 | 2021-08-09 11:00 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 11:00 | Eugene Water & Electric Board | August | PWS00287:SRC-AA | 2108105-01 | Source Water | 218 | 31 | 2021-08-09 11:00 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 11:52 | City of Pendleton | August | PWS00613:SRC-AM | 2108100-01 | Source Water | 218 | 31 | 2021-08-09 11:52 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 12:00 | City of Cottage Grove | August | PWS00236:SRC-BA | 2108095-01 | Source Water | 218 | 31 | 2021-08-09 12:00 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 12:07 | USFS Steamboat Work Center | August | PWS01091:SRC-AA | 2108125-01 | Source Water | 218 | 31 | 2021-08-09 12:07 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 13:00 | PP&L-Toketee Village | August | PWS01012:SRC-AA | 2108118-01 | Source Water | 218 | 31 | 2021-08-09 13:00 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 13:22 | City of Grants Pass | August | PWS00342:SRC-AA | 2108099-01 | Source Water | 218 | 31 | 2021-08-09 13:22 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 14:15 | City of Gates | August | PWS00317:SRC-AA | 2108097-01 | Source Water | 218 | 31 | 2021-08-09 14:15 | 2021-08-11 11:00 | 2021-08-13 17:21 | B21H107 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 14:15 | Milo Academy | August | PWS00250:SRC-AB | 2108131-01 | Source Water | 218 | 31 | 2021-08-09 14:15 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-09 14:20 | Canby Utility/Veolia Water | August | PWS00157:CH-A | 2108092-01 | Source Water | 218 | 31 | 2021-08-09 14:20 | 2021-08-10 12:45 | 2021-08-11 21:18 | B21H105 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-10 06:00 | City of Albany | August | PWS00012:SRC-BA | 2108094-01 | Source Water | 219 | 31 | 2021-08-10 06:00 | 2021-08-11 11:00 | 2021-08-13 17:21 | B21H107 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-11 09:33 | Douglas County Parks - Whistler's Bend | August | PWS93944:SRC-AB | 2108146-01 | Source Water | 220 | 31 | 2021-08-11 09:33 | 2021-08-12 12:00 | 2021-08-13 17:21 | B21H107 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 07:30 | Medford Water Commission | August | PWS00513:SRC-BA | 2108181-01 | Source Water | 225 | 32 | 2021-08-16 07:30 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 08:30 | City of Creswell | August | PWS00246:SRC-BK | 2108157-01 | Source Water | 225 | 32 | 2021-08-16 08:30 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 09:00 | Shangri La Water District | August | PWS00835:SRC-AA | 2108185-01 | Source Water | 225 | 32 | 2021-08-16 09:00 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 09:44 | South Fork Water Board - Oregon City | August | PWS00591:SRC-AA | 2108187-01 | Source Water | 225 | 32 | 2021-08-16 09:44 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 09:45 | Lyons Mehama Water District | August | PWS00493:SRC-AA | 2108179-01 | Source Water | 225 | 32 | 2021-08-16 09:45 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 10:00 | City of Monroe | August | PWS00540:SRC-BA | 2108166-01 | Source Water | 225 | 32 | 2021-08-16 10:00 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 10:05 | City of Lowell | August | PWS00492:SRC-BA | 2108165-01 | Source Water | 225 | 32 | 2021-08-16 10:05 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 10:30 | City of Myrtle Creek | August | PWS00550:SRC-AA | 2108167-01 | Source Water | 225 | 32 | 2021-08-16 10:30 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 10:30 | Umpqua Ranch Co-op | August | PWS00714:SRC-AH | 2108194-01 | Source Water | 225 | 32 | 2021-08-16 10:30 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.08 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-08-16 11:00 | City of Newport | August | PWS00566:CH-A | 2108168-01 | Source Water | 225 | 32 | 2021-08-16 11:00 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 11:10 | City of Rogue River | August | PWS00712:SRC-BA | 2108170-01 | Source Water | 225 | 32 | 2021-08-16 11:10 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 11:16 | City of Gold Hill | August | PWS00333:SRC-AA | 2108162-01 | Source Water | 225 | 32 | 2021-08-16 11:16 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.91 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-08-16 11:30 | Buell-Red Prairie Water Association | August | PWS01174:SRC-AA | 2108196-01 | Source Water | 225 | 32 | 2021-08-16 11:30 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.26 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2021-08-16 11:30 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2108177-01 | Source Water | 225 | 32 | 2021-08-16 11:30 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 11:30 | Roberts Creek Water District | August | PWS00717:SRC-AA | 2108183-01 | Source Water | 225 | 32 | 2021-08-16 11:30 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.19 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2021-08-16 12:00 | Josephine Co Pks Lake Selmac 1 | August | PWS90186:SRC-AA | 2108175-01 | Source Water | 225 | 32 | 2021-08-16 12:00 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 12:00 | Tri-City JW & SA | August | PWS00549:SRC-AA | 2108190-01 | Source Water | 225 | 32 | 2021-08-16 12:00 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 12:15 | City of Drain Shops | August | PWS00260:SRC-AA | 2108158-01 | Source Water | 225 | 32 | 2021-08-16 12:15 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 12:16 | City of Roseburg | August | PWS00720:SRC-AA | 2108172-01 | Source Water | 225 | 32 | 2021-08-16 12:16 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 12:30 | Ashland Water Department | August | PWS00047:SRC-AA | 2108156-01 | Source Water | 225 | 32 | 2021-08-16 12:30 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 13:00 | Stayton Water Supply | August | PWS00843:SRC-AA | 2108189-01 | Source Water | 225 | 32 | 2021-08-16 13:00 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 13:10 | Umpqua Basin Water Association | August | PWS00719:SRC-AA | 2108192-01 | Source Water | 225 | 32 | 2021-08-16 13:10 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 13:30 | Seaside Water Department | August | PWS00799:CH-A | 2108184-01 | Source Water | 225 | 32 | 2021-08-16 13:30 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-16 13:46 | Rainbow Water District | August | PWS00839:SRC-BB | 2108182-01 | Source Water | 225 | 32 | 2021-08-16 13:46 | 2021-08-17 12:45 | 2021-08-18 16:22 | B21H142 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-18 10:17 | Douglas County Parks - Whistler's Bend | August | PWS93944:SRC-AB | 2108173-01 | Source Water | 227 | 32 | 2021-08-18 10:17 | 2021-08-19 12:30 | 2021-08-23 14:47 | B21H249 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-18 10:30 | City of Jefferson | August | PWS00408:SRC-AA | 2108222-01 | Source Water | 227 | 32 | 2021-08-18 10:30 | 2021-08-19 12:30 | 2021-08-23 14:47 | B21H249 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-18 12:15 | City of Estacada | August | PWS00279:SRC-AA | 2108220-01 | Source Water | 227 | 32 | 2021-08-18 12:15 | 2021-08-19 12:30 | 2021-08-23 14:47 | B21H249 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 06:15 | Anglers Cove/SCHWC | August | PWS01483:SRC-AB | 2108233-01 | Source Water | 232 | 33 | 2021-08-23 06:15 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 06:30 | Hiland WC - Shady Cove | August | PWS01520:SRC-AA | 2108248-01 | Source Water | 232 | 33 | 2021-08-23 06:30 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 07:25 | City of Albany | August | PWS00012:SRC-BA | 2108237-01 | Source Water | 232 | 33 | 2021-08-23 07:25 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 07:30 | City of Pendleton | August | PWS00613:SRC-AM | 2108242-01 | Source Water | 232 | 33 | 2021-08-23 07:30 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 08:00 | Glide Water Association | August | PWS00326:SRC-AA | 2108247-01 | Source Water | 232 | 33 | 2021-08-23 08:00 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 08:00 | Winston-Dillard Water District | August | PWS00957:SRC-AA | 2108260-01 | Source Water | 232 | 33 | 2021-08-23 08:00 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 08:10 | Joint Water Commission | August | PWS00379:SRC-AA | 2108249-01 | Source Water | 232 | 33 | 2021-08-23 08:10 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 08:25 | Salem Public Works | August | PWS00731:SRC-AA | 2108255-01 | Source Water | 232 | 33 | 2021-08-23 08:25 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 09:10 | USFS Tiller Ranger Station | August | PWS01092:SRC-AA | 2108259-01 | Source Water | 232 | 33 | 2021-08-23 09:10 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 09:30 | City of Powers | August | PWS00672:SRC-AA | 2108243-01 | Source Water | 232 | 33 | 2021-08-23 09:30 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 09:30 | Springfield Utility Board | August | PWS00837:CH-A | 2108256-02 | Source Water | 232 | 33 | 2021-08-23 09:30 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 10:00 | Lake Oswego - Tigard Water Supply | August | PWS01542:SRC-AA | 2108250-01 | Source Water | 232 | 33 | 2021-08-23 10:00 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 10:05 | Springfield Utility Board | August | PWS00837:SRC-BB | 2108256-01 | Source Water | 232 | 33 | 2021-08-23 10:05 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 10:19 | Clarks Branch Water Association | August | PWS00548:SRC-AA | 2108245-01 | Source Water | 232 | 33 | 2021-08-23 10:19 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 10:40 | Camp Baker BSA | August | PWS91786:SRC-AA | 2108235-01 | Source Water | 232 | 33 | 2021-08-23 10:40 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 2.17 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-08-23 11:00 | Buell-Red Prairie Water Association | August | PWS01174:SRC-AA | 2108234-01RE1 | Source Water | 232 | 33 | 2021-08-23 11:00 | 2021-08-24 12:30 | 2021-08-27 14:18 | B21H301 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 11:20 | City of Wilsonville | August | PWS00954:SRC-HA | 2108244-01 | Source Water | 232 | 33 | 2021-08-23 11:20 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 11:30 | Eugene Water & Electric Board | August | PWS00287:SRC-AA | 2108246-01 | Source Water | 232 | 33 | 2021-08-23 11:30 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 11:40 | Roseburg Forest Products - Dillard | August | PWS94300:SRC-AA | 2108254-01 | Source Water | 232 | 33 | 2021-08-23 11:40 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.43 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-08-23 12:00 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2108261-01 | Source Water | 232 | 33 | 2021-08-23 12:00 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 12:11 | USFS Steamboat Work Center | August | PWS01091:SRC-AA | 2108258-01 | Source Water | 232 | 33 | 2021-08-23 12:11 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 12:30 | City of Cottage Grove | August | PWS00236:SRC-BA | 2108238-01 | Source Water | 232 | 33 | 2021-08-23 12:30 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 12:39 | USFS Horseshoe Bend CG | August | PWS94179:SRC-AA | 2108257-01 | Source Water | 232 | 33 | 2021-08-23 12:39 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 13:07 | PP&L-Toketee Village | August | PWS01012:SRC-AA | 2108253-01 | Source Water | 232 | 33 | 2021-08-23 13:07 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 13:30 | City of Gates | August | PWS00317:SRC-AA | 2108239-01 | Source Water | 232 | 33 | 2021-08-23 13:30 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 13:30 | City of Grants Pass | August | PWS00342:SRC-AA | 2108241-01 | Source Water | 232 | 33 | 2021-08-23 13:30 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-23 15:00 | Canby Utility/Veolia Water | August | PWS00157:CH-A | 2108236-01 | Source Water | 232 | 33 | 2021-08-23 15:00 | 2021-08-24 12:30 | 2021-08-25 13:47 | B21H285 | KMC
Microcystins ADDA, Total | 0.56 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-08-24 10:00 | Milo Academy | August | PWS00250:SRC-AB | 2108251-01 | Source Water | 233 | 33 | 2021-08-24 10:00 | 2021-08-26 10:45 | 2021-08-27 14:18 | B21H301 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-24 12:49 | North Clackamas County Water Commission | August | PWS00580:SRC-BA | 2108264-01 | Source Water | 233 | 33 | 2021-08-24 12:49 | 2021-08-25 16:00 | 2021-08-27 14:18 | B21H301 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 08:00 | Milo Academy | August | PWS00250:SRC-AB | 2108312-01 | Source Water | 239 | 34 | 2021-08-30 08:00 | 2021-08-31 13:45 | 2021-09-01 18:58 | B21H364 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 08:16 | Lyons Mehama Water District | August | PWS00493:SRC-AA | 2108296-01 | Source Water | 239 | 34 | 2021-08-30 08:16 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 08:20 | Medford Water Commission | August | PWS00513:SRC-BA | 2108297-01 | Source Water | 239 | 34 | 2021-08-30 08:20 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 7e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 08:36 | City of Jefferson | August | PWS00408:SRC-AA | 2108287-01 | Source Water | 239 | 34 | 2021-08-30 08:36 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.14 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2021-08-30 08:45 | City of Creswell | August | PWS00246:SRC-BK | 2108283-01 | Source Water | 239 | 34 | 2021-08-30 08:45 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 09:00 | Shangri La Water District | August | PWS00835:SRC-AA | 2108301-01 | Source Water | 239 | 34 | 2021-08-30 09:00 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 09:30 | Umpqua Ranch Co-op | August | PWS00714:SRC-AH | 2108306-01 | Source Water | 239 | 34 | 2021-08-30 09:30 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 09:45 | City of Monroe | August | PWS00540:SRC-BA | 2108289-01 | Source Water | 239 | 34 | 2021-08-30 09:45 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 09:45 | South Fork Water Board - Oregon City | August | PWS00591:SRC-AA | 2108302-01 | Source Water | 239 | 34 | 2021-08-30 09:45 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 3.28 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-08-30 10:00 | Buell-Red Prairie Water Association | August | PWS01174:SRC-AA | 2108313-01 | Source Water | 239 | 34 | 2021-08-30 10:00 | 2021-08-31 13:45 | 2021-09-01 18:58 | B21H364 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 10:10 | City of Gold Hill | August | PWS00333:SRC-AA | 2108286-01 | Source Water | 239 | 34 | 2021-08-30 10:10 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 10:30 | City of Estacada | August | PWS00279:SRC-AA | 2108285-01 | Source Water | 239 | 34 | 2021-08-30 10:30 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 10:30 | City of Myrtle Creek | August | PWS00550:SRC-AA | 2108290-01 | Source Water | 239 | 34 | 2021-08-30 10:30 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 10:40 | City of Rogue River | August | PWS00712:SRC-BA | 2108292-01 | Source Water | 239 | 34 | 2021-08-30 10:40 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 11:00 | City of Newport | August | PWS00566:CH-A | 2108291-01 | Source Water | 239 | 34 | 2021-08-30 11:00 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 11:20 | City of Drain Shops | August | PWS00260:SRC-AA | 2108284-01 | Source Water | 239 | 34 | 2021-08-30 11:20 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 11:20 | City of Lowell | August | PWS00492:SRC-BA | 2108288-01 | Source Water | 239 | 34 | 2021-08-30 11:20 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 12:00 | Tri-City JW & SA | August | PWS00549:SRC-AA | 2108304-01 | Source Water | 239 | 34 | 2021-08-30 12:00 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 12:10 | Umpqua Basin Water Association | August | PWS00719:SRC-AA | 2108305-01 | Source Water | 239 | 34 | 2021-08-30 12:10 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 12:20 | Ashland Water Department | August | PWS00047:SRC-AA | 2108282-01 | Source Water | 239 | 34 | 2021-08-30 12:20 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.38 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-08-30 12:30 | Josephine Co Pks Lake Selmac 1 | August | PWS90186:SRC-AA | 2108294-01 | Source Water | 239 | 34 | 2021-08-30 12:30 | 2021-08-31 13:45 | 2021-09-01 18:58 | B21H364 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 12:30 | Roberts Creek Water District | August | PWS00717:SRC-AA | 2108299-01 | Source Water | 239 | 34 | 2021-08-30 12:30 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 12:32 | City of Roseburg | August | PWS00720:SRC-AA | 2108293-01 | Source Water | 239 | 34 | 2021-08-30 12:32 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 12:42 | Douglas County Parks - Whistler's Bend | August | PWS93944:SRC-AB | 2108281-01 | Source Water | 239 | 34 | 2021-08-30 12:42 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.44 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-08-30 12:45 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2108295-01 | Source Water | 239 | 34 | 2021-08-30 12:45 | 2021-08-31 13:45 | 2021-09-01 18:58 | B21H364 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 13:00 | Stayton Water Supply | August | PWS00843:SRC-AA | 2108303-01 | Source Water | 239 | 34 | 2021-08-30 13:00 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 14:05 | Rainbow Water District | August | PWS00839:SRC-BB | 2108298-01 | Source Water | 239 | 34 | 2021-08-30 14:05 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-08-30 14:30 | Seaside Water Department | August | PWS00799:CH-A | 2108300-01 | Source Water | 239 | 34 | 2021-08-30 14:30 | 2021-08-31 14:30 | 2021-09-02 13:51 | B21H366 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 05:15 | Anglers Cove/SCHWC | September | PWS01483:SRC-AB | 2109035-01 | Source Water | 246 | 35 | 2021-09-07 05:15 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 05:30 | Hiland WC - Shady Cove | September | PWS01520:SRC-AA | 2109056-01 | Source Water | 246 | 35 | 2021-09-07 05:30 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 07:34 | City of Pendleton | September | PWS00613:SRC-AM | 2109048-01 | Source Water | 246 | 35 | 2021-09-07 07:34 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 08:00 | Roseburg Forest Products - Dillard | September | PWS94300:SRC-AA | 2109066-01 | Source Water | 246 | 35 | 2021-09-07 08:00 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 08:10 | Salem Public Works | September | PWS00731:SRC-AA | 2109067-01 | Source Water | 246 | 35 | 2021-09-07 08:10 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 08:14 | Joint Water Commission | September | PWS00379:SRC-AA | 2109058-01 | Source Water | 246 | 35 | 2021-09-07 08:14 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 08:15 | Glide Water Association | September | PWS00326:SRC-AA | 2109054-01 | Source Water | 246 | 35 | 2021-09-07 08:15 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 08:16 | City of Powers | September | PWS00672:SRC-AA | 2109049-01 | Source Water | 246 | 35 | 2021-09-07 08:16 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 08:30 | Springfield Utility Board | September | PWS00837:CH-A | 2109069-02 | Source Water | 246 | 35 | 2021-09-07 08:30 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 09:20 | Winston-Dillard Water District | September | PWS00957:SRC-AA | 2109074-01 | Source Water | 246 | 35 | 2021-09-07 09:20 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 09:24 | USFS Tiller Ranger Station | September | PWS01092:SRC-AA | 2109073-01 | Source Water | 246 | 35 | 2021-09-07 09:24 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 09:50 | Camp Baker BSA | September | PWS91786:SRC-AA | 2109038-01 | Source Water | 246 | 35 | 2021-09-07 09:50 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 10:23 | Clarks Branch Water Association | September | PWS00548:SRC-AA | 2109051-01 | Source Water | 246 | 35 | 2021-09-07 10:23 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 10:50 | Springfield Utility Board | September | PWS00837:SRC-BB | 2109069-01 | Source Water | 246 | 35 | 2021-09-07 10:50 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 2.86 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-09-07 11:00 | Buell-Red Prairie Water Association | September | PWS01174:SRC-AA | 2109037-01RE1 | Source Water | 246 | 35 | 2021-09-07 11:00 | 2021-09-08 16:30 | 2021-09-14 18:58 | B21I114 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 11:20 | Eugene Water & Electric Board | September | PWS00287:SRC-AA | 2109052-01 | Source Water | 246 | 35 | 2021-09-07 11:20 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 11:20 | Lake Oswego - Tigard Water Supply | September | PWS01542:SRC-AA | 2109059-01 | Source Water | 246 | 35 | 2021-09-07 11:20 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 11:35 | City of Wilsonville | September | PWS00954:SRC-HA | 2109050-01 | Source Water | 246 | 35 | 2021-09-07 11:35 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 11:41 | USFS Steamboat Work Center | September | PWS01091:SRC-AA | 2109071-01 | Source Water | 246 | 35 | 2021-09-07 11:41 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.19 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2021-09-07 12:00 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2109078-01RE1 | Source Water | 246 | 35 | 2021-09-07 12:00 | 2021-09-08 16:30 | 2021-09-14 18:58 | B21I114 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 12:23 | USFS Horseshoe Bend CG | September | PWS94179:SRC-AA | 2109070-01 | Source Water | 246 | 35 | 2021-09-07 12:23 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 12:30 | City of Cottage Grove | September | PWS00236:SRC-BA | 2109043-01 | Source Water | 246 | 35 | 2021-09-07 12:30 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.28 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2021-09-07 12:30 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2109077-01RE1 | Source Water | 246 | 35 | 2021-09-07 12:30 | 2021-09-08 16:30 | 2021-09-14 18:58 | B21I114 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 12:49 | PP&L-Toketee Village | September | PWS01012:SRC-AA | 2109064-01 | Source Water | 246 | 35 | 2021-09-07 12:49 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 13:45 | Canby Utility/Veolia Water | September | PWS00157:CH-A | 2109039-01 | Source Water | 246 | 35 | 2021-09-07 13:45 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-07 14:30 | City of Gates | September | PWS00317:SRC-AA | 2109044-01 | Source Water | 246 | 35 | 2021-09-07 14:30 | 2021-09-08 16:30 | 2021-09-13 14:24 | B21I083 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-08 08:30 | City of Albany | September | PWS00012:SRC-BA | 2109041-01 | Source Water | 247 | 35 | 2021-09-08 08:30 | 2021-09-09 12:00 | 2021-09-14 18:58 | B21I114 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-08 10:20 | North Clackamas County Water Commission | September | PWS00580:SRC-BA | 2109062-01 | Source Water | 247 | 35 | 2021-09-08 10:20 | 2021-09-09 12:00 | 2021-09-14 18:58 | B21I114 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-08 17:13 | City of Grants Pass | September | PWS00342:SRC-AA | 2109103-01 | Source Water | 247 | 35 | 2021-09-08 17:13 | 2021-09-10 13:30 | 2021-09-14 18:58 | B21I114 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-10 19:00 | Milo Academy | September | PWS00250:SRC-AB | 2109175-01 | Source Water | 249 | 35 | 2021-09-10 19:00 | 2021-09-15 14:30 | 2021-09-21 13:53 | B21I210 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 08:24 | Lyons Mehama Water District | September | PWS00493:SRC-AA | 2109149-01 | Source Water | 252 | 36 | 2021-09-13 08:24 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.62 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-09-13 08:45 | City of Creswell | September | PWS00246:SRC-BK | 2109136-01 | Source Water | 252 | 36 | 2021-09-13 08:45 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 08:51 | City of Jefferson | September | PWS00408:SRC-AA | 2109140-01 | Source Water | 252 | 36 | 2021-09-13 08:51 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 09:00 | Tri-City JW & SA | September | PWS00549:SRC-AA | 2109157-01 | Source Water | 252 | 36 | 2021-09-13 09:00 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 09:20 | Medford Water Commission | September | PWS00513:SRC-BA | 2109150-01 | Source Water | 252 | 36 | 2021-09-13 09:20 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 1.71 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-09-13 09:45 | Buell-Red Prairie Water Association | September | PWS01174:SRC-AA | 2109163-01 | Source Water | 252 | 36 | 2021-09-13 09:45 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 10:00 | South Fork Water Board - Oregon City | September | PWS00591:SRC-AA | 2109155-01 | Source Water | 252 | 36 | 2021-09-13 10:00 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 10:15 | City of Estacada | September | PWS00279:SRC-AA | 2109138-01 | Source Water | 252 | 36 | 2021-09-13 10:15 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 10:20 | City of Monroe | September | PWS00540:SRC-BA | 2109142-01 | Source Water | 252 | 36 | 2021-09-13 10:20 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 10:30 | City of Myrtle Creek | September | PWS00550:SRC-AA | 2109143-01 | Source Water | 252 | 36 | 2021-09-13 10:30 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 10:30 | Umpqua Basin Water Association | September | PWS00719:SRC-AA | 2109158-01 | Source Water | 252 | 36 | 2021-09-13 10:30 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 11:00 | City of Newport | September | PWS00566:CH-A | 2109144-01 | Source Water | 252 | 36 | 2021-09-13 11:00 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 11:05 | City of Rogue River | September | PWS00712:SRC-BA | 2109145-01 | Source Water | 252 | 36 | 2021-09-13 11:05 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 11:24 | Shangri La Water District | September | PWS00835:SRC-AA | 2109154-01 | Source Water | 252 | 36 | 2021-09-13 11:24 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.09 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-09-13 12:00 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2109147-01 | Source Water | 252 | 36 | 2021-09-13 12:00 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 12:08 | Roberts Creek Water District | September | PWS00717:SRC-AA | 2109152-01 | Source Water | 252 | 36 | 2021-09-13 12:08 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 12:15 | Ashland Water Department | September | PWS00047:SRC-AA | 2109135-01 | Source Water | 252 | 36 | 2021-09-13 12:15 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 12:25 | City of Drain Shops | September | PWS00260:SRC-AA | 2109137-01 | Source Water | 252 | 36 | 2021-09-13 12:25 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.26 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2021-09-13 12:30 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2109148-01 | Source Water | 252 | 36 | 2021-09-13 12:30 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 12:35 | Rainbow Water District | September | PWS00839:SRC-BB | 2109151-01 | Source Water | 252 | 36 | 2021-09-13 12:35 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 13:00 | City of Lowell | September | PWS00492:SRC-BA | 2109141-01 | Source Water | 252 | 36 | 2021-09-13 13:00 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 13:00 | Stayton Water Supply | September | PWS00843:SRC-AA | 2109156-01 | Source Water | 252 | 36 | 2021-09-13 13:00 | 2021-09-14 14:30 | 2021-09-15 15:01 | B21I171 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-13 19:00 | Milo Academy | September | PWS00250:SRC-AB | 2109175-03 | Source Water | 252 | 36 | 2021-09-13 19:00 | 2021-09-15 14:30 | 2021-09-21 13:53 | B21I210 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-14 15:00 | Seaside Water Department | September | PWS00799:CH-A | 2109153-01 | Source Water | 253 | 36 | 2021-09-14 15:00 | 2021-09-15 14:30 | 2021-09-21 13:53 | B21I210 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-15 08:55 | City of Gold Hill | September | PWS00333:SRC-AA | 2109139-01 | Source Water | 254 | 36 | 2021-09-15 08:55 | 2021-09-16 12:00 | 2021-09-21 13:53 | B21I210 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-15 09:30 | Umpqua Ranch Co-op | September | PWS00714:SRC-AH | 2109159-01 | Source Water | 254 | 36 | 2021-09-15 09:30 | 2021-09-16 12:00 | 2021-09-21 13:53 | B21I210 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-15 12:45 | City of Roseburg | September | PWS00720:SRC-AA | 2109186-01 | Source Water | 254 | 36 | 2021-09-15 12:45 | 2021-09-16 12:00 | 2021-09-21 13:53 | B21I210 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-16 12:37 | Douglas County Parks - Whistler's Bend | September | PWS93944:SRC-AB | 2109199-01 | Source Water | 255 | 36 | 2021-09-16 12:37 | 2021-09-17 11:00 | 2021-09-21 13:53 | B21I210 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 05:15 | Anglers Cove/SCHWC | September | PWS01483:SRC-AB | 2109200-01 | Source Water | 259 | 37 | 2021-09-20 05:15 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 5e-05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 06:47 | City of Albany | September | PWS00012:SRC-BA | 2109204-01 | Source Water | 259 | 37 | 2021-09-20 06:47 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 08:35 | Springfield Utility Board | September | PWS00837:CH-A | 2109222-02 | Source Water | 259 | 37 | 2021-09-20 08:35 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 08:45 | Salem Public Works | September | PWS00731:SRC-AA | 2109221-01 | Source Water | 259 | 37 | 2021-09-20 08:45 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.92 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-09-20 09:00 | City of Creswell | September | PWS00246:SRC-BK | 2109233-01 | Source Water | 259 | 37 | 2021-09-20 09:00 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 09:01 | Joint Water Commission | September | PWS00379:SRC-AA | 2109215-01 | Source Water | 259 | 37 | 2021-09-20 09:01 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 09:10 | North Clackamas County Water Commission | September | PWS00580:SRC-BA | 2109218-01 | Source Water | 259 | 37 | 2021-09-20 09:10 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 09:25 | Winston-Dillard Water District | September | PWS00957:SRC-AA | 2109226-01 | Source Water | 259 | 37 | 2021-09-20 09:25 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 09:40 | Camp Baker BSA | September | PWS91786:SRC-AA | 2109202-01 | Source Water | 259 | 37 | 2021-09-20 09:40 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 10:00 | Glide Water Association | September | PWS00326:SRC-AA | 2109213-01 | Source Water | 259 | 37 | 2021-09-20 10:00 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 10:16 | USFS Tiller Ranger Station | September | PWS01092:SRC-AA | 2109225-01 | Source Water | 259 | 37 | 2021-09-20 10:16 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 11:00 | Eugene Water & Electric Board | September | PWS00287:SRC-AA | 2109212-01 | Source Water | 259 | 37 | 2021-09-20 11:00 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 11:15 | City of Pendleton | September | PWS00613:SRC-AM | 2109208-01 | Source Water | 259 | 37 | 2021-09-20 11:15 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 11:16 | Clarks Branch Water Association | September | PWS00548:SRC-AA | 2109211-01 | Source Water | 259 | 37 | 2021-09-20 11:16 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.08 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-09-20 11:30 | City of Cottage Grove | September | PWS00236:SRC-BA | 2109205-01 | Source Water | 259 | 37 | 2021-09-20 11:30 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 12:20 | USFS Steamboat Work Center | September | PWS01091:SRC-AA | 2109224-01 | Source Water | 259 | 37 | 2021-09-20 12:20 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 12:50 | USFS Horseshoe Bend CG | September | PWS94179:SRC-AA | 2109223-01 | Source Water | 259 | 37 | 2021-09-20 12:50 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.08 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-09-20 13:00 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2109234-01 | Source Water | 259 | 37 | 2021-09-20 13:00 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 13:00 | Lake Oswego - Tigard Water Supply | September | PWS01542:SRC-AA | 2109216-01 | Source Water | 259 | 37 | 2021-09-20 13:00 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 13:10 | PP&L-Toketee Village | September | PWS01012:SRC-AA | 2109219-01 | Source Water | 259 | 37 | 2021-09-20 13:10 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 13:10 | Springfield Utility Board | September | PWS00837:SRC-BB | 2109222-01 | Source Water | 259 | 37 | 2021-09-20 13:10 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 13:30 | City of Gates | September | PWS00317:SRC-AA | 2109206-01 | Source Water | 259 | 37 | 2021-09-20 13:30 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.13 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2021-09-20 13:30 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2109235-01 | Source Water | 259 | 37 | 2021-09-20 13:30 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 13:35 | City of Grants Pass | September | PWS00342:SRC-AA | 2109207-01 | Source Water | 259 | 37 | 2021-09-20 13:35 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 14:30 | Canby Utility/Veolia Water | September | PWS00157:CH-A | 2109203-01 | Source Water | 259 | 37 | 2021-09-20 14:30 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-20 17:30 | Hiland WC - Shady Cove | September | PWS01520:SRC-AA | 2109214-01 | Source Water | 259 | 37 | 2021-09-20 17:30 | 2021-09-21 12:00 | 2021-09-22 14:33 | B21I253 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-21 09:10 | City of Powers | September | PWS00672:SRC-AA | 2109209-01 | Source Water | 260 | 37 | 2021-09-21 09:10 | 2021-09-22 12:30 | 2021-09-28 14:58 | B21I266 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-21 09:15 | Roseburg Forest Products - Dillard | September | PWS94300:SRC-AA | 2109220-01 | Source Water | 260 | 37 | 2021-09-21 09:15 | 2021-09-22 12:30 | 2021-09-28 14:58 | B21I266 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-21 13:20 | City of Wilsonville | September | PWS00954:SRC-HA | 2109247-01 | Source Water | 260 | 37 | 2021-09-21 13:20 | 2021-09-22 12:30 | 2021-09-28 14:58 | B21I266 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-22 10:30 | Milo Academy | September | PWS00250:SRC-AB | 2109217-01 | Source Water | 261 | 37 | 2021-09-22 10:30 | 2021-09-23 11:45 | 2021-09-28 14:58 | B21I266 | KMC
Microcystins ADDA, Total | 1.58 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-09-23 09:45 | Buell-Red Prairie Water Association | September | PWS01174:SRC-AA | 2109263-01 | Source Water | 262 | 37 | 2021-09-23 09:45 | 2021-09-24 10:45 | 2021-09-28 14:58 | B21I266 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 08:10 | City of Jefferson | September | PWS00408:SRC-AA | 2109273-01 | Source Water | 266 | 38 | 2021-09-27 08:10 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 09:00 | Tri-City JW & SA | September | PWS00549:SRC-AA | 2109290-01 | Source Water | 266 | 38 | 2021-09-27 09:00 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 09:19 | Douglas County Parks - Whistler's Bend | September | PWS93944:SRC-AB | 2109267-01 | Source Water | 266 | 38 | 2021-09-27 09:19 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 09:30 | Shangri La Water District | September | PWS00835:SRC-AA | 2109287-01 | Source Water | 266 | 38 | 2021-09-27 09:30 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.37 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-09-27 09:45 | City of Creswell | September | PWS00246:SRC-BK | 2109269-01 | Source Water | 266 | 38 | 2021-09-27 09:45 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 10:00 | Umpqua Ranch Co-op | September | PWS00714:SRC-AH | 2109292-01 | Source Water | 266 | 38 | 2021-09-27 10:00 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 10:05 | City of Estacada | September | PWS00279:SRC-AA | 2109271-01 | Source Water | 266 | 38 | 2021-09-27 10:05 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 10:15 | Lyons Mehama Water District | September | PWS00493:SRC-AA | 2109282-01 | Source Water | 266 | 38 | 2021-09-27 10:15 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 10:25 | City of Lowell | September | PWS00492:SRC-BA | 2109274-01 | Source Water | 266 | 38 | 2021-09-27 10:25 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 10:27 | South Fork Water Board - Oregon City | September | PWS00591:SRC-AA | 2109288-01 | Source Water | 266 | 38 | 2021-09-27 10:27 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 10:30 | City of Myrtle Creek | September | PWS00550:SRC-AA | 2109276-01 | Source Water | 266 | 38 | 2021-09-27 10:30 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 11:00 | City of Newport | September | PWS00566:CH-A | 2109277-01 | Source Water | 266 | 38 | 2021-09-27 11:00 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 11:00 | City of Rogue River | September | PWS00712:SRC-BA | 2109278-01 | Source Water | 266 | 38 | 2021-09-27 11:00 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 11:02 | City of Gold Hill | September | PWS00333:SRC-AA | 2109272-01 | Source Water | 266 | 38 | 2021-09-27 11:02 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 1.03 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2021-09-27 11:20 | Buell-Red Prairie Water Association | September | PWS01174:SRC-AA | 2109294-01 | Source Water | 266 | 38 | 2021-09-27 11:20 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 11:30 | Umpqua Basin Water Association | September | PWS00719:SRC-AA | 2109291-01 | Source Water | 266 | 38 | 2021-09-27 11:30 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.05 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 11:42 | City of Monroe | September | PWS00540:SRC-BA | 2109275-01 | Source Water | 266 | 38 | 2021-09-27 11:42 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 11:50 | Rainbow Water District | September | PWS00839:SRC-BB | 2109284-01 | Source Water | 266 | 38 | 2021-09-27 11:50 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 12:10 | City of Drain Shops | September | PWS00260:SRC-AA | 2109270-01 | Source Water | 266 | 38 | 2021-09-27 12:10 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.06 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 12:30 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2109280-01 | Source Water | 266 | 38 | 2021-09-27 12:30 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 12:30 | Roberts Creek Water District | September | PWS00717:SRC-AA | 2109285-01 | Source Water | 266 | 38 | 2021-09-27 12:30 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 12:45 | Ashland Water Department | September | PWS00047:SRC-AA | 2109268-01 | Source Water | 266 | 38 | 2021-09-27 12:45 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.14 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2021-09-27 13:00 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2109281-01 | Source Water | 266 | 38 | 2021-09-27 13:00 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 13:00 | Stayton Water Supply | September | PWS00843:SRC-AA | 2109289-01 | Source Water | 266 | 38 | 2021-09-27 13:00 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 13:37 | City of Roseburg | September | PWS00720:SRC-AA | 2109279-01 | Source Water | 266 | 38 | 2021-09-27 13:37 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 14:00 | Seaside Water Department | September | PWS00799:CH-A | 2109286-01 | Source Water | 266 | 38 | 2021-09-27 14:00 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-09-27 14:10 | Medford Water Commission | September | PWS00513:SRC-BA | 2109283-01 | Source Water | 266 | 38 | 2021-09-27 14:10 | 2021-09-28 13:40 | 2021-09-29 14:10 | B21I318 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 05:15 | Anglers Cove/SCHWC | October | PWS01483:SRC-AB | 2110011-01 | Source Water | 273 | 39 | 2021-10-04 05:15 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 05:30 | Hiland WC - Shady Cove | October | PWS01520:SRC-AA | 2110032-01 | Source Water | 273 | 39 | 2021-10-04 05:30 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 07:47 | Joint Water Commission | October | PWS00379:SRC-AA | 2110034-01 | Source Water | 273 | 39 | 2021-10-04 07:47 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 08:03 | Roseburg Forest Products - Dillard | October | PWS94300:SRC-AA | 2110042-01 | Source Water | 273 | 39 | 2021-10-04 08:03 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 08:10 | Salem Public Works | October | PWS00731:SRC-AA | 2110043-01 | Source Water | 273 | 39 | 2021-10-04 08:10 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 08:30 | Winston-Dillard Water District | October | PWS00957:SRC-AA | 2110050-01 | Source Water | 273 | 39 | 2021-10-04 08:30 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 08:40 | Springfield Utility Board | October | PWS00837:CH-A | 2110045-02 | Source Water | 273 | 39 | 2021-10-04 08:40 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 08:52 | City of Pendleton | October | PWS00613:SRC-AM | 2110024-01 | Source Water | 273 | 39 | 2021-10-04 08:52 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 09:20 | North Clackamas County Water Commission | October | PWS00580:SRC-BA | 2110038-01 | Source Water | 273 | 39 | 2021-10-04 09:20 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 10:00 | City of Powers | October | PWS00672:SRC-AA | 2110025-01 | Source Water | 273 | 39 | 2021-10-04 10:00 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 10:15 | Camp Baker BSA | October | PWS91786:SRC-AA | 2110014-01 | Source Water | 273 | 39 | 2021-10-04 10:15 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 10:30 | Glide Water Association | October | PWS00326:SRC-AA | 2110030-01 | Source Water | 273 | 39 | 2021-10-04 10:30 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 10:40 | Springfield Utility Board | October | PWS00837:SRC-BB | 2110045-01 | Source Water | 273 | 39 | 2021-10-04 10:40 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 11:19 | Eugene Water & Electric Board | October | PWS00287:SRC-AA | 2110028-01 | Source Water | 273 | 39 | 2021-10-04 11:19 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.07 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-10-04 12:00 | City of Cottage Grove | October | PWS00236:SRC-BA | 2110019-01 | Source Water | 273 | 39 | 2021-10-04 12:00 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 12:50 | Lake Oswego - Tigard Water Supply | October | PWS01542:SRC-AA | 2110035-01 | Source Water | 273 | 39 | 2021-10-04 12:50 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 13:00 | City of Wilsonville | October | PWS00954:SRC-HA | 2110026-01 | Source Water | 273 | 39 | 2021-10-04 13:00 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 13:15 | Canby Utility/Veolia Water | October | PWS00157:CH-A | 2110015-01 | Source Water | 273 | 39 | 2021-10-04 13:15 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 13:20 | City of Grants Pass | October | PWS00342:SRC-AA | 2110022-01 | Source Water | 273 | 39 | 2021-10-04 13:20 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 13:45 | City of Gates | October | PWS00317:SRC-AA | 2110020-01 | Source Water | 273 | 39 | 2021-10-04 13:45 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-04 16:15 | City of Albany | October | PWS00012:SRC-BA | 2110017-01 | Source Water | 273 | 39 | 2021-10-04 16:15 | 2021-10-06 12:30 | 2021-10-07 15:45 | B21J078 | KMC
Microcystins ADDA, Total | 0.1 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-10-05 08:30 | City of Creswell | October | PWS00246:SRC-BK | 2110061-01 | Source Water | 274 | 39 | 2021-10-05 08:30 | 2021-10-05 13:10 | 2021-10-06 13:17 | B21J060 | KMC
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-05 10:08 | USFS Tiller Ranger Station | October | PWS01092:SRC-AA | 2110049-01 | Source Water | 274 | 39 | 2021-10-05 10:08 | 2021-10-06 12:30 | 2021-10-07 15:45 | B21J078 | KMC
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-05 11:06 | Clarks Branch Water Association | October | PWS00548:SRC-AA | 2110027-01 | Source Water | 274 | 39 | 2021-10-05 11:06 | 2021-10-06 12:30 | 2021-10-07 15:45 | B21J078 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-05 12:42 | USFS Steamboat Work Center | October | PWS01091:SRC-AA | 2110047-01 | Source Water | 274 | 39 | 2021-10-05 12:42 | 2021-10-06 12:30 | 2021-10-07 15:45 | B21J078 | KMC
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-05 13:09 | PP&L-Toketee Village | October | PWS01012:SRC-AA | 2110040-01 | Source Water | 274 | 39 | 2021-10-05 13:09 | 2021-10-06 12:30 | 2021-10-07 15:45 | B21J078 | KMC
Microcystins ADDA, Total | 0.28 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2021-10-07 10:00 | Buell-Red Prairie Water Association | October | PWS01174:SRC-AA | 2110092-01RE1 | Source Water | 276 | 39 | 2021-10-07 10:00 | 2021-10-08 14:45 | 2021-10-14 14:27 | B21J175 | SD
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-07 10:30 | Milo Academy | October | PWS00250:SRC-AB | 2110037-01 | Source Water | 276 | 39 | 2021-10-07 10:30 | 2021-10-08 14:45 | 2021-10-12 18:05 | B21J130 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 08:38 | Lyons Mehama Water District | October | PWS00493:SRC-AA | 2110142-01 | Source Water | 280 | 40 | 2021-10-11 08:38 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.11 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-10-11 08:45 | City of Creswell | October | PWS00246:SRC-BK | 2110129-01 | Source Water | 280 | 40 | 2021-10-11 08:45 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 09:00 | Roberts Creek Water District | October | PWS00717:SRC-AA | 2110145-01 | Source Water | 280 | 40 | 2021-10-11 09:00 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 09:29 | City of Roseburg | October | PWS00720:SRC-AA | 2110139-01 | Source Water | 280 | 40 | 2021-10-11 09:29 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 09:30 | City of Monroe | October | PWS00540:SRC-BA | 2110135-01 | Source Water | 280 | 40 | 2021-10-11 09:30 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 09:30 | Shangri La Water District | October | PWS00835:SRC-AA | 2110147-01 | Source Water | 280 | 40 | 2021-10-11 09:30 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 09:50 | Douglas County Parks - Whistler's Bend | October | PWS93944:SRC-AB | 2110119-01 | Source Water | 280 | 40 | 2021-10-11 09:50 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 10:00 | City of Gold Hill | October | PWS00333:SRC-AA | 2110132-01 | Source Water | 280 | 40 | 2021-10-11 10:00 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 10:00 | City of Jefferson | October | PWS00408:SRC-AA | 2110133-01 | Source Water | 280 | 40 | 2021-10-11 10:00 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 10:00 | City of Newport | October | PWS00566:CH-A | 2110137-01 | Source Water | 280 | 40 | 2021-10-11 10:00 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 10:00 | Tri-City JW & SA | October | PWS00549:SRC-AA | 2110150-01 | Source Water | 280 | 40 | 2021-10-11 10:00 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 10:05 | City of Estacada | October | PWS00279:SRC-AA | 2110131-01 | Source Water | 280 | 40 | 2021-10-11 10:05 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.18 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2021-10-11 10:20 | Buell-Red Prairie Water Association | October | PWS01174:SRC-AA | 2110153-01 | Source Water | 280 | 40 | 2021-10-11 10:20 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 10:30 | City of Myrtle Creek | October | PWS00550:SRC-AA | 2110136-01 | Source Water | 280 | 40 | 2021-10-11 10:30 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 11:15 | Umpqua Basin Water Association | October | PWS00719:SRC-AA | 2110151-01 | Source Water | 280 | 40 | 2021-10-11 11:15 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 11:30 | City of Drain Shops | October | PWS00260:SRC-AA | 2110130-01 | Source Water | 280 | 40 | 2021-10-11 11:30 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 11:45 | City of Lowell | October | PWS00492:SRC-BA | 2110134-01 | Source Water | 280 | 40 | 2021-10-11 11:45 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 12:05 | City of Rogue River | October | PWS00712:SRC-BA | 2110138-01 | Source Water | 280 | 40 | 2021-10-11 12:05 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 12:05 | Medford Water Commission | October | PWS00513:SRC-BA | 2110143-01 | Source Water | 280 | 40 | 2021-10-11 12:05 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 12:30 | Ashland Water Department | October | PWS00047:SRC-AA | 2110128-01 | Source Water | 280 | 40 | 2021-10-11 12:30 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 12:30 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2110140-01 | Source Water | 280 | 40 | 2021-10-11 12:30 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 12:35 | South Fork Water Board - Oregon City | October | PWS00591:SRC-AA | 2110148-01 | Source Water | 280 | 40 | 2021-10-11 12:35 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.07 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-10-11 13:00 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2110141-01 | Source Water | 280 | 40 | 2021-10-11 13:00 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 13:00 | Stayton Water Supply | October | PWS00843:SRC-AA | 2110149-01 | Source Water | 280 | 40 | 2021-10-11 13:00 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-11 14:00 | Seaside Water Department | October | PWS00799:CH-A | 2110146-01 | Source Water | 280 | 40 | 2021-10-11 14:00 | 2021-10-12 13:30 | 2021-10-13 14:55 | B21J156 | SD
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-12 12:30 | Rainbow Water District | October | PWS00839:SRC-BB | 2110144-01 | Source Water | 281 | 40 | 2021-10-12 12:30 | 2021-10-13 13:45 | 2021-10-14 14:27 | B21J175 | SD
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 | City of Powers | October | PWS00672:SRC-AA | 2110213-01 | Source Water | 287 | 41 | 2021-10-18 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 05:15 | Anglers Cove/SCHWC | October | PWS01483:SRC-AB | 2110204-01 | Source Water | 287 | 41 | 2021-10-18 05:15 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 05:30 | Hiland WC - Shady Cove | October | PWS01520:SRC-AA | 2110218-01 | Source Water | 287 | 41 | 2021-10-18 05:30 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 06:30 | City of Albany | October | PWS00012:SRC-BA | 2110208-01 | Source Water | 287 | 41 | 2021-10-18 06:30 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 07:06 | City of Pendleton | October | PWS00613:SRC-AM | 2110212-01 | Source Water | 287 | 41 | 2021-10-18 07:06 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 08:25 | Springfield Utility Board | October | PWS00837:CH-A | 2110226-02 | Source Water | 287 | 41 | 2021-10-18 08:25 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 08:29 | Joint Water Commission | October | PWS00379:SRC-AA | 2110219-01 | Source Water | 287 | 41 | 2021-10-18 08:29 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 08:30 | Roseburg Forest Products - Dillard | October | PWS94300:SRC-AA | 2110224-01 | Source Water | 287 | 41 | 2021-10-18 08:30 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 08:30 | Winston-Dillard Water District | October | PWS00957:SRC-AA | 2110230-01 | Source Water | 287 | 41 | 2021-10-18 08:30 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 09:00 | North Clackamas County Water Commission | October | PWS00580:SRC-BA | 2110222-01 | Source Water | 287 | 41 | 2021-10-18 09:00 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.2 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2021-10-18 09:30 | Buell-Red Prairie Water Association | October | PWS01174:SRC-AA | 2110205-01 | Source Water | 287 | 41 | 2021-10-18 09:30 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 09:40 | Camp Baker BSA | October | PWS91786:SRC-AA | 2110206-01 | Source Water | 287 | 41 | 2021-10-18 09:40 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 10:05 | Glide Water Association | October | PWS00326:SRC-AA | 2110217-01 | Source Water | 287 | 41 | 2021-10-18 10:05 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 10:05 | USFS Tiller Ranger Station | October | PWS01092:SRC-AA | 2110229-01 | Source Water | 287 | 41 | 2021-10-18 10:05 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 10:30 | Milo Academy | October | PWS00250:SRC-AB | 2110221-01 | Source Water | 287 | 41 | 2021-10-18 10:30 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 10:32 | Salem Public Works | October | PWS00731:SRC-AA | 2110225-01 | Source Water | 287 | 41 | 2021-10-18 10:32 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 11:00 | Clarks Branch Water Association | October | PWS00548:SRC-AA | 2110215-01 | Source Water | 287 | 41 | 2021-10-18 11:00 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.09 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-10-18 11:00 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2110235-01 | Source Water | 287 | 41 | 2021-10-18 11:00 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.08 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-10-18 11:30 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2110234-01 | Source Water | 287 | 41 | 2021-10-18 11:30 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 11:30 | Springfield Utility Board | October | PWS00837:SRC-BB | 2110226-01 | Source Water | 287 | 41 | 2021-10-18 11:30 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.1 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2021-10-18 12:00 | City of Cottage Grove | October | PWS00236:SRC-BA | 2110209-01 | Source Water | 287 | 41 | 2021-10-18 12:00 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 12:28 | USFS Steamboat Work Center | October | PWS01091:SRC-AA | 2110228-01 | Source Water | 287 | 41 | 2021-10-18 12:28 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 12:30 | City of Wilsonville | October | PWS00954:SRC-HA | 2110214-01 | Source Water | 287 | 41 | 2021-10-18 12:30 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 12:50 | Lake Oswego - Tigard Water Supply | October | PWS01542:SRC-AA | 2110220-01 | Source Water | 287 | 41 | 2021-10-18 12:50 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 13:05 | PP&L-Toketee Village | October | PWS01012:SRC-AA | 2110223-01 | Source Water | 287 | 41 | 2021-10-18 13:05 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.03 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 13:35 | City of Grants Pass | October | PWS00342:SRC-AA | 2110211-01 | Source Water | 287 | 41 | 2021-10-18 13:35 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-18 14:10 | Canby Utility/Veolia Water | October | PWS00157:CH-A | 2110207-01 | Source Water | 287 | 41 | 2021-10-18 14:10 | 2021-10-19 13:05 | 2021-10-20 12:42 | B21J198 | SD
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-19 09:00 | City of Gates | October | PWS00317:SRC-AA | 2110210-01 | Source Water | 288 | 41 | 2021-10-19 09:00 | 2021-10-20 11:25 | 2021-10-21 10:14 | B21J266 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-19 13:38 | Eugene Water & Electric Board | October | PWS00287:SRC-AA | 2110254-01 | Source Water | 288 | 41 | 2021-10-19 13:38 | 2021-10-20 11:25 | 2021-10-21 10:14 | B21J266 | SD
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 07:00 | Umpqua Ranch Co-op | October | PWS00714:SRC-AH | 2110338-01 | Source Water | 294 | 42 | 2021-10-25 07:00 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 08:42 | South Fork Water Board - Oregon City | October | PWS00591:SRC-AA | 2110334-01 | Source Water | 294 | 42 | 2021-10-25 08:42 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 08:45 | City of Jefferson | October | PWS00408:SRC-AA | 2110319-01 | Source Water | 294 | 42 | 2021-10-25 08:45 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 08:45 | Lyons Mehama Water District | October | PWS00493:SRC-AA | 2110328-01 | Source Water | 294 | 42 | 2021-10-25 08:45 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.06 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 09:00 | City of Creswell | October | PWS00246:SRC-BK | 2110315-01 | Source Water | 294 | 42 | 2021-10-25 09:00 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 09:30 | Shangri La Water District | October | PWS00835:SRC-AA | 2110333-01 | Source Water | 294 | 42 | 2021-10-25 09:30 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 09:50 | City of Monroe | October | PWS00540:SRC-BA | 2110321-01 | Source Water | 294 | 42 | 2021-10-25 09:50 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 09:55 | City of Estacada | October | PWS00279:SRC-AA | 2110317-01 | Source Water | 294 | 42 | 2021-10-25 09:55 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 10:00 | City of Newport | October | PWS00566:CH-A | 2110323-01 | Source Water | 294 | 42 | 2021-10-25 10:00 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 10:07 | Douglas County Parks - Whistler's Bend | October | PWS93944:SRC-AB | 2110313-01 | Source Water | 294 | 42 | 2021-10-25 10:07 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 10:30 | City of Myrtle Creek | October | PWS00550:SRC-AA | 2110322-01 | Source Water | 294 | 42 | 2021-10-25 10:30 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 10:30 | Umpqua Basin Water Association | October | PWS00719:SRC-AA | 2110337-01 | Source Water | 294 | 42 | 2021-10-25 10:30 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 11:00 | City of Rogue River | October | PWS00712:SRC-BA | 2110324-01 | Source Water | 294 | 42 | 2021-10-25 11:00 | 2021-10-27 10:45 | 2021-10-28 14:33 | B21J340 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 11:00 | Tri-City JW & SA | October | PWS00549:SRC-AA | 2110336-01 | Source Water | 294 | 42 | 2021-10-25 11:00 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 11:30 | City of Drain Shops | October | PWS00260:SRC-AA | 2110316-01 | Source Water | 294 | 42 | 2021-10-25 11:30 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 12:00 | Roberts Creek Water District | October | PWS00717:SRC-AA | 2110331-01 | Source Water | 294 | 42 | 2021-10-25 12:00 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 12:30 | Ashland Water Department | October | PWS00047:SRC-AA | 2110344-01 | Source Water | 294 | 42 | 2021-10-25 12:30 | 2021-10-27 10:45 | 2021-10-28 14:33 | B21J340 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 13:00 | Stayton Water Supply | October | PWS00843:SRC-AA | 2110335-01 | Source Water | 294 | 42 | 2021-10-25 13:00 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 13:10 | Rainbow Water District | October | PWS00839:SRC-BB | 2110330-01 | Source Water | 294 | 42 | 2021-10-25 13:10 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 13:17 | City of Roseburg | October | PWS00720:SRC-AA | 2110325-01 | Source Water | 294 | 42 | 2021-10-25 13:17 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 14:00 | Seaside Water Department | October | PWS00799:CH-A | 2110332-01 | Source Water | 294 | 42 | 2021-10-25 14:00 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-25 14:15 | Medford Water Commission | October | PWS00513:SRC-BA | 2110329-01 | Source Water | 294 | 42 | 2021-10-25 14:15 | 2021-10-27 10:45 | 2021-10-28 14:33 | B21J340 | SD
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-26 10:00 | City of Lowell | October | PWS00492:SRC-BA | 2110320-01 | Source Water | 295 | 42 | 2021-10-26 10:00 | 2021-10-26 13:00 | 2021-10-27 11:30 | B21J329 | SD
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2021-10-26 14:45 | City of Gold Hill | October | PWS00333:SRC-AA | 2110352-01 | Source Water | 295 | 42 | 2021-10-26 14:45 | 2021-10-28 10:34 | 2021-10-29 11:23 | B21J363 | SD
Microcystins ADDA, Total | 0.0104 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 05:15 | Anglers Cove/SCHWC | May | PWS01483:SRC-AB | 2205003-01 | Source Water | 121 | 17 | 2022-05-02 05:15 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0056 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 05:30 | Hiland WC - Shady Cove | May | PWS01520:SRC-AA | 2205017-01 | Source Water | 121 | 17 | 2022-05-02 05:30 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0013 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 06:00 | USFS Tiller Ranger Station | May | PWS01092:SRC-AA | 2205029-01 | Source Water | 121 | 17 | 2022-05-02 06:00 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0039 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 07:20 | City of Albany | May | PWS00012:SRC-BA | 2205006-01 | Source Water | 121 | 17 | 2022-05-02 07:20 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 07:20 | Clarks Branch Water Association | May | PWS00548:SRC-AA | 2205014-01 | Source Water | 121 | 17 | 2022-05-02 07:20 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0061 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 07:45 | Springfield Utility Board | May | PWS00837:CH-A | 2205026-02 | Source Water | 121 | 17 | 2022-05-02 07:45 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0025 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 08:10 | Jackson Co Pks Emigrant Lake | May | PWS90730:CH-A | 2205018-01 | Source Water | 121 | 17 | 2022-05-02 08:10 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0046 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 08:10 | Roseburg Forest Products - Dillard | May | PWS94300:SRC-AA | 2205024-01 | Source Water | 121 | 17 | 2022-05-02 08:10 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0022 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 08:16 | Joint Water Commission | May | PWS00379:SRC-AA | 2205019-01 | Source Water | 121 | 17 | 2022-05-02 08:16 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0012 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 09:00 | North Clackamas County Water Commission | May | PWS00580:SRC-BA | 2205022-01 | Source Water | 121 | 17 | 2022-05-02 09:00 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0061 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 09:15 | Winston-Dillard Water District | May | PWS00957:SRC-AA | 2205030-01 | Source Water | 121 | 17 | 2022-05-02 09:15 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 09:25 | Springfield Utility Board | May | PWS00837:SRC-BB | 2205026-01 | Source Water | 121 | 17 | 2022-05-02 09:25 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0041 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 09:30 | Glide Water Association | May | PWS00326:SRC-AA | 2205016-01 | Source Water | 121 | 17 | 2022-05-02 09:30 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0028 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 10:10 | Salem Public Works | May | PWS00731:SRC-AA | 2205025-01 | Source Water | 121 | 17 | 2022-05-02 10:10 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0044 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 10:15 | Buell-Red Prairie Water Association | May | PWS01174:SRC-AA | 2205004-01 | Source Water | 121 | 17 | 2022-05-02 10:15 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0042 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 10:50 | City of Powers | May | PWS00672:SRC-AA | 2205011-01 | Source Water | 121 | 17 | 2022-05-02 10:50 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0014 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 11:00 | Lake Oswego - Tigard Water Supply | May | PWS01542:SRC-AA | 2205020-01 | Source Water | 121 | 17 | 2022-05-02 11:00 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0083 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 11:35 | City of Grants Pass | May | PWS00342:SRC-AA | 2205009-01 | Source Water | 121 | 17 | 2022-05-02 11:35 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0018 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 11:36 | City of Pendleton | May | PWS00613:SRC-AM | 2205010-01 | Source Water | 121 | 17 | 2022-05-02 11:36 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0029 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 11:40 | City of Wilsonville | May | PWS00954:SRC-HA | 2205012-01 | Source Water | 121 | 17 | 2022-05-02 11:40 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 12:00 | City of Cottage Grove | May | PWS00236:SRC-BA | 2205007-01 | Source Water | 121 | 17 | 2022-05-02 12:00 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0046 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 12:10 | Eugene Water & Electric Board | May | PWS00287:SRC-AA | 2205015-01 | Source Water | 121 | 17 | 2022-05-02 12:10 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 12:15 | USFS Steamboat Work Center | May | PWS01091:SRC-AA | 2205028-01 | Source Water | 121 | 17 | 2022-05-02 12:15 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0024 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 13:02 | USFS Horseshoe Bend CG | May | PWS94179:SRC-AA | 2205027-01 | Source Water | 121 | 17 | 2022-05-02 13:02 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 13:30 | Canby Utility/Veolia Water | May | PWS00157:CH-A | 2205005-01 | Source Water | 121 | 17 | 2022-05-02 13:30 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0012 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-02 13:50 | PP&L-Toketee Village | May | PWS01012:SRC-AA | 2205023-01 | Source Water | 121 | 17 | 2022-05-02 13:50 | 2022-05-03 12:00 | 2022-05-04 14:13 | B22E016 | KMC
Microcystins ADDA, Total | 0.0064 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-03 14:00 | City of Gates | May | PWS00317:SRC-AA | 2205008-01 | Source Water | 122 | 17 | 2022-05-03 14:00 | 2022-05-04 16:00 | 2022-05-09 14:58 | B22E027 | KMC
Microcystins ADDA, Total | 0.0062 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 08:45 | South Fork Water Board - Oregon City | May | PWS00591:SRC-AA | 2205084-01 | Source Water | 128 | 18 | 2022-05-09 08:45 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0302 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 08:48 | City of Jefferson | May | PWS00408:SRC-AA | 2205069-01 | Source Water | 128 | 18 | 2022-05-09 08:48 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0121 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 09:00 | City of Creswell | May | PWS00246:SRC-BK | 2205065-01 | Source Water | 128 | 18 | 2022-05-09 09:00 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0013 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 09:05 | Umpqua Basin Water Association | May | PWS00719:SRC-AA | 2205087-01 | Source Water | 128 | 18 | 2022-05-09 09:05 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.014 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 09:20 | City of Gold Hill | May | PWS00333:SRC-AA | 2205068-01 | Source Water | 128 | 18 | 2022-05-09 09:20 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 09:30 | Rainbow Water District | May | PWS00839:SRC-BB | 2205080-01 | Source Water | 128 | 18 | 2022-05-09 09:30 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 09:30 | Shangri La Water District | May | PWS00835:SRC-AA | 2205083-01 | Source Water | 128 | 18 | 2022-05-09 09:30 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0156 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 09:43 | City of Monroe | May | PWS00540:SRC-BA | 2205071-01 | Source Water | 128 | 18 | 2022-05-09 09:43 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0114 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 09:45 | Roberts Creek Water District | May | PWS00717:SRC-AA | 2205081-01RE1 | Source Water | 128 | 18 | 2022-05-09 09:45 | 2022-05-10 11:45 | 2022-05-12 13:21 | B22E072 | KMC
Microcystins ADDA, Total | 0.0114 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 09:45 | Roberts Creek Water District | May | PWS00717:SRC-AA | 2205081-01RE1 | Source Water | 128 | 18 | 2022-05-09 09:45 | 2022-05-10 11:45 | 2022-05-12 13:21 | B22E072 | KMC
Microcystins ADDA, Total | 0.0066 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 10:00 | City of Lowell | May | PWS00492:SRC-BA | 2205070-01 | Source Water | 128 | 18 | 2022-05-09 10:00 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0223 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 10:00 | City of Newport | May | PWS00566:CH-A | 2205073-01 | Source Water | 128 | 18 | 2022-05-09 10:00 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0152 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 10:30 | City of Myrtle Creek | May | PWS00550:SRC-AA | 2205072-01 | Source Water | 128 | 18 | 2022-05-09 10:30 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0106 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 10:30 | City of Rogue River | May | PWS00712:SRC-BA | 2205074-01 | Source Water | 128 | 18 | 2022-05-09 10:30 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0012 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 10:50 | City of Estacada | May | PWS00279:SRC-AA | 2205067-01 | Source Water | 128 | 18 | 2022-05-09 10:50 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0109 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 10:55 | Camp Baker BSA | May | PWS91786:SRC-AA | 2205064-01 | Source Water | 128 | 18 | 2022-05-09 10:55 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0083 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 11:00 | Stayton Water Supply | May | PWS00843:SRC-AA | 2205085-01 | Source Water | 128 | 18 | 2022-05-09 11:00 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 11:10 | Tri-City JW & SA | May | PWS00549:SRC-AA | 2205086-01 | Source Water | 128 | 18 | 2022-05-09 11:10 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 11:30 | Josephine Co Pks Lake Selmac 1 | May | PWS90186:SRC-AA | 2205076-01 | Source Water | 128 | 18 | 2022-05-09 11:30 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 11:30 | Lyons Mehama Water District | May | PWS00493:SRC-AA | 2205078-01 | Source Water | 128 | 18 | 2022-05-09 11:30 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0048 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 11:45 | Josephine Co Pks Lake Selmac 2 | May | PWS94645:SRC-AA | 2205077-01 | Source Water | 128 | 18 | 2022-05-09 11:45 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0357 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 11:48 | City of Roseburg | May | PWS00720:SRC-AA | 2205075-01 | Source Water | 128 | 18 | 2022-05-09 11:48 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0017 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 12:45 | Ashland Water Department | May | PWS00047:SRC-AA | 2205063-01 | Source Water | 128 | 18 | 2022-05-09 12:45 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0134 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 12:45 | Medford Water Commission | May | PWS00513:SRC-BA | 2205079-01 | Source Water | 128 | 18 | 2022-05-09 12:45 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 13:20 | City of Drain Shops | May | PWS00260:SRC-AA | 2205066-01 | Source Water | 128 | 18 | 2022-05-09 13:20 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-09 14:00 | Seaside Water Department | May | PWS00799:CH-A | 2205082-01 | Source Water | 128 | 18 | 2022-05-09 14:00 | 2022-05-10 11:45 | 2022-05-11 12:41 | B22E072 | KMC
Microcystins ADDA, Total | 0.0106 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-11 10:00 | Umpqua Ranch Co-op | May | PWS00714:SRC-AH | 2205088-01 | Source Water | 130 | 18 | 2022-05-11 10:00 | 2022-05-12 11:00 | 2022-05-16 09:17 | B22E110 | KMC
Microcystins ADDA, Total | 0.0205 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 05:15 | Anglers Cove/SCHWC | May | PWS01483:SRC-AB | 2205136-01 | Source Water | 135 | 19 | 2022-05-16 05:15 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0045 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 05:30 | Hiland WC - Shady Cove | May | PWS01520:SRC-AA | 2205149-01 | Source Water | 135 | 19 | 2022-05-16 05:30 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0052 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 07:00 | City of Albany | May | PWS00012:SRC-BA | 2205139-01 | Source Water | 135 | 19 | 2022-05-16 07:00 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0045 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 08:03 | Joint Water Commission | May | PWS00379:SRC-AA | 2205151-01 | Source Water | 135 | 19 | 2022-05-16 08:03 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0027 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 08:15 | Glide Water Association | May | PWS00326:SRC-AA | 2205148-01 | Source Water | 135 | 19 | 2022-05-16 08:15 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 08:20 | Springfield Utility Board | May | PWS00837:CH-A | 2205157-02 | Source Water | 135 | 19 | 2022-05-16 08:20 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0108 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 08:30 | Clarks Branch Water Association | May | PWS00548:SRC-AA | 2205146-01 | Source Water | 135 | 19 | 2022-05-16 08:30 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0115 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 08:40 | Jackson Co Pks Emigrant Lake | May | PWS90730:CH-A | 2205150-01 | Source Water | 135 | 19 | 2022-05-16 08:40 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 09:25 | USFS Tiller Ranger Station | May | PWS01092:SRC-AA | 2205160-01 | Source Water | 135 | 19 | 2022-05-16 09:25 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 09:30 | Salem Public Works | May | PWS00731:SRC-AA | 2205156-01 | Source Water | 135 | 19 | 2022-05-16 09:30 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0075 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 10:00 | Buell-Red Prairie Water Association | May | PWS01174:SRC-AA | 2205137-01 | Source Water | 135 | 19 | 2022-05-16 10:00 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0031 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 10:20 | Springfield Utility Board | May | PWS00837:SRC-BB | 2205157-01 | Source Water | 135 | 19 | 2022-05-16 10:20 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 10:30 | City of Powers | May | PWS00672:SRC-AA | 2205144-01 | Source Water | 135 | 19 | 2022-05-16 10:30 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0058 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 10:40 | North Clackamas County Water Commission | May | PWS00580:SRC-BA | 2205153-01 | Source Water | 135 | 19 | 2022-05-16 10:40 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 10:40 | Winston-Dillard Water District | May | PWS00957:SRC-AA | 2205161-01 | Source Water | 135 | 19 | 2022-05-16 10:40 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0107 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 10:45 | City of Gates | May | PWS00317:SRC-AA | 2205141-01 | Source Water | 135 | 19 | 2022-05-16 10:45 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0254 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 11:00 | City of Grants Pass | May | PWS00342:SRC-AA | 2205142-01 | Source Water | 135 | 19 | 2022-05-16 11:00 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 11:00 | Lake Oswego - Tigard Water Supply | May | PWS01542:SRC-AA | 2205152-01 | Source Water | 135 | 19 | 2022-05-16 11:00 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0067 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 11:04 | USFS Steamboat Work Center | May | PWS01091:SRC-AA | 2205159-01 | Source Water | 135 | 19 | 2022-05-16 11:04 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 11:33 | City of Pendleton | May | PWS00613:SRC-AM | 2205143-01 | Source Water | 135 | 19 | 2022-05-16 11:33 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0047 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 11:50 | PP&L-Toketee Village | May | PWS01012:SRC-AA | 2205154-01 | Source Water | 135 | 19 | 2022-05-16 11:50 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0102 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 12:10 | City of Cottage Grove | May | PWS00236:SRC-BA | 2205140-01 | Source Water | 135 | 19 | 2022-05-16 12:10 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 12:30 | USFS Horseshoe Bend CG | May | PWS94179:SRC-AA | 2205158-01 | Source Water | 135 | 19 | 2022-05-16 12:30 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0097 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 12:40 | Eugene Water & Electric Board | May | PWS00287:SRC-AA | 2205147-01 | Source Water | 135 | 19 | 2022-05-16 12:40 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0055 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 12:46 | City of Wilsonville | May | PWS00954:SRC-HA | 2205145-01 | Source Water | 135 | 19 | 2022-05-16 12:46 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0.0037 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-16 13:45 | Canby Utility/Veolia Water | May | PWS00157:CH-A | 2205138-01 | Source Water | 135 | 19 | 2022-05-16 13:45 | 2022-05-17 12:40 | 2022-05-18 12:17 | B22E162 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-18 07:38 | Roseburg Forest Products - Dillard | May | PWS94300:SRC-AA | 2205183-01RE1 | Source Water | 137 | 19 | 2022-05-18 07:38 | 2022-05-19 10:10 | 2022-05-20 11:21 | B22E206 | SD
Microcystins ADDA, Total | 0.0069 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 08:33 | Lyons Mehama Water District | May | PWS00493:SRC-AA | 2205210-01 | Source Water | 142 | 20 | 2022-05-23 08:33 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0029 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 09:00 | City of Creswell | May | PWS00246:SRC-BK | 2205197-01 | Source Water | 142 | 20 | 2022-05-23 09:00 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0038 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 09:10 | City of Estacada | May | PWS00279:SRC-AA | 2205199-01 | Source Water | 142 | 20 | 2022-05-23 09:10 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0054 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 09:16 | City of Jefferson | May | PWS00408:SRC-AA | 2205201-01 | Source Water | 142 | 20 | 2022-05-23 09:16 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 09:30 | Shangri La Water District | May | PWS00835:SRC-AA | 2205215-01 | Source Water | 142 | 20 | 2022-05-23 09:30 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0022 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 10:00 | South Fork Water Board - Oregon City | May | PWS00591:SRC-AA | 2205216-01 | Source Water | 142 | 20 | 2022-05-23 10:00 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 10:30 | City of Newport | May | PWS00566:CH-A | 2205205-01 | Source Water | 142 | 20 | 2022-05-23 10:30 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0018 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 10:30 | Tri-City JW & SA | May | PWS00549:SRC-AA | 2205218-01 | Source Water | 142 | 20 | 2022-05-23 10:30 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 10:40 | Rainbow Water District | May | PWS00839:SRC-BB | 2205212-01 | Source Water | 142 | 20 | 2022-05-23 10:40 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 10:40 | Umpqua Basin Water Association | May | PWS00719:SRC-AA | 2205219-01 | Source Water | 142 | 20 | 2022-05-23 10:40 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0033 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 10:45 | Camp Baker BSA | May | PWS91786:SRC-AA | 2205196-01 | Source Water | 142 | 20 | 2022-05-23 10:45 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 11:00 | Stayton Water Supply | May | PWS00843:SRC-AA | 2205217-01 | Source Water | 142 | 20 | 2022-05-23 11:00 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0055 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 11:30 | City of Lowell | May | PWS00492:SRC-BA | 2205202-01 | Source Water | 142 | 20 | 2022-05-23 11:30 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0035 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 11:30 | Roberts Creek Water District | May | PWS00717:SRC-AA | 2205213-01 | Source Water | 142 | 20 | 2022-05-23 11:30 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 11:40 | City of Drain Shops | May | PWS00260:SRC-AA | 2205198-01 | Source Water | 142 | 20 | 2022-05-23 11:40 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0085 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 11:48 | City of Roseburg | May | PWS00720:SRC-AA | 2205207-01 | Source Water | 142 | 20 | 2022-05-23 11:48 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.024 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 12:00 | Josephine Co Pks Lake Selmac 2 | May | PWS94645:SRC-AA | 2205209-01 | Source Water | 142 | 20 | 2022-05-23 12:00 | 2022-05-25 10:30 | 2022-05-26 11:51 | B22E254 | SD
Microcystins ADDA, Total | 0.0031 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 12:22 | City of Monroe | May | PWS00540:SRC-BA | 2205203-01 | Source Water | 142 | 20 | 2022-05-23 12:22 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0053 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 12:30 | Ashland Water Department | May | PWS00047:SRC-AA | 2205195-01 | Source Water | 142 | 20 | 2022-05-23 12:30 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0077 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 12:30 | Josephine Co Pks Lake Selmac 1 | May | PWS90186:SRC-AA | 2205208-01 | Source Water | 142 | 20 | 2022-05-23 12:30 | 2022-05-25 10:30 | 2022-05-26 11:51 | B22E254 | SD
Microcystins ADDA, Total | 0.0045 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 13:00 | Seaside Water Department | May | PWS00799:CH-A | 2205214-01 | Source Water | 142 | 20 | 2022-05-23 13:00 | 2022-05-25 10:30 | 2022-05-26 11:51 | B22E254 | SD
Microcystins ADDA, Total | 0.0073 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 14:00 | Medford Water Commission | May | PWS00513:SRC-BA | 2205211-01 | Source Water | 142 | 20 | 2022-05-23 14:00 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0054 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 14:05 | City of Rogue River | May | PWS00712:SRC-BA | 2205206-01 | Source Water | 142 | 20 | 2022-05-23 14:05 | 2022-05-24 11:15 | 2022-05-24 18:49 | B22E244 | KMC
Microcystins ADDA, Total | 0.0238 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-23 14:50 | City of Gold Hill | May | PWS00333:SRC-AA | 2205200-01 | Source Water | 142 | 20 | 2022-05-23 14:50 | 2022-05-25 10:30 | 2022-05-26 11:51 | B22E254 | SD
Microcystins ADDA, Total | 0.0182 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-24 10:30 | City of Myrtle Creek | May | PWS00550:SRC-AA | 2205204-01 | Source Water | 143 | 20 | 2022-05-24 10:30 | 2022-05-25 10:30 | 2022-05-26 11:51 | B22E254 | SD
Microcystins ADDA, Total | 0.0072 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-26 09:45 | Umpqua Ranch Co-op | May | PWS00714:SRC-AH | 2205254-01 | Source Water | 145 | 20 | 2022-05-26 09:45 | 2022-05-27 10:30 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0122 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 05:15 | Anglers Cove/SCHWC | May | PWS01483:SRC-AB | 2205256-01 | Source Water | 150 | 21 | 2022-05-31 05:15 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 05:30 | Hiland WC - Shady Cove | May | PWS01520:SRC-AA | 2205269-01 | Source Water | 150 | 21 | 2022-05-31 05:30 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0123 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 06:15 | City of Albany | May | PWS00012:SRC-BA | 2205259-01 | Source Water | 150 | 21 | 2022-05-31 06:15 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0029 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 07:30 | Winston-Dillard Water District | May | PWS00957:SRC-AA | 2205281-01 | Source Water | 150 | 21 | 2022-05-31 07:30 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0052 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 08:00 | Roseburg Forest Products - Dillard | May | PWS94300:SRC-AA | 2205275-01 | Source Water | 150 | 21 | 2022-05-31 08:00 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0072 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 08:02 | Joint Water Commission | May | PWS00379:SRC-AA | 2205271-01 | Source Water | 150 | 21 | 2022-05-31 08:02 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0051 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 08:20 | Salem Public Works | May | PWS00731:SRC-AA | 2205276-01 | Source Water | 150 | 21 | 2022-05-31 08:20 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 08:29 | USFS Tiller Ranger Station | May | PWS01092:SRC-AA | 2205280-01 | Source Water | 150 | 21 | 2022-05-31 08:29 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0084 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 08:30 | Glide Water Association | May | PWS00326:SRC-AA | 2205268-01 | Source Water | 150 | 21 | 2022-05-31 08:30 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 09:10 | Springfield Utility Board | May | PWS00837:CH-A | 2205277-02 | Source Water | 150 | 21 | 2022-05-31 09:10 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0089 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 09:20 | Jackson Co Pks Emigrant Lake | May | PWS90730:CH-A | 2205270-01 | Source Water | 150 | 21 | 2022-05-31 09:20 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0055 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 09:30 | Buell-Red Prairie Water Association | May | PWS01174:SRC-AA | 2205257-01 | Source Water | 150 | 21 | 2022-05-31 09:30 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 09:30 | North Clackamas County Water Commission | May | PWS00580:SRC-BA | 2205273-01 | Source Water | 150 | 21 | 2022-05-31 09:30 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 10:05 | Springfield Utility Board | May | PWS00837:SRC-BB | 2205277-01 | Source Water | 150 | 21 | 2022-05-31 10:05 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0036 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 10:10 | City of Powers | May | PWS00672:SRC-AA | 2205264-01 | Source Water | 150 | 21 | 2022-05-31 10:10 | 2022-06-02 10:08 | 2022-06-03 18:08 | B22F016 | KMC
Microcystins ADDA, Total | 0.0039 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 10:10 | Clarks Branch Water Association | May | PWS00548:SRC-AA | 2205266-01 | Source Water | 150 | 21 | 2022-05-31 10:10 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0112 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 10:20 | Eugene Water & Electric Board | May | PWS00287:SRC-AA | 2205267-01 | Source Water | 150 | 21 | 2022-05-31 10:20 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0049 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 10:30 | Lake Oswego - Tigard Water Supply | May | PWS01542:SRC-AA | 2205272-01 | Source Water | 150 | 21 | 2022-05-31 10:30 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0093 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 11:36 | City of Pendleton | May | PWS00613:SRC-AM | 2205263-01 | Source Water | 150 | 21 | 2022-05-31 11:36 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0079 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 11:45 | City of Gates | May | PWS00317:SRC-AA | 2205261-01 | Source Water | 150 | 21 | 2022-05-31 11:45 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0027 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 11:52 | PP&L-Toketee Village | May | PWS01012:SRC-AA | 2205274-01 | Source Water | 150 | 21 | 2022-05-31 11:52 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0016 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 12:00 | City of Cottage Grove | May | PWS00236:SRC-BA | 2205260-01 | Source Water | 150 | 21 | 2022-05-31 12:00 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0047 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 12:10 | USFS Horseshoe Bend CG | May | PWS94179:SRC-AA | 2205278-01 | Source Water | 150 | 21 | 2022-05-31 12:10 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0089 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 12:26 | USFS Steamboat Work Center | May | PWS01091:SRC-AA | 2205279-01 | Source Water | 150 | 21 | 2022-05-31 12:26 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 12:33 | City of Wilsonville | May | PWS00954:SRC-HA | 2205265-01 | Source Water | 150 | 21 | 2022-05-31 12:33 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0186 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 12:54 | City of Grants Pass | May | PWS00342:SRC-AA | 2205262-01 | Source Water | 150 | 21 | 2022-05-31 12:54 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0077 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-05-31 13:50 | Canby Utility/Veolia Water | May | PWS00157:CH-A | 2205258-01 | Source Water | 150 | 21 | 2022-05-31 13:50 | 2022-06-01 11:15 | 2022-06-01 19:48 | B22F008 | KMC
Microcystins ADDA, Total | 0.0015 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-01 10:24 | City of Monroe | June | PWS00540:SRC-BA | 2206012-01 | Source Water | 150 | 21 | 2022-06-01 10:24 | 2022-06-02 10:30 | 2022-06-03 18:08 | B22F016 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 07:30 | Umpqua Ranch Co-op | June | PWS00714:SRC-AH | 2206056-01 | Source Water | 155 | 22 | 2022-06-06 07:30 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0025 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 08:20 | South Fork Water Board - Oregon City | June | PWS00591:SRC-AA | 2206052-01 | Source Water | 155 | 22 | 2022-06-06 08:20 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0087 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 09:00 | City of Creswell | June | PWS00246:SRC-BK | 2206033-01 | Source Water | 155 | 22 | 2022-06-06 09:00 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 09:30 | Shangri La Water District | June | PWS00835:SRC-AA | 2206051-01 | Source Water | 155 | 22 | 2022-06-06 09:30 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 09:55 | Lyons Mehama Water District | June | PWS00493:SRC-AA | 2206046-01 | Source Water | 155 | 22 | 2022-06-06 09:55 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0022 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 10:00 | Roberts Creek Water District | June | PWS00717:SRC-AA | 2206049-01 | Source Water | 155 | 22 | 2022-06-06 10:00 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 10:30 | City of Jefferson | June | PWS00408:SRC-AA | 2206037-01 | Source Water | 155 | 22 | 2022-06-06 10:30 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0159 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 10:30 | City of Myrtle Creek | June | PWS00550:SRC-AA | 2206040-01 | Source Water | 155 | 22 | 2022-06-06 10:30 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0013 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 10:30 | City of Newport | June | PWS00566:CH-A | 2206041-01 | Source Water | 155 | 22 | 2022-06-06 10:30 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0083 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 10:40 | Camp Baker BSA | June | PWS91786:SRC-AA | 2206032-01 | Source Water | 155 | 22 | 2022-06-06 10:40 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0026 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 10:45 | Umpqua Basin Water Association | June | PWS00719:SRC-AA | 2206055-01 | Source Water | 155 | 22 | 2022-06-06 10:45 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0169 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 10:57 | City of Monroe | June | PWS00540:SRC-BA | 2206039-01 | Source Water | 155 | 22 | 2022-06-06 10:57 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0064 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 11:00 | City of Lowell | June | PWS00492:SRC-BA | 2206038-01 | Source Water | 155 | 22 | 2022-06-06 11:00 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 11:30 | Stayton Water Supply | June | PWS00843:SRC-AA | 2206053-01 | Source Water | 155 | 22 | 2022-06-06 11:30 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0069 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 11:40 | City of Roseburg | June | PWS00720:SRC-AA | 2206043-01 | Source Water | 155 | 22 | 2022-06-06 11:40 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0088 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 11:50 | City of Estacada | June | PWS00279:SRC-AA | 2206035-01 | Source Water | 155 | 22 | 2022-06-06 11:50 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0105 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 12:00 | Josephine Co Pks Lake Selmac 1 | June | PWS90186:SRC-AA | 2206044-01 | Source Water | 155 | 22 | 2022-06-06 12:00 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0102 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 12:30 | Ashland Water Department | June | PWS00047:SRC-AA | 2206031-01 | Source Water | 155 | 22 | 2022-06-06 12:30 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0162 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 12:30 | Josephine Co Pks Lake Selmac 2 | June | PWS94645:SRC-AA | 2206045-01 | Source Water | 155 | 22 | 2022-06-06 12:30 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 12:40 | Rainbow Water District | June | PWS00839:SRC-BB | 2206048-01 | Source Water | 155 | 22 | 2022-06-06 12:40 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0116 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 13:00 | Medford Water Commission | June | PWS00513:SRC-BA | 2206047-01 | Source Water | 155 | 22 | 2022-06-06 13:00 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0067 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 13:00 | Tri-City JW & SA | June | PWS00549:SRC-AA | 2206054-01 | Source Water | 155 | 22 | 2022-06-06 13:00 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0066 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-06 13:45 | Seaside Water Department | June | PWS00799:CH-A | 2206050-01 | Source Water | 155 | 22 | 2022-06-06 13:45 | 2022-06-07 11:30 | 2022-06-07 20:15 | B22F059 | KMC
Microcystins ADDA, Total | 0.0073 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-07 11:15 | City of Drain Shops | June | PWS00260:SRC-AA | 2206034-01RE1 | Source Water | 156 | 22 | 2022-06-07 11:15 | 2022-06-08 13:00 | 2022-06-10 11:55 | B22F109 | SD
Microcystins ADDA, Total | 0.021 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-07 11:20 | City of Rogue River | June | PWS00712:SRC-BA | 2206042-01RE1 | Source Water | 156 | 22 | 2022-06-07 11:20 | 2022-06-08 13:00 | 2022-06-10 11:55 | B22F109 | SD
Microcystins ADDA, Total | 0.0192 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-07 14:10 | City of Gold Hill | June | PWS00333:SRC-AA | 2206036-01RE1 | Source Water | 156 | 22 | 2022-06-07 14:10 | 2022-06-08 13:00 | 2022-06-10 11:55 | B22F109 | SD
Microcystins ADDA, Total | 0.0072 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 05:15 | Anglers Cove/SCHWC | June | PWS01483:SRC-AB | 2206107-01 | Source Water | 162 | 23 | 2022-06-13 05:15 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0236 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 05:30 | Hiland WC - Shady Cove | June | PWS01520:SRC-AA | 2206120-01 | Source Water | 162 | 23 | 2022-06-13 05:30 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 06:45 | City of Albany | June | PWS00012:SRC-BA | 2206110-01 | Source Water | 162 | 23 | 2022-06-13 06:45 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0149 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 07:45 | Springfield Utility Board | June | PWS00837:CH-A | 2206128-02 | Source Water | 162 | 23 | 2022-06-13 07:45 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0269 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 07:47 | Jackson Co Pks Emigrant Lake | June | PWS90730:CH-A | 2206121-01 | Source Water | 162 | 23 | 2022-06-13 07:47 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 08:18 | Roseburg Forest Products - Dillard | June | PWS94300:SRC-AA | 2206126-01 | Source Water | 162 | 23 | 2022-06-13 08:18 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0015 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 08:22 | Joint Water Commission | June | PWS00379:SRC-AA | 2206122-01 | Source Water | 162 | 23 | 2022-06-13 08:22 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0073 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 08:45 | Glide Water Association | June | PWS00326:SRC-AA | 2206119-01 | Source Water | 162 | 23 | 2022-06-13 08:45 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0139 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 09:10 | USFS Tiller Ranger Station | June | PWS01092:SRC-AA | 2206131-01 | Source Water | 162 | 23 | 2022-06-13 09:10 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.032 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 09:30 | Winston-Dillard Water District | June | PWS00957:SRC-AA | 2206132-01 | Source Water | 162 | 23 | 2022-06-13 09:30 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 09:40 | Springfield Utility Board | June | PWS00837:SRC-BB | 2206128-01 | Source Water | 162 | 23 | 2022-06-13 09:40 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0173 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 09:50 | North Clackamas County Water Commission | June | PWS00580:SRC-BA | 2206124-01 | Source Water | 162 | 23 | 2022-06-13 09:50 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0199 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 10:30 | Buell-Red Prairie Water Association | June | PWS01174:SRC-AA | 2206108-01 | Source Water | 162 | 23 | 2022-06-13 10:30 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0318 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 10:34 | Clarks Branch Water Association | June | PWS00548:SRC-AA | 2206117-01 | Source Water | 162 | 23 | 2022-06-13 10:34 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0094 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 10:40 | City of Powers | June | PWS00672:SRC-AA | 2206115-01 | Source Water | 162 | 23 | 2022-06-13 10:40 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0055 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 11:00 | Lake Oswego - Tigard Water Supply | June | PWS01542:SRC-AA | 2206123-01 | Source Water | 162 | 23 | 2022-06-13 11:00 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0145 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 11:15 | City of Grants Pass | June | PWS00342:SRC-AA | 2206113-01 | Source Water | 162 | 23 | 2022-06-13 11:15 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0027 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 11:20 | City of Pendleton | June | PWS00613:SRC-AM | 2206114-01 | Source Water | 162 | 23 | 2022-06-13 11:20 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0061 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 11:20 | Salem Public Works | June | PWS00731:SRC-AA | 2206127-01 | Source Water | 162 | 23 | 2022-06-13 11:20 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0081 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 12:05 | USFS Steamboat Work Center | June | PWS01091:SRC-AA | 2206130-01 | Source Water | 162 | 23 | 2022-06-13 12:05 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 12:10 | Eugene Water & Electric Board | June | PWS00287:SRC-AA | 2206118-01 | Source Water | 162 | 23 | 2022-06-13 12:10 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0076 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 12:15 | City of Cottage Grove | June | PWS00236:SRC-BA | 2206111-01 | Source Water | 162 | 23 | 2022-06-13 12:15 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0211 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 12:28 | City of Wilsonville | June | PWS00954:SRC-HA | 2206116-01 | Source Water | 162 | 23 | 2022-06-13 12:28 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0036 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 12:28 | USFS Horseshoe Bend CG | June | PWS94179:SRC-AA | 2206129-01 | Source Water | 162 | 23 | 2022-06-13 12:28 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.019 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 12:35 | PP&L-Toketee Village | June | PWS01012:SRC-AA | 2206125-01 | Source Water | 162 | 23 | 2022-06-13 12:35 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0024 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 13:45 | City of Gates | June | PWS00317:SRC-AA | 2206112-01 | Source Water | 162 | 23 | 2022-06-13 13:45 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0153 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-13 14:30 | Canby Utility/Veolia Water | June | PWS00157:CH-A | 2206109-01 | Source Water | 162 | 23 | 2022-06-13 14:30 | 2022-06-14 11:00 | 2022-06-14 20:12 | B22F152 | KMC
Microcystins ADDA, Total | 0.0153 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 08:26 | Lyons Mehama Water District | June | PWS00493:SRC-AA | 2206183-01 | Source Water | 169 | 24 | 2022-06-20 08:26 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 09:30 | Shangri La Water District | June | PWS00835:SRC-AA | 2206188-01 | Source Water | 169 | 24 | 2022-06-20 09:30 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0.0118 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 10:15 | City of Newport | June | PWS00566:CH-A | 2206178-01 | Source Water | 169 | 24 | 2022-06-20 10:15 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0.0052 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 10:15 | South Fork Water Board - Oregon City | June | PWS00591:SRC-AA | 2206189-01 | Source Water | 169 | 24 | 2022-06-20 10:15 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0.0142 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 11:00 | City of Lowell | June | PWS00492:SRC-BA | 2206175-01 | Source Water | 169 | 24 | 2022-06-20 11:00 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0.0018 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 11:00 | Roberts Creek Water District | June | PWS00717:SRC-AA | 2206186-01 | Source Water | 169 | 24 | 2022-06-20 11:00 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0.0083 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 11:50 | City of Drain Shops | June | PWS00260:SRC-AA | 2206171-01 | Source Water | 169 | 24 | 2022-06-20 11:50 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0.0051 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 11:57 | City of Roseburg | June | PWS00720:SRC-AA | 2206180-01 | Source Water | 169 | 24 | 2022-06-20 11:57 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0.0172 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 13:00 | Tri-City JW & SA | June | PWS00549:SRC-AA | 2206191-01 | Source Water | 169 | 24 | 2022-06-20 13:00 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0.0008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-20 13:00 | Umpqua Ranch Co-op | June | PWS00714:SRC-AH | 2206193-01 | Source Water | 169 | 24 | 2022-06-20 13:00 | 2022-06-21 13:00 | 2022-06-22 11:47 | B22F216 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 08:30 | City of Creswell | June | PWS00246:SRC-BK | 2206170-01 | Source Water | 170 | 24 | 2022-06-21 08:30 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 08:55 | Camp Baker BSA | June | PWS91786:SRC-AA | 2206169-01 | Source Water | 170 | 24 | 2022-06-21 08:55 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0054 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 10:20 | Umpqua Basin Water Association | June | PWS00719:SRC-AA | 2206192-01 | Source Water | 170 | 24 | 2022-06-21 10:20 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0015 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 10:29 | City of Monroe | June | PWS00540:SRC-BA | 2206176-01 | Source Water | 170 | 24 | 2022-06-21 10:29 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0088 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 10:30 | City of Myrtle Creek | June | PWS00550:SRC-AA | 2206177-01 | Source Water | 170 | 24 | 2022-06-21 10:30 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0186 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 10:50 | City of Jefferson | June | PWS00408:SRC-AA | 2206174-01 | Source Water | 170 | 24 | 2022-06-21 10:50 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 11:00 | Stayton Water Supply | June | PWS00843:SRC-AA | 2206190-01 | Source Water | 170 | 24 | 2022-06-21 11:00 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0018 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 11:05 | City of Estacada | June | PWS00279:SRC-AA | 2206172-01 | Source Water | 170 | 24 | 2022-06-21 11:05 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0081 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 12:00 | Josephine Co Pks Lake Selmac 1 | June | PWS90186:SRC-AA | 2206181-01 | Source Water | 170 | 24 | 2022-06-21 12:00 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0147 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 12:15 | Rainbow Water District | June | PWS00839:SRC-BB | 2206185-01 | Source Water | 170 | 24 | 2022-06-21 12:15 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 12:30 | Ashland Water Department | June | PWS00047:SRC-AA | 2206168-01 | Source Water | 170 | 24 | 2022-06-21 12:30 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-21 13:30 | Seaside Water Department | June | PWS00799:CH-A | 2206187-01 | Source Water | 170 | 24 | 2022-06-21 13:30 | 2022-06-22 13:30 | 2022-06-23 12:41 | B22F234 | SD
Microcystins ADDA, Total | 0.0154 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-22 13:00 | Medford Water Commission | June | PWS00513:SRC-BA | 2206208-01 | Source Water | 171 | 24 | 2022-06-22 13:00 | 2022-06-23 11:15 | 2022-06-28 12:38 | B22F251 | KMC
Microcystins ADDA, Total | 0.0046 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 05:15 | Anglers Cove/SCHWC | June | PWS01483:SRC-AB | 2206238-01 | Source Water | 176 | 25 | 2022-06-27 05:15 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0229 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 05:30 | Hiland WC - Shady Cove | June | PWS01520:SRC-AA | 2206251-01 | Source Water | 176 | 25 | 2022-06-27 05:30 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0048 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 06:15 | City of Albany | June | PWS00012:SRC-BA | 2206241-01 | Source Water | 176 | 25 | 2022-06-27 06:15 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.034 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 07:56 | Clarks Branch Water Association | June | PWS00548:SRC-AA | 2206248-01 | Source Water | 176 | 25 | 2022-06-27 07:56 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0182 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 08:00 | Salem Public Works | June | PWS00731:SRC-AA | 2206258-01 | Source Water | 176 | 25 | 2022-06-27 08:00 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0029 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 08:10 | City of Powers | June | PWS00672:SRC-AA | 2206246-01 | Source Water | 176 | 25 | 2022-06-27 08:10 | 2022-06-28 10:55 | 2022-06-29 10:58 | B22F286 | SD
Microcystins ADDA, Total | 0.0446 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 08:17 | Roseburg Forest Products - Dillard | June | PWS94300:SRC-AA | 2206257-01 | Source Water | 176 | 25 | 2022-06-27 08:17 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0011 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 08:30 | Springfield Utility Board | June | PWS00837:CH-A | 2206259-02 | Source Water | 176 | 25 | 2022-06-27 08:30 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 08:45 | North Clackamas County Water Commission | June | PWS00580:SRC-BA | 2206255-01 | Source Water | 176 | 25 | 2022-06-27 08:45 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0301 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 09:30 | Winston-Dillard Water District | June | PWS00957:SRC-AA | 2206263-01 | Source Water | 176 | 25 | 2022-06-27 09:30 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0202 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 09:35 | Jackson Co Pks Emigrant Lake | June | PWS90730:CH-A | 2206252-01 | Source Water | 176 | 25 | 2022-06-27 09:35 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0014 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 09:55 | PP&L-Toketee Village | June | PWS01012:SRC-AA | 2206256-01 | Source Water | 176 | 25 | 2022-06-27 09:55 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 2.6573 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-06-27 10:00 | Buell-Red Prairie Water Association | June | PWS01174:SRC-AA | 2206239-01 | Source Water | 176 | 25 | 2022-06-27 10:00 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0092 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 10:20 | USFS Horseshoe Bend CG | June | PWS94179:SRC-AA | 2206260-01 | Source Water | 176 | 25 | 2022-06-27 10:20 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0089 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 10:45 | USFS Steamboat Work Center | June | PWS01091:SRC-AA | 2206261-01 | Source Water | 176 | 25 | 2022-06-27 10:45 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0108 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 10:56 | City of Pendleton | June | PWS00613:SRC-AM | 2206245-01 | Source Water | 176 | 25 | 2022-06-27 10:56 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 11:00 | Springfield Utility Board | June | PWS00837:SRC-BB | 2206259-01 | Source Water | 176 | 25 | 2022-06-27 11:00 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 11:22 | USFS Tiller Ranger Station | June | PWS01092:SRC-AA | 2206262-01 | Source Water | 176 | 25 | 2022-06-27 11:22 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0062 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 12:00 | City of Cottage Grove | June | PWS00236:SRC-BA | 2206242-01 | Source Water | 176 | 25 | 2022-06-27 12:00 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0354 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 12:00 | Josephine Co Pks Lake Selmac 2 | June | PWS94645:SRC-AA | 2206266-01 | Source Water | 176 | 25 | 2022-06-27 12:00 | 2022-06-28 10:55 | 2022-06-29 10:58 | B22F286 | SD
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 13:00 | Lake Oswego - Tigard Water Supply | June | PWS01542:SRC-AA | 2206254-01 | Source Water | 176 | 25 | 2022-06-27 13:00 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0048 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 13:07 | City of Wilsonville | June | PWS00954:SRC-HA | 2206247-01 | Source Water | 176 | 25 | 2022-06-27 13:07 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0122 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 13:10 | Glide Water Association | June | PWS00326:SRC-AA | 2206250-01 | Source Water | 176 | 25 | 2022-06-27 13:10 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0265 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 13:25 | City of Grants Pass | June | PWS00342:SRC-AA | 2206244-01 | Source Water | 176 | 25 | 2022-06-27 13:25 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0141 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 14:00 | Canby Utility/Veolia Water | June | PWS00157:CH-A | 2206240-01 | Source Water | 176 | 25 | 2022-06-27 14:00 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0099 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-27 15:37 | Eugene Water & Electric Board | June | PWS00287:SRC-AA | 2206249-01 | Source Water | 176 | 25 | 2022-06-27 15:37 | 2022-06-29 12:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0023 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-28 09:46 | Joint Water Commission | June | PWS00379:SRC-AA | 2206253-01 | Source Water | 177 | 25 | 2022-06-28 09:46 | 2022-06-29 11:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0123 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-28 11:39 | City of Gold Hill | June | PWS00333:SRC-AA | 2206275-01 | Source Water | 177 | 25 | 2022-06-28 11:39 | 2022-06-29 11:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0133 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-28 12:00 | City of Rogue River | June | PWS00712:SRC-BA | 2206274-01 | Source Water | 177 | 25 | 2022-06-28 12:00 | 2022-06-29 11:30 | 2022-06-29 20:45 | B22F330 | KMC
Microcystins ADDA, Total | 0.0315 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-06-29 11:35 | City of Gates | June | PWS00317:SRC-AA | 2206279-01 | Source Water | 178 | 25 | 2022-06-29 11:35 | 2022-06-30 15:30 | 2022-07-01 12:02 | B22F339 | SD
Microcystins ADDA, Total | 0.0558 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 08:15 | Tri-City JW & SA | July | PWS00549:SRC-AA | 2207030-01 | Source Water | 184 | 26 | 2022-07-05 08:15 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0233 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 08:30 | Umpqua Basin Water Association | July | PWS00719:SRC-AA | 2207031-01 | Source Water | 184 | 26 | 2022-07-05 08:30 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0184 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 09:00 | Umpqua Ranch Co-op | July | PWS00714:SRC-AH | 2207032-01 | Source Water | 184 | 26 | 2022-07-05 09:00 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0149 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 09:15 | City of Newport | July | PWS00566:CH-A | 2207017-01 | Source Water | 184 | 26 | 2022-07-05 09:15 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0047 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 09:30 | City of Creswell | July | PWS00246:SRC-BK | 2207009-01 | Source Water | 184 | 26 | 2022-07-05 09:30 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0114 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 09:30 | Shangri La Water District | July | PWS00835:SRC-AA | 2207027-01 | Source Water | 184 | 26 | 2022-07-05 09:30 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0157 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 10:00 | Buell-Red Prairie Water Association | July | PWS01174:SRC-AA | 2207041-01 | Source Water | 184 | 26 | 2022-07-05 10:00 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0112 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 10:06 | South Fork Water Board - Oregon City | July | PWS00591:SRC-AA | 2207028-01 | Source Water | 184 | 26 | 2022-07-05 10:06 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0059 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 10:10 | City of Lowell | July | PWS00492:SRC-BA | 2207014-01 | Source Water | 184 | 26 | 2022-07-05 10:10 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.043 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 10:30 | Roberts Creek Water District | July | PWS00717:SRC-AA | 2207025-01 | Source Water | 184 | 26 | 2022-07-05 10:30 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0159 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 10:45 | City of Gold Hill | July | PWS00333:SRC-AA | 2207012-01 | Source Water | 184 | 26 | 2022-07-05 10:45 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.053 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 11:00 | City of Myrtle Creek | July | PWS00550:SRC-AA | 2207016-01 | Source Water | 184 | 26 | 2022-07-05 11:00 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0106 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 11:00 | Stayton Water Supply | July | PWS00843:SRC-AA | 2207029-01 | Source Water | 184 | 26 | 2022-07-05 11:00 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0091 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 11:12 | City of Monroe | July | PWS00540:SRC-BA | 2207015-01 | Source Water | 184 | 26 | 2022-07-05 11:12 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 11:15 | Camp Baker BSA | July | PWS91786:SRC-AA | 2207008-01 | Source Water | 184 | 26 | 2022-07-05 11:15 | 2022-07-07 11:00 | 2022-07-08 11:23 | B22G068 | SD
Microcystins ADDA, Total | 0.02 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 11:20 | City of Jefferson | July | PWS00408:SRC-AA | 2207013-01 | Source Water | 184 | 26 | 2022-07-05 11:20 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0105 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 11:25 | City of Rogue River | July | PWS00712:SRC-BA | 2207018-01 | Source Water | 184 | 26 | 2022-07-05 11:25 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0065 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 11:30 | City of Estacada | July | PWS00279:SRC-AA | 2207011-01 | Source Water | 184 | 26 | 2022-07-05 11:30 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0116 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 11:43 | City of Roseburg | July | PWS00720:SRC-AA | 2207019-01 | Source Water | 184 | 26 | 2022-07-05 11:43 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0156 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 11:50 | Rainbow Water District | July | PWS00839:SRC-BB | 2207024-01 | Source Water | 184 | 26 | 2022-07-05 11:50 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0084 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 12:00 | City of Drain Shops | July | PWS00260:SRC-AA | 2207010-01 | Source Water | 184 | 26 | 2022-07-05 12:00 | 2022-07-07 11:00 | 2022-07-08 11:23 | B22G068 | SD
Microcystins ADDA, Total | 0.0064 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 12:00 | Josephine Co Pks Lake Selmac 1 | July | PWS90186:SRC-AA | 2207020-01 | Source Water | 184 | 26 | 2022-07-05 12:00 | 2022-07-07 11:00 | 2022-07-08 11:23 | B22G068 | SD
Microcystins ADDA, Total | 0.0037 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 12:30 | Ashland Water Department | July | PWS00047:SRC-AA | 2207007-01 | Source Water | 184 | 26 | 2022-07-05 12:30 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0213 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 12:30 | Josephine Co Pks Lake Selmac 2 | July | PWS94645:SRC-AA | 2207021-01 | Source Water | 184 | 26 | 2022-07-05 12:30 | 2022-07-07 11:00 | 2022-07-08 11:23 | B22G068 | SD
Microcystins ADDA, Total | 0.0226 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 14:00 | Medford Water Commission | July | PWS00513:SRC-BA | 2207023-01 | Source Water | 184 | 26 | 2022-07-05 14:00 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0054 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-05 14:00 | Seaside Water Department | July | PWS00799:CH-A | 2207026-01 | Source Water | 184 | 26 | 2022-07-05 14:00 | 2022-07-06 12:00 | 2022-07-07 13:27 | B22G048 | KMC
Microcystins ADDA, Total | 0.0185 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-06 12:30 | Lyons Mehama Water District | July | PWS00493:SRC-AA | 2207022-01 | Source Water | 185 | 26 | 2022-07-06 12:30 | 2022-07-07 11:00 | 2022-07-08 11:23 | B22G068 | SD
Microcystins ADDA, Total | 0.0169 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 05:15 | Anglers Cove/SCHWC | July | PWS01483:SRC-AB | 2207056-01 | Source Water | 190 | 27 | 2022-07-11 05:15 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0229 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 05:30 | Hiland WC - Shady Cove | July | PWS01520:SRC-AA | 2207069-01 | Source Water | 190 | 27 | 2022-07-11 05:30 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0026 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 07:10 | City of Wilsonville | July | PWS00954:SRC-HA | 2207065-01 | Source Water | 190 | 27 | 2022-07-11 07:10 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0277 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 07:30 | Salem Public Works | July | PWS00731:SRC-AA | 2207076-01 | Source Water | 190 | 27 | 2022-07-11 07:30 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0134 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 08:45 | Springfield Utility Board | July | PWS00837:CH-A | 2207077-02 | Source Water | 190 | 27 | 2022-07-11 08:45 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 08:47 | Joint Water Commission | July | PWS00379:SRC-AA | 2207071-01 | Source Water | 190 | 27 | 2022-07-11 08:47 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0155 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 09:10 | Springfield Utility Board | July | PWS00837:SRC-BB | 2207077-01 | Source Water | 190 | 27 | 2022-07-11 09:10 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0116 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 09:10 | USFS Tiller Ranger Station | July | PWS01092:SRC-AA | 2207080-01 | Source Water | 190 | 27 | 2022-07-11 09:10 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0195 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 09:15 | North Clackamas County Water Commission | July | PWS00580:SRC-BA | 2207073-01 | Source Water | 190 | 27 | 2022-07-11 09:15 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0269 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 09:30 | Glide Water Association | July | PWS00326:SRC-AA | 2207068-01 | Source Water | 190 | 27 | 2022-07-11 09:30 | 2022-07-13 11:02 | 2022-07-14 11:05 | B22G127 | SD
Microcystins ADDA, Total | 0.0492 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 10:00 | Buell-Red Prairie Water Association | July | PWS01174:SRC-AA | 2207057-01 | Source Water | 190 | 27 | 2022-07-11 10:00 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0743 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-07-11 10:15 | Winston-Dillard Water District | July | PWS00957:SRC-AA | 2207081-01 | Source Water | 190 | 27 | 2022-07-11 10:15 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0269 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 10:31 | Clarks Branch Water Association | July | PWS00548:SRC-AA | 2207066-01 | Source Water | 190 | 27 | 2022-07-11 10:31 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0206 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 10:45 | City of Powers | July | PWS00672:SRC-AA | 2207064-01 | Source Water | 190 | 27 | 2022-07-11 10:45 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0407 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 11:05 | Roseburg Forest Products - Dillard | July | PWS94300:SRC-AA | 2207075-01 | Source Water | 190 | 27 | 2022-07-11 11:05 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 11:30 | Douglas County Parks - Whistler's Bend | July | PWS93944:SRC-AB | 2207087-01 | Source Water | 190 | 27 | 2022-07-11 11:30 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 11:36 | USFS Steamboat Work Center | July | PWS01091:SRC-AA | 2207079-01 | Source Water | 190 | 27 | 2022-07-11 11:36 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0023 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 11:47 | City of Pendleton | July | PWS00613:SRC-AM | 2207063-01 | Source Water | 190 | 27 | 2022-07-11 11:47 | 2022-07-13 11:02 | 2022-07-14 11:05 | B22G127 | SD
Microcystins ADDA, Total | 0.0037 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 11:59 | USFS Horseshoe Bend CG | July | PWS94179:SRC-AA | 2207078-01 | Source Water | 190 | 27 | 2022-07-11 11:59 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0124 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 12:19 | PP&L-Toketee Village | July | PWS01012:SRC-AA | 2207074-01 | Source Water | 190 | 27 | 2022-07-11 12:19 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0131 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 13:00 | City of Cottage Grove | July | PWS00236:SRC-BA | 2207060-01 | Source Water | 190 | 27 | 2022-07-11 13:00 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0247 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 13:00 | Eugene Water & Electric Board | July | PWS00287:SRC-AA | 2207067-01 | Source Water | 190 | 27 | 2022-07-11 13:00 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0042 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 14:00 | City of Albany | July | PWS00012:SRC-BA | 2207059-01 | Source Water | 190 | 27 | 2022-07-11 14:00 | 2022-07-13 11:02 | 2022-07-14 11:05 | B22G127 | SD
Microcystins ADDA, Total | 0.0514 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 14:00 | City of Gates | July | PWS00317:SRC-AA | 2207061-01 | Source Water | 190 | 27 | 2022-07-11 14:00 | 2022-07-13 11:02 | 2022-07-14 11:05 | B22G127 | SD
Microcystins ADDA, Total | 0.0084 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 14:45 | Canby Utility/Veolia Water | July | PWS00157:CH-A | 2207058-01 | Source Water | 190 | 27 | 2022-07-11 14:45 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0283 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 15:20 | City of Grants Pass | July | PWS00342:SRC-AA | 2207062-01 | Source Water | 190 | 27 | 2022-07-11 15:20 | 2022-07-13 11:02 | 2022-07-14 11:05 | B22G127 | SD
Microcystins ADDA, Total | 0.0003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-11 15:30 | Jackson Co Pks Emigrant Lake | July | PWS90730:CH-A | 2207070-01 | Source Water | 190 | 27 | 2022-07-11 15:30 | 2022-07-13 11:02 | 2022-07-14 11:05 | B22G127 | SD
Microcystins ADDA, Total | 0.0089 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-12 07:20 | Lake Oswego - Tigard Water Supply | July | PWS01542:SRC-AA | 2207072-01 | Source Water | 191 | 27 | 2022-07-12 07:20 | 2022-07-12 11:45 | 2022-07-12 20:04 | B22G105 | KMC
Microcystins ADDA, Total | 0.0054 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-17 16:00 | City of Creswell | July | PWS00246:SRC-BK | 2207138-01 | Source Water | 196 | 28 | 2022-07-17 16:00 | 2022-07-19 10:45 | 2022-07-19 18:52 | B22G190 | KMC
Microcystins ADDA, Total | 0.1185 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2022-07-18 07:50 | Tri-City JW & SA | July | PWS00549:SRC-AA | 2207159-01 | Source Water | 197 | 28 | 2022-07-18 07:50 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0154 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 08:29 | Lyons Mehama Water District | July | PWS00493:SRC-AA | 2207151-01 | Source Water | 197 | 28 | 2022-07-18 08:29 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 08:55 | Camp Baker BSA | July | PWS91786:SRC-AA | 2207137-01 | Source Water | 197 | 28 | 2022-07-18 08:55 | 2022-07-19 10:45 | 2022-07-19 18:52 | B22G190 | KMC
Microcystins ADDA, Total | 0.0194 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 09:00 | Josephine Co Pks Lake Selmac 1 | July | PWS90186:SRC-AA | 2207149-01 | Source Water | 197 | 28 | 2022-07-18 09:00 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0174 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 09:05 | City of Gold Hill | July | PWS00333:SRC-AA | 2207141-01 | Source Water | 197 | 28 | 2022-07-18 09:05 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.033 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 09:30 | Josephine Co Pks Lake Selmac 2 | July | PWS94645:SRC-AA | 2207150-01 | Source Water | 197 | 28 | 2022-07-18 09:30 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0041 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 09:30 | Shangri La Water District | July | PWS00835:SRC-AA | 2207156-01 | Source Water | 197 | 28 | 2022-07-18 09:30 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0099 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 09:34 | City of Roseburg | July | PWS00720:SRC-AA | 2207148-01 | Source Water | 197 | 28 | 2022-07-18 09:34 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0131 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 09:44 | City of Monroe | July | PWS00540:SRC-BA | 2207144-01 | Source Water | 197 | 28 | 2022-07-18 09:44 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.027 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 10:00 | City of Rogue River | July | PWS00712:SRC-BA | 2207147-01 | Source Water | 197 | 28 | 2022-07-18 10:00 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0291 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 10:20 | City of Jefferson | July | PWS00408:SRC-AA | 2207142-01 | Source Water | 197 | 28 | 2022-07-18 10:20 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0082 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 10:45 | Rainbow Water District | July | PWS00839:SRC-BB | 2207153-01 | Source Water | 197 | 28 | 2022-07-18 10:45 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0099 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 10:55 | Umpqua Basin Water Association | July | PWS00719:SRC-AA | 2207160-01 | Source Water | 197 | 28 | 2022-07-18 10:55 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.1 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-07-18 11:00 | City of Myrtle Creek | July | PWS00550:SRC-AA | 2207145-01 | Source Water | 197 | 28 | 2022-07-18 11:00 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.1101 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-07-18 11:00 | Roberts Creek Water District | July | PWS00717:SRC-AA | 2207154-01 | Source Water | 197 | 28 | 2022-07-18 11:00 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0438 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 11:00 | Stayton Water Supply | July | PWS00843:SRC-AA | 2207158-01 | Source Water | 197 | 28 | 2022-07-18 11:00 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0061 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 12:00 | South Fork Water Board - Oregon City | July | PWS00591:SRC-AA | 2207157-01 | Source Water | 197 | 28 | 2022-07-18 12:00 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0018 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 12:30 | Ashland Water Department | July | PWS00047:SRC-AA | 2207136-01 | Source Water | 197 | 28 | 2022-07-18 12:30 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0191 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 12:45 | City of Lowell | July | PWS00492:SRC-BA | 2207143-01 | Source Water | 197 | 28 | 2022-07-18 12:45 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0095 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 14:00 | Seaside Water Department | July | PWS00799:CH-A | 2207155-01 | Source Water | 197 | 28 | 2022-07-18 14:00 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0177 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 14:10 | City of Newport | July | PWS00566:CH-A | 2207146-01 | Source Water | 197 | 28 | 2022-07-18 14:10 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0194 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 14:40 | City of Drain Shops | July | PWS00260:SRC-AA | 2207139-01 | Source Water | 197 | 28 | 2022-07-18 14:40 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0168 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-18 17:00 | Medford Water Commission | July | PWS00513:SRC-BA | 2207152-01 | Source Water | 197 | 28 | 2022-07-18 17:00 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.016 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-19 07:30 | City of Estacada | July | PWS00279:SRC-AA | 2207140-01 | Source Water | 198 | 28 | 2022-07-19 07:30 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0111 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-19 10:30 | Umpqua Ranch Co-op | July | PWS00714:SRC-AH | 2207161-01 | Source Water | 198 | 28 | 2022-07-19 10:30 | 2022-07-20 12:00 | 2022-07-21 11:38 | B22G205 | SD
Microcystins ADDA, Total | 0.0121 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 05:15 | Anglers Cove/SCHWC | July | PWS01483:SRC-AB | 2207186-01 | Source Water | 204 | 29 | 2022-07-25 05:15 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0215 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 05:30 | Hiland WC - Shady Cove | July | PWS01520:SRC-AA | 2207199-01 | Source Water | 204 | 29 | 2022-07-25 05:30 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0065 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 07:00 | City of Albany | July | PWS00012:SRC-BA | 2207189-01 | Source Water | 204 | 29 | 2022-07-25 07:00 | 2022-07-27 13:50 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0188 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 07:50 | Lake Oswego - Tigard Water Supply | July | PWS01542:SRC-AA | 2207202-01 | Source Water | 204 | 29 | 2022-07-25 07:50 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0122 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 08:03 | Joint Water Commission | July | PWS00379:SRC-AA | 2207201-01 | Source Water | 204 | 29 | 2022-07-25 08:03 | 2022-07-27 07:30 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0173 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 08:10 | Springfield Utility Board | July | PWS00837:CH-A | 2207207-02 | Source Water | 204 | 29 | 2022-07-25 08:10 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0182 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 08:30 | Glide Water Association | July | PWS00326:SRC-AA | 2207198-01 | Source Water | 204 | 29 | 2022-07-25 08:30 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.1235 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2022-07-25 08:35 | Winston-Dillard Water District | July | PWS00957:SRC-AA | 2207211-01 | Source Water | 204 | 29 | 2022-07-25 08:35 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.1282 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2022-07-25 09:03 | Roseburg Forest Products - Dillard | July | PWS94300:SRC-AA | 2207205-01 | Source Water | 204 | 29 | 2022-07-25 09:03 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0087 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 09:06 | Jackson Co Pks Emigrant Lake | July | PWS90730:CH-A | 2207200-01 | Source Water | 204 | 29 | 2022-07-25 09:06 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0143 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 09:45 | North Clackamas County Water Commission | July | PWS00580:SRC-BA | 2207203-01 | Source Water | 204 | 29 | 2022-07-25 09:45 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0217 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 10:21 | Douglas County Parks - Whistler's Bend | July | PWS93944:SRC-AB | 2207224-01 | Source Water | 204 | 29 | 2022-07-25 10:21 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0238 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 10:24 | USFS Tiller Ranger Station | July | PWS01092:SRC-AA | 2207210-01 | Source Water | 204 | 29 | 2022-07-25 10:24 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.054 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 11:00 | Buell-Red Prairie Water Association | July | PWS01174:SRC-AA | 2207187-01 | Source Water | 204 | 29 | 2022-07-25 11:00 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0862 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-07-25 11:09 | Clarks Branch Water Association | July | PWS00548:SRC-AA | 2207196-01 | Source Water | 204 | 29 | 2022-07-25 11:09 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0282 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 11:23 | Salem Public Works | July | PWS00731:SRC-AA | 2207206-01 | Source Water | 204 | 29 | 2022-07-25 11:23 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0081 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 11:30 | City of Cottage Grove | July | PWS00236:SRC-BA | 2207190-01 | Source Water | 204 | 29 | 2022-07-25 11:30 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0239 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 12:13 | USFS Steamboat Work Center | July | PWS01091:SRC-AA | 2207209-01 | Source Water | 204 | 29 | 2022-07-25 12:13 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0151 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 12:40 | City of Wilsonville | July | PWS00954:SRC-HA | 2207195-01 | Source Water | 204 | 29 | 2022-07-25 12:40 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 12:49 | PP&L-Toketee Village | July | PWS01012:SRC-AA | 2207204-01 | Source Water | 204 | 29 | 2022-07-25 12:49 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0015 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 14:10 | Canby Utility/Veolia Water | July | PWS00157:CH-A | 2207188-01 | Source Water | 204 | 29 | 2022-07-25 14:10 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.013 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 14:10 | City of Grants Pass | July | PWS00342:SRC-AA | 2207192-01 | Source Water | 204 | 29 | 2022-07-25 14:10 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.063 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 15:45 | City of Gates | July | PWS00317:SRC-AA | 2207191-01 | Source Water | 204 | 29 | 2022-07-25 15:45 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0239 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 16:35 | Eugene Water & Electric Board | July | PWS00287:SRC-AA | 2207197-01 | Source Water | 204 | 29 | 2022-07-25 16:35 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0097 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 16:40 | City of Powers | July | PWS00672:SRC-AA | 2207194-01 | Source Water | 204 | 29 | 2022-07-25 16:40 | 2022-07-28 10:50 | 2022-07-29 10:27 | B22G280 | SD
Microcystins ADDA, Total | 0.0249 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-25 18:05 | Springfield Utility Board | July | PWS00837:SRC-BB | 2207207-01 | Source Water | 204 | 29 | 2022-07-25 18:05 | 2022-07-27 12:15 | 2022-07-28 15:49 | B22G266 | KMC
Microcystins ADDA, Total | 0.0191 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-26 11:56 | City of Pendleton | July | PWS00613:SRC-AM | 2207193-01 | Source Water | 205 | 29 | 2022-07-26 11:56 | 2022-07-28 10:50 | 2022-07-29 10:27 | B22G280 | SD
Microcystins ADDA, Total | 0.0214 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-28 11:35 | USFS Horseshoe Bend CG | July | PWS94179:SRC-AA | 2207246-01 | Source Water | 207 | 29 | 2022-07-28 11:35 | 2022-07-29 12:30 | 2022-08-01 13:53 | B22G300 | KMC
Microcystins ADDA, Total | 0.0077 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-07-31 13:30 | City of Creswell | July | PWS00246:SRC-BK | 2208005-01 | Source Water | 210 | 30 | 2022-07-31 13:30 | 2022-08-02 11:00 | 2022-08-03 12:15 | B22H021 | KMC
Microcystins ADDA, Total | 0.0141 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 08:28 | South Fork Water Board - Oregon City | August | PWS00591:SRC-AA | 2208024-01 | Source Water | 210 | 30 | 2022-08-01 08:28 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0092 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 08:30 | Shangri La Water District | August | PWS00835:SRC-AA | 2208023-01 | Source Water | 210 | 30 | 2022-08-01 08:30 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.1049 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-08-01 08:50 | Tri-City JW & SA | August | PWS00549:SRC-AA | 2208026-01 | Source Water | 210 | 30 | 2022-08-01 08:50 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0218 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 09:00 | Umpqua Basin Water Association | August | PWS00719:SRC-AA | 2208027-01 | Source Water | 210 | 30 | 2022-08-01 09:00 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0139 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 09:18 | City of Monroe | August | PWS00540:SRC-BA | 2208011-01 | Source Water | 210 | 30 | 2022-08-01 09:18 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0356 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 09:20 | City of Gold Hill | August | PWS00333:SRC-AA | 2208008-01 | Source Water | 210 | 30 | 2022-08-01 09:20 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0133 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 09:56 | Lyons Mehama Water District | August | PWS00493:SRC-AA | 2208018-01 | Source Water | 210 | 30 | 2022-08-01 09:56 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0617 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 10:00 | Roberts Creek Water District | August | PWS00717:SRC-AA | 2208021-01 | Source Water | 210 | 30 | 2022-08-01 10:00 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.1427 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2022-08-01 10:35 | City of Myrtle Creek | August | PWS00550:SRC-AA | 2208012-01 | Source Water | 210 | 30 | 2022-08-01 10:35 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.016 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 10:55 | City of Newport | August | PWS00566:CH-A | 2208013-01 | Source Water | 210 | 30 | 2022-08-01 10:55 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0257 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 11:00 | City of Rogue River | August | PWS00712:SRC-BA | 2208014-01 | Source Water | 210 | 30 | 2022-08-01 11:00 | 2022-08-02 11:00 | 2022-08-03 12:15 | B22H021 | KMC
Microcystins ADDA, Total | 0.0024 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 11:00 | City of Roseburg | August | PWS00720:SRC-AA | 2208015-01RE1 | Source Water | 210 | 30 | 2022-08-01 11:00 | 2022-08-03 12:50 | 2022-08-05 11:28 | B22H057 | SD
Microcystins ADDA, Total | 0.0176 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 11:00 | Stayton Water Supply | August | PWS00843:SRC-AA | 2208025-01RE1 | Source Water | 210 | 30 | 2022-08-01 11:00 | 2022-08-03 12:50 | 2022-08-05 11:28 | B22H057 | SD
Microcystins ADDA, Total | 0.0231 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 11:55 | City of Jefferson | August | PWS00408:SRC-AA | 2208009-01 | Source Water | 210 | 30 | 2022-08-01 11:55 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0572 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 12:15 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2208017-01 | Source Water | 210 | 30 | 2022-08-01 12:15 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0081 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 12:30 | Ashland Water Department | August | PWS00047:SRC-AA | 2208003-01 | Source Water | 210 | 30 | 2022-08-01 12:30 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0049 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 12:30 | City of Lowell | August | PWS00492:SRC-BA | 2208010-01 | Source Water | 210 | 30 | 2022-08-01 12:30 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0206 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 12:30 | Josephine Co Pks Lake Selmac 1 | August | PWS90186:SRC-AA | 2208016-01 | Source Water | 210 | 30 | 2022-08-01 12:30 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0069 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 12:50 | City of Estacada | August | PWS00279:SRC-AA | 2208007-01 | Source Water | 210 | 30 | 2022-08-01 12:50 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0055 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 13:15 | Camp Baker BSA | August | PWS91786:SRC-AA | 2208004-01 | Source Water | 210 | 30 | 2022-08-01 13:15 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0076 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 13:25 | Rainbow Water District | August | PWS00839:SRC-BB | 2208020-01 | Source Water | 210 | 30 | 2022-08-01 13:25 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0062 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-01 14:00 | Seaside Water Department | August | PWS00799:CH-A | 2208022-01 | Source Water | 210 | 30 | 2022-08-01 14:00 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0206 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-02 11:00 | Umpqua Ranch Co-op | August | PWS00714:SRC-AH | 2208028-01 | Source Water | 211 | 30 | 2022-08-02 11:00 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.0129 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-02 11:30 | Medford Water Commission | August | PWS00513:SRC-BA | 2208019-01 | Source Water | 211 | 30 | 2022-08-02 11:30 | 2022-08-03 12:50 | 2022-08-04 11:18 | B22H043 | SD
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-02 13:10 | City of Drain Shops | August | PWS00260:SRC-AA | 2208006-01 | Source Water | 211 | 30 | 2022-08-02 13:10 | 2022-08-04 12:30 | 2022-08-05 11:28 | B22H057 | SD
Microcystins ADDA, Total | 0.0261 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 05:15 | Anglers Cove/SCHWC | August | PWS01483:SRC-AB | 2208055-01 | Source Water | 217 | 31 | 2022-08-08 05:15 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0135 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 05:30 | Hiland WC - Shady Cove | August | PWS01520:SRC-AA | 2208069-01 | Source Water | 217 | 31 | 2022-08-08 05:30 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0137 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 08:15 | Springfield Utility Board | August | PWS00837:CH-A | 2208077-02 | Source Water | 217 | 31 | 2022-08-08 08:15 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0057 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 08:20 | Joint Water Commission | August | PWS00379:SRC-AA | 2208071-01 | Source Water | 217 | 31 | 2022-08-08 08:20 | 2022-08-09 13:00 | 2022-08-10 12:57 | B22H077 | KMC
Microcystins ADDA, Total | 0.0085 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 08:40 | Springfield Utility Board | August | PWS00837:SRC-BB | 2208077-01 | Source Water | 217 | 31 | 2022-08-08 08:40 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0231 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 09:10 | Glide Water Association | August | PWS00326:SRC-AA | 2208068-01 | Source Water | 217 | 31 | 2022-08-08 09:10 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0721 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-08-08 09:25 | Roseburg Forest Products - Dillard | August | PWS94300:SRC-AA | 2208075-01 | Source Water | 217 | 31 | 2022-08-08 09:25 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0882 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-08-08 09:40 | Winston-Dillard Water District | August | PWS00957:SRC-AA | 2208081-01 | Source Water | 217 | 31 | 2022-08-08 09:40 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0274 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 09:55 | Salem Public Works | August | PWS00731:SRC-AA | 2208076-01 | Source Water | 217 | 31 | 2022-08-08 09:55 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0284 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 10:00 | City of Gates | August | PWS00317:SRC-AA | 2208060-01 | Source Water | 217 | 31 | 2022-08-08 10:00 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0268 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 10:08 | USFS Tiller Ranger Station | August | PWS01092:SRC-AA | 2208080-01 | Source Water | 217 | 31 | 2022-08-08 10:08 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0096 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 10:20 | City of Powers | August | PWS00672:SRC-AA | 2208063-01 | Source Water | 217 | 31 | 2022-08-08 10:20 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0156 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 10:40 | Lake Oswego - Tigard Water Supply | August | PWS01542:SRC-AA | 2208072-01 | Source Water | 217 | 31 | 2022-08-08 10:40 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0074 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 11:46 | City of Pendleton | August | PWS00613:SRC-AM | 2208062-01 | Source Water | 217 | 31 | 2022-08-08 11:46 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0215 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 11:50 | City of Grants Pass | August | PWS00342:SRC-AA | 2208061-01 | Source Water | 217 | 31 | 2022-08-08 11:50 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0103 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 13:30 | City of Wilsonville | August | PWS00954:SRC-HA | 2208064-01 | Source Water | 217 | 31 | 2022-08-08 13:30 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0134 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 14:10 | Canby Utility/Veolia Water | August | PWS00157:CH-A | 2208057-01 | Source Water | 217 | 31 | 2022-08-08 14:10 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0102 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 15:15 | City of Cottage Grove | August | PWS00236:SRC-BA | 2208059-01 | Source Water | 217 | 31 | 2022-08-08 15:15 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0252 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 15:21 | USFS Steamboat Work Center | August | PWS01091:SRC-AA | 2208079-01 | Source Water | 217 | 31 | 2022-08-08 15:21 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0226 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 15:45 | Eugene Water & Electric Board | August | PWS00287:SRC-AA | 2208067-01 | Source Water | 217 | 31 | 2022-08-08 15:45 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0146 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 15:55 | Jackson Co Pks Emigrant Lake | August | PWS90730:CH-A | 2208070-01 | Source Water | 217 | 31 | 2022-08-08 15:55 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-08 16:15 | City of Albany | August | PWS00012:SRC-BA | 2208058-01 | Source Water | 217 | 31 | 2022-08-08 16:15 | 2022-08-10 11:45 | 2022-08-11 18:42 | B22H107 | KMC
Microcystins ADDA, Total | 0.0027 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-10 12:45 | North Clackamas County Water Commission | August | PWS00580:SRC-BA | 2208099-01 | Source Water | 219 | 31 | 2022-08-10 12:45 | 2022-08-11 14:00 | 2022-08-12 14:52 | B22H112 | SD
Microcystins ADDA, Total | 0.0054 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-11 09:30 | Douglas County Parks - Whistler's Bend | August | PWS93944:SRC-AB | 2208066-01 | Source Water | 220 | 31 | 2022-08-11 09:30 | 2022-08-15 10:30 | 2022-08-16 16:50 | B22H143 | KMC
Microcystins ADDA, Total | 0.0403 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-12 16:30 | Clarks Branch Water Association | August | PWS00548:SRC-AA | 2208153-01 | Source Water | 221 | 31 | 2022-08-12 16:30 | 2022-08-17 12:15 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-14 11:45 | City of Creswell | August | PWS00246:SRC-BK | 2208117-01 | Source Water | 223 | 31 | 2022-08-14 11:45 | 2022-08-16 10:45 | 2022-08-16 16:50 | B22H143 | KMC
Microcystins ADDA, Total | 0.0132 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 07:10 | Medford Water Commission | August | PWS00513:SRC-BA | 2208131-01 | Source Water | 224 | 32 | 2022-08-15 07:10 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0713 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-08-15 08:00 | Tri-City JW & SA | August | PWS00549:SRC-AA | 2208138-01 | Source Water | 224 | 32 | 2022-08-15 08:00 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0105 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 08:01 | Lyons Mehama Water District | August | PWS00493:SRC-AA | 2208130-01 | Source Water | 224 | 32 | 2022-08-15 08:01 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0322 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 08:30 | Roberts Creek Water District | August | PWS00717:SRC-AA | 2208133-01 | Source Water | 224 | 32 | 2022-08-15 08:30 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0079 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 09:08 | City of Gold Hill | August | PWS00333:SRC-AA | 2208120-01 | Source Water | 224 | 32 | 2022-08-15 09:08 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0024 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 09:30 | Shangri La Water District | August | PWS00835:SRC-AA | 2208135-01 | Source Water | 224 | 32 | 2022-08-15 09:30 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0126 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 09:45 | Umpqua Basin Water Association | August | PWS00719:SRC-AA | 2208139-01 | Source Water | 224 | 32 | 2022-08-15 09:45 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0357 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 10:30 | City of Myrtle Creek | August | PWS00550:SRC-AA | 2208124-01 | Source Water | 224 | 32 | 2022-08-15 10:30 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0058 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 10:40 | City of Jefferson | August | PWS00408:SRC-AA | 2208121-01 | Source Water | 224 | 32 | 2022-08-15 10:40 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0164 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 10:56 | City of Estacada | August | PWS00279:SRC-AA | 2208119-01 | Source Water | 224 | 32 | 2022-08-15 10:56 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0066 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 11:00 | Stayton Water Supply | August | PWS00843:SRC-AA | 2208137-01 | Source Water | 224 | 32 | 2022-08-15 11:00 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0871 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-08-15 11:30 | Josephine Co Pks Lake Selmac 1 | August | PWS90186:SRC-AA | 2208128-01 | Source Water | 224 | 32 | 2022-08-15 11:30 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 11:45 | Camp Baker BSA | August | PWS91786:SRC-AA | 2208116-01 | Source Water | 224 | 32 | 2022-08-15 11:45 | 2022-08-18 09:10 | 2022-08-18 17:51 | B22H186 | SD
Microcystins ADDA, Total | 0.0459 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 12:00 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2208129-01 | Source Water | 224 | 32 | 2022-08-15 12:00 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 12:13 | South Fork Water Board - Oregon City | August | PWS00591:SRC-AA | 2208136-01 | Source Water | 224 | 32 | 2022-08-15 12:13 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0067 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 12:18 | City of Roseburg | August | PWS00720:SRC-AA | 2208127-01 | Source Water | 224 | 32 | 2022-08-15 12:18 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0381 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 12:40 | City of Drain Shops | August | PWS00260:SRC-AA | 2208118-01 | Source Water | 224 | 32 | 2022-08-15 12:40 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0047 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 13:00 | Ashland Water Department | August | PWS00047:SRC-AA | 2208115-01 | Source Water | 224 | 32 | 2022-08-15 13:00 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 13:25 | Rainbow Water District | August | PWS00839:SRC-BB | 2208132-01 | Source Water | 224 | 32 | 2022-08-15 13:25 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0123 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 13:30 | City of Lowell | August | PWS00492:SRC-BA | 2208122-01 | Source Water | 224 | 32 | 2022-08-15 13:30 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0111 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 13:45 | Seaside Water Department | August | PWS00799:CH-A | 2208134-01 | Source Water | 224 | 32 | 2022-08-15 13:45 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0016 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 14:05 | PP&L-Toketee Village | August | PWS01012:SRC-AA | 2208155-01 | Source Water | 224 | 32 | 2022-08-15 14:05 | 2022-08-17 12:15 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0102 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 14:40 | City of Newport | August | PWS00566:CH-A | 2208125-01 | Source Water | 224 | 32 | 2022-08-15 14:40 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0206 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 14:42 | City of Monroe | August | PWS00540:SRC-BA | 2208123-01 | Source Water | 224 | 32 | 2022-08-15 14:42 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-15 14:52 | USFS Horseshoe Bend CG | August | PWS94179:SRC-AA | 2208152-01 | Source Water | 224 | 32 | 2022-08-15 14:52 | 2022-08-17 12:15 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0111 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-16 11:05 | City of Rogue River | August | PWS00712:SRC-BA | 2208126-01 | Source Water | 225 | 32 | 2022-08-16 11:05 | 2022-08-17 12:00 | 2022-08-18 11:15 | B22H177 | KMC
Microcystins ADDA, Total | 0.0073 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 05:15 | Anglers Cove/SCHWC | August | PWS01483:SRC-AB | 2208196-01 | Source Water | 231 | 33 | 2022-08-22 05:15 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0138 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 05:30 | Hiland WC - Shady Cove | August | PWS01520:SRC-AA | 2208210-01 | Source Water | 231 | 33 | 2022-08-22 05:30 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0079 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 06:45 | City of Wilsonville | August | PWS00954:SRC-HA | 2208205-01 | Source Water | 231 | 33 | 2022-08-22 06:45 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0113 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 07:40 | Springfield Utility Board | August | PWS00837:CH-A | 2208218-02 | Source Water | 231 | 33 | 2022-08-22 07:40 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0017 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 08:00 | Springfield Utility Board | August | PWS00837:SRC-BB | 2208218-01 | Source Water | 231 | 33 | 2022-08-22 08:00 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0024 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 08:14 | Joint Water Commission | August | PWS00379:SRC-AA | 2208212-01 | Source Water | 231 | 33 | 2022-08-22 08:14 | 2022-08-23 14:30 | 2022-08-24 18:24 | B22H234 | SD
Microcystins ADDA, Total | 0.0051 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 08:35 | Glide Water Association | August | PWS00326:SRC-AA | 2208209-01 | Source Water | 231 | 33 | 2022-08-22 08:35 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.057 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 08:45 | Winston-Dillard Water District | August | PWS00957:SRC-AA | 2208222-01 | Source Water | 231 | 33 | 2022-08-22 08:45 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0071 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 08:50 | Jackson Co Pks Emigrant Lake | August | PWS90730:CH-A | 2208211-01 | Source Water | 231 | 33 | 2022-08-22 08:50 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0136 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 09:00 | City of Gates | August | PWS00317:SRC-AA | 2208201-01 | Source Water | 231 | 33 | 2022-08-22 09:00 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0026 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 09:00 | North Clackamas County Water Commission | August | PWS00580:SRC-BA | 2208214-01 | Source Water | 231 | 33 | 2022-08-22 09:00 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0134 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 09:40 | Lake Oswego - Tigard Water Supply | August | PWS01542:SRC-AA | 2208213-01 | Source Water | 231 | 33 | 2022-08-22 09:40 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0138 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 10:00 | Douglas County Parks - Whistler's Bend | August | PWS93944:SRC-AB | 2208207-01 | Source Water | 231 | 33 | 2022-08-22 10:00 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0262 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 10:46 | Salem Public Works | August | PWS00731:SRC-AA | 2208217-01 | Source Water | 231 | 33 | 2022-08-22 10:46 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0463 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 10:47 | Roseburg Forest Products - Dillard | August | PWS94300:SRC-AA | 2208216-01 | Source Water | 231 | 33 | 2022-08-22 10:47 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0232 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 11:10 | City of Grants Pass | August | PWS00342:SRC-AA | 2208202-01 | Source Water | 231 | 33 | 2022-08-22 11:10 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.2467 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-08-22 11:30 | Buell-Red Prairie Water Association | August | PWS01174:SRC-AA | 2208197-01 | Source Water | 231 | 33 | 2022-08-22 11:30 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0139 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 11:30 | Umpqua Ranch Co-op | August | PWS00714:SRC-AH | 2208224-01 | Source Water | 231 | 33 | 2022-08-22 11:30 | 2022-08-23 14:30 | 2022-08-24 18:24 | B22H234 | SD
Microcystins ADDA, Total | 0.0117 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 11:50 | City of Pendleton | August | PWS00613:SRC-AM | 2208203-01 | Source Water | 231 | 33 | 2022-08-22 11:50 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0053 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 13:10 | City of Albany | August | PWS00012:SRC-BA | 2208199-01 | Source Water | 231 | 33 | 2022-08-22 13:10 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0123 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 14:30 | Canby Utility/Veolia Water | August | PWS00157:CH-A | 2208198-01 | Source Water | 231 | 33 | 2022-08-22 14:30 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0108 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 15:30 | City of Powers | August | PWS00672:SRC-AA | 2208204-01 | Source Water | 231 | 33 | 2022-08-22 15:30 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0235 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 16:15 | Eugene Water & Electric Board | August | PWS00287:SRC-AA | 2208208-01 | Source Water | 231 | 33 | 2022-08-22 16:15 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0127 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 16:28 | PP&L-Toketee Village | August | PWS01012:SRC-AA | 2208215-01 | Source Water | 231 | 33 | 2022-08-22 16:28 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0138 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 16:30 | City of Cottage Grove | August | PWS00236:SRC-BA | 2208200-01 | Source Water | 231 | 33 | 2022-08-22 16:30 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0119 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 17:25 | USFS Horseshoe Bend CG | August | PWS94179:SRC-AA | 2208219-01 | Source Water | 231 | 33 | 2022-08-22 17:25 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0187 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 17:48 | USFS Steamboat Work Center | August | PWS01091:SRC-AA | 2208220-01 | Source Water | 231 | 33 | 2022-08-22 17:48 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0029 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 18:23 | Clarks Branch Water Association | August | PWS00548:SRC-AA | 2208206-01RE1 | Source Water | 231 | 33 | 2022-08-22 18:23 | 2022-08-24 12:00 | 2022-08-26 13:29 | B22H271 | KMC
Microcystins ADDA, Total | 0.0375 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-22 20:23 | USFS Tiller Ranger Station | August | PWS01092:SRC-AA | 2208221-01 | Source Water | 231 | 33 | 2022-08-22 20:23 | 2022-08-24 12:00 | 2022-08-25 14:02 | B22H256 | KMC
Microcystins ADDA, Total | 0.0073 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-28 14:35 | City of Creswell | August | PWS00246:SRC-BK | 2208266-01 | Source Water | 237 | 33 | 2022-08-28 14:35 | 2022-08-30 11:00 | 2022-08-31 11:23 | B22H324 | SD
Microcystins ADDA, Total | 0.0666 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-08-29 07:42 | Milo Academy | August | PWS00250:SRC-AB | 2208292-01 | Source Water | 238 | 34 | 2022-08-29 07:42 | 2022-08-30 11:00 | 2022-08-31 11:23 | B22H324 | SD
Microcystins ADDA, Total | 0.0491 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 08:00 | Tri-City JW & SA | August | PWS00549:SRC-AA | 2208287-01 | Source Water | 238 | 34 | 2022-08-29 08:00 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0259 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 08:30 | City of Gold Hill | August | PWS00333:SRC-AA | 2208269-01 | Source Water | 238 | 34 | 2022-08-29 08:30 | 2022-08-30 11:00 | 2022-08-31 11:23 | B22H324 | SD
Microcystins ADDA, Total | 0.0036 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 09:30 | Shangri La Water District | August | PWS00835:SRC-AA | 2208284-01 | Source Water | 238 | 34 | 2022-08-29 09:30 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0413 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 09:40 | Roberts Creek Water District | August | PWS00717:SRC-AA | 2208282-01 | Source Water | 238 | 34 | 2022-08-29 09:40 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0151 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 09:45 | Medford Water Commission | August | PWS00513:SRC-BA | 2208280-01 | Source Water | 238 | 34 | 2022-08-29 09:45 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0072 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 10:00 | Umpqua Ranch Co-op | August | PWS00714:SRC-AH | 2208289-01 | Source Water | 238 | 34 | 2022-08-29 10:00 | 2022-08-30 11:00 | 2022-08-31 11:23 | B22H324 | SD
Microcystins ADDA, Total | 0.0225 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 11:00 | Stayton Water Supply | August | PWS00843:SRC-AA | 2208286-01 | Source Water | 238 | 34 | 2022-08-29 11:00 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 11:10 | Camp Baker BSA | August | PWS91786:SRC-AA | 2208265-01 | Source Water | 238 | 34 | 2022-08-29 11:10 | 2022-08-30 11:00 | 2022-08-31 11:23 | B22H324 | SD
Microcystins ADDA, Total | 0.0036 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 11:35 | City of Rogue River | August | PWS00712:SRC-BA | 2208275-01 | Source Water | 238 | 34 | 2022-08-29 11:35 | 2022-08-30 11:00 | 2022-08-31 11:23 | B22H324 | SD
Microcystins ADDA, Total | 0.0133 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 11:45 | City of Lowell | August | PWS00492:SRC-BA | 2208271-01 | Source Water | 238 | 34 | 2022-08-29 11:45 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0046 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 11:58 | Lyons Mehama Water District | August | PWS00493:SRC-AA | 2208279-01 | Source Water | 238 | 34 | 2022-08-29 11:58 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.1737 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-08-29 12:00 | Josephine Co Pks Lake Selmac 2 | August | PWS94645:SRC-AA | 2208278-01 | Source Water | 238 | 34 | 2022-08-29 12:00 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0144 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 12:29 | City of Roseburg | August | PWS00720:SRC-AA | 2208276-01 | Source Water | 238 | 34 | 2022-08-29 12:29 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.015 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 12:30 | Ashland Water Department | August | PWS00047:SRC-AA | 2208264-01 | Source Water | 238 | 34 | 2022-08-29 12:30 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.1267 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2022-08-29 12:30 | Josephine Co Pks Lake Selmac 1 | August | PWS90186:SRC-AA | 2208277-01 | Source Water | 238 | 34 | 2022-08-29 12:30 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0061 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 12:55 | Umpqua Basin Water Association | August | PWS00719:SRC-AA | 2208288-01 | Source Water | 238 | 34 | 2022-08-29 12:55 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0113 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 13:05 | City of Jefferson | August | PWS00408:SRC-AA | 2208270-01 | Source Water | 238 | 34 | 2022-08-29 13:05 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0099 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 13:20 | South Fork Water Board - Oregon City | August | PWS00591:SRC-AA | 2208285-01 | Source Water | 238 | 34 | 2022-08-29 13:20 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0155 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 14:22 | City of Monroe | August | PWS00540:SRC-BA | 2208272-01 | Source Water | 238 | 34 | 2022-08-29 14:22 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0175 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 14:45 | City of Drain Shops | August | PWS00260:SRC-AA | 2208267-01 | Source Water | 238 | 34 | 2022-08-29 14:45 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0161 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 15:35 | Rainbow Water District | August | PWS00839:SRC-BB | 2208281-01 | Source Water | 238 | 34 | 2022-08-29 15:35 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0055 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-29 16:00 | City of Newport | August | PWS00566:CH-A | 2208274-01 | Source Water | 238 | 34 | 2022-08-29 16:00 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0137 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-30 09:20 | City of Estacada | August | PWS00279:SRC-AA | 2208268-01RE1 | Source Water | 239 | 34 | 2022-08-30 09:20 | 2022-09-01 11:10 | 2022-09-06 12:40 | B22I032 | SD
Microcystins ADDA, Total | 0.0965 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-08-30 10:30 | Buell-Red Prairie Water Association | August | PWS01174:SRC-AA | 2208302-01 | Source Water | 239 | 34 | 2022-08-30 10:30 | 2022-08-31 11:30 | 2022-09-01 13:43 | B22H339 | KMC
Microcystins ADDA, Total | 0.0368 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-08-31 10:55 | City of Myrtle Creek | August | PWS00550:SRC-AA | 2208273-01RE1 | Source Water | 240 | 34 | 2022-08-31 10:55 | 2022-09-01 11:10 | 2022-09-06 12:40 | B22I032 | SD
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-02 14:00 | Seaside Water Department | September | PWS00799:CH-A | 2209051-01 | Source Water | 241 | 34 | 2022-09-02 14:00 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-05 09:00 | North Clackamas County Water Commission | September | PWS00580:SRC-BA | 2209035-01 | Source Water | 244 | 34 | 2022-09-05 09:00 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0091 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-05 09:30 | Jackson Co Pks Emigrant Lake | September | PWS90730:CH-A | 2209032-01 | Source Water | 244 | 34 | 2022-09-05 09:30 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0153 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-05 14:40 | City of Cottage Grove | September | PWS00236:SRC-BA | 2209021-01 | Source Water | 244 | 34 | 2022-09-05 14:40 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0134 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 07:35 | City of Wilsonville | September | PWS00954:SRC-HA | 2209026-01 | Source Water | 245 | 35 | 2022-09-06 07:35 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0147 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 07:48 | Salem Public Works | September | PWS00731:SRC-AA | 2209038-01 | Source Water | 245 | 35 | 2022-09-06 07:48 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0071 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 08:10 | USFS Tiller Ranger Station | September | PWS01092:SRC-AA | 2209042-01 | Source Water | 245 | 35 | 2022-09-06 08:10 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0195 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 08:20 | Roseburg Forest Products - Dillard | September | PWS94300:SRC-AA | 2209037-01 | Source Water | 245 | 35 | 2022-09-06 08:20 | 2022-09-08 11:30 | 2022-09-09 13:47 | B22I068 | KMC
Microcystins ADDA, Total | 0.0052 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 08:35 | Springfield Utility Board | September | PWS00837:CH-A | 2209039-02 | Source Water | 245 | 35 | 2022-09-06 08:35 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0028 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 08:36 | Joint Water Commission | September | PWS00379:SRC-AA | 2209033-01 | Source Water | 245 | 35 | 2022-09-06 08:36 | 2022-09-07 13:15 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.029 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 08:45 | Winston-Dillard Water District | September | PWS00957:SRC-AA | 2209043-01 | Source Water | 245 | 35 | 2022-09-06 08:45 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0117 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 08:55 | Lake Oswego - Tigard Water Supply | September | PWS01542:SRC-AA | 2209034-01 | Source Water | 245 | 35 | 2022-09-06 08:55 | 2022-09-08 11:30 | 2022-09-09 13:47 | B22I068 | KMC
Microcystins ADDA, Total | 0.0021 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 10:05 | Clarks Branch Water Association | September | PWS00548:SRC-AA | 2209027-01 | Source Water | 245 | 35 | 2022-09-06 10:05 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0013 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 10:50 | Springfield Utility Board | September | PWS00837:SRC-BB | 2209039-01 | Source Water | 245 | 35 | 2022-09-06 10:50 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0086 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 11:00 | City of Powers | September | PWS00672:SRC-AA | 2209025-01 | Source Water | 245 | 35 | 2022-09-06 11:00 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0089 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 11:12 | City of Pendleton | September | PWS00613:SRC-AM | 2209024-01 | Source Water | 245 | 35 | 2022-09-06 11:12 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.011 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 11:25 | Glide Water Association | September | PWS00326:SRC-AA | 2209030-01 | Source Water | 245 | 35 | 2022-09-06 11:25 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 11:50 | Eugene Water & Electric Board | September | PWS00287:SRC-AA | 2209029-01 | Source Water | 245 | 35 | 2022-09-06 11:50 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0141 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 12:41 | USFS Steamboat Work Center | September | PWS01091:SRC-AA | 2209041-01 | Source Water | 245 | 35 | 2022-09-06 12:41 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.013 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 12:45 | USFS Horseshoe Bend CG | September | PWS94179:SRC-AA | 2209040-01 | Source Water | 245 | 35 | 2022-09-06 12:45 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0074 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 13:25 | PP&L-Toketee Village | September | PWS01012:SRC-AA | 2209036-01 | Source Water | 245 | 35 | 2022-09-06 13:25 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0117 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 13:30 | City of Grants Pass | September | PWS00342:SRC-AA | 2209023-01 | Source Water | 245 | 35 | 2022-09-06 13:30 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0121 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 13:30 | Hiland WC - Shady Cove | September | PWS01520:SRC-AA | 2209031-01 | Source Water | 245 | 35 | 2022-09-06 13:30 | 2022-09-08 11:30 | 2022-09-09 13:47 | B22I068 | KMC
Microcystins ADDA, Total | 0.0095 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 13:45 | Anglers Cove/SCHWC | September | PWS01483:SRC-AB | 2209017-01 | Source Water | 245 | 35 | 2022-09-06 13:45 | 2022-09-08 11:30 | 2022-09-09 13:47 | B22I068 | KMC
Microcystins ADDA, Total | 0.0134 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 13:45 | City of Gates | September | PWS00317:SRC-AA | 2209022-01 | Source Water | 245 | 35 | 2022-09-06 13:45 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0091 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 15:00 | Canby Utility/Veolia Water | September | PWS00157:CH-A | 2209019-01 | Source Water | 245 | 35 | 2022-09-06 15:00 | 2022-09-07 11:00 | 2022-09-08 11:38 | B22I051 | SD
Microcystins ADDA, Total | 0.0083 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 15:00 | City of Albany | September | PWS00012:SRC-AA | 2209020-02 | Source Water | 245 | 35 | 2022-09-06 15:00 | 2022-09-08 11:30 | 2022-09-09 13:47 | B22I068 | KMC
Microcystins ADDA, Total | 0.0042 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-06 16:20 | City of Albany | September | PWS00012:SRC-BA | 2209020-01 | Source Water | 245 | 35 | 2022-09-06 16:20 | 2022-09-08 11:30 | 2022-09-09 13:47 | B22I068 | KMC
Microcystins ADDA, Total | 0.0048 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-10 13:00 | Douglas County Parks - Whistler's Bend | September | PWS93944:SRC-AB | 2209117-01 | Source Water | 249 | 35 | 2022-09-10 13:00 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-11 11:15 | City of Creswell | September | PWS00246:SRC-BK | 2209089-01 | Source Water | 250 | 35 | 2022-09-11 11:15 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0206 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 07:40 | Medford Water Commission | September | PWS00513:SRC-BA | 2209103-01 | Source Water | 251 | 35 | 2022-09-12 07:40 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0513 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 08:00 | Milo Academy | September | PWS00250:SRC-AB | 2209104-01 | Source Water | 251 | 35 | 2022-09-12 08:00 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 08:31 | City of Monroe | September | PWS00540:SRC-BA | 2209095-01 | Source Water | 251 | 35 | 2022-09-12 08:31 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0073 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 09:00 | South Fork Water Board - Oregon City | September | PWS00591:SRC-AA | 2209109-01 | Source Water | 251 | 35 | 2022-09-12 09:00 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0215 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 09:00 | Tri-City JW & SA | September | PWS00549:SRC-AA | 2209111-01 | Source Water | 251 | 35 | 2022-09-12 09:00 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0.0074 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 09:15 | City of Gold Hill | September | PWS00333:SRC-AA | 2209092-01 | Source Water | 251 | 35 | 2022-09-12 09:15 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0171 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 09:18 | Lyons Mehama Water District | September | PWS00493:SRC-AA | 2209102-01 | Source Water | 251 | 35 | 2022-09-12 09:18 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 09:30 | Shangri La Water District | September | PWS00835:SRC-AA | 2209108-01 | Source Water | 251 | 35 | 2022-09-12 09:30 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0283 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 10:00 | City of Estacada | September | PWS00279:SRC-AA | 2209091-01 | Source Water | 251 | 35 | 2022-09-12 10:00 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0.04 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 10:00 | Roberts Creek Water District | September | PWS00717:SRC-AA | 2209106-01 | Source Water | 251 | 35 | 2022-09-12 10:00 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 10:20 | City of Newport | September | PWS00566:CH-A | 2209097-01 | Source Water | 251 | 35 | 2022-09-12 10:20 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 10:25 | Camp Baker BSA | September | PWS91786:SRC-AA | 2209088-01 | Source Water | 251 | 35 | 2022-09-12 10:25 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0032 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 10:30 | Umpqua Basin Water Association | September | PWS00719:SRC-AA | 2209112-01 | Source Water | 251 | 35 | 2022-09-12 10:30 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0061 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 10:50 | City of Lowell | September | PWS00492:SRC-BA | 2209094-01 | Source Water | 251 | 35 | 2022-09-12 10:50 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0.013 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 11:00 | Stayton Water Supply | September | PWS00843:SRC-AA | 2209110-01 | Source Water | 251 | 35 | 2022-09-12 11:00 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0024 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 11:06 | City of Roseburg | September | PWS00720:SRC-AA | 2209099-01 | Source Water | 251 | 35 | 2022-09-12 11:06 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.027 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 11:30 | City of Myrtle Creek | September | PWS00550:SRC-AA | 2209096-01 | Source Water | 251 | 35 | 2022-09-12 11:30 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0.0014 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 11:40 | City of Rogue River | September | PWS00712:SRC-BA | 2209098-01 | Source Water | 251 | 35 | 2022-09-12 11:40 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0095 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 12:30 | Ashland Water Department | September | PWS00047:SRC-AA | 2209087-01 | Source Water | 251 | 35 | 2022-09-12 12:30 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0138 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 13:40 | City of Jefferson | September | PWS00408:SRC-AA | 2209093-01 | Source Water | 251 | 35 | 2022-09-12 13:40 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 13:45 | Seaside Water Department | September | PWS00799:CH-A | 2209107-01 | Source Water | 251 | 35 | 2022-09-12 13:45 | 2022-09-13 11:30 | 2022-09-13 20:26 | B22I101 | KMC
Microcystins ADDA, Total | 0.0302 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 14:10 | City of Drain Shops | September | PWS00260:SRC-AA | 2209090-01 | Source Water | 251 | 35 | 2022-09-12 14:10 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0.0153 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-12 14:25 | Rainbow Water District | September | PWS00839:SRC-BB | 2209105-01 | Source Water | 251 | 35 | 2022-09-12 14:25 | 2022-09-15 13:30 | 2022-09-16 12:50 | B22I128 | SD
Microcystins ADDA, Total | 0.0114 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-13 09:00 | Umpqua Ranch Co-op | September | PWS00714:SRC-AH | 2209113-01 | Source Water | 252 | 36 | 2022-09-13 09:00 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0.3041 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-09-13 12:00 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2209100-01 | Source Water | 252 | 36 | 2022-09-13 12:00 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0.3985 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-09-13 12:30 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2209101-01 | Source Water | 252 | 36 | 2022-09-13 12:30 | 2022-09-14 11:15 | 2022-09-15 10:32 | B22I110 | SD
Microcystins ADDA, Total | 0.4448 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-09-16 11:20 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2209191-01RE1 | Source Water | 255 | 36 | 2022-09-16 11:20 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.4782 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-09-16 11:35 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2209192-01RE1 | Source Water | 255 | 36 | 2022-09-16 11:35 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0253 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-18 13:10 | City of Grants Pass | September | PWS00342:SRC-AA | 2209170-01RE1 | Source Water | 257 | 36 | 2022-09-18 13:10 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0067 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 05:15 | Anglers Cove/SCHWC | September | PWS01483:SRC-AB | 2209164-01RE1 | Source Water | 258 | 36 | 2022-09-19 05:15 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0011 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 05:30 | Hiland WC - Shady Cove | September | PWS01520:SRC-AA | 2209178-01RE1 | Source Water | 258 | 36 | 2022-09-19 05:30 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0047 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 07:00 | North Clackamas County Water Commission | September | PWS00580:SRC-BA | 2209182-01 | Source Water | 258 | 36 | 2022-09-19 07:00 | 2022-09-21 11:45 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0035 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 07:40 | Springfield Utility Board | September | PWS00837:CH-A | 2209186-02RE1 | Source Water | 258 | 36 | 2022-09-19 07:40 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0096 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 07:53 | Joint Water Commission | September | PWS00379:SRC-AA | 2209180-01RE1 | Source Water | 258 | 36 | 2022-09-19 07:53 | 2022-09-20 15:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0138 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 08:15 | City of Albany | September | PWS00012:SRC-BA | 2209167-01RE1 | Source Water | 258 | 36 | 2022-09-19 08:15 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0273 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 08:30 | Winston-Dillard Water District | September | PWS00957:SRC-AA | 2209190-01RE1 | Source Water | 258 | 36 | 2022-09-19 08:30 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0393 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 08:31 | Roseburg Forest Products - Dillard | September | PWS94300:SRC-AA | 2209184-01RE1 | Source Water | 258 | 36 | 2022-09-19 08:31 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0185 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 08:40 | Salem Public Works | September | PWS00731:SRC-AA | 2209185-01RE1 | Source Water | 258 | 36 | 2022-09-19 08:40 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0155 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 09:00 | Glide Water Association | September | PWS00326:SRC-AA | 2209177-01RE1 | Source Water | 258 | 36 | 2022-09-19 09:00 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.001 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 09:25 | Springfield Utility Board | September | PWS00837:SRC-BB | 2209186-01RE1 | Source Water | 258 | 36 | 2022-09-19 09:25 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0214 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 09:40 | Douglas County Parks - Whistler's Bend | September | PWS93944:SRC-AB | 2209175-01 | Source Water | 258 | 36 | 2022-09-19 09:40 | 2022-09-21 11:45 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0037 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 10:35 | City of Powers | September | PWS00672:SRC-AA | 2209172-01RE1 | Source Water | 258 | 36 | 2022-09-19 10:35 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.3414 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-09-19 10:45 | Buell-Red Prairie Water Association | September | PWS01174:SRC-AA | 2209165-01RE1 | Source Water | 258 | 36 | 2022-09-19 10:45 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0149 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 11:20 | City of Wilsonville | September | PWS00954:SRC-HA | 2209173-01RE1 | Source Water | 258 | 36 | 2022-09-19 11:20 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0082 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 11:20 | USFS Tiller Ranger Station | September | PWS01092:SRC-AA | 2209189-01RE1 | Source Water | 258 | 36 | 2022-09-19 11:20 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0235 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 11:49 | City of Pendleton | September | PWS00613:SRC-AM | 2209171-01RE1 | Source Water | 258 | 36 | 2022-09-19 11:49 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0154 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 12:00 | City of Cottage Grove | September | PWS00236:SRC-BA | 2209168-01RE1 | Source Water | 258 | 36 | 2022-09-19 12:00 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0351 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 12:11 | Clarks Branch Water Association | September | PWS00548:SRC-AA | 2209174-01RE1 | Source Water | 258 | 36 | 2022-09-19 12:11 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0185 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 12:15 | Eugene Water & Electric Board | September | PWS00287:SRC-AA | 2209176-01RE1 | Source Water | 258 | 36 | 2022-09-19 12:15 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0078 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 12:23 | Jackson Co Pks Emigrant Lake | September | PWS90730:CH-A | 2209179-01 | Source Water | 258 | 36 | 2022-09-19 12:23 | 2022-09-21 11:45 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0152 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 13:30 | City of Gates | September | PWS00317:SRC-AA | 2209169-01RE1 | Source Water | 258 | 36 | 2022-09-19 13:30 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0122 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 13:46 | USFS Steamboat Work Center | September | PWS01091:SRC-AA | 2209188-01RE1 | Source Water | 258 | 36 | 2022-09-19 13:46 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0126 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 14:00 | Canby Utility/Veolia Water | September | PWS00157:CH-A | 2209166-01RE1 | Source Water | 258 | 36 | 2022-09-19 14:00 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0107 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 14:00 | USFS Horseshoe Bend CG | September | PWS94179:SRC-AA | 2209187-01RE1 | Source Water | 258 | 36 | 2022-09-19 14:00 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.0065 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 14:14 | PP&L-Toketee Village | September | PWS01012:SRC-AA | 2209183-01RE1 | Source Water | 258 | 36 | 2022-09-19 14:14 | 2022-09-20 12:00 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-19 14:30 | Lake Oswego - Tigard Water Supply | September | PWS01542:SRC-AA | 2209181-01 | Source Water | 258 | 36 | 2022-09-19 14:30 | 2022-09-21 11:45 | 2022-09-22 14:00 | B22I195 | SD
Microcystins ADDA, Total | 0.4729 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-09-21 10:30 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2209219-01 | Source Water | 260 | 37 | 2022-09-21 10:30 | 2022-09-22 11:30 | 2022-09-22 18:57 | B22I217 | SD
Microcystins ADDA, Total | 0.2137 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-09-21 10:50 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2209218-01 | Source Water | 260 | 37 | 2022-09-21 10:50 | 2022-09-22 11:30 | 2022-09-22 18:57 | B22I217 | SD
Microcystins ADDA, Total | 0.0154 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 08:10 | Medford Water Commission | September | PWS00513:SRC-BA | 2209253-01 | Source Water | 265 | 37 | 2022-09-26 08:10 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.0197 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 08:30 | City of Estacada | September | PWS00279:SRC-AA | 2209241-01 | Source Water | 265 | 37 | 2022-09-26 08:30 | 2022-09-28 10:45 | 2022-09-28 18:01 | B22I325 | SD
Microcystins ADDA, Total | 0.0011 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 08:51 | Lyons Mehama Water District | September | PWS00493:SRC-AA | 2209252-01 | Source Water | 265 | 37 | 2022-09-26 08:51 | 2022-09-28 10:45 | 2022-09-28 18:01 | B22I325 | SD
Microcystins ADDA, Total | 0.0465 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 09:00 | City of Creswell | September | PWS00246:SRC-BK | 2209239-01 | Source Water | 265 | 37 | 2022-09-26 09:00 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.0109 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 09:14 | City of Monroe | September | PWS00540:SRC-BA | 2209245-01 | Source Water | 265 | 37 | 2022-09-26 09:14 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 09:30 | Shangri La Water District | September | PWS00835:SRC-AA | 2209258-01 | Source Water | 265 | 37 | 2022-09-26 09:30 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.0295 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 09:45 | City of Drain Shops | September | PWS00260:SRC-AA | 2209240-01 | Source Water | 265 | 37 | 2022-09-26 09:45 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 09:50 | South Fork Water Board - Oregon City | September | PWS00591:SRC-AA | 2209259-01 | Source Water | 265 | 37 | 2022-09-26 09:50 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.0233 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 10:00 | City of Myrtle Creek | September | PWS00550:SRC-AA | 2209246-01 | Source Water | 265 | 37 | 2022-09-26 10:00 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 10:00 | Umpqua Ranch Co-op | September | PWS00714:SRC-AH | 2209263-01 | Source Water | 265 | 37 | 2022-09-26 10:00 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 10:15 | Camp Baker BSA | September | PWS91786:SRC-AA | 2209238-01 | Source Water | 265 | 37 | 2022-09-26 10:15 | 2022-09-28 10:45 | 2022-09-28 18:01 | B22I325 | SD
Microcystins ADDA, Total | 0.038 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 10:15 | Roberts Creek Water District | September | PWS00717:SRC-AA | 2209256-01 | Source Water | 265 | 37 | 2022-09-26 10:15 | 2022-09-28 10:45 | 2022-09-28 18:01 | B22I325 | SD
Microcystins ADDA, Total | 0.2175 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-09-26 10:30 | Buell-Red Prairie Water Association | September | PWS01174:SRC-AA | 2209264-01 | Source Water | 265 | 37 | 2022-09-26 10:30 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 10:40 | City of Newport | September | PWS00566:CH-A | 2209247-01 | Source Water | 265 | 37 | 2022-09-26 10:40 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.0031 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 10:55 | City of Roseburg | September | PWS00720:SRC-AA | 2209249-01 | Source Water | 265 | 37 | 2022-09-26 10:55 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.01 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 11:00 | Stayton Water Supply | September | PWS00843:SRC-AA | 2209260-01 | Source Water | 265 | 37 | 2022-09-26 11:00 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.0242 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 11:05 | City of Lowell | September | PWS00492:SRC-BA | 2209244-01 | Source Water | 265 | 37 | 2022-09-26 11:05 | 2022-09-28 10:45 | 2022-09-28 18:01 | B22I325 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 11:15 | Umpqua Basin Water Association | September | PWS00719:SRC-AA | 2209262-01 | Source Water | 265 | 37 | 2022-09-26 11:15 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 11:55 | City of Rogue River | September | PWS00712:SRC-BA | 2209248-01 | Source Water | 265 | 37 | 2022-09-26 11:55 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.0005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 12:15 | City of Gold Hill | September | PWS00333:SRC-AA | 2209242-01 | Source Water | 265 | 37 | 2022-09-26 12:15 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 12:30 | Ashland Water Department | September | PWS00047:SRC-AA | 2209237-01 | Source Water | 265 | 37 | 2022-09-26 12:30 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.0066 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 12:35 | Rainbow Water District | September | PWS00839:SRC-BB | 2209255-01 | Source Water | 265 | 37 | 2022-09-26 12:35 | 2022-09-28 10:45 | 2022-09-28 18:01 | B22I325 | SD
Microcystins ADDA, Total | 0.0006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 12:50 | City of Jefferson | September | PWS00408:SRC-AA | 2209243-01 | Source Water | 265 | 37 | 2022-09-26 12:50 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 14:00 | Seaside Water Department | September | PWS00799:CH-A | 2209257-01 | Source Water | 265 | 37 | 2022-09-26 14:00 | 2022-09-27 11:15 | 2022-09-27 19:45 | B22I309 | KMC
Microcystins ADDA, Total | 0.0237 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-26 15:00 | Tri-City JW & SA | September | PWS00549:SRC-AA | 2209261-01 | Source Water | 265 | 37 | 2022-09-26 15:00 | 2022-09-28 10:45 | 2022-09-28 18:01 | B22I325 | SD
Microcystins ADDA, Total | 0.3559 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-09-28 10:50 | Josephine Co Pks Lake Selmac 1 | September | PWS90186:SRC-AA | 2209250-01 | Source Water | 267 | 38 | 2022-09-28 10:50 | 2022-09-29 11:00 | 2022-09-29 19:39 | B22I343 | KMC
Microcystins ADDA, Total | 0.5368 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-09-28 11:15 | Josephine Co Pks Lake Selmac 2 | September | PWS94645:SRC-AA | 2209251-01 | Source Water | 267 | 38 | 2022-09-28 11:15 | 2022-09-29 11:00 | 2022-09-29 19:39 | B22I343 | KMC
Microcystins ADDA, Total | 0.0538 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-09-28 11:15 | Milo Academy | September | PWS00250:SRC-AB | 2209254-01 | Source Water | 267 | 38 | 2022-09-28 11:15 | 2022-09-29 11:00 | 2022-09-29 19:39 | B22I343 | KMC
Microcystins ADDA, Total | 0.0045 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-02 14:25 | City of Grants Pass | October | PWS00342:SRC-AA | 2210007-01 | Source Water | 271 | 38 | 2022-10-02 14:25 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0088 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 05:15 | Anglers Cove/SCHWC | October | PWS01483:SRC-AB | 2210001-01 | Source Water | 272 | 38 | 2022-10-03 05:15 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0011 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 05:30 | Hiland WC - Shady Cove | October | PWS01520:SRC-AA | 2210015-01 | Source Water | 272 | 38 | 2022-10-03 05:30 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0043 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 06:00 | City of Albany | October | PWS00012:SRC-BA | 2210004-01 | Source Water | 272 | 38 | 2022-10-03 06:00 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0078 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 08:00 | North Clackamas County Water Commission | October | PWS00580:SRC-BA | 2210019-01 | Source Water | 272 | 38 | 2022-10-03 08:00 | 2022-10-05 12:00 | 2022-10-06 11:15 | B22J033 | SD
Microcystins ADDA, Total | 0.012 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 08:09 | Joint Water Commission | October | PWS00379:SRC-AA | 2210017-01 | Source Water | 272 | 38 | 2022-10-03 08:09 | 2022-10-05 12:00 | 2022-10-06 11:15 | B22J033 | SD
Microcystins ADDA, Total | 0.0021 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 08:10 | Springfield Utility Board | October | PWS00837:CH-A | 2210023-02 | Source Water | 272 | 38 | 2022-10-03 08:10 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0098 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 08:20 | Jackson Co Pks Emigrant Lake | October | PWS90730:CH-A | 2210016-01 | Source Water | 272 | 38 | 2022-10-03 08:20 | 2022-10-05 12:00 | 2022-10-06 11:15 | B22J033 | SD
Microcystins ADDA, Total | 0.0086 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 09:00 | USFS Tiller Ranger Station | October | PWS01092:SRC-AA | 2210026-01 | Source Water | 272 | 38 | 2022-10-03 09:00 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0117 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 09:36 | City of Pendleton | October | PWS00613:SRC-AM | 2210008-01 | Source Water | 272 | 38 | 2022-10-03 09:36 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0064 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 09:40 | Lake Oswego - Tigard Water Supply | October | PWS01542:SRC-AA | 2210018-01 | Source Water | 272 | 38 | 2022-10-03 09:40 | 2022-10-05 12:00 | 2022-10-06 11:15 | B22J033 | SD
Microcystins ADDA, Total | 0.0062 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 09:50 | Clarks Branch Water Association | October | PWS00548:SRC-AA | 2210011-01 | Source Water | 272 | 38 | 2022-10-03 09:50 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.014 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 10:00 | City of Wilsonville | October | PWS00954:SRC-HA | 2210010-01 | Source Water | 272 | 38 | 2022-10-03 10:00 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0132 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 10:00 | Salem Public Works | October | PWS00731:SRC-AA | 2210022-01 | Source Water | 272 | 38 | 2022-10-03 10:00 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0029 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 10:00 | Springfield Utility Board | October | PWS00837:SRC-BB | 2210023-01 | Source Water | 272 | 38 | 2022-10-03 10:00 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0072 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 10:10 | City of Powers | October | PWS00672:SRC-AA | 2210009-01 | Source Water | 272 | 38 | 2022-10-03 10:10 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.1976 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-10-03 10:30 | Buell-Red Prairie Water Association | October | PWS01174:SRC-AA | 2210002-01 | Source Water | 272 | 38 | 2022-10-03 10:30 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.2609 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-10-03 10:30 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2210037-01 | Source Water | 272 | 38 | 2022-10-03 10:30 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0221 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 10:45 | Winston-Dillard Water District | October | PWS00957:SRC-AA | 2210027-01 | Source Water | 272 | 38 | 2022-10-03 10:45 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.1219 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2022-10-03 11:15 | City of Cottage Grove | October | PWS00236:SRC-BA | 2210005-01 | Source Water | 272 | 38 | 2022-10-03 11:15 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0091 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 11:15 | Eugene Water & Electric Board | October | PWS00287:SRC-AA | 2210013-01 | Source Water | 272 | 38 | 2022-10-03 11:15 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.6626 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-10-03 11:15 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2210038-01 | Source Water | 272 | 38 | 2022-10-03 11:15 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0081 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 11:20 | Glide Water Association | October | PWS00326:SRC-AA | 2210014-01 | Source Water | 272 | 38 | 2022-10-03 11:20 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0036 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 11:55 | USFS Steamboat Work Center | October | PWS01091:SRC-AA | 2210025-01 | Source Water | 272 | 38 | 2022-10-03 11:55 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0041 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 12:55 | PP&L-Toketee Village | October | PWS01012:SRC-AA | 2210020-01 | Source Water | 272 | 38 | 2022-10-03 12:55 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0032 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 13:45 | City of Gates | October | PWS00317:SRC-AA | 2210006-01 | Source Water | 272 | 38 | 2022-10-03 13:45 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0034 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-03 14:25 | Canby Utility/Veolia Water | October | PWS00157:CH-A | 2210003-01 | Source Water | 272 | 38 | 2022-10-03 14:25 | 2022-10-04 11:15 | 2022-10-04 19:49 | B22J029 | KMC
Microcystins ADDA, Total | 0.0136 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-05 13:00 | Roseburg Forest Products - Dillard | October | PWS94300:SRC-AA | 2210063-01 | Source Water | 274 | 39 | 2022-10-05 13:00 | 2022-10-06 12:00 | 2022-10-07 11:04 | B22J067 | SD
Microcystins ADDA, Total | 0.0163 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 07:50 | Medford Water Commission | October | PWS00513:SRC-BA | 2210083-01 | Source Water | 279 | 39 | 2022-10-10 07:50 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0121 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 08:44 | Lyons Mehama Water District | October | PWS00493:SRC-AA | 2210082-01 | Source Water | 279 | 39 | 2022-10-10 08:44 | 2022-10-12 11:45 | 2022-10-13 13:20 | B22J133 | KMC
Microcystins ADDA, Total | 0.2202 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-10-10 09:00 | City of Creswell | October | PWS00246:SRC-BK | 2210069-01 | Source Water | 279 | 39 | 2022-10-10 09:00 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0182 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 09:45 | City of Gold Hill | October | PWS00333:SRC-AA | 2210072-01 | Source Water | 279 | 39 | 2022-10-10 09:45 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0376 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 09:47 | Milo Academy | October | PWS00250:SRC-AB | 2210084-01 | Source Water | 279 | 39 | 2022-10-10 09:47 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0064 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 09:50 | Umpqua Basin Water Association | October | PWS00719:SRC-AA | 2210092-01 | Source Water | 279 | 39 | 2022-10-10 09:50 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0018 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 10:00 | Shangri La Water District | October | PWS00835:SRC-AA | 2210088-01 | Source Water | 279 | 39 | 2022-10-10 10:00 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0269 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 10:15 | City of Myrtle Creek | October | PWS00550:SRC-AA | 2210076-01 | Source Water | 279 | 39 | 2022-10-10 10:15 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 10:17 | South Fork Water Board - Oregon City | October | PWS00591:SRC-AA | 2210089-01 | Source Water | 279 | 39 | 2022-10-10 10:17 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0098 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 10:20 | City of Newport | October | PWS00566:CH-A | 2210077-01 | Source Water | 279 | 39 | 2022-10-10 10:20 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0241 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 10:20 | Rainbow Water District | October | PWS00839:SRC-BB | 2210085-01 | Source Water | 279 | 39 | 2022-10-10 10:20 | 2022-10-12 11:45 | 2022-10-13 13:20 | B22J133 | KMC
Microcystins ADDA, Total | 0.0065 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 10:30 | Camp Baker BSA | October | PWS91786:SRC-AA | 2210068-01 | Source Water | 279 | 39 | 2022-10-10 10:30 | 2022-10-12 11:45 | 2022-10-13 13:20 | B22J133 | KMC
Microcystins ADDA, Total | 0.0188 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 10:30 | City of Estacada | October | PWS00279:SRC-AA | 2210071-01 | Source Water | 279 | 39 | 2022-10-10 10:30 | 2022-10-12 11:45 | 2022-10-13 13:20 | B22J133 | KMC
Microcystins ADDA, Total | 0.0659 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-10-10 10:30 | Roberts Creek Water District | October | PWS00717:SRC-AA | 2210086-01 | Source Water | 279 | 39 | 2022-10-10 10:30 | 2022-10-12 11:45 | 2022-10-13 13:20 | B22J133 | KMC
Microcystins ADDA, Total | 0.6998 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-10-10 10:40 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2210080-01 | Source Water | 279 | 39 | 2022-10-10 10:40 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0122 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 10:54 | City of Monroe | October | PWS00540:SRC-BA | 2210075-01 | Source Water | 279 | 39 | 2022-10-10 10:54 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0121 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 11:00 | Stayton Water Supply | October | PWS00843:SRC-AA | 2210090-01 | Source Water | 279 | 39 | 2022-10-10 11:00 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.5925 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-10-10 11:25 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2210081-01 | Source Water | 279 | 39 | 2022-10-10 11:25 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0073 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 11:58 | City of Roseburg | October | PWS00720:SRC-AA | 2210079-01 | Source Water | 279 | 39 | 2022-10-10 11:58 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0041 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 12:30 | Ashland Water Department | October | PWS00047:SRC-AA | 2210067-01 | Source Water | 279 | 39 | 2022-10-10 12:30 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0135 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 12:35 | City of Rogue River | October | PWS00712:SRC-BA | 2210078-01 | Source Water | 279 | 39 | 2022-10-10 12:35 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 13:11 | City of Jefferson | October | PWS00408:SRC-AA | 2210073-01 | Source Water | 279 | 39 | 2022-10-10 13:11 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0022 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 14:00 | Seaside Water Department | October | PWS00799:CH-A | 2210087-01 | Source Water | 279 | 39 | 2022-10-10 14:00 | 2022-10-11 11:15 | 2022-10-12 11:25 | B22J118 | SD
Microcystins ADDA, Total | 0.0133 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 14:15 | City of Lowell | October | PWS00492:SRC-BA | 2210074-01 | Source Water | 279 | 39 | 2022-10-10 14:15 | 2022-10-12 11:45 | 2022-10-13 13:20 | B22J133 | KMC
Microcystins ADDA, Total | 0.0523 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-10 14:20 | Tri-City JW & SA | October | PWS00549:SRC-AA | 2210091-01 | Source Water | 279 | 39 | 2022-10-10 14:20 | 2022-10-12 11:45 | 2022-10-13 13:20 | B22J133 | KMC
Microcystins ADDA, Total | 0.0422 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-11 12:15 | City of Drain Shops | October | PWS00260:SRC-AA | 2210070-01 | Source Water | 280 | 40 | 2022-10-11 12:15 | 2022-10-12 11:45 | 2022-10-13 13:20 | B22J133 | KMC
Microcystins ADDA, Total | 0.004 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-11 18:30 | Umpqua Ranch Co-op | October | PWS00714:SRC-AH | 2210129-01 | Source Water | 280 | 40 | 2022-10-11 18:30 | 2022-10-13 11:30 | 2022-10-14 16:25 | B22J145 | KMC
Microcystins ADDA, Total | 0.0033 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-16 11:35 | City of Grants Pass | October | PWS00342:SRC-AA | 2210187-01 | Source Water | 285 | 40 | 2022-10-16 11:35 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0173 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 05:15 | Anglers Cove/SCHWC | October | PWS01483:SRC-AB | 2210181-01 | Source Water | 286 | 40 | 2022-10-17 05:15 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0153 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 05:30 | Hiland WC - Shady Cove | October | PWS01520:SRC-AA | 2210194-01 | Source Water | 286 | 40 | 2022-10-17 05:30 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0045 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 06:30 | City of Albany | October | PWS00012:SRC-BA | 2210184-01 | Source Water | 286 | 40 | 2022-10-17 06:30 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0165 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 08:00 | Glide Water Association | October | PWS00326:SRC-AA | 2210193-01 | Source Water | 286 | 40 | 2022-10-17 08:00 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0154 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 08:08 | Joint Water Commission | October | PWS00379:SRC-AA | 2210196-01 | Source Water | 286 | 40 | 2022-10-17 08:08 | 2022-10-18 14:45 | 2022-10-20 10:44 | B22J204 | SD
Microcystins ADDA, Total | 0.0151 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 08:30 | Salem Public Works | October | PWS00731:SRC-AA | 2210200-01 | Source Water | 286 | 40 | 2022-10-17 08:30 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0154 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 08:30 | Springfield Utility Board | October | PWS00837:CH-A | 2210201-02 | Source Water | 286 | 40 | 2022-10-17 08:30 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0313 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 08:45 | Winston-Dillard Water District | October | PWS00957:SRC-AA | 2210204-01 | Source Water | 286 | 40 | 2022-10-17 08:45 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0264 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 09:17 | City of Wilsonville | October | PWS00954:SRC-HA | 2210190-01 | Source Water | 286 | 40 | 2022-10-17 09:17 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0078 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 09:30 | Springfield Utility Board | October | PWS00837:SRC-BB | 2210201-01 | Source Water | 286 | 40 | 2022-10-17 09:30 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0052 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 09:30 | USFS Tiller Ranger Station | October | PWS01092:SRC-AA | 2210203-01 | Source Water | 286 | 40 | 2022-10-17 09:30 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0122 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 09:45 | North Clackamas County Water Commission | October | PWS00580:SRC-BA | 2210198-01 | Source Water | 286 | 40 | 2022-10-17 09:45 | 2022-10-19 10:30 | 2022-10-20 10:44 | B22J204 | SD
Microcystins ADDA, Total | 0.0123 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 10:30 | City of Powers | October | PWS00672:SRC-AA | 2210189-01 | Source Water | 286 | 40 | 2022-10-17 10:30 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0103 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 10:41 | Clarks Branch Water Association | October | PWS00548:SRC-AA | 2210191-01 | Source Water | 286 | 40 | 2022-10-17 10:41 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0076 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 10:51 | Eugene Water & Electric Board | October | PWS00287:SRC-AA | 2210192-01 | Source Water | 286 | 40 | 2022-10-17 10:51 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.2107 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-10-17 11:00 | Buell-Red Prairie Water Association | October | PWS01174:SRC-AA | 2210182-01 | Source Water | 286 | 40 | 2022-10-17 11:00 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0023 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 11:40 | Lake Oswego - Tigard Water Supply | October | PWS01542:SRC-AA | 2210197-01 | Source Water | 286 | 40 | 2022-10-17 11:40 | 2022-10-19 10:30 | 2022-10-20 10:44 | B22J204 | SD
Microcystins ADDA, Total | 0.006 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 11:46 | City of Pendleton | October | PWS00613:SRC-AM | 2210188-01 | Source Water | 286 | 40 | 2022-10-17 11:46 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.3447 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-10-17 11:50 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2210206-01 | Source Water | 286 | 40 | 2022-10-17 11:50 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 12:05 | PP&L-Toketee Village | October | PWS01012:SRC-AA | 2210199-01 | Source Water | 286 | 40 | 2022-10-17 12:05 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.4456 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-10-17 12:20 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2210205-01 | Source Water | 286 | 40 | 2022-10-17 12:20 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.021 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 13:00 | Roseburg Forest Products - Dillard | October | PWS94300:SRC-AA | 2210180-01 | Source Water | 286 | 40 | 2022-10-17 13:00 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.1811 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-10-17 13:15 | City of Cottage Grove | October | PWS00236:SRC-BA | 2210185-01 | Source Water | 286 | 40 | 2022-10-17 13:15 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0229 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 13:30 | City of Gates | October | PWS00317:SRC-AA | 2210186-01 | Source Water | 286 | 40 | 2022-10-17 13:30 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 14:00 | USFS Steamboat Work Center | October | PWS01091:SRC-AA | 2210202-01 | Source Water | 286 | 40 | 2022-10-17 14:00 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0014 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 14:15 | Canby Utility/Veolia Water | October | PWS00157:CH-A | 2210183-01 | Source Water | 286 | 40 | 2022-10-17 14:15 | 2022-10-18 11:20 | 2022-10-18 18:08 | B22J199 | SD
Microcystins ADDA, Total | 0.0049 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-17 14:25 | Jackson Co Pks Emigrant Lake | October | PWS90730:CH-A | 2210195-01 | Source Water | 286 | 40 | 2022-10-17 14:25 | 2022-10-19 10:30 | 2022-10-20 10:44 | B22J204 | SD
Microcystins ADDA, Total | 0.0243 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 | City of Myrtle Creek | October | PWS00550:SRC-AA | 2210156-01 | Source Water | 293 | 41 | 2022-10-24 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0062 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 07:07 | Medford Water Commission | October | PWS00513:SRC-BA | 2210163-01 | Source Water | 293 | 41 | 2022-10-24 07:07 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0012 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 07:48 | Lyons Mehama Water District | October | PWS00493:SRC-AA | 2210162-01 | Source Water | 293 | 41 | 2022-10-24 07:48 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.0037 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 08:30 | Roberts Creek Water District | October | PWS00717:SRC-AA | 2210166-01 | Source Water | 293 | 41 | 2022-10-24 08:30 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.0009 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 08:30 | South Fork Water Board - Oregon City | October | PWS00591:SRC-AA | 2210169-01 | Source Water | 293 | 41 | 2022-10-24 08:30 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0112 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 08:34 | Milo Academy | October | PWS00250:SRC-AB | 2210164-01 | Source Water | 293 | 41 | 2022-10-24 08:34 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.1257 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2022-10-24 09:00 | City of Creswell | October | PWS00246:SRC-BK | 2210149-01 | Source Water | 293 | 41 | 2022-10-24 09:00 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.008 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 10:00 | Shangri La Water District | October | PWS00835:SRC-AA | 2210168-01 | Source Water | 293 | 41 | 2022-10-24 10:00 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0382 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 10:30 | City of Lowell | October | PWS00492:SRC-BA | 2210154-01 | Source Water | 293 | 41 | 2022-10-24 10:30 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.0136 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 10:35 | Camp Baker BSA | October | PWS91786:SRC-AA | 2210148-01 | Source Water | 293 | 41 | 2022-10-24 10:35 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.0022 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 10:35 | City of Estacada | October | PWS00279:SRC-AA | 2210151-01 | Source Water | 293 | 41 | 2022-10-24 10:35 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0109 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 11:00 | City of Newport | October | PWS00566:CH-A | 2210157-01 | Source Water | 293 | 41 | 2022-10-24 11:00 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0097 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 11:00 | Stayton Water Supply | October | PWS00843:SRC-AA | 2210170-01 | Source Water | 293 | 41 | 2022-10-24 11:00 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0131 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 11:16 | City of Roseburg | October | PWS00720:SRC-AA | 2210159-01 | Source Water | 293 | 41 | 2022-10-24 11:16 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.013 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 11:25 | Umpqua Basin Water Association | October | PWS00719:SRC-AA | 2210172-01 | Source Water | 293 | 41 | 2022-10-24 11:25 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0062 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 11:30 | City of Jefferson | October | PWS00408:SRC-AA | 2210153-01 | Source Water | 293 | 41 | 2022-10-24 11:30 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.022 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 12:21 | City of Monroe | October | PWS00540:SRC-BA | 2210155-01 | Source Water | 293 | 41 | 2022-10-24 12:21 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0022 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 12:30 | Ashland Water Department | October | PWS00047:SRC-AA | 2210147-01 | Source Water | 293 | 41 | 2022-10-24 12:30 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0249 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 12:50 | City of Drain Shops | October | PWS00260:SRC-AA | 2210150-01 | Source Water | 293 | 41 | 2022-10-24 12:50 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0113 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 14:00 | Seaside Water Department | October | PWS00799:CH-A | 2210167-01 | Source Water | 293 | 41 | 2022-10-24 14:00 | 2022-10-25 11:00 | 2022-10-25 18:26 | B22J286 | KMC
Microcystins ADDA, Total | 0.0124 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 14:00 | Tri-City JW & SA | October | PWS00549:SRC-AA | 2210171-01 | Source Water | 293 | 41 | 2022-10-24 14:00 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-24 15:55 | Rainbow Water District | October | PWS00839:SRC-BB | 2210165-01 | Source Water | 293 | 41 | 2022-10-24 15:55 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.0174 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-25 09:15 | City of Gold Hill | October | PWS00333:SRC-AA | 2210152-01 | Source Water | 294 | 42 | 2022-10-25 09:15 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.4376 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-10-25 09:40 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2210161-01 | Source Water | 294 | 42 | 2022-10-25 09:40 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.0038 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-25 10:15 | City of Rogue River | October | PWS00712:SRC-BA | 2210158-01 | Source Water | 294 | 42 | 2022-10-25 10:15 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.9457 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-10-25 10:35 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2210160-01 | Source Water | 294 | 42 | 2022-10-25 10:35 | 2022-10-26 12:45 | 2022-10-27 11:37 | B22J310 | SD
Microcystins ADDA, Total | 0.0258 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-27 09:00 | Umpqua Ranch Co-op | October | PWS00714:SRC-AH | 2210280-01 | Source Water | 296 | 42 | 2022-10-27 09:00 | 2022-10-28 10:30 | 2022-10-31 13:08 | B22J321 | KMC
Microcystins ADDA, Total | 0.0035 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-30 11:45 | City of Grants Pass | October | PWS00342:SRC-AA | 2210293-01 | Source Water | 299 | 42 | 2022-10-30 11:45 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0072 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 05:15 | Anglers Cove/SCHWC | October | PWS01483:SRC-AB | 2210287-01 | Source Water | 300 | 42 | 2022-10-31 05:15 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0056 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 05:30 | Hiland WC - Shady Cove | October | PWS01520:SRC-AA | 2210300-01 | Source Water | 300 | 42 | 2022-10-31 05:30 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0128 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 06:00 | City of Albany | October | PWS00012:SRC-BA | 2210290-01 | Source Water | 300 | 42 | 2022-10-31 06:00 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0025 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 07:45 | USFS Tiller Ranger Station | October | PWS01092:SRC-AA | 2210310-01 | Source Water | 300 | 42 | 2022-10-31 07:45 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0037 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 07:54 | Roseburg Forest Products - Dillard | October | PWS94300:SRC-AA | 2210306-01 | Source Water | 300 | 42 | 2022-10-31 07:54 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0024 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 08:00 | Glide Water Association | October | PWS00326:SRC-AA | 2210299-01 | Source Water | 300 | 42 | 2022-10-31 08:00 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0308 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 08:02 | Joint Water Commission | October | PWS00379:SRC-AA | 2210302-01RE1 | Source Water | 300 | 42 | 2022-10-31 08:02 | 2022-11-01 14:30 | 2022-11-03 10:32 | B22K032 | SD
Microcystins ADDA, Total | 0.0173 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 08:15 | Winston-Dillard Water District | October | PWS00957:SRC-AA | 2210311-01 | Source Water | 300 | 42 | 2022-10-31 08:15 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0057 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 08:50 | Clarks Branch Water Association | October | PWS00548:SRC-AA | 2210297-01 | Source Water | 300 | 42 | 2022-10-31 08:50 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0182 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 09:00 | Salem Public Works | October | PWS00731:SRC-AA | 2210307-01 | Source Water | 300 | 42 | 2022-10-31 09:00 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.005 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 09:15 | Springfield Utility Board | October | PWS00837:SRC-BB | 2210308-01 | Source Water | 300 | 42 | 2022-10-31 09:15 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0831 | 0.064 | 0.115 | 0.15 | 0.3 | 1 |  |  |  | µg/L | 2022-10-31 09:30 | City of Cottage Grove | October | PWS00236:SRC-BA | 2210291-01 | Source Water | 300 | 42 | 2022-10-31 09:30 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0034 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 10:00 | Springfield Utility Board | October | PWS00837:CH-A | 2210308-02 | Source Water | 300 | 42 | 2022-10-31 10:00 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.003 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 10:00 | USFS Steamboat Work Center | October | PWS01091:SRC-AA | 2210309-01 | Source Water | 300 | 42 | 2022-10-31 10:00 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.007 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 10:33 | City of Pendleton | October | PWS00613:SRC-AM | 2210294-01 | Source Water | 300 | 42 | 2022-10-31 10:33 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0083 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 10:35 | PP&L-Toketee Village | October | PWS01012:SRC-AA | 2210305-01 | Source Water | 300 | 42 | 2022-10-31 10:35 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.002 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 10:40 | Lake Oswego - Tigard Water Supply | October | PWS01542:SRC-AA | 2210303-01RE1 | Source Water | 300 | 42 | 2022-10-31 10:40 | 2022-11-02 10:30 | 2022-11-03 10:32 | B22K032 | SD
Microcystins ADDA, Total | 0.0114 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 10:45 | City of Powers | October | PWS00672:SRC-AA | 2210295-01 | Source Water | 300 | 42 | 2022-10-31 10:45 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.2799 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-10-31 11:30 | Josephine Co Pks Lake Selmac 2 | October | PWS94645:SRC-AA | 2211007-01RE1 | Source Water | 300 | 42 | 2022-10-31 11:30 | 2022-11-02 10:30 | 2022-11-03 10:32 | B22K032 | SD
Microcystins ADDA, Total | 0.0095 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 11:30 | North Clackamas County Water Commission | October | PWS00580:SRC-BA | 2210304-01 | Source Water | 300 | 42 | 2022-10-31 11:30 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0367 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 11:45 | City of Gates | October | PWS00317:SRC-AA | 2210292-01 | Source Water | 300 | 42 | 2022-10-31 11:45 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.3944 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-10-31 12:00 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2211006-01RE1 | Source Water | 300 | 42 | 2022-10-31 12:00 | 2022-11-02 10:30 | 2022-11-03 10:32 | B22K032 | SD
Microcystins ADDA, Total | 0.3944 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 | 1 | µg/L | 2022-10-31 12:00 | Josephine Co Pks Lake Selmac 1 | October | PWS90186:SRC-AA | 2211006-01RE1 | Source Water | 300 | 42 | 2022-10-31 12:00 | 2022-11-02 10:30 | 2022-11-03 10:32 | B22K032 | SD
Microcystins ADDA, Total | 0.0108 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 12:07 | Eugene Water & Electric Board | October | PWS00287:SRC-AA | 2210298-01 | Source Water | 300 | 42 | 2022-10-31 12:07 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0239 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 12:40 | City of Wilsonville | October | PWS00954:SRC-HA | 2210296-01 | Source Water | 300 | 42 | 2022-10-31 12:40 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.0131 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-10-31 14:50 | Canby Utility/Veolia Water | October | PWS00157:CH-A | 2210289-01 | Source Water | 300 | 42 | 2022-10-31 14:50 | 2022-11-01 11:00 | 2022-11-02 12:30 | B22K005 | KMC
Microcystins ADDA, Total | 0.1823 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-11-01 11:00 | Buell-Red Prairie Water Association | November | PWS01174:SRC-AA | 2210288-01RE1 | Source Water | 300 | 42 | 2022-11-01 11:00 | 2022-11-02 10:30 | 2022-11-03 10:32 | B22K032 | SD
Microcystins ADDA, Total | 0.0074 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-11-03 08:06 | Jackson Co Pks Emigrant Lake | November | PWS90730:CH-A | 2211018-01RE1 | Source Water | 302 | 43 | 2022-11-03 08:06 | 2022-11-04 10:30 | 2022-11-09 17:41 | B22K087 | SD
Microcystins ADDA, Total | 0.2223 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-11-08 11:30 | Josephine Co Pks Lake Selmac 2 | November | PWS94645:SRC-AA | 2211031-01 | Source Water | 307 | 43 | 2022-11-08 11:30 | 2022-11-09 11:20 | 2022-11-09 17:41 | B22K087 | SD
Microcystins ADDA, Total | 0.2576 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-11-08 12:15 | Josephine Co Pks Lake Selmac 1 | November | PWS90186:SRC-AA | 2211030-01 | Source Water | 307 | 43 | 2022-11-08 12:15 | 2022-11-09 11:20 | 2022-11-09 17:41 | B22K087 | SD
Microcystins ADDA, Total | 0.2365 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-11-14 12:00 | Josephine Co Pks Lake Selmac 2 | November | PWS94645:SRC-AA | 2211051-01 | Source Water | 313 | 44 | 2022-11-14 12:00 | 2022-11-15 12:40 | 2022-11-16 12:06 | B22K153 | SD
Microcystins ADDA, Total | 0.0497 | 0.064 | 0.115 | 0.15 | 0.3 |  |  |  |  | µg/L | 2022-11-14 12:10 | Josephine Co Pks Lake Selmac 1 | November | PWS90186:SRC-AA | 2211050-01 | Source Water | 313 | 44 | 2022-11-14 12:10 | 2022-11-15 12:40 | 2022-11-16 12:06 | B22K153 | SD
Microcystins ADDA, Total | 0.245 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 | 1 |  | µg/L | 2022-11-21 11:30 | Josephine Co Pks Lake Selmac 2 | November | PWS94645:SRC-AA | 2211082-01 | Source Water | 320 | 45 | 2022-11-21 11:30 | 2022-11-22 10:45 | 2022-11-22 18:06 | B22K243 | KMC
Microcystins ADDA, Total | 0.1376 | 0.064 | 0.115 | 0.15 | 0.3 | 1 | 1 |  |  | µg/L | 2022-11-21 11:50 | Josephine Co Pks Lake Selmac 1 | November | PWS90186:SRC-AA | 2211081-01 | Source Water | 320 | 45 | 2022-11-21 11:50 | 2022-11-22 10:45 | 2022-11-22 18:06 | B22K243 | KMC
Note on G1: Kale Clauson: Over Detection Limit
Note on H1: Kale Clauson: Over Reporting Limit
Note on I1: Kale Clauson: Over Trigger Level
Note on J1: Kale Clauson: Over Health Advisory Limit

## Fig2_Tabulation
| 2020 |  |  | 2021 |  |  | 2022 |  |  | Detections
CLIENT | >DL (0.064 µg/L) | >RL (0.115 µg/L) | >HAL (0.3 µg/L) | >DL (0.064 µg/L) | >RL (0.115 µg/L) | >HAL (0.3 µg/L) | >DL (0.064 µg/L) | >RL (0.115 µg/L) | >HAL (0.3 µg/L) | 2020 | 2021 | 2022
Anglers Cove/SCHWC | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ashland Water Department | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
Buell-Red Prairie Water Association | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1
Camp Baker BSA | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Canby Utility/Veolia Water | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Albany | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Cottage Grove | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1
City of Creswell | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1
City of Drain Shops | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Estacada | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Gates | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
City of Gold Hill | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Grants Pass | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Jefferson | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Lowell | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Monroe | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Myrtle Creek | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1
City of Newport | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0
City of Pendleton | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Powers | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Rogue River | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Roseburg | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
City of Wilsonville | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Clarks Branch Water Association | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1
Douglas County Parks - Whistler's Bend | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Eugene Water & Electric Board | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Glide Water Association | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hiland WC - Shady Cove | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Jackson Co Pks Emigrant Lake | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
Joint Water Commission | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Josephine Co Pks Lake Selmac 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1
Josephine Co Pks Lake Selmac 2 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1
Lake Oswego - Tigard Water Supply | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Lyons Mehama Water District | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Medford Water Commission | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Milo Academy | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1
North Clackamas County Water Commission | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PP&L-Toketee Village | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Rainbow Water District | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Roberts Creek Water District | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1
Roseburg Forest Products - Dillard | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1
Salem Public Works | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
Seaside Water Department | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Shangri La Water District | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
South Fork Water Board - Oregon City | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Springfield Utility Board | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Stayton Water Supply | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
Tri-City JW & SA | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1
Umpqua Basin Water Association | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Umpqua Ranch Co-op | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
USFS Horseshoe Bend CG |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
USFS Steamboat Work Center | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
USFS Tiller Ranger Station | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Winston-Dillard Water District | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1
Grand Total | 2 | 2 | 4 | 2 | 0 | 5 | 3 | 6 | 3 | 8 | 7 | 12