## 091624ICP_TO178-477 Sample Data
09162024_ICP-OES_TO 178-477 Sample Data
ANALYSIS TECHNIUE: |  | ICP-OES |  |  |  | Calibration date: |  | 2024-09-16
ANALYSIS METHOD: |  | 30M |  |  |  | Calibration Standards: |  | 2024-09-04
ANALYSIS BATCH: |  | TO 178-475, 476 & 477 + TO 212-22 |  |  |  | Task Order# |  | 68HERC22F0178
Task Order: |  | TO 178 |  |  |  | Instrument : |  | Agilent 5900 SVDV ICP-OES (S/N: MY2138CP13) with Agilent SPS 4 Autosampler (S/N: AU213011544)
UNITS: |  | mg/L (ppm) |  |  |  | Analyst: |  | Mahendranath Arambewela
Comments: | [5 empty] | SOP# |  | K-LRTD-SOP-1185-2
Sample results are reported as mg/L (ppm)
Worksheet exported from: C:\Users\USEPA\Documents\Agilent\ICP Expert\My Results\09162024_ICP-OES_TO 178-475 476 & 477 + TO 212-22.esws
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Solution Label | Type | Date Time | Ag 328.068 C ppm | Al 396.152 C ppm | As 188.980 C ppm | B 249.678 C ppm | Ba 455.403 C ppm | Be 313.107 C ppm | Ca 370.602 C ppm | Cd 214.439 C ppm | Co 238.892 C ppm | Cr 267.716 C ppm | Cu 324.754 C ppm | Fe 259.940 C ppm | K 766.491 C ppm | Li 670.783 C ppm | Mg 279.078 C ppm | Mn 257.610 C ppm | Mo 202.032 C ppm | Na 588.995 C ppm | Ni 231.604 C ppm | P 178.222 C ppm | Pb 220.353 C ppm | S 181.972 C ppm | Sb 217.582 C ppm | Se 196.026 C ppm | Si 251.611 C ppm | Sr 421.552 C ppm | Ti 334.941 C ppm | V 292.401 C ppm | Zn 206.200 C ppm
10x 0.5% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:12 | -0.0002 | 16.7305 | 0.0047 | 0.0161 | 0.0027 | -0.0001 | 0.2078 | -0.0006 | -0.0008 | 0.0305 | 0.002 | 165.192 | -0.0097 | 0.185 | 0.8543 | 1.1406 | 0.0135 | 5.1539 | 0.6358 | 1.1378 | -0.0013 | 62.0289 | 0.112 | 0.0028 | 0.1715 | 0.0005 | 0.099 | 0.0081 | 0.6053
10x 1% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:14 | -0.0003 | 27.958 | 0.0007 | 0.0364 | 0.0028 | -0.0002 | 0.365 | -0.0008 | -0.0009 | 0.0487 | 0.0036 | 232.2671 | 0.0034 | 0.2207 | 1.4973 | 1.9398 | 0.0086 | 14.4739 | 0.9754 | 1.4515 | 0.0011 | 120.8814 | 0.1134 | -0.0035 | 0.2665 | 0.0008 | 0.1934 | 0.0151 | 0.9293
10x 5% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:17 | -0.0004 | 12.8167 | 0.0022 | 0.1787 | 0.003 | -0.0002 | 1.8491 | 0.0003 | 0.0117 | 0.1818 | 0.009 | 942.0008 | 0.0787 | 0.2548 | 7.4451 | 9.3401 | 0.0109 | 33.0439 | 1.7583 | 4.5448 | 0.0171 | 610.9164 | 0.1236 | -0.0188 | 0.5968 | 0.0037 | 0.9988 | 0.052 | 5.1065
10x 10% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:19 | -0.0001 | 1.1084 | 0.0044 | 0.3716 | 0.0049 | -0.0001 | 3.46 | 0.0006 | -0.0201 | 0.3857 | 0.0269 | 1803.3566 | 0.1637 | -0.0001 | 14.3827 | 17.8596 | 0.0145 | 64.8239 | 0.8456 | 0.1584 | 0.0536 | 1161.3306 | 0.1428 | -0.0442 | 0.7729 | 0.0067 | 2.032 | 0.1281 | 12.3307
10x 15% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:22 | 0.0008 | 1.6543 | 0.0024 | 0.5727 | 0.0069 | -0.0001 | 5.0803 | 0.0044 | -0.0217 | 0.574 | 0.0429 | 2533.4012 | 0.3017 | 0.0193 | 21.0202 | 26.9829 | 0.0205 | 97.0954 | 1.0375 | 0.241 | 0.0866 | 1657.7279 | 0.1884 | -0.0706 | 1.1423 | 0.0106 | 2.9999 | 0.2017 | 18.3706
10x 0.5% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:24 | -0.0004 | 9.2607 | 0.0001 | 0.0195 | 0.0008 | -0.0002 | 0.2165 | -0.0007 | -0.0005 | 0.0335 | 0.0048 | 166.8406 | -0.0054 | 0.0205 | 0.8947 | 0.9895 | 0.0023 | 3.9598 | 1.2588 | 0.4528 | -0.0015 | 64.8784 | 0.1198 | -0.0011 | 0.0991 | 0.0005 | 0.1114 | 0.0075 | 0.4791
10x 1% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:27 | -0.0005 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | -0.0005 | 0.003 | 0.0598 | 0.0075 | 252.3585 | -0.005 | 0.0271 | 1.6867 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.34 | -0.0011 | 126.2896 | 0.1232 | -0.0045 | 0.1505 | 0.0008 | 0.2255 | 0.0103 | 0.8442
10x 5% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:29 | -0.0007 | 10.346 | 0.0036 | 0.2226 | 0.0016 | -0.0002 | 1.9549 | -0.0004 | 0.0045 | 0.2301 | 0.0089 | 995.943 | 0.0957 | 0.0909 | 7.9403 | 9.5788 | 0.0086 | 33.5317 | 2.5013 | 0.4545 | 0.0218 | 641.9245 | 0.1276 | -0.0256 | 0.469 | 0.0038 | 1.1323 | 0.0654 | 4.2772
10x 10% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:32 | -0.0005 | 1.1497 | 0.0007 | 0.4192 | 0.0033 | -0.0002 | 3.8504 | 0.0052 | -0.0184 | 0.4303 | 0.0286 | 1867.6644 | 0.246 | -0.0007 | 14.6679 | 18.5866 | 0.0157 | 62.795 | 2.4582 | 0.4111 | 0.0513 | 1190.7709 | 0.1454 | -0.0374 | 0.8201 | 0.0073 | 2.1431 | 0.1329 | 9.1637
10x 15% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:34 | 0.0002 | 2.6273 | 0.0027 | 0.6048 | 0.0045 | -0.0001 | 5.4141 | 0.0036 | -0.024 | 0.6136 | 0.0459 | 2550.4615 | 0.3147 | 0.0679 | 20.89 | 27.0173 | 0.021 | 91.9079 | 2.762 | 0.504 | 0.0881 | 1668.2033 | 0.1881 | -0.076 | 1.163 | 0.0105 | 3.0669 | 0.1928 | 14.6004
09032024 MB | Sample | 2024-09-16 15:44 | 0.0001 | 0.0065 | 0.0046 | -0.0008 | 0.0003 | 0 | 0.0054 | 0 | 0 | -0.0003 | 0.0001 | 0.0112 | -0.0022 | -0.0025 | 0.0165 | 0.0002 | 0.0108 | 0.0161 | -0.0001 | 0.007 | -0.0047 | 0.0203 | 0.1349 | 0.0211 | 0.0511 | 0 | 0.0008 | 0.0002 | 0.0026
10x 09032024 MB | Sample | 2024-09-16 15:47 | -0.0001 | 0.0021 | 0.0012 | -0.0022 | -0.0001 | -0.0001 | 0.0692 | 0 | -0.0002 | 0.0001 | 0.0095 | 0.0095 | 0.0043 | -0.0024 | 0.0092 | 0.0008 | 0.0038 | 0.0017 | -0.0019 | 0.0041 | -0.0038 | 0.0121 | 0.1222 | 0.014 | 0.0395 | 0.0001 | 0.0004 | -0.0002 | 0.0043
10x 09032024 MFB | Sample | 2024-09-16 15:50 | 0.01 | 0.4842 | 0.2424 | 0.0482 | 0.2575 | 0.0488 | 0.5556 | 0.0972 | 0.5035 | 0.2469 | 0.2494 | 0.5049 | 0.4666 | -0.0026 | 0.481 | 0.1007 | 0.0969 | 0.5096 | 0.2426 | 0.1056 | 0.2554 | 0.2545 | 0.2459 | 0.2492 | 0.1389 | 0.1021 | 0.0984 | 0.2517 | 0.0959
10x 1% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:27 | -0.0005 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | -0.0005 | 0.003 | 0.0598 | 0.0075 | 252.3585 | -0.005 | 0.0271 | 1.6867 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.34 | -0.0011 | 126.2896 | 0.1232 | -0.0045 | 0.1505 | 0.0008 | 0.2255 | 0.0103 | 0.8442
10x 1% FeSO4 NMC Battery-TM -MS -Analytical Duplicate | Sample | 2024-09-16 15:52 | -0.0006 | 9.926 | 0.0023 | 0.0414 | 0.0005 | -0.0001 | 0.4151 | -0.0006 | 0.001 | 0.0595 | 0.0074 | 250.9591 | 0.0009 | 0.0273 | 1.6778 | 1.866 | 0.0054 | 7.6281 | 1.3942 | 0.339 | -0.0019 | 125.1867 | 0.107 | 0.0002 | 0.1468 | 0.0008 | 0.2242 | 0.0141 | 0.837
10x 1% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:27 | -0.0005 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | -0.0005 | 0.003 | 0.0598 | 0.0075 | 252.3585 | -0.005 | 0.0271 | 1.6867 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.34 | -0.0011 | 126.2896 | 0.1232 | -0.0045 | 0.1505 | 0.0008 | 0.2255 | 0.0103 | 0.8442
10x 1% FeSO4 NMC Battery-TM - Lab Duplicate | Sample | 2024-09-16 15:55 | -0.0005 | 10.1338 | 0.0009 | 0.0418 | 0.0005 | -0.0001 | 0.4152 | -0.0004 | 0.0019 | 0.0602 | 0.007 | 254.5683 | 0.0132 | 0.0276 | 1.702 | 1.8836 | 0.0041 | 7.7345 | 1.4203 | 0.3435 | 0.0011 | 127.3304 | 0.1219 | -0.0028 | 0.1507 | 0.0008 | 0.2256 | 0.0104 | 0.8588
10x 1% FeSO4- NMC Battery-TM - MS | Sample | 2024-09-16 15:57 | 0.0086 | 9.0598 | 0.2474 | 0.0856 | 0.2559 | 0.0501 | 0.8491 | 0.097 | 0.4421 | 0.2973 | 0.264 | 226.1028 | 0.5031 | 0.0246 | 1.9952 | 1.7875 | 0.0984 | 7.4719 | 1.4852 | 0.4004 | 0.2582 | 113.0942 | 0.2897 | 0.2235 | 0.2466 | 0.1025 | 0.2923 | 0.2657 | 0.8221
10x 1% FeSO4- NMC Battery-TM - MS - Analytical Duplicate | Sample | 2024-09-16 16:00 | 0.0088 | 9.0013 | 0.2404 | 0.0851 | 0.2538 | 0.0495 | 0.8404 | 0.0964 | 0.4283 | 0.2947 | 0.2623 | 223.522 | 0.5033 | 0.0244 | 1.9777 | 1.7773 | 0.0995 | 7.4075 | 1.4732 | 0.4017 | 0.2545 | 112.3137 | 0.2914 | 0.2225 | 0.2436 | 0.1017 | 0.2906 | 0.2634 | 0.8162
09162024-Lab Blank2 | Sample | 2024-09-16 16:10 | -0.0002 | 0.0002 | 0.0028 | -0.002 | -0.0002 | -0.0001 | -0.0004 | 0 | 0 | -0.0001 | -0.0001 | -0.0006 | -0.0006 | -0.0028 | 0.0089 | 0.0003 | 0.0099 | 0.0006 | -0.0008 | 0.0029 | -0.0038 | 0.0026 | 0.1404 | 0.0175 | 0.0405 | 0 | 0.0005 | 0.0001 | -0.0003
10x 09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:12 | 0.0096 | 0.4748 | 0.2451 | 0.0481 | 0.2503 | 0.0477 | 0.4792 | 0.0958 | 0.4968 | 0.2431 | 0.2458 | 0.4895 | 0.4567 | -0.0025 | 0.4718 | 0.099 | 0.0977 | 0.4998 | 0.2393 | 0.1027 | 0.2517 | 0.249 | 0.2292 | 0.2467 | 0.138 | 0.0992 | 0.0975 | 0.2482 | 0.0947
10x 09162024-Lab Blank Spike2-Analytical Duplicate | Sample | 2024-09-16 16:15 | 0.0094 | 0.4731 | 0.2451 | 0.0479 | 0.2495 | 0.0479 | 0.4755 | 0.095 | 0.4928 | 0.2416 | 0.244 | 0.4898 | 0.4622 | -0.0025 | 0.4708 | 0.0983 | 0.0975 | 0.4962 | 0.239 | 0.1038 | 0.2495 | 0.2474 | 0.2699 | 0.2515 | 0.1385 | 0.099 | 0.0971 | 0.2466 | 0.0941
5x 09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:17 | 0.0191 | 0.9536 | 0.4833 | 0.0956 | 0.4984 | 0.0957 | 0.9469 | 0.1902 | 0.9818 | 0.4803 | 0.4857 | 0.9772 | 0.9523 | -0.004 | 0.9318 | 0.1966 | 0.1898 | 0.995 | 0.4745 | 0.2007 | 0.4971 | 0.4885 | 0.4249 | 0.4916 | 0.2326 | 0.1967 | 0.1917 | 0.4906 | 0.1871
09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:20 | 0.0982 | 4.8272 | 2.5113 | 0.496 | 2.5174 | 0.4893 | 4.835 | 0.9661 | 4.9875 | 2.4259 | 2.4646 | 4.9443 | 4.9978 | -0.0021 | 4.793 | 0.9926 | 0.9949 | 5.0132 | 2.4267 | 1.0326 | 2.5598 | 2.5024 | 2.2541 | 2.5109 | 1.0178 | 0.9985 | 0.9771 | 2.4875 | 0.9598
09162024-Lab Blank Spike2-Analytical Duplicate | Sample | 2024-09-16 16:22 | 0.0983 | 4.7796 | 2.53 | 0.496 | 2.5 | 0.4859 | 4.8357 | 0.9665 | 4.9703 | 2.4279 | 2.4548 | 4.9051 | 4.9558 | -0.0021 | 4.7911 | 0.9789 | 1.0113 | 4.9854 | 2.4277 | 1.0324 | 2.5647 | 2.5009 | 2.295 | 2.5114 | 1.0228 | 0.9917 | 0.9779 | 2.4727 | 0.9611
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Lowest Cal Std |  |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Highest Cal Std |  |  | 20 | 200 | 20 | 20 | 50 | 20 | 2000 | 20 | 20 | 50 | 100 | 1000 | 100 | 100 | 2000 | 50 | 20 | 100 | 50 | 200 | 200 | 100 | 50 | 20 | 100 | 20 | 50 | 20 | 100
Correlation |  |  | 1 | 0.99998 | 0.99998 | 1 | 1 | 0.99999 | 0.99996 | 1 | 0.99999 | 0.99996 | 0.99998 | 1 | 0.99998 | 0.99999 | 0.99999 | 0.99999 | 0.99976 | 0.99998 | 0.99998 | 0.99999 | 0.99981 | 0.99997 | 0.99994 | 0.99998 | 0.99985 | 0.99998 | 1 | 1 | 1
ICP-OES MDL |  |  | 0.0045598 | 0.0048715 | 0.0060725 | 0.0016216 | 0.00011879 | 0.000168 | 0.0038865 | 0.00024729 | 0.00024729 | 0.0007 | 0.00076374 | 0.0031683 | 0.032873 | 0.0015923 | 0.0080477 | 0.0008057 | 0.00050865 | 0.001351 | 0.0024937 | 0.0072662 | 0.0049325 | 0.0058515 | 0.0056274 | 0.0064431 | 0.002358 | 0.000168 | 0.00011879 | 0.00093287 | 0.00059397
ICP-OES MRL/PRDL |  |  | 0.045598 | 0.048715 | 0.060725 | 0.016216 | 0.0011879 | 0.00168 | 0.038865 | 0.0024729 | 0.0024729 | 0.0074 | 0.0076374 | 0.031683 | 0.32873 | 0.015923 | 0.080477 | 0.008057 | 0.0050865 | 0.01351 | 0.024937 | 0.072662 | 0.049325 | 0.058515 | 0.056274 | 0.064431 | 0.02358 | 0.00168 | 0.0011879 | 0.0093287 | 0.0059397
Where
Calculated/Estimated MDLs & MRLs
Dilution required
* | Sampling Error
** | Different concentration than the normal
Mislabeled

## 091624ICP-OES_TO178-477Data
09162024_ICP-OES_TO 178-477 Data Report
ANALYSIS TECHNIUE: |  | ICP-OES |  |  |  | Calibration date: |  | 2024-09-16
ANALYSIS METHOD: |  | 30M |  |  |  | Calibration Standards: |  | 2024-09-04
ANALYSIS BATCH: |  | TO 178-475, 476 & 477 + TO 212-22 |  |  |  | Task Order# |  | 68HERC22F0178
Task Order: |  | TO 178 |  |  |  | Instrument : |  | Agilent 5900 SVDV ICP-OES (S/N: MY2138CP13) with Agilent SPS 4 Autosampler (S/N: AU213011544)
UNITS: |  | mg/L (ppm) |  |  |  | Analyst: |  | Mahendranath Arambewela
Comments: | [5 empty] | SOP# |  | K-LRTD-SOP-1185-2
Sample results are reported as mg/L (ppm)
Worksheet exported from: C:\Users\USEPA\Documents\Agilent\ICP Expert\My Results\09162024_ICP-OES_TO 178-475 476 & 477 + TO 212-22.esws
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
|  |  | Ag (328.068 nm) ppm | Al (396.152 nm) ppm | As (188.980 nm) ppm | B (249.678 nm) ppm | Ba (455.403 nm) ppm | Be (313.107 nm) ppm | Ca (370.602 nm) ppm | Cd (214.439 nm) ppm | Co (238.892 nm) ppm | Cr (267.716 nm) ppm | Cu (324.754 nm) ppm | Fe (259.940 nm) ppm | K (766.491 nm) ppm | Li (670.783 nm) ppm | Mg (279.078 nm) ppm | Mn (257.610 nm) ppm | Mo (202.032 nm) ppm | Na (588.995 nm) ppm | Ni (231.604 nm) ppm | P (178.222 nm) ppm | Pb (220.353 nm) ppm | S (181.972 nm) ppm | Sb (217.582 nm) ppm | Se (196.026 nm) ppm | Si (251.611 nm) ppm | Sr (421.552 nm) ppm | Ti (334.941 nm) ppm | V (292.401 nm) ppm | Zn (206.200 nm) ppm
IECS1 (Al 1000 Int.) | IECS |  | Al: 1000
IECS2 (Fe 1000 Int.) | IECS |  | Fe: 1000
IECS3 (Mn 100 Int.) | IECS |  | Mn: 100
IECS4 (Ca Mg 1000 Int.) | IECS |  | Ca: 1000; Mg: 1000
IECS5 (Ti 20 Int.) | IECS |  | Ti: 20
ICV 0.5ppm All | QC |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
ICV 5ppm All | QC |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
CCS1 (SS 5ppm) | QC |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
CCV2 5ppm All | QC |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
ICV 1ppm All | QC | bar: Ag to Zn
|  | MDL | 0.0045598 | 0.0048715 | 0.0060725 | 0.0016216 | 0.00011879 | 0.000168 | 0.0038865 | 0.00024729 | 0.00024729 | 0.00038187 | 0.00076374 | 0.0031683 | 0.032873 | 0.0015923 | 0.0080477 | 0.0008057 | 0.00050865 | 0.001351 | 0.0024937 | 0.0072662 | 0.0049325 | 0.0058515 | 0.0056274 | 0.0064431 | 0.002358 | 0.000168 | 0.00011879 | 0.00093287 | 0.00059397
|  | MRL | 0.045598 | 0.048715 | 0.060725 | 0.016216 | 0.0011879 | 0.00168 | 0.038865 | 0.0024729 | 0.0024729 | 0.0038187 | 0.0076374 | 0.031683 | 0.32873 | 0.015923 | 0.080477 | 0.008057 | 0.0050865 | 0.01351 | 0.024937 | 0.072662 | 0.049325 | 0.058515 | 0.056274 | 0.064431 | 0.02358 | 0.00168 | 0.0011879 | 0.0093287 | 0.0059397
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Solution Label | Type | Date Time | Ag 328.068 C ppm | Al 396.152 C ppm | As 188.980 C ppm | B 249.678 C ppm | Ba 455.403 C ppm | Be 313.107 C ppm | Ca 370.602 C ppm | Cd 214.439 C ppm | Co 238.892 C ppm | Cr 267.716 C ppm | Cu 324.754 C ppm | Fe 259.940 C ppm | K 766.491 C ppm | Li 670.783 C ppm | Mg 279.078 C ppm | Mn 257.610 C ppm | Mo 202.032 C ppm | Na 588.995 C ppm | Ni 231.604 C ppm | P 178.222 C ppm | Pb 220.353 C ppm | S 181.972 C ppm | Sb 217.582 C ppm | Se 196.026 C ppm | Si 251.611 C ppm | Sr 421.552 C ppm | Ti 334.941 C ppm | V 292.401 C ppm | Zn 206.200 C ppm
Calibration Blank | Blank | 2024-09-16 09:25 | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 !
Calibration Standard 0.5ppm | Standard | 2024-09-16 09:27 | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 !
Calibration Standard 1ppm | Standard | 2024-09-16 09:30 | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 !
Calibration Standard 5ppm | Standard | 2024-09-16 09:32 | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 !
Calibration Standard 10ppm | Standard | 2024-09-16 09:35 | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 !
Calibration Standard 20ppm | Standard | 2024-09-16 09:37 | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 !
Calibration Standard 50ppm | Standard | 2024-09-16 09:40 |  | 50.0000 ! |  |  | 50.0000 ! |  | 50.0000 ! |  |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! |  | 50.0000 ! |  | 50.0000 !
Calibration Standard-A 100ppm | Standard | 2024-09-16 09:42 | Al: 100.0000 !; Ca: 100.0000 !; Cu: 100.0000 !; Fe: 100.0000 !; K: 100.0000 !; Li: 100.0000 !; Mg: 100.0000 !; Na: 100.0000 !; Pb: 100.0000 !; Si: 100.0000 !; Zn: 100.0000 !
Calibration Standard-B 100ppm | Standard | 2024-09-16 09:45 | P: 100.0000 !; S: 100.0000 !
Calibration Standard Al+Fe+Ca+Mg200ppm | Standard | 2024-09-16 09:47 | Al: 200.0000 !; Ca: 200.0000 !; Fe: 200.0000 !; Mg: 200.0000 !
Calibration Standard Ca+Mg500ppm | Standard | 2024-09-16 09:50 | Ca: 500.0000 !; Mg: 500.0000 !
Calibration Standard Ca+Mg1000ppm | Standard | 2024-09-16 09:52 | Ca: 1000.0000 !; Mg: 1000.0000 !
Calibration Standard Ca+Mg2000ppm | Standard | 2024-09-16 09:55 | Ca: 2000.0000 !; Mg: 2000.0000 !
Calibration Standard P+S200ppm | Standard | 2024-09-16 09:58 | P: 200.0000 !; S: 200.0000 !
Calibration Standard Pb200ppm | Standard | 2024-09-16 10:00 | Pb: 200.0000 !
Calibration Standard Fe 1000ppm | Standard | 2024-09-16 10:03 | Fe: 1000.0000 !
Rinse1 | QC | 2024-09-16 10:23 | -0.0001 | 0.0029 | 0.0003 | 0.0028 | 0 | -0.0001 | 0.0074 | 0 | 0 | -0.0001 | 0.0003 | -0.0007 | -0.0027 | -0.0012 | 0.0136 | 0.0001 | -0.0003 | 0.0028 | -0.0012 | 0.0039 | -0.0018 | 0 | 0.1434 | 0.014 | 0.0227 | 0 | 0 | -0.0003 | 0.0003
Rinse | QC | 2024-09-16 10:26 | -0.0002 | -0.001 | 0.0017 | 0.0023 | 0 | 0 | 0.0047 | 0.0001 | 0 | 0 | 0 | 0 | -0.0035 | -0.0014 | 0.0135 | 0.0001 | 0.0006 | 0.0009 | -0.0013 | 0.0039 | -0.0028 | 0.0035 | 0.1455 | 0.0165 | 0.0216 | 0 | 0.0001 | -0.0002 | 0.0004
IECS1 (Al 1000 Int.) | QC | 2024-09-16 10:28 | 0 | 1039.4457 | 0.0001 | 0.0051 | 0.0097 | 0 | 0.0094 | 0 | -0.0002 | 0.0002 | 0.0024 | 0.0036 | -0.0048 | -0.0013 | 0.0228 | 0.0003 | 0.0031 | 0.0078 | -0.0011 | 0.0094 | 0.0259 | -0.0037 | 0.3079 | 0.005 | 0.0381 | 0.0001 | 0.0002 | 0.0004 | 0.0046
% Recovery |  |  | Al: 103.9%
Rinse1 | QC | 2024-09-16 10:31 | 0.0001 | 0.0145 | 0.002 | 0.0023 | -0.0001 | -0.0001 | 0.0048 | 0 | 0.0002 | -0.0001 | 0.0004 | 0.0007 | -0.003 | -0.0007 | 0.0116 | 0.0001 | 0.0003 | 0.0015 | -0.0013 | 0.0041 | -0.003 | -0.0004 | 0.1456 | 0.0156 | 0.0232 | 0 | 0.0001 | 0 | 0.0002
Rinse | QC | 2024-09-16 10:33 | -0.0001 | 0.0052 | 0.0005 | 0.0016 | 0.0001 | 0 | 0.004 | -0.0001 | 0.0002 | 0 | 0.0001 | 0.0001 | -0.0072 | -0.0016 | 0.0092 | 0.0001 | -0.0005 | 0.0007 | -0.0013 | 0.0011 | -0.0016 | 0.0011 | 0.1465 | 0.0135 | 0.0227 | 0 | 0.0001 | 0.0002 | 0.0005
IECS2 (Fe 1000 Int.) | QC | 2024-09-16 10:36 | -0.0003 | 0.0337 | 0.0003 | -0.0057 | -0.0001 | 0 | -0.0055 | 0.0016 | -0.0042 | 0.0045 | -0.0002 | 1008.1799 | -0.006 | -0.0014 | 1.5117 K | 0.009 | -0.0006 | 0.0017 | 0.0065 | 0.0006 | 0.0143 | 0.0065 | 0.1359 | -0.021 | 0.0163 | 0.0004 | 0 | -0.0022 | 0.0031
% Recovery |  |  | Fe: 100.8%
Rinse1 | QC | 2024-09-16 10:38 | 0.0002 | 0.0052 | 0.0005 | 0.0011 | -0.0001 | -0.0001 | 0.005 | 0 | -0.0002 | 0 | 0.0002 | 0.021 | -0.0077 | -0.0015 | 0.0119 | 0.0003 | 0 | 0.0008 | -0.0017 | 0.0011 | -0.0011 | 0.0005 | 0.149 | 0.0127 | 0.0249 | 0 | 0 | -0.0003 | 0.0003
Rinse | QC | 2024-09-16 10:41 | 0.0002 | 0.0024 | 0.001 | 0.0008 | -0.0001 | 0 | 0.0046 | 0.0001 | -0.0001 | 0 | 0.0005 | 0.0062 | 0.0017 | -0.0009 | 0.0122 | 0.0003 | 0.0004 | 0.0004 | -0.0018 | 0.002 | -0.0032 | 0.0032 | 0.1449 | 0.0104 | 0.0255 | 0 | 0 | 0 | 0.0005
IECS3 (Mn 100 Int.) | QC | 2024-09-16 10:43 | -0.0025 | 0.0031 | 0.0005 | 0.0002 | -0.0001 | 0 | 0.6990 K | 0 | -0.0028 | -0.0194 | -0.0003 | 0.0014 | -0.0031 | -0.0007 | 0.0146 | 99.8516 | -0.0003 | 0.0003 | -0.0018 | 0.0028 | -0.0033 | -0.0985 | 0.1538 | 0.0221 | 0.0259 | 0 | 0 | 0.0005 | 0.0001
% Recovery |  |  | Mn: 99.85%
IECS4 (Ca Mg 1000 Int.) | QC | 2024-09-16 10:46 | -0.0001 | 0.1027 | 0.0003 | 0.0044 | 0.002 | 0.0001 | 989.1619 | 0.0082 | 0.0002 | 0.0012 | 0.0029 | 0.0042 | 0.0027 | -0.0024 | 1023.732 | 0.0042 | 0.0007 | 0.0176 | 0.0022 | 0.0042 | 0.0047 | 0.0481 | 0.1568 | 0.0145 | 0.0547 | 0.0079 | -0.0004 | 0.0007 | 0.0119
% Recovery |  |  | Ca: 98.92%; Mg: 102.4%
Rinse1 | QC | 2024-09-16 10:48 | 0.0001 | 0.0011 | 0.0011 | 0.0004 | 0 | 0 | 0.013 | 0.0001 | 0.0003 | 0.0001 | 0 | 0.0008 | -0.0054 | -0.0016 | 0.0204 | 0.0006 | -0.0005 | 0.0032 | -0.0001 | 0.0015 | -0.0043 | 0.0036 | 0.1509 | 0.0133 | 0.026 | 0 | 0.0001 | 0.0001 | 0.0004
Rinse | QC | 2024-09-16 10:51 | 0.0001 | -0.0007 | 0.0006 | 0.0003 | 0 | -0.0001 | 0.0084 | 0 | 0.0002 | -0.0001 | 0.0003 | 0.0003 | -0.0008 | -0.0016 | 0.017 | 0.0001 | 0 | 0.0003 | -0.0005 | -0.0008 | -0.0026 | 0.003 | 0.1514 | 0.0127 | 0.0269 | 0 | 0.0001 | -0.0003 | 0.0001
IECS5 (Ti 20 Int.) | QC | 2024-09-16 10:53 | 0.0002 | 0.0021 | 0.0013 | 0.0017 | -0.0003 | -0.0003 | 0.0932 | 0 | -0.0001 | -0.0002 | -0.0007 | 0.0013 | -0.0044 | -0.0017 | 0.0181 | 0 | -0.0002 | 0.0009 | -0.0011 | -0.0005 | -0.005 | 0.001 | 0.1475 | 0.011 | 0.0664 | 0 | 20.1663 | -0.0053 | -0.0002
% Recovery |  |  | Ti: 100.8%
Rinse1 | QC | 2024-09-16 10:56 | 0.0001 | 0 | 0.0007 | 0.0004 | -0.0001 | -0.0001 | 0.0078 | 0.0001 | 0.0001 | -0.0004 | 0.0001 | 0.0007 | -0.0044 | -0.0018 | 0.0126 | 0.0001 | -0.0007 | 0.0032 | -0.002 | 0.002 | -0.0018 | 0.0003 | 0.1487 | 0.0133 | 0.0295 | 0 | 0.0045 | -0.0004 | 0.0003
Rinse | QC | 2024-09-16 10:58 | 0.0001 | -0.0004 | -0.0017 | 0.0005 | -0.0001 | 0 | 0.0045 | 0.0002 | 0.0003 | 0 | 0.0003 | 0.0006 | -0.0073 | -0.0016 | 0.0154 | 0 | -0.0003 | 0.0011 | -0.0013 | 0.0002 | -0.0032 | 0.0033 | 0.1499 | 0.0173 | 0.0284 | 0 | 0.0015 | -0.0003 | 0.0001
ICB | QC | 2024-09-16 11:01 | -0.0001 | -0.0003 | 0.0006 | 0.0003 | -0.0001 | 0 | 0.0018 | -0.0001 | 0 | 0 | 0 | 0.0004 | -0.0121 | -0.0011 | 0.0144 | 0.0003 | 0 | -0.0008 | -0.0008 | -0.0008 | -0.0025 | 0 | 0.1443 | 0.0121 | 0.0278 | 0 | 0.0008 | -0.0002 | 0.0001
ICV (SS 5ppm All) | QC | 2024-09-16 11:04 | 5.0775 | 4.972 | 5.1286 | 4.9878 | 5.0499 | 5.0092 | 5.0185 | 4.9925 | 5.0763 | 4.9767 | 4.9806 | 4.9897 | 5.0416 | 5.0421 | 4.9082 | 4.966 | 5.2977 | 5.0765 | 4.9849 | 5.0175 | 5.2863 | 4.9194 | 4.8948 | 5.04 | 4.8172 | 5.0491 | 5.0118 | 5.0597 | 4.9762
% Recovery |  |  | 101.5% | 99.44% | 102.6% | 99.76% | 101% | 100.2% | 100.4% | 99.85% | 101.5% | 99.53% | 99.61% | 99.79% | 100.8% | 100.8% | 98.16% | 99.32% | 106% | 101.5% | 99.7% | 100.4% | 105.7% | 98.39% | 97.9% | 100.8% | 96.34% | 101% | 100.2% | 101.2% | 99.52%
Rinse | QC | 2024-09-16 11:06 | -0.0001 | 0.0023 | 0.0152 | 0.0027 | 0 | 0.0001 | 0.0045 | 0.0002 | 0 | -0.0001 | 0.0002 | 0.0013 | -0.0002 | -0.0019 | 0.0119 | 0.0002 | 0.0447 | 0.0006 | -0.0018 | 0.002 | -0.0008 | 0.004 | 0.1531 | 0.0316 | 0.0308 | 0.0001 | 0.0019 | 0.0001 | 0.0005
CCB | QC | 2024-09-16 11:09 | 0.0001 | 0.0004 | 0.0035 | 0.0013 | 0 | -0.0001 | -0.0006 | 0 | -0.0001 | -0.0004 | 0.0002 | -0.0001 | -0.0051 | -0.0016 | 0.0102 | 0.0002 | 0.0124 | 0.0005 | -0.0016 | 0.0026 | -0.0007 | 0.0008 | 0.1446 | 0.0153 | 0.03 | 0 | 0.0008 | 0 | 0.0001
0.5ppm All (LLQC) | QC | 2024-09-16 11:11 | 0.4927 | 0.4909 | 0.4909 | 0.4928 | 0.5017 | 0.4915 | 0.5271 | 0.4928 | 0.5017 | 0.495 | 0.4901 | 0.4967 | 0.4768 | 0.5007 | 0.4789 | 0.4968 | 0.4952 | 0.5129 | 0.4935 | 0.493 | 0.5168 | 0.4814 | 0.4738 | 0.4904 | 0.5186 | 0.5 | 0.4889 | 0.4975 | 0.5029
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-09-16 11:17 | 0.4919 | 0.4943 | 0.4841 | 0.4933 | 0.5079 | 0.4937 | 0.5257 | 0.4921 | 0.5017 | 0.4942 | 0.4882 | 0.5023 | 0.481 | 0.5035 | 0.4779 | 0.5008 | 0.495 | 0.5172 | 0.4926 | 0.4859 | 0.5157 | 0.4794 | 0.5206 | 0.4937 | 0.519 | 0.504 | 0.4879 | 0.4967 | 0.5032
RPD |  |  | 0.1625% | 0.6902% | 1.395% | 0.1014% | 1.228% | 0.4466% | 0.266% | 0.1421% | 0% | 0.1617% | 0.3884% | 1.121% | 0.877% | 1.626% | 0.209% | 0.8019% | 0.0404% | 0.8349% | 0.1825% | 1.451% | 0.2131% | 0.4163% | 9.413% | 0.6707% | 0.0771% | 0.7968% | 0.2048% | 0.1609% | 0.05964%
0.5ppm All (LLQC) | QC | 2024-09-16 11:11 | 0.4927 | 0.4909 | 0.4909 | 0.4928 | 0.5017 | 0.4915 | 0.5271 | 0.4928 | 0.5017 | 0.495 | 0.4901 | 0.4967 | 0.4768 | 0.5007 | 0.4789 | 0.4968 | 0.4952 | 0.5129 | 0.4935 | 0.493 | 0.5168 | 0.4814 | 0.4738 | 0.4904 | 0.5186 | 0.5 | 0.4889 | 0.4975 | 0.5029
% Recovery |  |  | 98.54% | 98.18% | 98.18% | 98.56% | 100.3% | 98.3% | 105.4% | 98.56% | 100.3% | 99% | 98.02% | 99.34% | 95.36% | 100.1% | 95.78% | 99.36% | 99.04% | 102.6% | 98.7% | 98.6% | 103.4% | 96.28% | 94.76% | 98.08% | 103.7% | 100% | 97.78% | 99.5% | 100.6%
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-09-16 11:17 | 0.4919 | 0.4943 | 0.4841 | 0.4933 | 0.5079 | 0.4937 | 0.5257 | 0.4921 | 0.5017 | 0.4942 | 0.4882 | 0.5023 | 0.481 | 0.5035 | 0.4779 | 0.5008 | 0.495 | 0.5172 | 0.4926 | 0.4859 | 0.5157 | 0.4794 | 0.5206 | 0.4937 | 0.519 | 0.504 | 0.4879 | 0.4967 | 0.5032
% Recovery |  |  | 98.38% | 98.86% | 96.82% | 98.66% | 101.6% | 98.74% | 105.1% | 98.42% | 100.3% | 98.84% | 97.64% | 100.5% | 96.2% | 100.7% | 95.58% | 100.2% | 99% | 103.4% | 98.52% | 97.18% | 103.1% | 95.88% | 104.1% | 98.74% | 103.8% | 100.8% | 97.58% | 99.34% | 100.6%
CCV (5ppm All) | QC | 2024-09-16 11:20 | 4.9564 | 4.9015 | 4.9868 | 4.9646 | 5.0263 | 4.9731 | 4.9525 | 4.9397 | 5.0485 | 4.9404 | 4.9298 | 4.9523 | 4.9989 | 5.0673 | 4.7836 | 4.9289 | 5.1843 | 4.9983 | 4.9356 | 4.9243 | 5.2052 | 4.8513 | 4.9682 | 4.9647 | 4.9199 | 4.9792 | 4.9033 | 4.9818 | 4.939
% Recovery |  |  | 99.13% | 98.03% | 99.74% | 99.29% | 100.5% | 99.46% | 99.05% | 98.79% | 101% | 98.81% | 98.6% | 99.05% | 99.98% | 101.3% | 95.67% | 98.58% | 103.7% | 99.97% | 98.71% | 98.49% | 104.1% | 97.03% | 99.36% | 99.29% | 98.4% | 99.58% | 98.07% | 99.64% | 98.78%
CCB | QC | 2024-09-16 11:23 | 0 | -0.0017 | 0.0093 | 0.0025 | 0 | 0 | -0.0022 | 0.0001 | 0.0003 | -0.0001 | 0 | -0.0019 | 0.0052 | -0.002 | 0.0119 | 0.0003 | 0.0317 | 0.0009 | -0.0011 | -0.0009 | -0.004 | -0.0029 | 0.159 | 0.0246 | 0.0315 | 0 | 0.0013 | 0 | 0
Rinse | QC | 2024-09-16 14:43 | 0 | -0.0006 | 0.0011 | -0.0022 | 0 | 0 | 0.0054 | -0.0001 | -0.0003 | 0 | 0.0005 | 0.0044 | -0.0013 | -0.001 | 0.011 | 0.0005 | 0.0001 | 0.0009 | -0.0012 | 0.003 | -0.0036 | 0.0203 | 0.1431 | 0.0103 | 0.0401 | 0 | 0.0001 | -0.0002 | 0.002
CCV (5ppm All) | QC | 2024-09-16 14:46 | 4.9704 | 4.942 | 4.9863 | 4.9527 | 5.0236 | 4.9202 | 4.9103 | 4.8662 | 4.9798 | 4.913 | 4.9843 | 4.9533 | 4.9402 | 5.0763 | 4.7555 | 4.9187 | 5.1652 | 4.9884 | 4.8957 | 5.0625 | 5.1848 | 4.9282 | 4.7384 | 4.9697 | 4.942 | 4.9761 | 4.8704 | 4.977 | 4.8546
% Recovery |  |  | 99.41% | 98.84% | 99.73% | 99.05% | 100.5% | 98.4% | 98.21% | 97.32% | 99.6% | 98.26% | 99.69% | 99.07% | 98.8% | 101.5% | 95.11% | 98.37% | 103.3% | 99.77% | 97.91% | 101.2% | 103.7% | 98.56% | 94.77% | 99.39% | 98.84% | 99.52% | 97.41% | 99.54% | 97.09%
CCB | QC | 2024-09-16 14:49 | -0.0001 | 0.0022 | 0.0076 | -0.0014 | -0.0001 | 0 | -0.0007 | 0 | 0 | -0.0001 | 0 | -0.0005 | -0.0105 | -0.0022 | 0.0108 | 0 | 0.0246 | 0 | -0.001 | 0.0003 | -0.0034 | 0.0021 | 0.1465 | 0.0187 | 0.0396 | 0.0001 | 0.001 | -0.0001 | 0.0006
0.5ppm All (LLQC) | QC | 2024-09-16 14:52 | 0.4957 | 0.4911 | 0.4896 | 0.4913 | 0.5041 | 0.4921 | 0.5221 | 0.4862 | 0.501 | 0.4925 | 0.497 | 0.4995 | 0.4568 | 0.5029 | 0.4729 | 0.496 | 0.5004 | 0.514 | 0.49 | 0.508 | 0.5129 | 0.4879 | 0.4594 | 0.4944 | 0.5295 | 0.5008 | 0.4879 | 0.4982 | 0.497
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-09-16 14:54 | 0.4922 | 0.4921 | 0.4931 | 0.4885 | 0.5044 | 0.4919 | 0.519 | 0.4827 | 0.4965 | 0.4891 | 0.4923 | 0.4949 | 0.4695 | 0.5014 | 0.4737 | 0.4961 | 0.5024 | 0.515 | 0.4873 | 0.5002 | 0.5089 | 0.4868 | 0.4837 | 0.4931 | 0.5265 | 0.5014 | 0.4842 | 0.495 | 0.4932
RPD |  |  | 0.7086% | 0.2034% | 0.7123% | 0.5715% | 0.05949% | 0.04065% | 0.5955% | 0.7225% | 0.9023% | 0.6927% | 0.9502% | 0.9252% | 2.742% | 1.626% | 0.169% | 0.02016% | 0.3989% | 0.1944% | 0.5525% | 1.547% | 0.7829% | 0.2257% | 5.153% | 0.2633% | 0.5682% | 0.1197% | 0.7612% | 0.6444% | 0.7675%
0.5ppm All (LLQC) | QC | 2024-09-16 14:52 | 0.4957 | 0.4911 | 0.4896 | 0.4913 | 0.5041 | 0.4921 | 0.5221 | 0.4862 | 0.501 | 0.4925 | 0.497 | 0.4995 | 0.4568 | 0.5029 | 0.4729 | 0.496 | 0.5004 | 0.514 | 0.49 | 0.508 | 0.5129 | 0.4879 | 0.4594 | 0.4944 | 0.5295 | 0.5008 | 0.4879 | 0.4982 | 0.497
% Recovery |  |  | 99.14% | 98.22% | 97.92% | 98.26% | 100.8% | 98.42% | 104.4% | 97.24% | 100.2% | 98.5% | 99.4% | 99.9% | 91.36% | 100.6% | 94.58% | 99.2% | 100.1% | 102.8% | 98% | 101.6% | 102.6% | 97.58% | 91.88% | 98.88% | 105.9% | 100.2% | 97.58% | 99.64% | 99.4%
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-09-16 14:54 | 0.4922 | 0.4921 | 0.4931 | 0.4885 | 0.5044 | 0.4919 | 0.519 | 0.4827 | 0.4965 | 0.4891 | 0.4923 | 0.4949 | 0.4695 | 0.5014 | 0.4737 | 0.4961 | 0.5024 | 0.515 | 0.4873 | 0.5002 | 0.5089 | 0.4868 | 0.4837 | 0.4931 | 0.5265 | 0.5014 | 0.4842 | 0.495 | 0.4932
% Recovery |  |  | 98.44% | 98.42% | 98.62% | 97.7% | 100.9% | 98.38% | 103.8% | 96.54% | 99.3% | 97.82% | 98.46% | 98.98% | 93.9% | 100.3% | 94.74% | 99.22% | 100.5% | 103% | 97.46% | 100% | 101.8% | 97.36% | 96.74% | 98.62% | 105.3% | 100.3% | 96.84% | 99% | 98.64%
CCS1 (SS 5ppm) | QC | 2024-09-16 14:57 | 5.0583 | 5.0199 | 5.0967 | 4.9366 | 5.0908 | 5.0007 | 4.9225 | 4.8641 | 4.9749 | 4.9013 | 5.01 | 5.0044 | 5.0233 | 5.0326 | 4.8132 | 5.0134 | 5.2946 | 5.06 | 4.8958 | 5.1232 | 5.2159 | 4.9565 | 4.7614 | 5.0131 | 4.8179 | 5.0737 | 4.9428 | 5.0139 | 4.8218
% Recovery |  |  | 101.2% | 100.4% | 101.9% | 98.73% | 101.8% | 100% | 98.45% | 97.28% | 99.5% | 98.03% | 100.2% | 100.1% | 100.5% | 100.7% | 96.26% | 100.3% | 105.9% | 101.2% | 97.92% | 102.5% | 104.3% | 99.13% | 95.23% | 100.3% | 96.36% | 101.5% | 98.86% | 100.3% | 96.44%
CCB | QC | 2024-09-16 14:59 | -0.0002 | 0.0009 | 0.0153 | -0.0005 | 0 | 0.0001 | -0.0008 | 0.0001 | 0.0001 | 0 | -0.0001 | 0.0013 | -0.0026 | -0.0017 | 0.0098 | 0.0003 | 0.0449 | 0.001 | -0.0018 | 0.0001 | -0.0044 | 0.003 | 0.1442 | 0.0248 | 0.0399 | 0.0001 | 0.0016 | -0.0004 | 0.0001
MQ-Blank-02 | Sample | 2024-09-16 15:02 | 0 | -0.0016 | 0.0028 | -0.002 | -0.0001 | -0.0001 | 0.0011 | -0.0001 | -0.0002 | -0.0003 | -0.0001 | -0.0003 | -0.0034 | -0.0019 | 0.0088 | 0.0002 | 0.0026 | 0.0017 | -0.0016 | -0.0003 | -0.0027 | 0 | 0.1378 | 0.0145 | 0.0397 | 0 | 0 | 0.0001 | -0.0003
Rinse | QC | 2024-09-16 15:04 | -0.0003 | 0.0008 | 0.0021 | -0.0015 | -0.0001 | -0.0001 | 0.0025 | 0 | -0.0006 | -0.0001 | 0.0001 | 0.0006 | -0.0065 | -0.0023 | 0.0107 | 0.0001 | 0.0075 | 0.0011 | -0.0017 | -0.0013 | -0.0023 | -0.0017 | 0.1299 | 0.018 | 0.0402 | 0.0001 | 0.0004 | 0 | 0.0004
CCV (5ppm All) | QC | 2024-09-16 15:07 | 5.0249 | 4.9705 | 5.0374 | 5.0051 | 5.0447 | 4.9725 | 4.9481 | 4.8925 | 5.0257 | 4.9467 | 5.0461 | 4.9693 | 4.9999 | 5.0903 | 4.7898 | 4.9404 | 5.2175 | 5.012 | 4.9324 | 5.1309 | 5.2282 | 4.9765 | 4.6955 | 5.0156 | 4.9991 | 5.0029 | 4.92 | 5.0155 | 4.8589
% Recovery |  |  | 100.5% | 99.41% | 100.7% | 100.1% | 100.9% | 99.45% | 98.96% | 97.85% | 100.5% | 98.93% | 100.9% | 99.39% | 100% | 101.8% | 95.8% | 98.81% | 104.4% | 100.2% | 98.65% | 102.6% | 104.6% | 99.53% | 93.91% | 100.3% | 99.98% | 100.1% | 98.4% | 100.3% | 97.18%
CCB | QC | 2024-09-16 15:09 | -0.0001 | 0.0004 | 0.0116 | -0.0001 | -0.0001 | -0.0001 | 0.0003 | 0 | 0 | -0.0001 | 0.0001 | -0.0003 | -0.0008 | -0.0027 | 0.0091 | 0.0002 | 0.0447 | 0.0016 | -0.0009 | 0.0002 | -0.0049 | 0.0014 | 0.1595 | 0.0289 | 0.0399 | 0.0001 | 0.0018 | 0 | -0.0002
10x 0.5% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:12 | -0.0002 | 16.7305 | 0.0047 | 0.0161 | 0.0027 | -0.0001 | 0.2078 | -0.0006 | -0.0008 | 0.0305 | 0.002 | 165.192 | -0.0097 | 0.185 | 0.8543 | 1.1406 | 0.0135 | 5.1539 | 0.6358 | 1.1378 | -0.0013 | 62.0289 | 0.112 | 0.0028 | 0.1715 | 0.0005 | 0.099 | 0.0081 | 0.6053
10x 1% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:14 | -0.0003 | 27.958 | 0.0007 | 0.0364 | 0.0028 | -0.0002 | 0.365 | -0.0008 | -0.0009 | 0.0487 | 0.0036 | 232.2671 | 0.0034 | 0.2207 | 1.4973 | 1.9398 | 0.0086 | 14.4739 | 0.9754 | 1.4515 | 0.0011 | 120.8814 | 0.1134 | -0.0035 | 0.2665 | 0.0008 | 0.1934 | 0.0151 | 0.9293
10x 5% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:17 | -0.0004 | 12.8167 | 0.0022 | 0.1787 | 0.003 | -0.0002 | 1.8491 | 0.0003 | 0.0117 | 0.1818 | 0.009 | 942.0008 | 0.0787 | 0.2548 | 7.4451 | 9.3401 | 0.0109 | 33.0439 | 1.7583 | 4.5448 | 0.0171 | 610.9164 | 0.1236 | -0.0188 | 0.5968 | 0.0037 | 0.9988 | 0.052 | 5.1065
10x 10% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:19 | -0.0001 | 1.1084 | 0.0044 | 0.3716 | 0.0049 | -0.0001 | 3.46 | 0.0006 | -0.0201 | 0.3857 | 0.0269 | 1803.3566 | 0.1637 | -0.0001 | 14.3827 | 17.8596 | 0.0145 | 64.8239 | 0.8456 | 0.1584 | 0.0536 | 1161.3306 | 0.1428 | -0.0442 | 0.7729 | 0.0067 | 2.032 | 0.1281 | 12.3307
10x 15% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:22 | 0.0008 | 1.6543 | 0.0024 | 0.5727 | 0.0069 | -0.0001 | 5.0803 | 0.0044 | -0.0217 | 0.574 | 0.0429 | 2533.4012 | 0.3017 | 0.0193 | 21.0202 | 26.9829 | 0.0205 | 97.0954 | 1.0375 | 0.241 | 0.0866 | 1657.7279 | 0.1884 | -0.0706 | 1.1423 | 0.0106 | 2.9999 | 0.2017 | 18.3706
10x 0.5% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:24 | -0.0004 | 9.2607 | 0.0001 | 0.0195 | 0.0008 | -0.0002 | 0.2165 | -0.0007 | -0.0005 | 0.0335 | 0.0048 | 166.8406 | -0.0054 | 0.0205 | 0.8947 | 0.9895 | 0.0023 | 3.9598 | 1.2588 | 0.4528 | -0.0015 | 64.8784 | 0.1198 | -0.0011 | 0.0991 | 0.0005 | 0.1114 | 0.0075 | 0.4791
10x 1% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:27 | -0.0005 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | -0.0005 | 0.003 | 0.0598 | 0.0075 | 252.3585 | -0.005 | 0.0271 | 1.6867 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.34 | -0.0011 | 126.2896 | 0.1232 | -0.0045 | 0.1505 | 0.0008 | 0.2255 | 0.0103 | 0.8442
10x 5% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:29 | -0.0007 | 10.346 | 0.0036 | 0.2226 | 0.0016 | -0.0002 | 1.9549 | -0.0004 | 0.0045 | 0.2301 | 0.0089 | 995.943 | 0.0957 | 0.0909 | 7.9403 | 9.5788 | 0.0086 | 33.5317 | 2.5013 | 0.4545 | 0.0218 | 641.9245 | 0.1276 | -0.0256 | 0.469 | 0.0038 | 1.1323 | 0.0654 | 4.2772
10x 10% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:32 | -0.0005 | 1.1497 | 0.0007 | 0.4192 | 0.0033 | -0.0002 | 3.8504 | 0.0052 | -0.0184 | 0.4303 | 0.0286 | 1867.6644 | 0.246 | -0.0007 | 14.6679 | 18.5866 | 0.0157 | 62.795 | 2.4582 | 0.4111 | 0.0513 | 1190.7709 | 0.1454 | -0.0374 | 0.8201 | 0.0073 | 2.1431 | 0.1329 | 9.1637
10x 15% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:34 | 0.0002 | 2.6273 | 0.0027 | 0.6048 | 0.0045 | -0.0001 | 5.4141 | 0.0036 | -0.024 | 0.6136 | 0.0459 | 2550.4615 | 0.3147 | 0.0679 | 20.89 | 27.0173 | 0.021 | 91.9079 | 2.762 | 0.504 | 0.0881 | 1668.2033 | 0.1881 | -0.076 | 1.163 | 0.0105 | 3.0669 | 0.1928 | 14.6004
Rinse | QC | 2024-09-16 15:37 | -0.0002 | 0.0007 | 0.0001 | -0.002 | -0.0001 | 0 | 0.0036 | 0 | -0.0002 | -0.0003 | -0.0003 | 0.0573 | -0.0003 | -0.0023 | 0.0078 | 0.0008 | 0.0003 | 0.0014 | -0.0017 | 0.0012 | -0.0048 | 0.06 | 0.1348 | 0.0109 | 0.0388 | 0 | 0.0004 | 0.0001 | 0.0004
CCV (5ppm All) | QC | 2024-09-16 15:39 | 4.9966 | 4.9582 | 5.0108 | 4.9704 | 5.0439 | 4.9613 | 4.9149 | 4.8515 | 4.9861 | 4.9138 | 5.0218 | 4.9991 | 4.9974 | 5.0844 | 4.7607 | 4.9498 | 5.1802 | 5.0091 | 4.8985 | 5.1151 | 5.1963 | 4.9881 | 4.6549 | 4.9943 | 4.9613 | 5.0072 | 4.8773 | 4.9782 | 4.8263
% Recovery |  |  | 99.93% | 99.16% | 100.2% | 99.41% | 100.9% | 99.23% | 98.3% | 97.03% | 99.72% | 98.28% | 100.4% | 99.98% | 99.95% | 101.7% | 95.21% | 99% | 103.6% | 100.2% | 97.97% | 102.3% | 103.9% | 99.76% | 93.1% | 99.89% | 99.23% | 100.1% | 97.55% | 99.56% | 96.53%
CCB | QC | 2024-09-16 15:42 | -0.0003 | 0.0012 | 0.0113 | -0.0006 | 0 | -0.0001 | -0.002 | 0.0001 | -0.0002 | 0 | -0.0002 | 0.0115 | 0.0061 | -0.0027 | 0.0087 | 0.0003 | 0.0396 | -0.0005 | -0.0017 | 0.0021 | -0.0049 | 0.0214 | 0.1544 | 0.0297 | 0.0408 | 0.0001 | 0.0021 | 0.0002 | 0.0002
09032024 MB | Sample | 2024-09-16 15:44 | 0.0001 | 0.0065 | 0.0046 | -0.0008 | 0.0003 | 0 | 0.0054 | 0 | 0 | -0.0003 | 0.0001 | 0.0112 | -0.0022 | -0.0025 | 0.0165 | 0.0002 | 0.0108 | 0.0161 | -0.0001 | 0.007 | -0.0047 | 0.0203 | 0.1349 | 0.0211 | 0.0511 | 0 | 0.0008 | 0.0002 | 0.0026
10x 09032024 MB | Sample | 2024-09-16 15:47 | -0.0001 | 0.0021 | 0.0012 | -0.0022 | -0.0001 | -0.0001 | 0.0692 | 0 | -0.0002 | 0.0001 | 0.0095 | 0.0095 | 0.0043 | -0.0024 | 0.0092 | 0.0008 | 0.0038 | 0.0017 | -0.0019 | 0.0041 | -0.0038 | 0.0121 | 0.1222 | 0.014 | 0.0395 | 0.0001 | 0.0004 | -0.0002 | 0.0043
10x 09032024 MFB | Sample | 2024-09-16 15:50 | 0.01 | 0.4842 | 0.2424 | 0.0482 | 0.2575 | 0.0488 | 0.5556 | 0.0972 | 0.5035 | 0.2469 | 0.2494 | 0.5049 | 0.4666 | -0.0026 | 0.481 | 0.1007 | 0.0969 | 0.5096 | 0.2426 | 0.1056 | 0.2554 | 0.2545 | 0.2459 | 0.2492 | 0.1389 | 0.1021 | 0.0984 | 0.2517 | 0.0959
R% |  |  | 101% | 96.42% | 96.48% | 100.8% | 103% | 97.8% | 97.28% | 97.2% | 100.7% | 98.72% | 95.96% | 99.08% | 92.46% |  | 94.36% | 99.9% | 93.1% | 101.6% | 97.8% | 101.5% | 103.7% | 96.96% | 49.48% | 94.08% | 99.4% | 102% | 98% | 100.8% | 91.6%
10x 1% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:27 | -0.0005 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | -0.0005 | 0.003 | 0.0598 | 0.0075 | 252.3585 | -0.005 | 0.0271 | 1.6867 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.34 | -0.0011 | 126.2896 | 0.1232 | -0.0045 | 0.1505 | 0.0008 | 0.2255 | 0.0103 | 0.8442
10x 1% FeSO4 NMC Battery-TM -MS -Analytical Duplicate | Sample | 2024-09-16 15:52 | -0.0006 | 9.926 | 0.0023 | 0.0414 | 0.0005 | -0.0001 | 0.4151 | -0.0006 | 0.001 | 0.0595 | 0.0074 | 250.9591 | 0.0009 | 0.0273 | 1.6778 | 1.866 | 0.0054 | 7.6281 | 1.3942 | 0.339 | -0.0019 | 125.1867 | 0.107 | 0.0002 | 0.1468 | 0.0008 | 0.2242 | 0.0141 | 0.837
RPD |  |  | 18.18% | 0.4293% | 51.61% | 0.9709% | 0% | 0% | 0.3378% | 18.18% | 100% | 0.5029% | 1.342% | 0.5561% | 287.8% | 1.626% | 0.5291% | 0.3851% | 51.16% | 1.217% | 0.4937% | 0.2946% | 53.33% | 0.8771% | 14.07% | 218.6% | 2.489% | 0% | 0.5782% | 31.15% | 0.8565%
10x 1% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:27 | -0.0005 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | -0.0005 | 0.003 | 0.0598 | 0.0075 | 252.3585 | -0.005 | 0.0271 | 1.6867 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.34 | -0.0011 | 126.2896 | 0.1232 | -0.0045 | 0.1505 | 0.0008 | 0.2255 | 0.0103 | 0.8442
10x 1% FeSO4 NMC Battery-TM - Lab Duplicate | Sample | 2024-09-16 15:55 | -0.0005 | 10.1338 | 0.0009 | 0.0418 | 0.0005 | -0.0001 | 0.4152 | -0.0004 | 0.0019 | 0.0602 | 0.007 | 254.5683 | 0.0132 | 0.0276 | 1.702 | 1.8836 | 0.0041 | 7.7345 | 1.4203 | 0.3435 | 0.0011 | 127.3304 | 0.1219 | -0.0028 | 0.1507 | 0.0008 | 0.2256 | 0.0104 | 0.8588
RPD |  |  | 0% | 1.643% | 125% | 1.932% | 0% | 0% | 0.3619% | 22.22% | 44.9% | 0.6667% | 6.897% | 0.8718% | 443.9% | 1.626% | 0.903% | 0.5537% | 24.66% | 0.1682% | 1.361% | 1.024% | #DIV/0! | 0.8208% | 1.061% | 46.58% | 0.1328% | 0% | 0.04434% | 0.9662% | 1.715%
10x 1% FeSO4- NMC Battery-TM - MS | Sample | 2024-09-16 15:57 | 0.0086 | 9.0598 | 0.2474 | 0.0856 | 0.2559 | 0.0501 | 0.8491 | 0.097 | 0.4421 | 0.2973 | 0.264 | 226.1028 | 0.5031 | 0.0246 | 1.9952 | 1.7875 | 0.0984 | 7.4719 | 1.4852 | 0.4004 | 0.2582 | 113.0942 | 0.2897 | 0.2235 | 0.2466 | 0.1025 | 0.2923 | 0.2657 | 0.8221
10x 1% FeSO4- NMC Battery-TM - MS - Analytical Duplicate | Sample | 2024-09-16 16:00 | 0.0088 | 9.0013 | 0.2404 | 0.0851 | 0.2538 | 0.0495 | 0.8404 | 0.0964 | 0.4283 | 0.2947 | 0.2623 | 223.522 | 0.5033 | 0.0244 | 1.9777 | 1.7773 | 0.0995 | 7.4075 | 1.4732 | 0.4017 | 0.2545 | 112.3137 | 0.2914 | 0.2225 | 0.2436 | 0.1017 | 0.2906 | 0.2634 | 0.8162
RPD |  |  | 2.299% | 0.6478% | 2.87% | 0.5858% | 0.824% | 1.205% | 1.03% | 0.6205% | 3.171% | 0.8784% | 0.646% | 1.148% | 0.03975% | 1.626% | 0.881% | 0.5723% | 1.112% | 0.8656% | 0.8112% | 0.3241% | 1.443% | 0.6925% | 0.5851% | 0.4484% | 1.224% | 0.7835% | 0.5833% | 0.8694% | 0.7203%
10x 1% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:27 | -0.0005 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | -0.0005 | 0.003 | 0.0598 | 0.0075 | 252.3585 | -0.005 | 0.0271 | 1.6867 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.34 | -0.0011 | 126.2896 | 0.1232 | -0.0045 | 0.1505 | 0.0008 | 0.2255 | 0.0103 | 0.8442
10x 1% FeSO4- NMC Battery-TM - MS | Sample | 2024-09-16 15:57 | 0.0086 | 9.0598 | 0.2474 | 0.0856 | 0.2559 | 0.0501 | 0.8491 | 0.097 | 0.4421 | 0.2973 | 0.264 | 226.1028 | 0.5031 | 0.0246 | 1.9952 | 1.7875 | 0.0984 | 7.4719 | 1.4852 | 0.4004 | 0.2582 | 113.0942 | 0.2897 | 0.2235 | 0.2466 | 0.1025 | 0.2923 | 0.2657 | 0.8221
R% |  |  | 91% | -181.8% | 97.4% | 89.2% | 102.2% | 100.4% | 87.08% | 97.5% | 87.82% | 95% | 102.6% | -5251% | 101.6% |  | 61.7% | -85.7% | 95.2% | -49.92% | 33.64% | 60.4% | 103.7% | -5278% | 66.6% | 91.2% | 96.1% | 101.7% | 66.8% | 102.2% | -22.1%
10x 1% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:27 | -0.0005 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | -0.0005 | 0.003 | 0.0598 | 0.0075 | 252.3585 | -0.005 | 0.0271 | 1.6867 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.34 | -0.0011 | 126.2896 | 0.1232 | -0.0045 | 0.1505 | 0.0008 | 0.2255 | 0.0103 | 0.8442
10x 1% FeSO4- NMC Battery-TM - MS - Analytical Duplicate | Sample | 2024-09-16 16:00 | 0.0088 | 9.0013 | 0.2404 | 0.0851 | 0.2538 | 0.0495 | 0.8404 | 0.0964 | 0.4283 | 0.2947 | 0.2623 | 223.522 | 0.5033 | 0.0244 | 1.9777 | 1.7773 | 0.0995 | 7.4075 | 1.4732 | 0.4017 | 0.2545 | 112.3137 | 0.2914 | 0.2225 | 0.2436 | 0.1017 | 0.2906 | 0.2634 | 0.8162
R% |  |  | 93% | -193.5% | 94.6% | 88.2% | 101.3% | 99.2% | 85.34% | 96.9% | 85.06% | 93.96% | 101.9% | -5767% | 101.7% |  | 58.2% | -95.9% | 96.3% | -62.8% | 28.84% | 61.7% | 102.2% | -5590% | 67.28% | 90.8% | 93.1% | 100.9% | 65.1% | 101.2% | -28%
Rinse | QC | 2024-09-16 16:02 | -0.0001 | -0.0005 | -0.0005 | -0.0027 | 0 | -0.0001 | 0.0025 | -0.0001 | -0.0001 | -0.0002 | 0.0001 | 0.0083 | -0.0045 | -0.0025 | 0.0091 | 0.0004 | 0.0023 | 0.0005 | -0.0009 | 0.0021 | -0.0031 | 0.011 | 0.1363 | 0.0118 | 0.0397 | 0 | 0.0001 | -0.0002 | -0.0002
CCV (5ppm All) | QC | 2024-09-16 16:05 | 4.9856 | 4.9762 | 4.993 | 4.9644 | 5.066 | 4.9907 | 4.9075 | 4.8468 | 4.9857 | 4.9056 | 5.0109 | 5.0163 | 5.0064 | 5.0822 | 4.7434 | 4.9688 | 5.1709 | 5.0046 | 4.8942 | 5.0949 | 5.1957 | 4.9602 | 4.6236 | 4.9812 | 4.9395 | 5.0144 | 4.8748 | 4.9748 | 4.818
% Recovery |  |  | 99.71% | 99.52% | 99.86% | 99.29% | 101.3% | 99.81% | 98.15% | 96.94% | 99.71% | 98.11% | 100.2% | 100.3% | 100.1% | 101.6% | 94.87% | 99.38% | 103.4% | 100.1% | 97.88% | 101.9% | 103.9% | 99.2% | 92.47% | 99.62% | 98.79% | 100.3% | 97.5% | 99.5% | 96.36%
CCB | QC | 2024-09-16 16:07 | -0.0003 | 0 | 0.0141 | -0.0005 | -0.0001 | 0.0002 | -0.0022 | 0 | -0.0002 | -0.0002 | 0 | 0.0025 | -0.0046 | -0.0025 | 0.0091 | 0 | 0.0441 | -0.0004 | -0.0012 | -0.0001 | -0.0041 | 0.0084 | 0.1721 | 0.0292 | 0.0406 | 0.0001 | 0.0017 | 0.0001 | 0
09162024-Lab Blank2 | Sample | 2024-09-16 16:10 | -0.0002 | 0.0002 | 0.0028 | -0.002 | -0.0002 | -0.0001 | -0.0004 | 0 | 0 | -0.0001 | -0.0001 | -0.0006 | -0.0006 | -0.0028 | 0.0089 | 0.0003 | 0.0099 | 0.0006 | -0.0008 | 0.0029 | -0.0038 | 0.0026 | 0.1404 | 0.0175 | 0.0405 | 0 | 0.0005 | 0.0001 | -0.0003
10x 09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:12 | 0.0096 | 0.4748 | 0.2451 | 0.0481 | 0.2503 | 0.0477 | 0.4792 | 0.0958 | 0.4968 | 0.2431 | 0.2458 | 0.4895 | 0.4567 | -0.0025 | 0.4718 | 0.099 | 0.0977 | 0.4998 | 0.2393 | 0.1027 | 0.2517 | 0.249 | 0.2292 | 0.2467 | 0.138 | 0.0992 | 0.0975 | 0.2482 | 0.0947
10x 09162024-Lab Blank Spike2-Analytical Duplicate | Sample | 2024-09-16 16:15 | 0.0094 | 0.4731 | 0.2451 | 0.0479 | 0.2495 | 0.0479 | 0.4755 | 0.095 | 0.4928 | 0.2416 | 0.244 | 0.4898 | 0.4622 | -0.0025 | 0.4708 | 0.0983 | 0.0975 | 0.4962 | 0.239 | 0.1038 | 0.2495 | 0.2474 | 0.2699 | 0.2515 | 0.1385 | 0.099 | 0.0971 | 0.2466 | 0.0941
RPD |  |  | 2.105% | 0.3587% | 0% | 0.4167% | 0.3201% | 0.4184% | 0.7751% | 0.8386% | 0.8084% | 0.6189% | 0.735% | 0.06127% | 1.197% | 1.626% | 0.2122% | 0.7096% | 0.2049% | 0.7229% | 0.1254% | 1.065% | 0.8779% | 0.6446% | 16.31% | 1.927% | 0.3617% | 0.2018% | 0.4111% | 0.6467% | 0.6356%
09162024-Lab Blank2 | Sample | 2024-09-16 16:10 | -0.0002 | 0.0002 | 0.0028 | -0.002 | -0.0002 | -0.0001 | -0.0004 | 0 | 0 | -0.0001 | -0.0001 | -0.0006 | -0.0006 | -0.0028 | 0.0089 | 0.0003 | 0.0099 | 0.0006 | -0.0008 | 0.0029 | -0.0038 | 0.0026 | 0.1404 | 0.0175 | 0.0405 | 0 | 0.0005 | 0.0001 | -0.0003
10x 09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:12 | 0.0096 | 0.4748 | 0.2451 | 0.0481 | 0.2503 | 0.0477 | 0.4792 | 0.0958 | 0.4968 | 0.2431 | 0.2458 | 0.4895 | 0.4567 | -0.0025 | 0.4718 | 0.099 | 0.0977 | 0.4998 | 0.2393 | 0.1027 | 0.2517 | 0.249 | 0.2292 | 0.2467 | 0.138 | 0.0992 | 0.0975 | 0.2482 | 0.0947
R% |  |  | 98% | 94.92% | 96.92% | 100.2% | 100.2% | 95.6% | 95.92% | 95.8% | 99.36% | 97.28% | 98.36% | 98.02% | 91.46% |  | 92.58% | 98.7% | 87.8% | 99.84% | 96.04% | 99.8% | 102.2% | 98.56% | 35.52% | 91.68% | 97.5% | 99.2% | 97% | 99.24% | 95%
09162024-Lab Blank2 | Sample | 2024-09-16 16:10 | -0.0002 | 0.0002 | 0.0028 | -0.002 | -0.0002 | -0.0001 | -0.0004 | 0 | 0 | -0.0001 | -0.0001 | -0.0006 | -0.0006 | -0.0028 | 0.0089 | 0.0003 | 0.0099 | 0.0006 | -0.0008 | 0.0029 | -0.0038 | 0.0026 | 0.1404 | 0.0175 | 0.0405 | 0 | 0.0005 | 0.0001 | -0.0003
10x 09162024-Lab Blank Spike2-Analytical Duplicate | Sample | 2024-09-16 16:15 | 0.0094 | 0.4731 | 0.2451 | 0.0479 | 0.2495 | 0.0479 | 0.4755 | 0.095 | 0.4928 | 0.2416 | 0.244 | 0.4898 | 0.4622 | -0.0025 | 0.4708 | 0.0983 | 0.0975 | 0.4962 | 0.239 | 0.1038 | 0.2495 | 0.2474 | 0.2699 | 0.2515 | 0.1385 | 0.099 | 0.0971 | 0.2466 | 0.0941
R% |  |  | 96% | 94.58% | 96.92% | 99.8% | 99.88% | 96% | 95.18% | 95% | 98.56% | 96.68% | 97.64% | 98.08% | 92.56% |  | 92.38% | 98% | 87.6% | 99.12% | 95.92% | 100.9% | 101.3% | 97.92% | 51.8% | 93.6% | 98% | 99% | 96.6% | 98.6% | 94.4%
09162024-Lab Blank2 | Sample | 2024-09-16 16:10 | -0.0002 | 0.0002 | 0.0028 | -0.002 | -0.0002 | -0.0001 | -0.0004 | 0 | 0 | -0.0001 | -0.0001 | -0.0006 | -0.0006 | -0.0028 | 0.0089 | 0.0003 | 0.0099 | 0.0006 | -0.0008 | 0.0029 | -0.0038 | 0.0026 | 0.1404 | 0.0175 | 0.0405 | 0 | 0.0005 | 0.0001 | -0.0003
5x 09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:17 | 0.0191 | 0.9536 | 0.4833 | 0.0956 | 0.4984 | 0.0957 | 0.9469 | 0.1902 | 0.9818 | 0.4803 | 0.4857 | 0.9772 | 0.9523 | -0.004 | 0.9318 | 0.1966 | 0.1898 | 0.995 | 0.4745 | 0.2007 | 0.4971 | 0.4885 | 0.4249 | 0.4916 | 0.2326 | 0.1967 | 0.1917 | 0.4906 | 0.1871
R% |  |  | 96.5% | 95.34% | 96.1% | 97.6% | 99.72% | 95.8% | 94.73% | 95.1% | 98.18% | 96.08% | 97.16% | 97.78% | 95.29% |  | 92.29% | 98.15% | 89.95% | 99.44% | 95.06% | 98.9% | 100.2% | 97.18% | 56.9% | 94.82% | 96.05% | 98.35% | 95.6% | 98.1% | 93.7%
09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:20 | 0.0982 | 4.8272 | 2.5113 | 0.496 | 2.5174 | 0.4893 | 4.835 | 0.9661 | 4.9875 | 2.4259 | 2.4646 | 4.9443 | 4.9978 | -0.0021 | 4.793 | 0.9926 | 0.9949 | 5.0132 | 2.4267 | 1.0326 | 2.5598 | 2.5024 | 2.2541 | 2.5109 | 1.0178 | 0.9985 | 0.9771 | 2.4875 | 0.9598
09162024-Lab Blank Spike2-Analytical Duplicate | Sample | 2024-09-16 16:22 | 0.0983 | 4.7796 | 2.53 | 0.496 | 2.5 | 0.4859 | 4.8357 | 0.9665 | 4.9703 | 2.4279 | 2.4548 | 4.9051 | 4.9558 | -0.0021 | 4.7911 | 0.9789 | 1.0113 | 4.9854 | 2.4277 | 1.0324 | 2.5647 | 2.5009 | 2.295 | 2.5114 | 1.0228 | 0.9917 | 0.9779 | 2.4727 | 0.9611
RPD |  |  | 0.1018% | 0.991% | 0.7419% | 0% | 0.6936% | 0.6973% | 0.01448% | 0.0414% | 0.3455% | 0.08241% | 0.3984% | 0.796% | 0.8439% | 1.626% | 0.03965% | 1.39% | 1.635% | 0.5561% | 0.0412% | 0.01937% | 0.1912% | 0.05996% | 1.798% | 0.01991% | 0.4901% | 0.6833% | 0.08184% | 0.5968% | 0.1354%
09162024-Lab Blank2 | Sample | 2024-09-16 16:10 | -0.0002 | 0.0002 | 0.0028 | -0.002 | -0.0002 | -0.0001 | -0.0004 | 0 | 0 | -0.0001 | -0.0001 | -0.0006 | -0.0006 | -0.0028 | 0.0089 | 0.0003 | 0.0099 | 0.0006 | -0.0008 | 0.0029 | -0.0038 | 0.0026 | 0.1404 | 0.0175 | 0.0405 | 0 | 0.0005 | 0.0001 | -0.0003
09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:20 | 0.0982 | 4.8272 | 2.5113 | 0.496 | 2.5174 | 0.4893 | 4.835 | 0.9661 | 4.9875 | 2.4259 | 2.4646 | 4.9443 | 4.9978 | -0.0021 | 4.793 | 0.9926 | 0.9949 | 5.0132 | 2.4267 | 1.0326 | 2.5598 | 2.5024 | 2.2541 | 2.5109 | 1.0178 | 0.9985 | 0.9771 | 2.4875 | 0.9598
R% |  |  | 98.4% | 96.54% | 100.3% | 99.6% | 100.7% | 97.88% | 96.71% | 96.61% | 99.75% | 97.04% | 98.59% | 98.9% | 99.97% |  | 95.68% | 99.23% | 98.5% | 100.3% | 97.1% | 103.3% | 102.5% | 99.99% | 84.55% | 99.74% | 101.8% | 99.85% | 97.66% | 99.5% | 96.01%
09162024-Lab Blank2 | Sample | 2024-09-16 16:10 | -0.0002 | 0.0002 | 0.0028 | -0.002 | -0.0002 | -0.0001 | -0.0004 | 0 | 0 | -0.0001 | -0.0001 | -0.0006 | -0.0006 | -0.0028 | 0.0089 | 0.0003 | 0.0099 | 0.0006 | -0.0008 | 0.0029 | -0.0038 | 0.0026 | 0.1404 | 0.0175 | 0.0405 | 0 | 0.0005 | 0.0001 | -0.0003
09162024-Lab Blank Spike2-Analytical Duplicate | Sample | 2024-09-16 16:22 | 0.0983 | 4.7796 | 2.53 | 0.496 | 2.5 | 0.4859 | 4.8357 | 0.9665 | 4.9703 | 2.4279 | 2.4548 | 4.9051 | 4.9558 | -0.0021 | 4.7911 | 0.9789 | 1.0113 | 4.9854 | 2.4277 | 1.0324 | 2.5647 | 2.5009 | 2.295 | 2.5114 | 1.0228 | 0.9917 | 0.9779 | 2.4727 | 0.9611
R% |  |  | 98.5% | 95.59% | 101.1% | 99.6% | 100% | 97.2% | 96.72% | 96.65% | 99.41% | 97.12% | 98.2% | 98.11% | 99.13% |  | 95.64% | 97.86% | 100.1% | 99.7% | 97.14% | 103.2% | 102.7% | 99.93% | 86.18% | 99.76% | 102.3% | 99.17% | 97.74% | 98.9% | 96.14%
Rinse | QC | 2024-09-16 16:25 | -0.0002 | -0.0015 | 0.0134 | -0.002 | 0.0001 | 0 | 0.0036 | 0 | 0.0002 | -0.0002 | 0 | 0.0002 | -0.0098 | -0.0024 | 0.0111 | 0.0002 | 0.0246 | -0.0001 | -0.0017 | 0.0007 | -0.0013 | 0.0072 | 0.1759 | 0.0233 | 0.0404 | 0.0001 | 0.0005 | 0.0001 | 0.0004
0.5ppm All (LLQC) | QC | 2024-09-16 16:27 | 0.4942 | 0.4939 | 0.4907 | 0.4903 | 0.5025 | 0.4924 | 0.5195 | 0.4821 | 0.4982 | 0.4896 | 0.4962 | 0.4955 | 0.4766 | 0.505 | 0.4721 | 0.4941 | 0.4989 | 0.5167 | 0.4874 | 0.5087 | 0.5118 | 0.4843 | 0.4226 | 0.4922 | 0.5305 | 0.4991 | 0.4848 | 0.4956 | 0.4915
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-09-16 16:30 | 0.4927 | 0.4921 | 0.4896 | 0.4892 | 0.5029 | 0.4909 | 0.5145 | 0.4803 | 0.4961 | 0.488 | 0.4955 | 0.493 | 0.4742 | 0.5003 | 0.4707 | 0.4976 | 0.5035 | 0.5128 | 0.4864 | 0.5046 | 0.5095 | 0.4885 | 0.4436 | 0.4985 | 0.5269 | 0.4996 | 0.4834 | 0.494 | 0.4898
RPD |  |  | 0.304% | 0.3651% | 0.2244% | 0.2246% | 0.07957% | 0.3051% | 0.9671% | 0.3741% | 0.4224% | 0.3273% | 0.1412% | 0.5058% | 0.5048% | 1.626% | 0.297% | 0.7059% | 0.9178% | 0.7576% | 0.2054% | 0.8092% | 0.4504% | 0.8635% | 4.849% | 1.272% | 0.6809% | 0.1001% | 0.2892% | 0.3234% | 0.3465%
0.5ppm All (LLQC) | QC | 2024-09-16 16:27 | 0.4942 | 0.4939 | 0.4907 | 0.4903 | 0.5025 | 0.4924 | 0.5195 | 0.4821 | 0.4982 | 0.4896 | 0.4962 | 0.4955 | 0.4766 | 0.505 | 0.4721 | 0.4941 | 0.4989 | 0.5167 | 0.4874 | 0.5087 | 0.5118 | 0.4843 | 0.4226 | 0.4922 | 0.5305 | 0.4991 | 0.4848 | 0.4956 | 0.4915
% Recovery |  |  | 98.84% | 98.78% | 98.14% | 98.06% | 100.5% | 98.48% | 103.9% | 96.42% | 99.64% | 97.92% | 99.24% | 99.1% | 95.32% | 101% | 94.42% | 98.82% | 99.78% | 103.3% | 97.48% | 101.7% | 102.4% | 96.86% | 84.52% | 98.44% | 106.1% | 99.82% | 96.96% | 99.12% | 98.3%
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-09-16 16:30 | 0.4927 | 0.4921 | 0.4896 | 0.4892 | 0.5029 | 0.4909 | 0.5145 | 0.4803 | 0.4961 | 0.488 | 0.4955 | 0.493 | 0.4742 | 0.5003 | 0.4707 | 0.4976 | 0.5035 | 0.5128 | 0.4864 | 0.5046 | 0.5095 | 0.4885 | 0.4436 | 0.4985 | 0.5269 | 0.4996 | 0.4834 | 0.494 | 0.4898
% Recovery |  |  | 98.54% | 98.42% | 97.92% | 97.84% | 100.6% | 98.18% | 102.9% | 96.06% | 99.22% | 97.6% | 99.1% | 98.6% | 94.84% | 100.1% | 94.14% | 99.52% | 100.7% | 102.6% | 97.28% | 100.9% | 101.9% | 97.7% | 88.72% | 99.7% | 105.4% | 99.92% | 96.68% | 98.8% | 97.96%
CCS1 (SS 5ppm) | QC | 2024-09-16 16:32 | 5.1043 | 5.0418 | 5.15 | 4.9823 | 5.109 | 5.0224 | 4.9623 | 4.8963 | 5.0266 | 4.9387 | 5.0618 | 5.0121 | 5.054 | 5.02 | 4.8546 | 5.0096 | 5.3433 | 5.0553 | 4.9368 | 5.2082 | 5.27 | 5.0062 | 4.6884 | 5.058 | 4.8473 | 5.0967 | 4.9774 | 5.0616 | 4.8537
% Recovery |  |  | 102.1% | 100.8% | 103% | 99.65% | 102.2% | 100.4% | 99.25% | 97.93% | 100.5% | 98.77% | 101.2% | 100.2% | 101.1% | 100.4% | 97.09% | 100.2% | 106.9% | 101.1% | 98.74% | 104.2% | 105.4% | 100.1% | 93.77% | 101.2% | 96.95% | 101.9% | 99.55% | 101.2% | 97.07%
CCB | QC | 2024-09-16 16:35 | -0.0001 | 0.0004 | 0.0159 | -0.0009 | 0 | 0 | 0.0012 | 0 | -0.0002 | -0.0002 | 0 | -0.0005 | -0.002 | -0.0024 | 0.0092 | 0.0001 | 0.0426 | -0.0002 | -0.0011 | 0.0001 | -0.0021 | 0.0012 | 0.15 | 0.0258 | 0.0414 | 0.0001 | 0.0016 | -0.0002 | -0.0001
MQ-Blank-03 | Sample | 2024-09-16 16:37 | 0.0001 | 0.001 | -0.0005 | -0.0027 | -0.0001 | 0 | -0.0001 | -0.0002 | 0 | -0.0003 | -0.0004 | 0.0012 | -0.0011 | -0.002 | 0.0098 | 0.0002 | 0.0015 | 0.0001 | -0.0022 | 0.0008 | -0.0028 | -0.0005 | 0.1433 | 0.0124 | 0.039 | 0 | 0 | -0.0003 | -0.0002
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Lowest Cal Std |  |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Highest Cal Std |  |  | 20 | 200 | 20 | 20 | 50 | 20 | 2000 | 20 | 20 | 50 | 100 | 1000 | 100 | 100 | 2000 | 50 | 20 | 100 | 50 | 200 | 200 | 100 | 50 | 20 | 100 | 20 | 50 | 20 | 100
Correlation |  |  | 1 | 0.99998 | 0.99998 | 1 | 1 | 0.99999 | 0.99996 | 1 | 0.99999 | 0.99996 | 0.99998 | 1 | 0.99998 | 0.99999 | 0.99999 | 0.99999 | 0.99976 | 0.99998 | 0.99998 | 0.99999 | 0.99981 | 0.99997 | 0.99994 | 0.99998 | 0.99985 | 0.99998 | 1 | 1 | 1
ICP-OES MDL |  |  | 0.0045598 | 0.0048715 | 0.0060725 | 0.0016216 | 0.00011879 | 0.000168 | 0.0038865 | 0.00024729 | 0.00024729 | 0.0007 | 0.00076374 | 0.0031683 | 0.032873 | 0.0015923 | 0.0080477 | 0.0008057 | 0.00050865 | 0.001351 | 0.0024937 | 0.0072662 | 0.0049325 | 0.0058515 | 0.0056274 | 0.0064431 | 0.002358 | 0.000168 | 0.00011879 | 0.00093287 | 0.00059397
ICP-OES MRL/PRDL |  |  | 0.045598 | 0.048715 | 0.060725 | 0.016216 | 0.0011879 | 0.00168 | 0.038865 | 0.0024729 | 0.0024729 | 0.0074 | 0.0076374 | 0.031683 | 0.32873 | 0.015923 | 0.080477 | 0.008057 | 0.0050865 | 0.01351 | 0.024937 | 0.072662 | 0.049325 | 0.058515 | 0.056274 | 0.064431 | 0.02358 | 0.00168 | 0.0011879 | 0.0093287 | 0.0059397
Where |  |  | Ba: QC Recoveries
| Calibration Standards |  | Ba: 100%; Be: 90<R%>110
| IEC Check Standards |  | Ba: 80%; Be: 70<R%>80
| CCV & Secondary source QC check standards |  | Ba: 70%; Be: 70%>R%>120%
| Low Level QC chek standards |  | Ba: 60%; Be: 65%>R%>135%
| Calculated/Estimated MDLs & MRLs |  | Ba: 100%; Be: 65<R%>135
| Dilution required |  | Ba: 20%; Be: RPD<20%
* | Sampling Error |  | Ba: 21%; Be: RPD>20%
** | Different concentration than the normal |  | Ba: 20; Be: RSD<20%
| Mislabeled |  | Ba: 21; Be: RSD>20%
[7 empty] | 20 | STDEV<20

## 09162024ICP-OES_TO178-477RD
Worksheet exported from: C:\Users\USEPA\Documents\Agilent\ICP Expert\My Results\09162024_ICP-OES_TO 178-475 476 & 477 + TO 212-22.esws
Solution Label | Type | Date Time | LIMS ID | Customer | Customer Ref. | Sample Date Time | Sample Site | Description | Ag (328.068 nm) ppm | Al (396.152 nm) ppm | As (188.980 nm) ppm | B (249.678 nm) ppm | Ba (455.403 nm) ppm | Be (313.107 nm) ppm | Ca (370.602 nm) ppm | Cd (214.439 nm) ppm | Co (238.892 nm) ppm | Cr (267.716 nm) ppm | Cu (324.754 nm) ppm | Fe (259.940 nm) ppm | K (766.491 nm) ppm | Li (670.783 nm) ppm | Mg (279.078 nm) ppm | Mn (257.610 nm) ppm | Mo (202.032 nm) ppm | Na (588.995 nm) ppm | Ni (231.604 nm) ppm | P (178.222 nm) ppm | Pb (220.353 nm) ppm | S (181.972 nm) ppm | Sb (217.582 nm) ppm | Se (196.026 nm) ppm | Si (251.611 nm) ppm | Sn (189.925 nm) ppm | Sr (421.552 nm) ppm | Ti (334.941 nm) ppm | V (292.401 nm) ppm | Zn (206.200 nm) ppm | Y (371.029 nm) Ratio | Y (371.029 nm) Ratio 1 | Y (488.368 nm) Ratio | Ag 328.068 C ppm | Al 396.152 C ppm | As 188.980 C ppm | B 249.678 C ppm | Ba 455.403 C ppm | Be 313.107 C ppm | Ca 370.602 C ppm | Cd 214.439 C ppm | Co 238.892 C ppm | Cr 267.716 C ppm | Cu 324.754 C ppm | Fe 259.940 C ppm | K 766.491 C ppm | Li 670.783 C ppm | Mg 279.078 C ppm | Mn 257.610 C ppm | Mo 202.032 C ppm | Na 588.995 C ppm | Ni 231.604 C ppm | P 178.222 C ppm | Pb 220.353 C ppm | S 181.972 C ppm | Sb 217.582 C ppm | Se 196.026 C ppm | Si 251.611 C ppm | Sr 421.552 C ppm | Ti 334.941 C ppm | V 292.401 C ppm | Zn 206.200 C ppm
Calibration Blank | Blank | 2024-09-16 09:25 | [6 empty] | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 1.00 ! | 1.00 ! | 1.00 !
Calibration Standard 0.5ppm | Standard | 2024-09-16 09:27 | [6 empty] | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 1.01 ! | 1.01 ! | 1.01 !
Calibration Standard 1ppm | Standard | 2024-09-16 09:30 | [6 empty] | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.01 ! | 1.00 ! | 1.00 !
Calibration Standard 5ppm | Standard | 2024-09-16 09:32 | [6 empty] | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 1.00 ! | 1.00 ! | 0.99 !
Calibration Standard 10ppm | Standard | 2024-09-16 09:35 | [6 empty] | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 1.01 ! | 0.99 ! | 0.99 !
Calibration Standard 20ppm | Standard | 2024-09-16 09:37 | [6 empty] | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 1.00 ! | 0.97 ! | 0.97 !
Calibration Standard 50ppm | Standard | 2024-09-16 09:40 | [7 empty] | 50.0000 ! |  |  | 50.0000 ! |  | 50.0000 ! |  |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! |  |  | 50.0000 ! |  | 50.0000 ! | 0.98 ! | 0.94 ! | 0.93 !
Calibration Standard-A 100ppm | Standard | 2024-09-16 09:42 | [7 empty] | 100.0000 ! |  |  |  |  | 100.0000 ! |  |  |  | 100.0000 ! | 100.0000 ! | 100.0000 ! | 100.0000 ! | 100.0000 ! |  |  | 100.0000 ! |  |  | 100.0000 ! |  |  |  | 100.0000 ! |  |  |  |  | 100.0000 ! | 0.94 ! | 0.90 ! | 0.91 !
Calibration Standard-B 100ppm | Standard | 2024-09-16 09:45 | [25 empty] | 100.0000 ! |  | 100.0000 ! | [8 empty] | 1.02 ! | 1.00 ! | 0.99 !
Calibration Standard Al+Fe+Ca+Mg200ppm | Standard | 2024-09-16 09:47 | [7 empty] | 200.0000 ! |  |  |  |  | 200.0000 ! |  |  |  |  | 200.0000 ! |  |  | 200.0000 ! | [15 empty] | 0.97 ! | 0.95 ! | 0.95 !
Calibration Standard Ca+Mg500ppm | Standard | 2024-09-16 09:50 | [12 empty] | 500.0000 ! | [7 empty] | 500.0000 ! | [15 empty] | 0.95 ! | 0.92 ! | 0.92 !
Calibration Standard Ca+Mg1000ppm | Standard | 2024-09-16 09:52 | [12 empty] | 1000.0000 ! | [7 empty] | 1000.0000 ! | [15 empty] | 0.91 ! | 0.88 ! | 0.88 !
Calibration Standard Ca+Mg2000ppm | Standard | 2024-09-16 09:55 | [12 empty] | 2000.0000 ! | [7 empty] | 2000.0000 ! | [15 empty] | 0.87 ! | 0.83 ! | 0.83 !
Calibration Standard P+S200ppm | Standard | 2024-09-16 09:58 | [25 empty] | 200.0000 ! |  | 200.0000 ! | [8 empty] | 1.04 ! | 1.01 ! | 1.01 !
Calibration Standard Pb200ppm | Standard | 2024-09-16 10:00 | [26 empty] | 200.0000 ! | [9 empty] | 1.03 ! | 1.00 ! | 1.00 !
Calibration Standard Fe 1000ppm | Standard | 2024-09-16 10:03 | [17 empty] | 1000.0000 ! | [18 empty] | 0.99 ! | 0.97 ! | 0.97 !
Rinse1 | QC | 2024-09-16 10:23 | [6 empty] | -0.0001 ! | 0.0029 ! | 0.0003 ! | 0.0028 ! | 0.0000 ! | -0.0001 ! | 0.0074 ! | 0.0000 ! | 0.0000 ! | -0.0001 ! | 0.0003 ! | -0.0007 ! | -0.0027 ! | -0.0012 ! | 0.0136 ! | 0.0001 ! | -0.0003 ! | 0.0028 ! | -0.0012 ! | 0.0039 ! | -0.0018 ! | 0.0000 ! | 0.1434 ! | 0.0140 ! | 0.0227 ! | -0.0007 ! | 0.0000 ! | 0.0000 ! | -0.0003 ! | 0.0003 ! | 1.02 ! | 1.01 ! | 1.00 ! | -0.0001 | 0.0029 | 0.0003 | 0.0028 | 0 | -0.0001 | 0.0074 | 0 | 0 | -0.0001 | 0.0003 | -0.0007 | -0.0027 | -0.0012 | 0.0136 | 0.0001 | -0.0003 | 0.0028 | -0.0012 | 0.0039 | -0.0018 | 0 | 0.1434 | 0.014 | 0.0227 | 0 | 0 | -0.0003 | 0.0003
Rinse | QC | 2024-09-16 10:26 | [6 empty] | -0.0002 ! | -0.0010 ! | 0.0017 ! | 0.0023 ! | 0.0000 ! | 0.0000 ! | 0.0047 ! | 0.0001 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | -0.0035 ! | -0.0014 ! | 0.0135 ! | 0.0001 ! | 0.0006 ! | 0.0009 ! | -0.0013 ! | 0.0039 ! | -0.0028 ! | 0.0035 ! | 0.1455 ! | 0.0165 ! | 0.0216 ! | -0.0014 ! | 0.0000 ! | 0.0001 ! | -0.0002 ! | 0.0004 ! | 1.03 ! | 1.01 ! | 1.01 ! | -0.0002 | -0.001 | 0.0017 | 0.0023 | 0 | 0 | 0.0047 | 0.0001 | 0 | 0 | 0 | 0 | -0.0035 | -0.0014 | 0.0135 | 0.0001 | 0.0006 | 0.0009 | -0.0013 | 0.0039 | -0.0028 | 0.0035 | 0.1455 | 0.0165 | 0.0216 | 0 | 0.0001 | -0.0002 | 0.0004
IECS1 (Al 1000 Int.) | QC | 2024-09-16 10:28 | [6 empty] | 0.0000 ! | 1039.4457 !o | 0.0001 ! | 0.0051 ! | 0.0097 ! | 0.0000 ! | 0.0094 ! | 0.0000 ! | -0.0002 ! | 0.0002 ! | 0.0024 ! | 0.0036 ! | -0.0048 ! | -0.0013 ! | 0.0228 ! | 0.0003 ! | 0.0031 ! | 0.0078 ! | -0.0011 ! | 0.0094 ! | 0.0259 ! | -0.0037 ! | 0.3079 ! | 0.0154 ! | 0.0381 ! | -0.0006 ! | 0.0001 ! | 0.0002 ! | 0.0004 ! | 0.0046 ! | 0.95 ! | 0.98 ! | 0.98 ! | 0 | 1039.4457 | 0.0001 | 0.0051 | 0.0097 | 0 | 0.0094 | 0 | -0.0002 | 0.0002 | 0.0024 | 0.0036 | -0.0048 | -0.0013 | 0.0228 | 0.0003 | 0.0031 | 0.0078 | -0.0011 | 0.0094 | 0.0259 | -0.0037 | 0.3079 | 0.005 | 0.0381 | 0.0001 | 0.0002 | 0.0004 | 0.0046
Rinse1 | QC | 2024-09-16 10:31 | [6 empty] | 0.0001 | 0.0145 | 0.002 | 0.0023 | -0.0001 | -0.0001 | 0.0048 | 0 | 0.0002 | -0.0001 | 0.0004 | 0.0007 | -0.003 | -0.0007 | 0.0116 | 0.0001 | 0.0003 | 0.0015 | -0.0013 | 0.0041 | -0.003 | -0.0004 | 0.1456 | 0.0156 | 0.0232 | -0.0003 | 0 | 0.0001 | 0 | 0.0002 | 1.02 | 1.01 | 1.01 | 0.0001 | 0.0145 | 0.002 | 0.0023 | -0.0001 | -0.0001 | 0.0048 | 0 | 0.0002 | -0.0001 | 0.0004 | 0.0007 | -0.003 | -0.0007 | 0.0116 | 0.0001 | 0.0003 | 0.0015 | -0.0013 | 0.0041 | -0.003 | -0.0004 | 0.1456 | 0.0156 | 0.0232 | 0 | 0.0001 | 0 | 0.0002
Rinse | QC | 2024-09-16 10:33 | [6 empty] | -0.0001 | 0.0052 | 0.0005 | 0.0016 | 0.0001 | 0 | 0.004 | -0.0001 | 0.0002 | 0 | 0.0001 | 0.0001 | -0.0072 | -0.0016 | 0.0092 | 0.0001 | -0.0005 | 0.0007 | -0.0013 | 0.0011 | -0.0016 | 0.0011 | 0.1465 | 0.0135 | 0.0227 | -0.0008 | 0 | 0.0001 | 0.0002 | 0.0005 | 1.03 | 1.02 | 1.01 | -0.0001 | 0.0052 | 0.0005 | 0.0016 | 0.0001 | 0 | 0.004 | -0.0001 | 0.0002 | 0 | 0.0001 | 0.0001 | -0.0072 | -0.0016 | 0.0092 | 0.0001 | -0.0005 | 0.0007 | -0.0013 | 0.0011 | -0.0016 | 0.0011 | 0.1465 | 0.0135 | 0.0227 | 0 | 0.0001 | 0.0002 | 0.0005
IECS2 (Fe 1000 Int.) | QC | 2024-09-16 10:36 | [6 empty] | -0.0003 | 0.0337 | 0.0003 | -0.0057 | -0.0001 | 0 | -0.0055 | 0.0167 | 0.0109 | 0.0045 | -0.0113 | 1008.1799 | -0.006 | -0.0014 | 1.6831 | 0.009 | -0.0006 | 0.0017 | 0.0065 | 0.0157 | 0.0324 | 0.0066 | 0.1258 | -0.0028 | 0.0163 | 0.0044 | 0.0004 | 0 | 0.0038 | 0.0031 | 1 | 0.97 | 0.97 | -0.0003 | 0.0337 | 0.0003 | -0.0057 | -0.0001 | 0 | -0.0055 | 0.0016 | -0.0042 | 0.0045 | -0.0002 | 1008.1799 | -0.006 | -0.0014 | 1.5117 K | 0.009 | -0.0006 | 0.0017 | 0.0065 | 0.0006 | 0.0143 | 0.0065 | 0.1359 | -0.021 | 0.0163 | 0.0004 | 0 | -0.0022 | 0.0031
Rinse1 | QC | 2024-09-16 10:38 | [6 empty] | 0.0002 | 0.0052 | 0.0005 | 0.0011 | -0.0001 | -0.0001 | 0.005 | 0 | -0.0002 | 0 | 0.0002 | 0.021 | -0.0077 | -0.0015 | 0.0119 | 0.0003 | 0 | 0.0008 | -0.0017 | 0.0011 | -0.0011 | 0.0005 | 0.149 | 0.0127 | 0.0249 | -0.0007 | 0 | 0 | -0.0003 | 0.0003 | 1.02 | 1.02 | 1.01 | 0.0002 | 0.0052 | 0.0005 | 0.0011 | -0.0001 | -0.0001 | 0.005 | 0 | -0.0002 | 0 | 0.0002 | 0.021 | -0.0077 | -0.0015 | 0.0119 | 0.0003 | 0 | 0.0008 | -0.0017 | 0.0011 | -0.0011 | 0.0005 | 0.149 | 0.0127 | 0.0249 | 0 | 0 | -0.0003 | 0.0003
Rinse | QC | 2024-09-16 10:41 | [6 empty] | 0.0002 | 0.0024 | 0.001 | 0.0008 | -0.0001 | 0 | 0.0046 | 0.0001 | -0.0001 | 0 | 0.0005 | 0.0062 | 0.0017 | -0.0009 | 0.0122 | 0.0003 | 0.0004 | 0.0004 | -0.0018 | 0.002 | -0.0032 | 0.0032 | 0.1449 | 0.0104 | 0.0255 | -0.0004 | 0 | 0 | 0 | 0.0005 | 1.02 | 1.01 | 1.01 | 0.0002 | 0.0024 | 0.001 | 0.0008 | -0.0001 | 0 | 0.0046 | 0.0001 | -0.0001 | 0 | 0.0005 | 0.0062 | 0.0017 | -0.0009 | 0.0122 | 0.0003 | 0.0004 | 0.0004 | -0.0018 | 0.002 | -0.0032 | 0.0032 | 0.1449 | 0.0104 | 0.0255 | 0 | 0 | 0 | 0.0005
IECS3 (Mn 100 Int.) | QC | 2024-09-16 10:43 | [6 empty] | 0.0074 | 0.0031 | 0.0005 | 0.0002 | -0.0001 | 0 | 0.699 | 0 | 0.0022 | 0.0086 | -0.0003 | 0.0014 | -0.0031 | -0.0007 | 0.0146 | 99.8516 o | -0.0003 | 0.0003 | -0.0018 | 0.0028 | 0.0097 | 0.4407 | 0.1538 | 0.0791 | 0.0638 | -0.0017 | 0 | 0 | 0.0005 | 0.0001 | 1.02 | 1 | 1 | -0.0025 | 0.0031 | 0.0005 | 0.0002 | -0.0001 | 0 | 0.6990 K | 0 | -0.0028 | -0.0194 | -0.0003 | 0.0014 | -0.0031 | -0.0007 | 0.0146 | 99.8516 | -0.0003 | 0.0003 | -0.0018 | 0.0028 | -0.0033 | -0.0985 | 0.1538 | 0.0221 | 0.0259 | 0 | 0 | 0.0005 | 0.0001
IECS4 (Ca Mg 1000 Int.) | QC | 2024-09-16 10:46 | [6 empty] | -0.0001 | 0.1027 | 0.0003 | 0.0044 | 0.002 | 0.0001 | 989.1619 | 0.0082 | 0.0002 | 0.0012 | 0.0029 | 0.0042 | 0.0027 | -0.0024 | 1023.732 | 0.0042 | 0.0007 | 0.0176 | 0.0022 | 0.0042 | 0.0047 | 0.0481 | 0.1568 | 0.0145 | 0.0547 | 0.0001 | 0.0079 | -0.0004 | 0.0007 | 0.0119 | 0.92 | 0.88 | 0.88 | -0.0001 | 0.1027 | 0.0003 | 0.0044 | 0.002 | 0.0001 | 989.1619 | 0.0082 | 0.0002 | 0.0012 | 0.0029 | 0.0042 | 0.0027 | -0.0024 | 1023.732 | 0.0042 | 0.0007 | 0.0176 | 0.0022 | 0.0042 | 0.0047 | 0.0481 | 0.1568 | 0.0145 | 0.0547 | 0.0079 | -0.0004 | 0.0007 | 0.0119
Rinse1 | QC | 2024-09-16 10:48 | [6 empty] | 0.0001 | 0.0011 | 0.0011 | 0.0004 | 0 | 0 | 0.013 | 0.0001 | 0.0003 | 0.0001 | 0 | 0.0008 | -0.0054 | -0.0016 | 0.0204 | 0.0006 | -0.0005 | 0.0032 | -0.0001 | 0.0015 | -0.0043 | 0.0036 | 0.1509 | 0.0133 | 0.026 | -0.0015 | 0 | 0.0001 | 0.0001 | 0.0004 | 1.02 | 1.01 | 1.01 | 0.0001 | 0.0011 | 0.0011 | 0.0004 | 0 | 0 | 0.013 | 0.0001 | 0.0003 | 0.0001 | 0 | 0.0008 | -0.0054 | -0.0016 | 0.0204 | 0.0006 | -0.0005 | 0.0032 | -0.0001 | 0.0015 | -0.0043 | 0.0036 | 0.1509 | 0.0133 | 0.026 | 0 | 0.0001 | 0.0001 | 0.0004
Rinse | QC | 2024-09-16 10:51 | [6 empty] | 0.0001 | -0.0007 | 0.0006 | 0.0003 | 0 | -0.0001 | 0.0084 | 0 | 0.0002 | -0.0001 | 0.0003 | 0.0003 | -0.0008 | -0.0016 | 0.017 | 0.0001 | 0 | 0.0003 | -0.0005 | -0.0008 | -0.0026 | 0.003 | 0.1514 | 0.0127 | 0.0269 | -0.0007 | 0 | 0.0001 | -0.0003 | 0.0001 | 1.02 | 1.01 | 1 | 0.0001 | -0.0007 | 0.0006 | 0.0003 | 0 | -0.0001 | 0.0084 | 0 | 0.0002 | -0.0001 | 0.0003 | 0.0003 | -0.0008 | -0.0016 | 0.017 | 0.0001 | 0 | 0.0003 | -0.0005 | -0.0008 | -0.0026 | 0.003 | 0.1514 | 0.0127 | 0.0269 | 0 | 0.0001 | -0.0003 | 0.0001
IECS5 (Ti 20 Int.) | QC | 2024-09-16 10:53 | [6 empty] | 0.0002 | 0.0021 | 0.0013 | 0.0017 | -0.0003 | -0.0003 | 0.0932 | 0 | -0.0001 | -0.0002 | -0.0007 | 0.0013 | -0.0044 | -0.0017 | 0.0181 | 0 | -0.0002 | 0.0009 | -0.0011 | -0.0005 | -0.005 | 0.001 | 0.1475 | 0.011 | 0.0664 | -0.0027 | 0 | 20.1663 | 0.0128 | -0.0002 | 1.03 | 1.01 | 1 | 0.0002 | 0.0021 | 0.0013 | 0.0017 | -0.0003 | -0.0003 | 0.0932 | 0 | -0.0001 | -0.0002 | -0.0007 | 0.0013 | -0.0044 | -0.0017 | 0.0181 | 0 | -0.0002 | 0.0009 | -0.0011 | -0.0005 | -0.005 | 0.001 | 0.1475 | 0.011 | 0.0664 | 0 | 20.1663 K | -0.0053 | -0.0002
Rinse1 | QC | 2024-09-16 10:56 | [6 empty] | 0.0001 | 0 | 0.0007 | 0.0004 | -0.0001 | -0.0001 | 0.0078 | 0.0001 | 0.0001 | -0.0004 | 0.0001 | 0.0007 | -0.0044 | -0.0018 | 0.0126 | 0.0001 | -0.0007 | 0.0032 | -0.002 | 0.002 | -0.0018 | 0.0003 | 0.1487 | 0.0133 | 0.0295 | -0.0012 | 0 | 0.0045 | -0.0004 | 0.0003 | 1.02 | 1.01 | 1 | 0.0001 | 0 | 0.0007 | 0.0004 | -0.0001 | -0.0001 | 0.0078 | 0.0001 | 0.0001 | -0.0004 | 0.0001 | 0.0007 | -0.0044 | -0.0018 | 0.0126 | 0.0001 | -0.0007 | 0.0032 | -0.002 | 0.002 | -0.0018 | 0.0003 | 0.1487 | 0.0133 | 0.0295 | 0 | 0.0045 | -0.0004 | 0.0003
Rinse | QC | 2024-09-16 10:58 | [6 empty] | 0.0001 | -0.0004 | -0.0017 | 0.0005 | -0.0001 | 0 | 0.0045 | 0.0002 | 0.0003 | 0 | 0.0003 | 0.0006 | -0.0073 | -0.0016 | 0.0154 | 0 | -0.0003 | 0.0011 | -0.0013 | 0.0002 | -0.0032 | 0.0033 | 0.1499 | 0.0173 | 0.0284 | -0.0013 | 0 | 0.0015 | -0.0003 | 0.0001 | 1.02 | 1.01 | 1 | 0.0001 | -0.0004 | -0.0017 | 0.0005 | -0.0001 | 0 | 0.0045 | 0.0002 | 0.0003 | 0 | 0.0003 | 0.0006 | -0.0073 | -0.0016 | 0.0154 | 0 | -0.0003 | 0.0011 | -0.0013 | 0.0002 | -0.0032 | 0.0033 | 0.1499 | 0.0173 | 0.0284 | 0 | 0.0015 | -0.0003 | 0.0001
ICB | QC | 2024-09-16 11:01 | [6 empty] | -0.0001 | -0.0003 | 0.0006 | 0.0003 | -0.0001 | 0 | 0.0018 | -0.0001 | 0 | 0 | 0 | 0.0004 | -0.0121 | -0.0011 | 0.0144 | 0.0003 | 0 | -0.0008 | -0.0008 | -0.0008 | -0.0025 | 0 | 0.1443 | 0.0121 | 0.0278 | -0.0005 | 0 | 0.0008 | -0.0002 | 0.0001 | 1.02 | 1 | 1 | -0.0001 | -0.0003 | 0.0006 | 0.0003 | -0.0001 | 0 | 0.0018 | -0.0001 | 0 | 0 | 0 | 0.0004 | -0.0121 | -0.0011 | 0.0144 | 0.0003 | 0 | -0.0008 | -0.0008 | -0.0008 | -0.0025 | 0 | 0.1443 | 0.0121 | 0.0278 | 0 | 0.0008 | -0.0002 | 0.0001
ICV (SS 5ppm All) | QC | 2024-09-16 11:04 | [6 empty] | 5.078 | 4.972 | 5.1286 | 4.9878 | 5.0499 | 5.0092 | 5.0185 | 4.9925 | 5.0767 | 4.9781 | 4.9805 | 4.9897 | 5.0416 | 5.0421 | 4.9091 | 4.966 | 5.2977 | 5.0765 | 4.9849 | 5.0176 | 5.2871 | 4.9463 | 4.8947 | 5.0429 | 4.8191 | 5.0969 | 5.0491 | 5.0118 | 5.0643 | 4.9762 | 1.02 | 1 | 0.99 | 5.0775 | 4.972 | 5.1286 | 4.9878 | 5.0499 | 5.0092 | 5.0185 | 4.9925 | 5.0763 | 4.9767 | 4.9806 | 4.9897 | 5.0416 | 5.0421 | 4.9082 | 4.966 | 5.2977 | 5.0765 | 4.9849 | 5.0175 | 5.2863 | 4.9194 | 4.8948 | 5.04 | 4.8172 | 5.0491 | 5.0118 | 5.0597 | 4.9762
Rinse | QC | 2024-09-16 11:06 | [6 empty] | -0.0001 | 0.0023 | 0.0152 | 0.0027 | 0 | 0.0001 | 0.0045 | 0.0002 | 0 | -0.0001 | 0.0002 | 0.0013 | -0.0002 | -0.0019 | 0.0119 | 0.0002 | 0.0447 | 0.0006 | -0.0018 | 0.002 | -0.0008 | 0.004 | 0.1531 | 0.0316 | 0.0308 | 0.0039 | 0.0001 | 0.0019 | 0.0001 | 0.0005 | 1.03 | 1.02 | 1.01 | -0.0001 | 0.0023 | 0.0152 | 0.0027 | 0 | 0.0001 | 0.0045 | 0.0002 | 0 | -0.0001 | 0.0002 | 0.0013 | -0.0002 | -0.0019 | 0.0119 | 0.0002 | 0.0447 | 0.0006 | -0.0018 | 0.002 | -0.0008 | 0.004 | 0.1531 | 0.0316 | 0.0308 | 0.0001 | 0.0019 | 0.0001 | 0.0005
CCB | QC | 2024-09-16 11:09 | [6 empty] | 0.0001 | 0.0004 | 0.0035 | 0.0013 | 0 | -0.0001 | -0.0006 | 0 | -0.0001 | -0.0004 | 0.0002 | -0.0001 | -0.0051 | -0.0016 | 0.0102 | 0.0002 | 0.0124 | 0.0005 | -0.0016 | 0.0026 | -0.0007 | 0.0008 | 0.1446 | 0.0153 | 0.03 | 0.0001 | 0 | 0.0008 | 0 | 0.0001 | 1.03 | 1.01 | 1.01 | 0.0001 | 0.0004 | 0.0035 | 0.0013 | 0 | -0.0001 | -0.0006 | 0 | -0.0001 | -0.0004 | 0.0002 | -0.0001 | -0.0051 | -0.0016 | 0.0102 | 0.0002 | 0.0124 | 0.0005 | -0.0016 | 0.0026 | -0.0007 | 0.0008 | 0.1446 | 0.0153 | 0.03 | 0 | 0.0008 | 0 | 0.0001
0.5ppm All (LLQC) | QC | 2024-09-16 11:11 | [6 empty] | 0.4928 | 0.4909 | 0.4909 | 0.4928 | 0.5017 | 0.4915 | 0.5271 | 0.4928 | 0.5017 | 0.4951 | 0.4901 | 0.4967 | 0.4768 | 0.5007 | 0.479 | 0.4968 | 0.4952 | 0.5129 | 0.4935 | 0.493 | 0.5169 | 0.4841 | 0.4738 | 0.4907 | 0.5188 | 0.4899 | 0.5 | 0.4889 | 0.498 | 0.5029 | 1.04 | 1.01 | 1.01 | 0.4927 | 0.4909 | 0.4909 | 0.4928 | 0.5017 | 0.4915 | 0.5271 | 0.4928 | 0.5017 | 0.495 | 0.4901 | 0.4967 | 0.4768 | 0.5007 | 0.4789 | 0.4968 | 0.4952 | 0.5129 | 0.4935 | 0.493 | 0.5168 | 0.4814 | 0.4738 | 0.4904 | 0.5186 | 0.5 | 0.4889 | 0.4975 | 0.5029
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-09-16 11:17 | [6 empty] | 0.492 | 0.4943 | 0.4841 | 0.4933 | 0.5079 | 0.4937 | 0.5257 | 0.4921 | 0.5017 | 0.4943 | 0.4882 | 0.5023 | 0.481 | 0.5035 | 0.4779 | 0.5008 | 0.495 | 0.5172 | 0.4926 | 0.486 | 0.5158 | 0.4821 | 0.5206 | 0.494 | 0.5192 | 0.4903 | 0.504 | 0.4879 | 0.4971 | 0.5032 | 1.04 | 1.01 | 1.01 | 0.4919 | 0.4943 | 0.4841 | 0.4933 | 0.5079 | 0.4937 | 0.5257 | 0.4921 | 0.5017 | 0.4942 | 0.4882 | 0.5023 | 0.481 | 0.5035 | 0.4779 | 0.5008 | 0.495 | 0.5172 | 0.4926 | 0.4859 | 0.5157 | 0.4794 | 0.5206 | 0.4937 | 0.519 | 0.504 | 0.4879 | 0.4967 | 0.5032
CCV (5ppm All) | QC | 2024-09-16 11:20 | [6 empty] | 4.9569 | 4.9015 | 4.9868 | 4.9646 | 5.0263 | 4.9731 | 4.9525 | 4.9397 | 5.0489 | 4.9418 | 4.9297 | 4.9523 | 4.9989 | 5.0673 | 4.7845 | 4.9289 | 5.1843 | 4.9983 | 4.9356 | 4.9244 | 5.2059 | 4.8779 | 4.9682 | 4.9677 | 4.9218 | 4.9662 | 4.9792 | 4.9033 | 4.9862 | 4.939 | 1.04 | 1 | 0.99 | 4.9564 | 4.9015 | 4.9868 | 4.9646 | 5.0263 | 4.9731 | 4.9525 | 4.9397 | 5.0485 | 4.9404 | 4.9298 | 4.9523 | 4.9989 | 5.0673 | 4.7836 | 4.9289 | 5.1843 | 4.9983 | 4.9356 | 4.9243 | 5.2052 | 4.8513 | 4.9682 | 4.9647 | 4.9199 | 4.9792 | 4.9033 | 4.9818 | 4.939
CCB | QC | 2024-09-16 11:23 | [6 empty] | 0 | -0.0017 | 0.0093 | 0.0025 | 0 | 0 | -0.0022 | 0.0001 | 0.0003 | -0.0001 | 0 | -0.0019 | 0.0052 | -0.002 | 0.0119 | 0.0003 | 0.0317 | 0.0009 | -0.0011 | -0.0009 | -0.004 | -0.0029 | 0.159 | 0.0246 | 0.0315 | 0.0024 | 0 | 0.0013 | 0 | 0 | 1.04 | 1.01 | 1.01 | 0 | -0.0017 | 0.0093 | 0.0025 | 0 | 0 | -0.0022 | 0.0001 | 0.0003 | -0.0001 | 0 | -0.0019 | 0.0052 | -0.002 | 0.0119 | 0.0003 | 0.0317 | 0.0009 | -0.0011 | -0.0009 | -0.004 | -0.0029 | 0.159 | 0.0246 | 0.0315 | 0 | 0.0013 | 0 | 0
Rinse | QC | 2024-09-16 14:43 | [6 empty] | 0 | -0.0006 | 0.0011 | -0.0022 | 0 | 0 | 0.0054 | -0.0001 | -0.0003 | 0 | 0.0005 | 0.0044 | -0.0013 | -0.001 | 0.011 | 0.0005 | 0.0001 | 0.0009 | -0.0012 | 0.003 | -0.0036 | 0.0203 | 0.1431 | 0.0103 | 0.0401 | 0.0002 | 0 | 0.0001 | -0.0002 | 0.002 | 1.08 | 1.01 | 1.01 | 0 | -0.0006 | 0.0011 | -0.0022 | 0 | 0 | 0.0054 | -0.0001 | -0.0003 | 0 | 0.0005 | 0.0044 | -0.0013 | -0.001 | 0.011 | 0.0005 | 0.0001 | 0.0009 | -0.0012 | 0.003 | -0.0036 | 0.0203 | 0.1431 | 0.0103 | 0.0401 | 0 | 0.0001 | -0.0002 | 0.002
CCV (5ppm All) | QC | 2024-09-16 14:46 | [6 empty] | 4.9709 | 4.942 | 4.9863 | 4.9527 | 5.0236 | 4.9202 | 4.9103 | 4.8663 | 4.9801 | 4.9144 | 4.9842 | 4.9533 | 4.9402 | 5.0763 | 4.7564 | 4.9187 | 5.1652 | 4.9884 | 4.8957 | 5.0625 | 5.1856 | 4.9547 | 4.7383 | 4.9727 | 4.9438 | 4.9363 | 4.9761 | 4.8704 | 4.9814 | 4.8546 | 1.07 | 1 | 1 | 4.9704 | 4.942 | 4.9863 | 4.9527 | 5.0236 | 4.9202 | 4.9103 | 4.8662 | 4.9798 | 4.913 | 4.9843 | 4.9533 | 4.9402 | 5.0763 | 4.7555 | 4.9187 | 5.1652 | 4.9884 | 4.8957 | 5.0625 | 5.1848 | 4.9282 | 4.7384 | 4.9697 | 4.942 | 4.9761 | 4.8704 | 4.977 | 4.8546
CCB | QC | 2024-09-16 14:49 | [6 empty] | -0.0001 | 0.0022 | 0.0076 | -0.0014 | -0.0001 | 0 | -0.0007 | 0 | 0 | -0.0001 | 0 | -0.0005 | -0.0105 | -0.0022 | 0.0108 | 0 | 0.0246 | 0 | -0.001 | 0.0003 | -0.0034 | 0.0021 | 0.1465 | 0.0187 | 0.0396 | 0.0015 | 0.0001 | 0.001 | -0.0001 | 0.0006 | 1.08 | 1.01 | 1.01 | -0.0001 | 0.0022 | 0.0076 | -0.0014 | -0.0001 | 0 | -0.0007 | 0 | 0 | -0.0001 | 0 | -0.0005 | -0.0105 | -0.0022 | 0.0108 | 0 | 0.0246 | 0 | -0.001 | 0.0003 | -0.0034 | 0.0021 | 0.1465 | 0.0187 | 0.0396 | 0.0001 | 0.001 | -0.0001 | 0.0006
0.5ppm All (LLQC) | QC | 2024-09-16 14:52 | [6 empty] | 0.4957 | 0.4911 | 0.4896 | 0.4913 | 0.5041 | 0.4921 | 0.5221 | 0.4862 | 0.501 | 0.4926 | 0.497 | 0.4995 | 0.4568 | 0.5029 | 0.473 | 0.496 | 0.5004 | 0.514 | 0.49 | 0.508 | 0.513 | 0.4906 | 0.4594 | 0.4947 | 0.5297 | 0.4877 | 0.5008 | 0.4879 | 0.4986 | 0.497 | 1.08 | 1.01 | 1.01 | 0.4957 | 0.4911 | 0.4896 | 0.4913 | 0.5041 | 0.4921 | 0.5221 | 0.4862 | 0.501 | 0.4925 | 0.497 | 0.4995 | 0.4568 | 0.5029 | 0.4729 | 0.496 | 0.5004 | 0.514 | 0.49 | 0.508 | 0.5129 | 0.4879 | 0.4594 | 0.4944 | 0.5295 | 0.5008 | 0.4879 | 0.4982 | 0.497
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-09-16 14:54 | [6 empty] | 0.4923 | 0.4921 | 0.4931 | 0.4885 | 0.5044 | 0.4919 | 0.519 | 0.4827 | 0.4965 | 0.4892 | 0.4923 | 0.4949 | 0.4695 | 0.5014 | 0.4738 | 0.4961 | 0.5024 | 0.515 | 0.4873 | 0.5002 | 0.509 | 0.4895 | 0.4837 | 0.4934 | 0.5267 | 0.4873 | 0.5014 | 0.4842 | 0.4954 | 0.4932 | 1.08 | 1.01 | 1.02 | 0.4922 | 0.4921 | 0.4931 | 0.4885 | 0.5044 | 0.4919 | 0.519 | 0.4827 | 0.4965 | 0.4891 | 0.4923 | 0.4949 | 0.4695 | 0.5014 | 0.4737 | 0.4961 | 0.5024 | 0.515 | 0.4873 | 0.5002 | 0.5089 | 0.4868 | 0.4837 | 0.4931 | 0.5265 | 0.5014 | 0.4842 | 0.495 | 0.4932
CCS1 (SS 5ppm) | QC | 2024-09-16 14:57 | [6 empty] | 5.0588 | 5.0199 | 5.0967 | 4.9366 | 5.0908 | 5.0007 | 4.9225 | 4.8642 | 4.9752 | 4.9027 | 5.0099 | 5.0044 | 5.0233 | 5.0326 | 4.8141 | 5.0134 | 5.2946 | 5.06 | 4.8958 | 5.1232 | 5.2166 | 4.9836 | 4.7613 | 5.0161 | 4.8198 | 5.0181 | 5.0737 | 4.9428 | 5.0184 | 4.8218 | 1.08 | 0.99 | 1 | 5.0583 | 5.0199 | 5.0967 | 4.9366 | 5.0908 | 5.0007 | 4.9225 | 4.8641 | 4.9749 | 4.9013 | 5.01 | 5.0044 | 5.0233 | 5.0326 | 4.8132 | 5.0134 | 5.2946 | 5.06 | 4.8958 | 5.1232 | 5.2159 | 4.9565 | 4.7614 | 5.0131 | 4.8179 | 5.0737 | 4.9428 | 5.0139 | 4.8218
CCB | QC | 2024-09-16 14:59 | [6 empty] | -0.0002 | 0.0009 | 0.0153 | -0.0005 | 0 | 0.0001 | -0.0008 | 0.0001 | 0.0001 | 0 | -0.0001 | 0.0013 | -0.0026 | -0.0017 | 0.0098 | 0.0003 | 0.0449 | 0.001 | -0.0018 | 0.0001 | -0.0044 | 0.003 | 0.1442 | 0.0248 | 0.0399 | 0.0043 | 0.0001 | 0.0016 | -0.0004 | 0.0001 | 1.08 | 1 | 1 | -0.0002 | 0.0009 | 0.0153 | -0.0005 | 0 | 0.0001 | -0.0008 | 0.0001 | 0.0001 | 0 | -0.0001 | 0.0013 | -0.0026 | -0.0017 | 0.0098 | 0.0003 | 0.0449 | 0.001 | -0.0018 | 0.0001 | -0.0044 | 0.003 | 0.1442 | 0.0248 | 0.0399 | 0.0001 | 0.0016 | -0.0004 | 0.0001
MQ-Blank-02 | Sample | 2024-09-16 15:02 | [6 empty] | 0 | -0.0016 | 0.0028 | -0.002 | -0.0001 | -0.0001 | 0.0011 | -0.0001 | -0.0002 | -0.0003 | -0.0001 | -0.0003 | -0.0034 | -0.0019 | 0.0088 | 0.0002 | 0.0026 | 0.0017 | -0.0016 | -0.0003 | -0.0027 | 0 | 0.1378 | 0.0145 | 0.0397 | -0.0012 | 0 | 0 | 0.0001 | -0.0003 | 1.08 | 1.02 | 1.02 | 0 | -0.0016 | 0.0028 | -0.002 | -0.0001 | -0.0001 | 0.0011 | -0.0001 | -0.0002 | -0.0003 | -0.0001 | -0.0003 | -0.0034 | -0.0019 | 0.0088 | 0.0002 | 0.0026 | 0.0017 | -0.0016 | -0.0003 | -0.0027 | 0 | 0.1378 | 0.0145 | 0.0397 | 0 | 0 | 0.0001 | -0.0003
Rinse | QC | 2024-09-16 15:04 | [6 empty] | -0.0003 | 0.0008 | 0.0021 | -0.0015 | -0.0001 | -0.0001 | 0.0025 | 0 | -0.0006 | -0.0001 | 0.0001 | 0.0006 | -0.0065 | -0.0023 | 0.0107 | 0.0001 | 0.0075 | 0.0011 | -0.0017 | -0.0013 | -0.0023 | -0.0017 | 0.1299 | 0.018 | 0.0402 | 0.0012 | 0.0001 | 0.0004 | 0 | 0.0004 | 1.07 | 1.01 | 1.01 | -0.0003 | 0.0008 | 0.0021 | -0.0015 | -0.0001 | -0.0001 | 0.0025 | 0 | -0.0006 | -0.0001 | 0.0001 | 0.0006 | -0.0065 | -0.0023 | 0.0107 | 0.0001 | 0.0075 | 0.0011 | -0.0017 | -0.0013 | -0.0023 | -0.0017 | 0.1299 | 0.018 | 0.0402 | 0.0001 | 0.0004 | 0 | 0.0004
CCV (5ppm All) | QC | 2024-09-16 15:07 | [6 empty] | 5.0254 | 4.9705 | 5.0374 | 5.0051 | 5.0447 | 4.9725 | 4.9481 | 4.8926 | 5.026 | 4.9481 | 5.0461 | 4.9693 | 4.9999 | 5.0903 | 4.7906 | 4.9404 | 5.2175 | 5.012 | 4.9324 | 5.131 | 5.2289 | 5.0032 | 4.6955 | 5.0185 | 5.001 | 4.9794 | 5.0029 | 4.92 | 5.0199 | 4.8589 | 1.06 | 0.99 | 1 | 5.0249 | 4.9705 | 5.0374 | 5.0051 | 5.0447 | 4.9725 | 4.9481 | 4.8925 | 5.0257 | 4.9467 | 5.0461 | 4.9693 | 4.9999 | 5.0903 | 4.7898 | 4.9404 | 5.2175 | 5.012 | 4.9324 | 5.1309 | 5.2282 | 4.9765 | 4.6955 | 5.0156 | 4.9991 | 5.0029 | 4.92 | 5.0155 | 4.8589
CCB | QC | 2024-09-16 15:09 | [6 empty] | -0.0001 | 0.0004 | 0.0116 | -0.0001 | -0.0001 | -0.0001 | 0.0003 | 0 | 0 | -0.0001 | 0.0001 | -0.0003 | -0.0008 | -0.0027 | 0.0091 | 0.0002 | 0.0447 | 0.0016 | -0.0009 | 0.0002 | -0.0049 | 0.0014 | 0.1595 | 0.0289 | 0.0399 | 0.002 | 0.0001 | 0.0018 | 0 | -0.0002 | 1.08 | 1.01 | 1.01 | -0.0001 | 0.0004 | 0.0116 | -0.0001 | -0.0001 | -0.0001 | 0.0003 | 0 | 0 | -0.0001 | 0.0001 | -0.0003 | -0.0008 | -0.0027 | 0.0091 | 0.0002 | 0.0447 | 0.0016 | -0.0009 | 0.0002 | -0.0049 | 0.0014 | 0.1595 | 0.0289 | 0.0399 | 0.0001 | 0.0018 | 0 | -0.0002
10x 0.5% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:12 | [6 empty] | -0.0001 | 16.7305 | 0.0047 | 0.0161 | 0.0027 | -0.0001 | 0.2078 | 0.0019 | 0.0018 | 0.0309 | 0.0002 | 165.192 | -0.0097 | 0.185 | 0.8824 | 1.1406 | 0.0135 | 5.1539 | 0.6358 | 1.1403 | 0.0018 | 62.0351 | 0.1104 | 0.0066 | 0.1719 | 0.0034 | 0.0005 | 0.099 | 0.0092 | 0.6053 | 1.08 | 1 | 1 | -0.0002 | 16.7305 | 0.0047 | 0.0161 | 0.0027 | -0.0001 | 0.2078 | -0.0006 | -0.0008 | 0.0305 | 0.002 | 165.192 | -0.0097 | 0.185 | 0.8543 | 1.1406 | 0.0135 | 5.1539 | 0.6358 | 1.1378 | -0.0013 | 62.0289 | 0.112 | 0.0028 | 0.1715 | 0.0005 | 0.099 | 0.0081 | 0.6053
10x 1% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:14 | [6 empty] | -0.0001 | 27.958 | 0.0007 | 0.0364 | 0.0028 | -0.0002 | 0.365 | 0.0027 | 0.0027 | 0.0493 | 0.001 | 232.2671 | 0.0034 | 0.2207 | 1.5367 | 1.9398 | 0.0086 | 14.4739 | 0.9754 | 1.455 | 0.0055 | 120.8919 | 0.1111 | 0.0021 | 0.2672 | 0.0029 | 0.0008 | 0.1934 | 0.0167 | 0.9293 | 1.07 | 0.99 | 1 | -0.0003 | 27.958 | 0.0007 | 0.0364 | 0.0028 | -0.0002 | 0.365 | -0.0008 | -0.0009 | 0.0487 | 0.0036 | 232.2671 | 0.0034 | 0.2207 | 1.4973 | 1.9398 | 0.0086 | 14.4739 | 0.9754 | 1.4515 | 0.0011 | 120.8814 | 0.1134 | -0.0035 | 0.2665 | 0.0008 | 0.1934 | 0.0151 | 0.9293
10x 5% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:17 | [6 empty] | 0.0005 | 12.8167 | 0.0022 | 0.1787 | 0.003 | -0.0002 | 1.8491 | 0.0144 | 0.0263 | 0.1844 | -0.0013 | 942.0008 | 0.0787 | 0.2548 | 7.6052 | 9.3401 | 0.0109 | 33.0439 | 1.7583 | 4.5589 | 0.0353 | 610.9668 o | 0.1142 | 0.0036 | 0.6003 | 0.0162 | 0.0037 | 0.9988 | 0.0585 | 5.1065 | 1.05 | 0.96 | 0.97 | -0.0004 | 12.8167 | 0.0022 | 0.1787 | 0.003 | -0.0002 | 1.8491 | 0.0003 | 0.0117 | 0.1818 | 0.009 | 942.0008 | 0.0787 | 0.2548 | 7.4451 | 9.3401 | 0.0109 | 33.0439 | 1.7583 | 4.5448 | 0.0171 | 610.9164 | 0.1236 | -0.0188 | 0.5968 | 0.0037 | 0.9988 | 0.052 | 5.1065
10x 10% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:19 | [6 empty] | 0.0017 | 1.1084 | 0.0044 | 0.3716 | 0.0049 | -0.0001 | 3.46 | 0.0276 | 0.0079 | 0.3907 | 0.0071 | 1803.3566 o | 0.1637 | -0.0001 | 14.6893 | 17.8596 | 0.0145 | 64.8239 | 0.8456 | 0.1855 | 0.0884 | 1161.4271 o | 0.1248 | -0.0016 | 0.7796 | 0.0339 | 0.0067 | 2.032 | 0.1408 | 12.3307 | 1.02 | 0.94 | 0.94 | -0.0001 | 1.1084 | 0.0044 | 0.3716 | 0.0049 | -0.0001 | 3.46 | 0.0006 | -0.0201 | 0.3857 | 0.0269 | 1803.3566 | 0.1637 | -0.0001 | 14.3827 | 17.8596 | 0.0145 | 64.8239 | 0.8456 | 0.1584 | 0.0536 | 1161.3306 | 0.1428 | -0.0442 | 0.7729 | 0.0067 | 2.032 | 0.1281 | 12.3307
10x 15% FeSO4 LFP Battery-TM | Sample | 2024-09-16 15:22 | [6 empty] | 0.0035 | 1.6543 | 0.0024 | 0.5727 | 0.0069 | -0.0001 | 5.0803 | 0.0424 | 0.0176 | 0.5815 | 0.015 | 2533.4012 o | 0.3017 | 0.0193 | 21.4508 | 26.9829 | 0.0205 | 97.0954 | 1.0375 | 0.279 | 0.1357 | 1657.8736 o | 0.163 | -0.0096 | 1.1526 | 0.0474 | 0.0106 | 2.9999 | 0.2196 | 18.3706 | 1 | 0.91 | 0.91 | 0.0008 | 1.6543 | 0.0024 | 0.5727 | 0.0069 | -0.0001 | 5.0803 | 0.0044 | -0.0217 | 0.574 | 0.0429 | 2533.4012 | 0.3017 | 0.0193 | 21.0202 | 26.9829 | 0.0205 | 97.0954 | 1.0375 | 0.241 | 0.0866 | 1657.7279 | 0.1884 | -0.0706 | 1.1423 | 0.0106 | 2.9999 | 0.2017 | 18.3706
10x 0.5% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:24 | [6 empty] | -0.0003 | 9.2607 | 0.0001 | 0.0195 | 0.0008 | -0.0002 | 0.2165 | 0.0018 | 0.0021 | 0.0338 | 0.003 | 166.8406 | -0.0054 | 0.0205 | 0.923 | 0.9895 | 0.0023 | 3.9598 | 1.2588 | 0.4553 | 0.0016 | 64.8838 | 0.1182 | 0.0026 | 0.0995 | 0.0027 | 0.0005 | 0.1114 | 0.0086 | 0.4791 | 1.08 | 0.99 | 1 | -0.0004 | 9.2607 | 0.0001 | 0.0195 | 0.0008 | -0.0002 | 0.2165 | -0.0007 | -0.0005 | 0.0335 | 0.0048 | 166.8406 | -0.0054 | 0.0205 | 0.8947 | 0.9895 | 0.0023 | 3.9598 | 1.2588 | 0.4528 | -0.0015 | 64.8784 | 0.1198 | -0.0011 | 0.0991 | 0.0005 | 0.1114 | 0.0075 | 0.4791
10x 1% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:27 | [6 empty] | -0.0003 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | 0.0033 | 0.0069 | 0.0603 | 0.0048 | 252.3585 | -0.005 | 0.0271 | 1.7296 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.3438 | 0.0037 | 126.2997 | 0.1207 | 0.0012 | 0.1512 | 0.0034 | 0.0008 | 0.2255 | 0.012 | 0.8442 | 1.07 | 0.99 | 1 | -0.0005 | 9.9687 | 0.0039 | 0.041 | 0.0005 | -0.0001 | 0.4137 | -0.0005 | 0.003 | 0.0598 | 0.0075 | 252.3585 | -0.005 | 0.0271 | 1.6867 | 1.8732 | 0.0032 | 7.7215 | 1.4011 | 0.34 | -0.0011 | 126.2896 | 0.1232 | -0.0045 | 0.1505 | 0.0008 | 0.2255 | 0.0103 | 0.8442
10x 5% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:29 | [6 empty] | 0.0003 | 10.346 | 0.0036 | 0.2226 | 0.0016 | -0.0002 | 1.9549 | 0.0146 | 0.0199 | 0.2327 | -0.0021 | 995.943 | 0.0957 | 0.0909 | 8.1096 | 9.5788 | 0.0086 | 33.5317 | 2.5013 | 0.4694 | 0.0409 | 641.9763 o | 0.1176 | -0.0021 | 0.4727 | 0.0192 | 0.0038 | 1.1323 | 0.0724 | 4.2772 | 1.05 | 0.96 | 0.97 | -0.0007 | 10.346 | 0.0036 | 0.2226 | 0.0016 | -0.0002 | 1.9549 | -0.0004 | 0.0045 | 0.2301 | 0.0089 | 995.943 | 0.0957 | 0.0909 | 7.9403 | 9.5788 | 0.0086 | 33.5317 | 2.5013 | 0.4545 | 0.0218 | 641.9245 | 0.1276 | -0.0256 | 0.469 | 0.0038 | 1.1323 | 0.0654 | 4.2772
10x 10% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:32 | [6 empty] | 0.0014 | 1.1497 | 0.0007 | 0.4192 | 0.0033 | -0.0002 | 3.8504 | 0.0332 | 0.0105 | 0.4355 | 0.0081 | 1867.6644 o | 0.246 | -0.0007 | 14.9854 | 18.5866 | 0.0157 | 62.795 | 2.4582 | 0.4391 | 0.0873 | 1190.8712 o | 0.1267 | 0.0068 | 0.8271 | 0.0356 | 0.0073 | 2.1431 | 0.146 | 9.1637 | 1.02 | 0.93 | 0.94 | -0.0005 | 1.1497 | 0.0007 | 0.4192 | 0.0033 | -0.0002 | 3.8504 | 0.0052 | -0.0184 | 0.4303 | 0.0286 | 1867.6644 | 0.246 | -0.0007 | 14.6679 | 18.5866 | 0.0157 | 62.795 | 2.4582 | 0.4111 | 0.0513 | 1190.7709 | 0.1454 | -0.0374 | 0.8201 | 0.0073 | 2.1431 | 0.1329 | 9.1637
10x 15% FeSO4 NMC Battery-TM | Sample | 2024-09-16 15:34 | [6 empty] | 0.0029 | 2.6273 | 0.0027 | 0.6048 | 0.0045 | -0.0001 | 5.4141 | 0.0419 | 0.0156 | 0.6211 | 0.0178 | 2550.4615 o | 0.3147 | 0.0679 | 21.3236 | 27.0173 | 0.021 | 91.9079 | 2.762 | 0.5423 | 0.1375 | 1668.3492 o | 0.1626 | -0.0146 | 1.1733 | 0.05 | 0.0105 | 3.0669 | 0.2109 | 14.6004 | 1 | 0.91 | 0.91 | 0.0002 | 2.6273 | 0.0027 | 0.6048 | 0.0045 | -0.0001 | 5.4141 | 0.0036 | -0.024 | 0.6136 | 0.0459 | 2550.4615 | 0.3147 | 0.0679 | 20.89 | 27.0173 | 0.021 | 91.9079 | 2.762 | 0.504 | 0.0881 | 1668.2033 | 0.1881 | -0.076 | 1.163 | 0.0105 | 3.0669 | 0.1928 | 14.6004
Rinse | QC | 2024-09-16 15:37 | [6 empty] | -0.0002 | 0.0007 | 0.0001 | -0.002 | -0.0001 | 0 | 0.0036 | 0 | -0.0002 | -0.0003 | -0.0003 | 0.0573 | -0.0003 | -0.0023 | 0.0078 | 0.0008 | 0.0003 | 0.0014 | -0.0017 | 0.0012 | -0.0048 | 0.06 | 0.1348 | 0.0109 | 0.0388 | 0.0003 | 0 | 0.0004 | 0.0001 | 0.0004 | 1.08 | 1.01 | 1.02 | -0.0002 | 0.0007 | 0.0001 | -0.002 | -0.0001 | 0 | 0.0036 | 0 | -0.0002 | -0.0003 | -0.0003 | 0.0573 | -0.0003 | -0.0023 | 0.0078 | 0.0008 | 0.0003 | 0.0014 | -0.0017 | 0.0012 | -0.0048 | 0.06 | 0.1348 | 0.0109 | 0.0388 | 0 | 0.0004 | 0.0001 | 0.0004
CCV (5ppm All) | QC | 2024-09-16 15:39 | [6 empty] | 4.9971 | 4.9582 | 5.0108 | 4.9704 | 5.0439 | 4.9613 | 4.9149 | 4.8515 | 4.9864 | 4.9152 | 5.0218 | 4.9991 | 4.9974 | 5.0844 | 4.7616 | 4.9498 | 5.1802 | 5.0091 | 4.8985 | 5.1152 | 5.197 | 5.0148 | 4.6549 | 4.9973 | 4.9632 | 4.9399 | 5.0072 | 4.8773 | 4.9826 | 4.8263 | 1.07 | 1 | 1 | 4.9966 | 4.9582 | 5.0108 | 4.9704 | 5.0439 | 4.9613 | 4.9149 | 4.8515 | 4.9861 | 4.9138 | 5.0218 | 4.9991 | 4.9974 | 5.0844 | 4.7607 | 4.9498 | 5.1802 | 5.0091 | 4.8985 | 5.1151 | 5.1963 | 4.9881 | 4.6549 | 4.9943 | 4.9613 | 5.0072 | 4.8773 | 4.9782 | 4.8263
CCB | QC | 2024-09-16 15:42 | [6 empty] | -0.0003 | 0.0012 | 0.0113 | -0.0006 | 0 | -0.0001 | -0.002 | 0.0001 | -0.0002 | 0 | -0.0002 | 0.0115 | 0.0061 | -0.0027 | 0.0087 | 0.0003 | 0.0396 | -0.0005 | -0.0017 | 0.0021 | -0.0049 | 0.0214 | 0.1544 | 0.0297 | 0.0408 | 0.002 | 0.0001 | 0.0021 | 0.0002 | 0.0002 | 1.08 | 1 | 1.01 | -0.0003 | 0.0012 | 0.0113 | -0.0006 | 0 | -0.0001 | -0.002 | 0.0001 | -0.0002 | 0 | -0.0002 | 0.0115 | 0.0061 | -0.0027 | 0.0087 | 0.0003 | 0.0396 | -0.0005 | -0.0017 | 0.0021 | -0.0049 | 0.0214 | 0.1544 | 0.0297 | 0.0408 | 0.0001 | 0.0021 | 0.0002 | 0.0002
09032024 MB | Sample | 2024-09-16 15:44 | [6 empty] | 0.0001 | 0.0065 | 0.0046 | -0.0008 | 0.0003 | 0 | 0.0054 | 0 | 0 | -0.0003 | 0.0001 | 0.0112 | -0.0022 | -0.0025 | 0.0165 | 0.0002 | 0.0108 | 0.0161 | -0.0001 | 0.007 | -0.0047 | 0.0203 | 0.1349 | 0.0211 | 0.0511 | 0.0004 | 0 | 0.0008 | 0.0002 | 0.0026 | 1.07 | 1 | 1 | 0.0001 | 0.0065 | 0.0046 | -0.0008 | 0.0003 | 0 | 0.0054 | 0 | 0 | -0.0003 | 0.0001 | 0.0112 | -0.0022 | -0.0025 | 0.0165 | 0.0002 | 0.0108 | 0.0161 | -0.0001 | 0.007 | -0.0047 | 0.0203 | 0.1349 | 0.0211 | 0.0511 | 0 | 0.0008 | 0.0002 | 0.0026
10x 09032024 MB | Sample | 2024-09-16 15:47 | [6 empty] | -0.0001 | 0.0021 | 0.0012 | -0.0022 | -0.0001 | -0.0001 | 0.0692 | 0 | -0.0002 | 0.0001 | 0.0095 | 0.0095 | 0.0043 | -0.0024 | 0.0092 | 0.0008 | 0.0038 | 0.0017 | -0.0019 | 0.0041 | -0.0038 | 0.0121 | 0.1222 | 0.014 | 0.0395 | -0.0008 | 0.0001 | 0.0004 | -0.0002 | 0.0043 | 1.08 | 1.01 | 1.01 | -0.0001 | 0.0021 | 0.0012 | -0.0022 | -0.0001 | -0.0001 | 0.0692 | 0 | -0.0002 | 0.0001 | 0.0095 | 0.0095 | 0.0043 | -0.0024 | 0.0092 | 0.0008 | 0.0038 | 0.0017 | -0.0019 | 0.0041 | -0.0038 | 0.0121 | 0.1222 | 0.014 | 0.0395 | 0.0001 | 0.0004 | -0.0002 | 0.0043
10x 09032024 MFB | Sample | 2024-09-16 15:50 | [6 empty] | 0.01 | 0.4842 | 0.2424 | 0.0482 | 0.2575 | 0.0488 | 0.5556 | 0.0972 | 0.5035 | 0.2469 | 0.2494 | 0.5049 | 0.4666 | -0.0026 | 0.481 | 0.1007 | 0.0969 | 0.5096 | 0.2426 | 0.1056 | 0.2554 | 0.255 | 0.2459 | 0.2493 | 0.1389 | 0.0985 | 0.1021 | 0.0984 | 0.2518 | 0.0959 | 1.08 | 1 | 1.02 | 0.01 | 0.4842 | 0.2424 | 0.0482 | 0.2575 | 0.0488 | 0.5556 | 0.0972 | 0.5035 | 0.2469 | 0.2494 | 0.5049 | 0.4666 | -0.0026 | 0.481 | 0.1007 | 0.0969 | 0.5096 | 0.2426 | 0.1056 | 0.2554 | 0.2545 | 0.2459 | 0.2492 | 0.1389 | 0.1021 | 0.0984 | 0.2517 | 0.0959
10x 1% FeSO4 NMC Battery-TM -MS -Analytical Duplicate | Sample | 2024-09-16 15:52 | [6 empty] | -0.0005 | 9.926 | 0.0023 | 0.0414 | 0.0005 | -0.0001 | 0.4151 | 0.0032 | 0.0048 | 0.06 | 0.0047 | 250.9591 | 0.0009 | 0.0273 | 1.7205 | 1.866 | 0.0054 | 7.6281 | 1.3942 | 0.3428 | 0.0029 | 125.1968 | 0.1045 | 0.0059 | 0.1475 | 0.0051 | 0.0008 | 0.2242 | 0.0158 | 0.837 | 1.08 | 1 | 1.01 | -0.0006 | 9.926 | 0.0023 | 0.0414 | 0.0005 | -0.0001 | 0.4151 | -0.0006 | 0.001 | 0.0595 | 0.0074 | 250.9591 | 0.0009 | 0.0273 | 1.6778 | 1.866 | 0.0054 | 7.6281 | 1.3942 | 0.339 | -0.0019 | 125.1867 | 0.107 | 0.0002 | 0.1468 | 0.0008 | 0.2242 | 0.0141 | 0.837
10x 1% FeSO4 NMC Battery-TM - Lab Duplicate | Sample | 2024-09-16 15:55 | [6 empty] | -0.0003 | 10.1338 | 0.0009 | 0.0418 | 0.0005 | -0.0001 | 0.4152 | 0.0035 | 0.0058 | 0.0607 | 0.0042 | 254.5683 | 0.0132 | 0.0276 | 1.7453 | 1.8836 | 0.0041 | 7.7345 | 1.4203 | 0.3473 | 0.0059 | 127.3406 | 0.1194 | 0.003 | 0.1514 | 0.0044 | 0.0008 | 0.2256 | 0.0122 | 0.8588 | 1.07 | 1 | 1.01 | -0.0005 | 10.1338 | 0.0009 | 0.0418 | 0.0005 | -0.0001 | 0.4152 | -0.0004 | 0.0019 | 0.0602 | 0.007 | 254.5683 | 0.0132 | 0.0276 | 1.702 | 1.8836 | 0.0041 | 7.7345 | 1.4203 | 0.3435 | 0.0011 | 127.3304 | 0.1219 | -0.0028 | 0.1507 | 0.0008 | 0.2256 | 0.0104 | 0.8588
10x 1% FeSO4- NMC Battery-TM - MS | Sample | 2024-09-16 15:57 | [6 empty] | 0.0088 | 9.0598 | 0.2474 | 0.0856 | 0.2559 | 0.0501 | 0.8491 | 0.1004 | 0.4455 | 0.2978 | 0.2616 | 226.1028 | 0.5031 | 0.0246 | 2.0337 | 1.7875 | 0.0984 | 7.4719 | 1.4852 | 0.4037 | 0.2625 | 113.1039 | 0.2874 | 0.2286 | 0.2473 | 0.1017 | 0.1025 | 0.2923 | 0.2674 | 0.8221 | 1.07 | 0.99 | 1 | 0.0086 | 9.0598 | 0.2474 | 0.0856 | 0.2559 | 0.0501 | 0.8491 | 0.097 | 0.4421 | 0.2973 | 0.264 | 226.1028 | 0.5031 | 0.0246 | 1.9952 | 1.7875 | 0.0984 | 7.4719 | 1.4852 | 0.4004 | 0.2582 | 113.0942 | 0.2897 | 0.2235 | 0.2466 | 0.1025 | 0.2923 | 0.2657 | 0.8221
10x 1% FeSO4- NMC Battery-TM - MS - Analytical Duplicate | Sample | 2024-09-16 16:00 | [6 empty] | 0.009 | 9.0013 | 0.2404 | 0.0851 | 0.2538 | 0.0495 | 0.8404 | 0.0997 | 0.4318 | 0.2952 | 0.2598 | 223.522 | 0.5033 | 0.0244 | 2.0157 | 1.7773 | 0.0995 | 7.4075 | 1.4732 | 0.405 | 0.2587 | 112.3233 | 0.2891 | 0.2276 | 0.2443 | 0.1011 | 0.1017 | 0.2906 | 0.265 | 0.8162 | 1.08 | 1 | 1 | 0.0088 | 9.0013 | 0.2404 | 0.0851 | 0.2538 | 0.0495 | 0.8404 | 0.0964 | 0.4283 | 0.2947 | 0.2623 | 223.522 | 0.5033 | 0.0244 | 1.9777 | 1.7773 | 0.0995 | 7.4075 | 1.4732 | 0.4017 | 0.2545 | 112.3137 | 0.2914 | 0.2225 | 0.2436 | 0.1017 | 0.2906 | 0.2634 | 0.8162
Rinse | QC | 2024-09-16 16:02 | [6 empty] | -0.0001 | -0.0005 | -0.0005 | -0.0027 | 0 | -0.0001 | 0.0025 | -0.0001 | -0.0001 | -0.0002 | 0.0001 | 0.0083 | -0.0045 | -0.0025 | 0.0091 | 0.0004 | 0.0023 | 0.0005 | -0.0009 | 0.0021 | -0.0031 | 0.011 | 0.1363 | 0.0118 | 0.0397 | 0.0004 | 0 | 0.0001 | -0.0002 | -0.0002 | 1.08 | 1.01 | 1.01 | -0.0001 | -0.0005 | -0.0005 | -0.0027 | 0 | -0.0001 | 0.0025 | -0.0001 | -0.0001 | -0.0002 | 0.0001 | 0.0083 | -0.0045 | -0.0025 | 0.0091 | 0.0004 | 0.0023 | 0.0005 | -0.0009 | 0.0021 | -0.0031 | 0.011 | 0.1363 | 0.0118 | 0.0397 | 0 | 0.0001 | -0.0002 | -0.0002
CCV (5ppm All) | QC | 2024-09-16 16:05 | [6 empty] | 4.9861 | 4.9762 | 4.993 | 4.9644 | 5.066 | 4.9907 | 4.9075 | 4.8469 | 4.986 | 4.907 | 5.0109 | 5.0163 | 5.0064 | 5.0822 | 4.7443 | 4.9688 | 5.1709 | 5.0046 | 4.8942 | 5.095 | 5.1965 | 4.9871 | 4.6235 | 4.9841 | 4.9414 | 4.9351 | 5.0144 | 4.8748 | 4.9793 | 4.818 | 1.08 | 0.99 | 1 | 4.9856 | 4.9762 | 4.993 | 4.9644 | 5.066 | 4.9907 | 4.9075 | 4.8468 | 4.9857 | 4.9056 | 5.0109 | 5.0163 | 5.0064 | 5.0822 | 4.7434 | 4.9688 | 5.1709 | 5.0046 | 4.8942 | 5.0949 | 5.1957 | 4.9602 | 4.6236 | 4.9812 | 4.9395 | 5.0144 | 4.8748 | 4.9748 | 4.818
CCB | QC | 2024-09-16 16:07 | [6 empty] | -0.0003 | 0 | 0.0141 | -0.0005 | -0.0001 | 0.0002 | -0.0022 | 0 | -0.0002 | -0.0002 | 0 | 0.0025 | -0.0046 | -0.0025 | 0.0091 | 0 | 0.0441 | -0.0004 | -0.0012 | -0.0001 | -0.0041 | 0.0084 | 0.1721 | 0.0292 | 0.0406 | 0.0009 | 0.0001 | 0.0017 | 0.0001 | 0 | 1.07 | 1.02 | 1.02 | -0.0003 | 0 | 0.0141 | -0.0005 | -0.0001 | 0.0002 | -0.0022 | 0 | -0.0002 | -0.0002 | 0 | 0.0025 | -0.0046 | -0.0025 | 0.0091 | 0 | 0.0441 | -0.0004 | -0.0012 | -0.0001 | -0.0041 | 0.0084 | 0.1721 | 0.0292 | 0.0406 | 0.0001 | 0.0017 | 0.0001 | 0
09162024-Lab Blank2 | Sample | 2024-09-16 16:10 | [6 empty] | -0.0002 | 0.0002 | 0.0028 | -0.002 | -0.0002 | -0.0001 | -0.0004 | 0 | 0 | -0.0001 | -0.0001 | -0.0006 | -0.0006 | -0.0028 | 0.0089 | 0.0003 | 0.0099 | 0.0006 | -0.0008 | 0.0029 | -0.0038 | 0.0026 | 0.1404 | 0.0175 | 0.0405 | -0.0004 | 0 | 0.0005 | 0.0001 | -0.0003 | 1.08 | 1.01 | 1.02 | -0.0002 | 0.0002 | 0.0028 | -0.002 | -0.0002 | -0.0001 | -0.0004 | 0 | 0 | -0.0001 | -0.0001 | -0.0006 | -0.0006 | -0.0028 | 0.0089 | 0.0003 | 0.0099 | 0.0006 | -0.0008 | 0.0029 | -0.0038 | 0.0026 | 0.1404 | 0.0175 | 0.0405 | 0 | 0.0005 | 0.0001 | -0.0003
10x 09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:12 | [6 empty] | 0.0096 | 0.4748 | 0.2451 | 0.0481 | 0.2503 | 0.0477 | 0.4792 | 0.0958 | 0.4968 | 0.2431 | 0.2458 | 0.4895 | 0.4567 | -0.0025 | 0.4719 | 0.099 | 0.0977 | 0.4998 | 0.2393 | 0.1027 | 0.2518 | 0.2495 | 0.2292 | 0.2467 | 0.1381 | 0.0965 | 0.0992 | 0.0975 | 0.2482 | 0.0947 | 1.08 | 1.01 | 1.02 | 0.0096 | 0.4748 | 0.2451 | 0.0481 | 0.2503 | 0.0477 | 0.4792 | 0.0958 | 0.4968 | 0.2431 | 0.2458 | 0.4895 | 0.4567 | -0.0025 | 0.4718 | 0.099 | 0.0977 | 0.4998 | 0.2393 | 0.1027 | 0.2517 | 0.249 | 0.2292 | 0.2467 | 0.138 | 0.0992 | 0.0975 | 0.2482 | 0.0947
10x 09162024-Lab Blank Spike2-Analytical Duplicate | Sample | 2024-09-16 16:15 | [6 empty] | 0.0095 | 0.4731 | 0.2451 | 0.0479 | 0.2495 | 0.0479 | 0.4755 | 0.095 | 0.4928 | 0.2416 | 0.244 | 0.4898 | 0.4622 | -0.0025 | 0.4709 | 0.0983 | 0.0975 | 0.4962 | 0.239 | 0.1038 | 0.2495 | 0.248 | 0.2699 | 0.2515 | 0.1385 | 0.0967 | 0.099 | 0.0971 | 0.2467 | 0.0941 | 1.08 | 1.01 | 1.02 | 0.0094 | 0.4731 | 0.2451 | 0.0479 | 0.2495 | 0.0479 | 0.4755 | 0.095 | 0.4928 | 0.2416 | 0.244 | 0.4898 | 0.4622 | -0.0025 | 0.4708 | 0.0983 | 0.0975 | 0.4962 | 0.239 | 0.1038 | 0.2495 | 0.2474 | 0.2699 | 0.2515 | 0.1385 | 0.099 | 0.0971 | 0.2466 | 0.0941
5x 09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:17 | [6 empty] | 0.0191 | 0.9536 | 0.4833 | 0.0956 | 0.4984 | 0.0957 | 0.9469 | 0.1902 | 0.9818 | 0.4804 | 0.4857 | 0.9772 | 0.9523 | -0.004 | 0.932 | 0.1966 | 0.1898 | 0.995 | 0.4745 | 0.2007 | 0.4971 | 0.4895 | 0.4248 | 0.4918 | 0.2327 | 0.1929 | 0.1967 | 0.1917 | 0.4908 | 0.1871 | 1.09 | 1.01 | 1.01 | 0.0191 | 0.9536 | 0.4833 | 0.0956 | 0.4984 | 0.0957 | 0.9469 | 0.1902 | 0.9818 | 0.4803 | 0.4857 | 0.9772 | 0.9523 | -0.004 | 0.9318 | 0.1966 | 0.1898 | 0.995 | 0.4745 | 0.2007 | 0.4971 | 0.4885 | 0.4249 | 0.4916 | 0.2326 | 0.1967 | 0.1917 | 0.4906 | 0.1871
09162024-Lab Blank Spike2 | Sample | 2024-09-16 16:20 | [6 empty] | 0.0983 | 4.8272 | 2.5113 | 0.496 | 2.5174 | 0.4893 | 4.835 | 0.9662 | 4.9876 | 2.4261 | 2.4645 | 4.9443 | 4.9978 | -0.0021 | 4.7939 | 0.9926 | 0.9949 | 5.0132 | 2.4267 | 1.0327 | 2.5601 | 2.5077 | 2.2541 | 2.5116 | 1.0182 | 0.997 | 0.9985 | 0.9771 | 2.4884 | 0.9598 | 1.08 | 1 | 1.01 | 0.0982 | 4.8272 | 2.5113 | 0.496 | 2.5174 | 0.4893 | 4.835 | 0.9661 | 4.9875 | 2.4259 | 2.4646 | 4.9443 | 4.9978 | -0.0021 | 4.793 | 0.9926 | 0.9949 | 5.0132 | 2.4267 | 1.0326 | 2.5598 | 2.5024 | 2.2541 | 2.5109 | 1.0178 | 0.9985 | 0.9771 | 2.4875 | 0.9598
09162024-Lab Blank Spike2-Analytical Duplicate | Sample | 2024-09-16 16:22 | [6 empty] | 0.0984 | 4.7796 | 2.53 | 0.496 | 2.5 | 0.4859 | 4.8357 | 0.9666 | 4.9704 | 2.4282 | 2.4547 | 4.9051 | 4.9558 | -0.0021 | 4.7919 | 0.9789 | 1.0113 | 4.9854 | 2.4277 | 1.0325 | 2.5649 | 2.5062 | 2.295 | 2.5121 | 1.0232 | 0.9934 | 0.9917 | 0.9779 | 2.4737 | 0.9611 | 1.07 | 1.01 | 1.01 | 0.0983 | 4.7796 | 2.53 | 0.496 | 2.5 | 0.4859 | 4.8357 | 0.9665 | 4.9703 | 2.4279 | 2.4548 | 4.9051 | 4.9558 | -0.0021 | 4.7911 | 0.9789 | 1.0113 | 4.9854 | 2.4277 | 1.0324 | 2.5647 | 2.5009 | 2.295 | 2.5114 | 1.0228 | 0.9917 | 0.9779 | 2.4727 | 0.9611
Rinse | QC | 2024-09-16 16:25 | [6 empty] | -0.0002 | -0.0015 | 0.0134 | -0.002 | 0.0001 | 0 | 0.0036 | 0 | 0.0002 | -0.0002 | 0 | 0.0002 | -0.0098 | -0.0024 | 0.0111 | 0.0002 | 0.0246 | -0.0001 | -0.0017 | 0.0007 | -0.0013 | 0.0072 | 0.1759 | 0.0233 | 0.0404 | 0.0024 | 0.0001 | 0.0005 | 0.0001 | 0.0004 | 1.07 | 1 | 1.01 | -0.0002 | -0.0015 | 0.0134 | -0.002 | 0.0001 | 0 | 0.0036 | 0 | 0.0002 | -0.0002 | 0 | 0.0002 | -0.0098 | -0.0024 | 0.0111 | 0.0002 | 0.0246 | -0.0001 | -0.0017 | 0.0007 | -0.0013 | 0.0072 | 0.1759 | 0.0233 | 0.0404 | 0.0001 | 0.0005 | 0.0001 | 0.0004
0.5ppm All (LLQC) | QC | 2024-09-16 16:27 | [6 empty] | 0.4943 | 0.4939 | 0.4907 | 0.4903 | 0.5025 | 0.4924 | 0.5195 | 0.4821 | 0.4982 | 0.4898 | 0.4962 | 0.4955 | 0.4766 | 0.505 | 0.4722 | 0.4941 | 0.4989 | 0.5167 | 0.4874 | 0.5087 | 0.5119 | 0.4869 | 0.4226 | 0.4925 | 0.5307 | 0.487 | 0.4991 | 0.4848 | 0.4961 | 0.4915 | 1.08 | 1.01 | 1.02 | 0.4942 | 0.4939 | 0.4907 | 0.4903 | 0.5025 | 0.4924 | 0.5195 | 0.4821 | 0.4982 | 0.4896 | 0.4962 | 0.4955 | 0.4766 | 0.505 | 0.4721 | 0.4941 | 0.4989 | 0.5167 | 0.4874 | 0.5087 | 0.5118 | 0.4843 | 0.4226 Q | 0.4922 | 0.5305 | 0.4991 | 0.4848 | 0.4956 | 0.4915
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-09-16 16:30 | [6 empty] | 0.4928 | 0.4921 | 0.4896 | 0.4892 | 0.5029 | 0.4909 | 0.5145 | 0.4803 | 0.4961 | 0.4882 | 0.4955 | 0.493 | 0.4742 | 0.5003 | 0.4708 | 0.4976 | 0.5035 | 0.5128 | 0.4864 | 0.5046 | 0.5096 | 0.4912 | 0.4436 Q | 0.4988 | 0.5271 | 0.4852 | 0.4996 | 0.4834 | 0.4944 | 0.4898 | 1.09 | 1.01 | 1.01 | 0.4927 | 0.4921 | 0.4896 | 0.4892 | 0.5029 | 0.4909 | 0.5145 | 0.4803 | 0.4961 | 0.488 | 0.4955 | 0.493 | 0.4742 | 0.5003 | 0.4707 | 0.4976 | 0.5035 | 0.5128 | 0.4864 | 0.5046 | 0.5095 | 0.4885 | 0.4436 | 0.4985 | 0.5269 | 0.4996 | 0.4834 | 0.494 | 0.4898
CCS1 (SS 5ppm) | QC | 2024-09-16 16:32 | [6 empty] | 5.1048 | 5.0418 | 5.15 | 4.9823 | 5.109 | 5.0224 | 4.9623 | 4.8964 | 5.0269 | 4.9401 | 5.0617 | 5.0121 | 5.054 | 5.02 | 4.8555 | 5.0096 | 5.3433 | 5.0553 | 4.9368 | 5.2083 | 5.2707 | 5.0333 | 4.6883 | 5.061 | 4.8492 | 5.0673 | 5.0967 | 4.9774 | 5.0661 | 4.8537 | 1.08 | 0.99 | 1 | 5.1043 | 5.0418 | 5.15 | 4.9823 | 5.109 | 5.0224 | 4.9623 | 4.8963 | 5.0266 | 4.9387 | 5.0618 | 5.0121 | 5.054 | 5.02 | 4.8546 | 5.0096 | 5.3433 | 5.0553 | 4.9368 | 5.2082 | 5.27 | 5.0062 | 4.6884 | 5.058 | 4.8473 | 5.0967 | 4.9774 | 5.0616 | 4.8537
CCB | QC | 2024-09-16 16:35 | [6 empty] | -0.0001 | 0.0004 | 0.0159 | -0.0009 | 0 | 0 | 0.0012 | 0 | -0.0002 | -0.0002 | 0 | -0.0005 | -0.002 | -0.0024 | 0.0092 | 0.0001 | 0.0426 | -0.0002 | -0.0011 | 0.0001 | -0.0021 | 0.0012 | 0.15 | 0.0258 | 0.0414 | 0.0051 | 0.0001 | 0.0016 | -0.0002 | -0.0001 | 1.08 | 1 | 1.01 | -0.0001 | 0.0004 | 0.0159 | -0.0009 | 0 | 0 | 0.0012 | 0 | -0.0002 | -0.0002 | 0 | -0.0005 | -0.002 | -0.0024 | 0.0092 | 0.0001 | 0.0426 | -0.0002 | -0.0011 | 0.0001 | -0.0021 | 0.0012 | 0.15 | 0.0258 | 0.0414 | 0.0001 | 0.0016 | -0.0002 | -0.0001
MQ-Blank-03 | Sample | 2024-09-16 16:37 | [6 empty] | 0.0001 | 0.001 | -0.0005 | -0.0027 | -0.0001 | 0 | -0.0001 | -0.0002 | 0 | -0.0003 | -0.0004 | 0.0012 | -0.0011 | -0.002 | 0.0098 | 0.0002 | 0.0015 | 0.0001 | -0.0022 | 0.0008 | -0.0028 | -0.0005 | 0.1433 | 0.0124 | 0.039 | -0.0018 | 0 | 0 | -0.0003 | -0.0002 | 1.1 | 1.02 | 1.03 | 0.0001 | 0.001 | -0.0005 | -0.0027 | -0.0001 | 0 | -0.0001 | -0.0002 | 0 | -0.0003 | -0.0004 | 0.0012 | -0.0011 | -0.002 | 0.0098 | 0.0002 | 0.0015 | 0.0001 | -0.0022 | 0.0008 | -0.0028 | -0.0005 | 0.1433 | 0.0124 | 0.039 | 0 | 0 | -0.0003 | -0.0002