## 081224ICP_TO178-444 Sample Data
08122024_ICP-OES_TO 178-444 Sample Data
ANALYSIS TECHNIUE: |  | ICP-OES |  |  |  | Calibration date: |  | 2024-08-12
ANALYSIS METHOD: |  | 30M |  |  |  | Calibration Standards: |  | 2024-08-08
ANALYSIS BATCH: |  | TO 058-16 + TO 178-444 |  |  |  | Task Order# |  | 68HERC22F0178
Task Order: |  | TO 178 |  |  |  | Instrument : |  | Agilent 5900 SVDV ICP-OES (S/N: MY2138CP13) with Agilent SPS 4 Autosampler (S/N: AU213011544)
UNITS: |  | mg/L (ppm) |  |  |  | Analyst: |  | Mahendranath Arambewela
Comments: | [5 empty] | SOP# |  | K-LRTD-SOP-1185-2
Sample results are reported as mg/L (ppm)
Worksheet exported from: C:\Users\USEPA\Documents\Agilent\ICP Expert\My Results\08122024_ICP-OES_TO 058-16 + TO 178-444.esws
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Solution Label | Type | Date Time | Ag 328.068 C ppm | Al 396.152 C ppm | As 188.980 C ppm | B 249.678 C ppm | Ba 455.403 C ppm | Be 313.107 C ppm | Ca 370.602 C ppm | Cd 214.439 C ppm | Co 238.892 C ppm | Cr 267.716 C ppm | Cu 324.754 C ppm | Fe 259.940 C ppm | K 766.491 C ppm | Li 670.783 C ppm | Mg 279.078 C ppm | Mn 257.610 C ppm | Mo 202.032 C ppm | Na 588.995 C ppm | Ni 231.604 C ppm | P 178.222 C ppm | Pb 220.353 C ppm | S 181.972 C ppm | Sb 217.582 C ppm | Se 196.026 C ppm | Si 251.611 C ppm | Sr 421.552 C ppm | Ti 334.941 C ppm | V 292.401 C ppm | Zn 206.200 C ppm
08122024 LB | Sample | 2024-08-12 13:15 | -0.0003 | -0.0028 | 0.0055 | 0.0004 | 0 | 0.0001 | 0.0117 | 0.0001 | 0.0001 | 0 | 0 | 0.0036 | -0.0031 | -0.0121 | -0.0096 | 0.0002 | 0.0163 | 0.0001 | -0.0029 | 0.0057 | -0.0038 | 0.0131 | 0.2045 | 0.0121 | 0.0053 | 0 | 0.0005 | -0.0004 | 0.0024
10x 08122024 LCS | Sample | 2024-08-12 13:18 | 0.0091 | 0.4808 | 0.2381 | 0.049 | 0.2464 | 0.0484 | 0.4819 | 0.097 | 0.4828 | 0.2427 | 0.2364 | 0.5098 | 0.4665 | -0.0031 | 0.4748 | 0.1016 | 0.09 | 0.5027 | 0.2372 | 0.105 | 0.2364 | 0.2413 | 0.2395 | 0.2451 | 0.0953 | 0.099 | 0.096 | 0.24 | 0.0955
10x 08122024 LCS- Analytical Duplicate | Sample | 2024-08-12 13:20 | 0.0096 | 0.4706 | 0.2391 | 0.0478 | 0.2449 | 0.0478 | 0.4823 | 0.0967 | 0.4839 | 0.2431 | 0.2358 | 0.5085 | 0.4641 | -0.0035 | 0.4727 | 0.1003 | 0.0919 | 0.4959 | 0.238 | 0.1112 | 0.2349 | 0.2455 | 0.272 | 0.2403 | 0.0951 | 0.0981 | 0.0959 | 0.2402 | 0.096
5x 08122024 LCS | Sample | 2024-08-12 13:23 | 0.0188 | 0.9633 | 0.4861 | 0.0982 | 0.4967 | 0.0972 | 0.9646 | 0.196 | 0.9805 | 0.4875 | 0.4754 | 1.0228 | 0.9475 | -0.0056 | 0.9728 | 0.2039 | 0.1772 | 1.0081 | 0.4832 | 0.2067 | 0.4828 | 0.4827 | 0.4015 | 0.481 | 0.1909 | 0.1984 | 0.1928 | 0.4842 | 0.1926
08122024 LCS | Sample | 2024-08-12 13:25 | 0.0951 | 4.7787 | 2.4618 | 0.4929 | 2.4434 | 0.4802 | 4.8454 | 0.985 | 4.9104 | 2.4231 | 2.3844 | 5.0502 | 4.9365 | -0.0039 | 4.9593 | 1.0057 | 0.9308 | 4.9235 | 2.437 | 1.0304 | 2.4452 | 2.418 | 2.2314 | 2.4796 | 0.9748 | 0.9806 | 0.9708 | 2.435 | 0.9729
08122024 LCS - Analytical Duplicate | Sample | 2024-08-12 13:28 | 0.094 | 4.779 | 2.4543 | 0.4892 | 2.4386 | 0.4797 | 4.8193 | 0.9718 | 4.8742 | 2.4055 | 2.3732 | 5.0517 | 4.8974 | -0.0045 | 4.9228 | 1.0044 | 0.9468 | 4.9058 | 2.421 | 1.0196 | 2.4296 | 2.4032 | 1.8238 | 2.4647 | 0.9709 | 0.978 | 0.9649 | 2.4248 | 0.967
0.5% NaCl LFP Battery-TM | Sample | 2024-08-12 13:58 | -0.0012 | 557.7988 | 0.0258 | 0.0046 | 0.0923 | -0.0001 | 2.6321 | 0.0019 | 0.0114 | 0.143 | 0.0216 | 1314.2552 | 0.6112 | 2.4354 | 1.5195 | 5.0189 | 0.037 | 1471.3982 | 17.2464 | 57.286 | 0.0063 | 2.5667 | 0.2012 | -0.0521 | 3.8912 | 0.0362 | 0.139 | 0.0832 | 0.2196
1% NaCl LFP Battery-TM | Sample | 2024-08-12 14:00 | -0.0017 | 467.9349 | 0.0215 | 0.011 | 0.0733 | -0.0001 | 4.6387 | 0.0032 | 0.0006 | 0.1437 | 0.0259 | 989.5368 | 1.0757 | 5.0366 | 1.2138 | 4.4929 | 0.0272 |  | 13.531 | 111.2263 | 0.0068 | 3.4047 | 0.2015 | -0.0365 | 5.4906 | 0.0503 | 0.1286 | 0.0682 | 0.1464
5% NaCl LFP Battery-TM | Sample | 2024-08-12 14:03 | -0.0016 | 814.1376 | 0.0251 | 0.0328 | 0.0692 | -0.0002 | 18.1943 | 0.003 | 0.0032 | 0.2392 | 0.0786 | 858.3883 | 5.3255 | 23.4352 | 1.4624 | 3.4276 | 0.0422 |  | 13.4173 | 196.4268 | 0.0057 | 15.3393 | 0.205 | -0.0417 | 9.7089 | 0.2306 | 0.2101 | 0.1067 | 0.0789
10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:05 | -0.0017 | 1293.5693 | 0.0355 | 0.0604 | 0.0904 | -0.0002 | 33.5257 | -0.0002 | 0.0034 | 0.3419 | 0.1374 | 1055.2288 | 13.6907 | 32.0557 | 2.2245 | 3.8145 | 0.0519 |  | 17.1599 | 159.2554 | 0.0041 | 29.3639 | 0.302 | -0.0609 | 10.0706 | 0.4585 | 0.318 | 0.1462 | 0.1659
15% NaCl LFP Battery-TM | Sample | 2024-08-12 14:08 | -0.0014 | 963.621 | 0.0262 | 0.1545 | 0.0046 | -0.0003 | 46.0159 | 0.0015 | 0.0008 | 0.3604 | 0.1651 | 753.6171 | 24.654 | 14.3553 | 4.7087 | 0.7849 | 0.0206 |  | 10.7765 | 16.5137 | -0.0045 | 41.614 | 0.2007 | -0.0331 | 0.7093 | 0.677 | 0.2669 | 0.0657 | 0.0409
0.5% NaCl MMC Battery-TM | Sample | 2024-08-12 14:10 | -0.0007 | 456.0864 | 0.0159 | 0.0858 | 0.0079 | -0.0002 | 2.7328 | 0.004 | -0.0061 | 0.2566 | 0.0556 | 1202.9261 | 0.7907 | 1.2715 | 3.2923 | 1.0584 | 0.0114 | 1666.8729 | 13.7941 | 18.9552 | 0.0044 | 1.7617 | 0.2311 | -0.055 | 0.5751 | 0.0324 | 0.1634 | 0.0449 | 0.0646
1% NaCl MMC Battery-TM | Sample | 2024-08-12 14:13 | -0.0004 | 656.2899 | 0.0243 | 0.1069 | 0.0083 | -0.0001 | 4.0848 | 0.0036 | 0.0092 | 0.3919 | 0.1357 | 1271.6013 | 1.1073 | 1.3916 | 4.9223 | 1.1126 | 0.0174 |  | 15.8358 | 22.503 | 0.0029 | 2.9429 | 0.2691 | -0.0593 | 0.6418 | 0.0542 | 0.2544 | 0.0666 | 0.0698
5% NaCl MMC Battery-TM | Sample | 2024-08-12 14:15 | -0.0015 | 793.9437 | 0.0196 | 0.1193 | 0.006 | -0.0002 | 15.8604 | 0.004 | 0.185 | 0.4084 | 0.1636 | 971.2853 | 7.1104 | 8.1921 | 4.9497 | 0.8952 | 0.0178 |  | 13.0225 | 23.9881 | -0.0011 | 13.5099 | 0.249 | -0.0527 | 0.6662 | 0.2582 | 0.2724 | 0.0683 | 0.0548
10% NaCl MMC Battery-TM | Sample | 2024-08-12 14:18 | -0.0014 | 1049.6128 | 0.0364 | 0.1497 | 0.0045 | -0.0002 | 37.1117 | 0.0025 | 0.0143 | 0.4244 | 0.1993 | 900.7167 | 17.0243 | 15.7913 | 5.1979 | 0.8891 | 0.0191 |  | 12.539 | 20.2247 | -0.0034 | 33.0031 | 0.2312 | -0.0555 | 0.7378 | 0.5309 | 0.3155 | 0.0769 | 0.0485
15% NaCl MMC Battery-TM | Sample | 2024-08-12 14:20 | -0.0016 | 958.1671 | 0.0176 | 0.0771 | 0.0719 | -0.0001 | 45.0598 | 0.005 | 0.0053 | 0.1875 | 0.0846 | 730.2844 | 24.2192 | 30.2284 | 1.8596 | 3.2229 | 0.0328 |  | 11.9989 | 102.5226 | -0.0074 | 40.7388 | 0.2175 | -0.0551 | 8.7413 | 0.6679 | 0.2083 | 0.0876 | 0.064
08092024 MB | Sample | 2024-08-12 14:33 | -0.0005 | 0.0018 | 0.0079 | -0.0001 | 0 | 0 | -0.0091 | 0.0002 | 0.0001 | -0.0001 | 0.0005 | 0.0072 | -0.001 | -0.0037 | -0.0098 | 0.0002 | 0.014 | 0.4276 | -0.0023 | 0.0084 | -0.0027 | 0.007 | 0.1698 | 0.0075 | 0.0144 | 0 | 0.0005 | -0.0001 | 0.0011
1008092024 MFB | Sample | 2024-08-12 14:35 | 0.0955 | 4.8215 | 2.4388 | 0.4803 | 2.4493 | 0.4779 | 4.8377 | 0.9756 | 4.9165 | 2.4226 | 2.4159 | 5.0866 | 4.9341 | -0.0121 | 4.9412 | 1.0065 | 0.9871 | 5.5818 | 2.4371 | 1.0079 | 2.4329 | 2.3612 | 2.1057 | 2.3829 | 1.0051 | 0.9825 | 0.9772 | 2.4441 | 0.9651
08092024 MFB - Analyticsal Duplicate | Sample | 2024-08-12 14:38 | 0.0948 | 4.7914 | 2.4428 | 0.4807 | 2.4307 | 0.4767 | 4.8348 | 0.9699 | 4.8704 | 2.4239 | 2.4111 | 5.0631 | 4.8962 | -0.004 | 4.9375 | 1.0018 | 0.9995 | 5.5088 | 2.4317 | 1.0115 | 2.432 | 2.3437 | 2.1615 | 2.3907 | 1.0066 | 0.9764 | 0.9766 | 2.4348 | 0.9642
5x 08092024 MFB | Sample | 2024-08-12 14:40 | 0.0188 | 0.9539 | 0.4915 | 0.0947 | 0.4846 | 0.0947 | 0.9759 | 0.193 | 0.9692 | 0.4822 | 0.4712 | 1.0115 | 0.9427 | -0.007 | 0.9587 | 0.1998 | 0.1988 | 1.2827 | 0.4788 | 0.2066 | 0.4784 | 0.471 | 0.459 | 0.4883 | 0.2018 | 0.1935 | 0.1914 | 0.4791 | 0.1898
5x 10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:43 | -0.0008 | 215.6443 | 0.0122 | 0.0078 | 0.0176 | 0 | 7.4977 | 0.0001 | 0.0015 | 0.079 | 0.0257 | 243.4776 | 1.5368 | 3.0259 | 0.4649 | 0.8884 | 0.0221 |  | 4.1748 | 35.9493 | -0.0002 | 6.7007 | 0.153 | -0.0125 | 1.1647 | 0.0894 | 0.0671 | 0.0324 | 0.0436
5x 10% NaCl LFP Battery-TM - MS | Sample | 2024-08-12 14:45 | 0.0176 | 189.7969 | 0.5183 | 0.106 | 0.4979 | 0.0908 | 7.6228 | 0.1866 | 0.8229 | 0.5387 | 0.572 | 215.9512 | 2.5462 | 2.6037 | 1.3208 | 0.971 | 0.2052 |  | 4.0047 | 32.264 | 0.4631 | 6.4754 | 0.5718 | 0.4763 | 1.3654 | 0.2731 | 0.2521 | 0.5194 | 0.2224
10% NaCl LFP Battery-TM - Analytical Duplicate | Sample | 2024-08-12 14:48 | -0.0018 | 1291.7171 | 0.0276 | 0.0566 | 0.0902 | -0.0003 | 33.4245 | 0.0036 | 0.0035 | 0.3409 | 0.1377 | 1058.1629 | 13.5663 | 31.5973 | 2.2124 | 3.8145 | 0.052 |  | 17.0984 | 158.9981 | 0.011 | 29.428 | 0.2068 | -0.0515 | 9.1782 | 0.4576 | 0.3173 | 0.152 | 0.1655
10% NaCl LFP Battery-TM - Lab Duplicate | Sample | 2024-08-12 14:50 | -0.0017 | 1307.6594 | 0.0302 | 0.056 | 0.0905 | -0.0003 | 33.3458 | 0.004 | -0.0001 | 0.3441 | 0.137 | 1063.5783 | 13.9791 | 31.9609 | 2.2103 | 3.8301 | 0.047 | 2435.7052 | 16.914 | 158.4421 | 0.0061 | 29.3356 | 0.2646 | -0.0649 | 8.6894 | 0.459 | 0.3177 | 0.1492 | 0.1639
10% NaCl LFP Battery-TM - MS | Sample | 2024-08-12 14:53 | 0.1095 | 1125.5413 | 2.3395 | 0.5103 | 2.5982 | 0.406 | 34.1408 | 0.7543 | 3.4841 | 2.354 | 3.0146 | 952.7033 | 20.8246 | 26.5031 | 5.5548 | 4.2462 | 0.926 |  | 16.6532 | 143.6606 | 1.8626 | 28.9446 | 2.3307 | 2.0901 | 9.5887 | 1.4172 | 1.2005 | 2.3862 | 0.8689
10% NaCl LFP Battery-TM - MS - Analytical Duplicate | Sample | 2024-08-12 14:55 | 0.11 | 1121.6721 | 2.3542 | 0.5118 | 2.5932 | 0.4086 | 34.1838 | 0.7539 | 3.4883 | 2.3576 | 3.0374 | 954.4947 | 20.6754 | 26.3993 | 5.5434 | 4.238 | 0.9317 |  | 16.7073 | 143.7686 | 1.8719 | 28.9841 | 2.3846 | 2.102 | 9.5116 | 1.4139 | 1.2022 | 2.3926 | 0.868
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Lowest Cal Std |  |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Highest Cal Std |  |  | 20 | 200 | 20 | 20 | 50 | 20 | 2000 | 20 | 20 | 50 | 100 | 1000 | 100 | 100 | 2000 | 50 | 20 | 100 | 50 | 200 | 200 | 100 | 50 | 20 | 100 | 20 | 50 | 20 | 100
Correlation |  |  | 0.99997 | 0.99997 | 0.99997 | 1 | 1 | 0.99999 | 0.99987 | 1 | 0.99998 | 0.99999 | 0.99999 | 0.99998 | 0.99998 | 0.99999 | 0.99996 | 0.99988 | 0.99973 | 0.99998 | 0.99999 | 0.9999 | 0.99974 | 0.99964 | 0.9998 | 0.99992 | 0.99933 | 0.99999 | 0.99996 | 0.99999 | 0.99999
ICP-OES MDL |  |  | 0.0045598 | 0.0048715 | 0.0060725 | 0.0016216 | 0.00011879 | 0.000168 | 0.0038865 | 0.00024729 | 0.00024729 | 0.0007 | 0.00076374 | 0.0031683 | 0.032873 | 0.0015923 | 0.0080477 | 0.0008057 | 0.00050865 | 0.001351 | 0.0024937 | 0.0072662 | 0.0049325 | 0.0058515 | 0.0056274 | 0.0064431 | 0.002358 | 0.000168 | 0.00011879 | 0.00093287 | 0.00059397
ICP-OES MRL/PRDL |  |  | 0.045598 | 0.048715 | 0.060725 | 0.016216 | 0.0011879 | 0.00168 | 0.038865 | 0.0024729 | 0.0024729 | 0.0074 | 0.0076374 | 0.031683 | 0.32873 | 0.015923 | 0.080477 | 0.008057 | 0.0050865 | 0.01351 | 0.024937 | 0.072662 | 0.049325 | 0.058515 | 0.056274 | 0.064431 | 0.02358 | 0.00168 | 0.0011879 | 0.0093287 | 0.0059397
Where
Calculated/Estimated MDLs & MRLs
Dilution required
* | Sampling Error
** | Different concentration than the normal
Mislabeled

## Lithium vs Phosphorus
| Lithium | Phosphorus
0.5% NaCl LFP Battery-TM | 2.4354
1% NaCl LFP Battery-TM | 5.0366
5% NaCl LFP Battery-TM | 23.4352
10% NaCl LFP Battery-TM | 32.0557
15% NaCl LFP Battery-TM | 14.3553
0.5% NaCl MMC Battery-TM | 1.2715 | 18.9552
1% NaCl MMC Battery-TM | 1.3916 | 22.503
5% NaCl MMC Battery-TM | 8.1921 | 23.9881
10% NaCl MMC Battery-TM | 15.7913 | 20.2247
15% NaCl MMC Battery-TM | 30.2284 | 102.5226
0.5% NaOH LFP Battery-TM | 4.02 | 16.946
1% NaOH LFP Battery-TM | 15.259 | 46.786
5% NaOH LFP Battery-TM | 27.127 | 72.243
10% NaOH LFP Battery-TM | 35.499 | 51.269
15% NaOH LFP Battery-TM | 32.944 | 48.455
0.5% NaOH NMC Battery-TM | 5.141 | 1.9
1% NaOH NMC Battery-TM | 12.131 | 11.518
5% NaOH NMC Battery-TM | 23.011 | 17.791
10% NaOH NMC Battery-TM | 27.686 | 24.112
15% NaOH NMC Battery-TM | 26.426 | 21.233
0.5% FeSO4 LFP Battery-TM | 0.1939 | 10.9318 |  | Lithium | Phosphorus
1% FeSO4 LFP Battery-TM | 0.2957 | 14.1726 | 0.5% NaCl MMC Battery-TM | 1.2715 | 18.9552
5% FeSO4 LFP Battery-TM | 1.1755 | 43.2472 | 1% NaCl MMC Battery-TM | 1.3916 | 22.503
10% FeSO4 LFP Battery-TM | 2.5841 | 1.5305 | 5% NaCl MMC Battery-TM | 8.1921 | 23.9881
15% FeSO4 LFP Battery-TM | 3.2673 | 2.2225 | 10% NaCl MMC Battery-TM | 15.7913 | 20.2247
0.5% FeSO4 MMC Battery-TM | 0.1521 | 3.9465 | 15% NaCl MMC Battery-TM | 30.2284 | 102.5226
1% FeSO4 MMC Battery-TM | 0.2966 | 3.0423 | 0.5% NaOH NMC Battery-TM | 5.141 | 1.9
5% FeSO4 MMC Battery-TM | 1.37 | 4.3408 | 1% NaOH NMC Battery-TM | 12.131 | 11.518
10% FeSO4 MMC Battery-TM | 3.3313 | 3.8218 | 5% NaOH NMC Battery-TM | 23.011 | 17.791
15% FeSO4 MMC Battery-TM | 4.3804 | 4.6762 | 10% NaOH NMC Battery-TM | 27.686 | 24.112
|  |  | 15% NaOH NMC Battery-TM | 26.426 | 21.233
|  |  | 0.5% FeSO4 MMC Battery-TM | 0.1521 | 3.9465
|  |  | 1% FeSO4 MMC Battery-TM | 0.2966 | 3.0423
|  |  | 5% FeSO4 MMC Battery-TM | 1.37 | 4.3408
|  |  | 10% FeSO4 MMC Battery-TM | 3.3313 | 3.8218
|  |  | 15% FeSO4 MMC Battery-TM | 4.3804 | 4.6762

## 081224ICP-OES_TO178-444 Data
08122024_ICP-OES_TO 178-444 Data Report
ANALYSIS TECHNIUE: |  | ICP-OES |  |  |  | Calibration date: |  | 2024-08-12
ANALYSIS METHOD: |  | 30M |  |  |  | Calibration Standards: |  | 2024-08-08
ANALYSIS BATCH: |  | TO 058-16 + TO 178-444 |  |  |  | Task Order# |  | 68HERC22F0178
Task Order: |  | TO 178 |  |  |  | Instrument : |  | Agilent 5900 SVDV ICP-OES (S/N: MY2138CP13) with Agilent SPS 4 Autosampler (S/N: AU213011544)
UNITS: |  | mg/L (ppm) |  |  |  | Analyst: |  | Mahendranath Arambewela
Comments: | [5 empty] | SOP# |  | K-LRTD-SOP-1185-2
Sample results are reported as mg/L (ppm)
Worksheet exported from: C:\Users\USEPA\Documents\Agilent\ICP Expert\My Results\08122024_ICP-OES_TO 058-16 + TO 178-444.esws
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
|  |  | Ag (328.068 nm) ppm | Al (396.152 nm) ppm | As (188.980 nm) ppm | B (249.678 nm) ppm | Ba (455.403 nm) ppm | Be (313.107 nm) ppm | Ca (370.602 nm) ppm | Cd (214.439 nm) ppm | Co (238.892 nm) ppm | Cr (267.716 nm) ppm | Cu (324.754 nm) ppm | Fe (259.940 nm) ppm | K (766.491 nm) ppm | Li (670.783 nm) ppm | Mg (279.078 nm) ppm | Mn (257.610 nm) ppm | Mo (202.032 nm) ppm | Na (588.995 nm) ppm | Ni (231.604 nm) ppm | P (178.222 nm) ppm | Pb (220.353 nm) ppm | S (181.972 nm) ppm | Sb (217.582 nm) ppm | Se (196.026 nm) ppm | Si (251.611 nm) ppm | Sr (421.552 nm) ppm | Ti (334.941 nm) ppm | V (292.401 nm) ppm | Zn (206.200 nm) ppm
IECS1 (Al 1000 Int.) | IECS |  | Al: 1000
IECS2 (Fe 1000 Int.) | IECS |  | Fe: 1000
IECS3 (Mn 100 Int.) | IECS |  | Mn: 100
IECS4 (Ca Mg 1000 Int.) | IECS |  | Ca: 1000; Mg: 1000
IECS5 (Ti 20 Int.) | IECS |  | Ti: 20
ICV 0.5ppm All | QC |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
ICV 5ppm All | QC |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
CCS1 (SS 5ppm) | QC |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
CCV2 5ppm All | QC |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
ICV 1ppm All | QC | bar: Ag to Zn
|  | MDL | 0.0045598 | 0.0048715 | 0.0060725 | 0.0016216 | 0.00011879 | 0.000168 | 0.0038865 | 0.00024729 | 0.00024729 | 0.00038187 | 0.00076374 | 0.0031683 | 0.032873 | 0.0015923 | 0.0080477 | 0.0008057 | 0.00050865 | 0.001351 | 0.0024937 | 0.0072662 | 0.0049325 | 0.0058515 | 0.0056274 | 0.0064431 | 0.002358 | 0.000168 | 0.00011879 | 0.00093287 | 0.00059397
|  | MRL | 0.045598 | 0.048715 | 0.060725 | 0.016216 | 0.0011879 | 0.00168 | 0.038865 | 0.0024729 | 0.0024729 | 0.0038187 | 0.0076374 | 0.031683 | 0.32873 | 0.015923 | 0.080477 | 0.008057 | 0.0050865 | 0.01351 | 0.024937 | 0.072662 | 0.049325 | 0.058515 | 0.056274 | 0.064431 | 0.02358 | 0.00168 | 0.0011879 | 0.0093287 | 0.0059397
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Solution Label | Type | Date Time | Ag 328.068 C ppm | Al 396.152 C ppm | As 188.980 C ppm | B 249.678 C ppm | Ba 455.403 C ppm | Be 313.107 C ppm | Ca 370.602 C ppm | Cd 214.439 C ppm | Co 238.892 C ppm | Cr 267.716 C ppm | Cu 324.754 C ppm | Fe 259.940 C ppm | K 766.491 C ppm | Li 670.783 C ppm | Mg 279.078 C ppm | Mn 257.610 C ppm | Mo 202.032 C ppm | Na 588.995 C ppm | Ni 231.604 C ppm | P 178.222 C ppm | Pb 220.353 C ppm | S 181.972 C ppm | Sb 217.582 C ppm | Se 196.026 C ppm | Si 251.611 C ppm | Sr 421.552 C ppm | Ti 334.941 C ppm | V 292.401 C ppm | Zn 206.200 C ppm
Calibration Blank | Blank | 2024-08-12 09:19 | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 !
Calibration Standard 0.5ppm | Standard | 2024-08-12 09:21 | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 !
Calibration Standard 1ppm | Standard | 2024-08-12 09:24 | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 !
Calibration Standard 5ppm | Standard | 2024-08-12 09:26 | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 !
Calibration Standard 10ppm | Standard | 2024-08-12 09:29 | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 !
Calibration Standard 20ppm | Standard | 2024-08-12 09:31 | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 !
Calibration Standard 50ppm | Standard | 2024-08-12 09:34 |  | 50.0000 ! |  |  | 50.0000 ! |  | 50.0000 ! |  |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! |  | 50.0000 ! |  | 50.0000 !
Calibration Standard-A 100ppm | Standard | 2024-08-12 09:36 | Al: 100.0000 !; Ca: 100.0000 !; Cu: 100.0000 !; Fe: 100.0000 !; K: 100.0000 !; Li: 100.0000 !; Mg: 100.0000 !; Na: 100.0000 !; Pb: 100.0000 !; Si: 100.0000 !; Zn: 100.0000 !
Calibration Standard-B 100ppm | Standard | 2024-08-12 09:39 | P: 100.0000 !; S: 100.0000 !
Calibration Standard Al+Fe+Ca+Mg200ppm | Standard | 2024-08-12 09:41 | Al: 200.0000 !; Ca: 200.0000 !; Fe: 200.0000 !; Mg: 200.0000 !
Calibration Standard Ca+Mg500ppm | Standard | 2024-08-12 09:44 | Ca: 500.0000 !; Mg: 500.0000 !
Calibration Standard Ca+Mg1000ppm | Standard | 2024-08-12 09:47 | Ca: 1000.0000 !; Mg: 1000.0000 !
Calibration Standard Ca+Mg2000ppm | Standard | 2024-08-12 09:49 | Ca: 2000.0000 !; Mg: 2000.0000 !
Calibration Standard P+S200ppm | Standard | 2024-08-12 09:52 | P: 200.0000 !; S: 200.0000 !
Calibration Standard Pb200ppm | Standard | 2024-08-12 09:54 | Pb: 200.0000 !
Calibration Standard Fe 1000ppm | Standard | 2024-08-12 09:57 | Fe: 1000.0000 !
Rinse1 | QC | 2024-08-12 09:59 | -0.0001 | -0.0031 | 0.0024 | 0.0095 | -0.0001 | 0 | -0.0063 | 0.0002 | -0.0003 | 0.0003 | -0.0003 | 0.0239 | -0.0058 | -0.0047 | -0.0035 | 0.0002 | 0.004 | -0.0083 | -0.0016 | 0.0057 | -0.003 | 0.013 | 0.2725 | 0.0064 | 0.0037 | 0 | 0.0001 | -0.0002 | 0.0003
Rinse | QC | 2024-08-12 10:02 | -0.0002 | -0.0039 | 0.0018 | 0.0084 | 0 | 0 | -0.0065 | 0.0002 | 0.0003 | -0.0002 | 0 | 0.005 | 0.0081 | -0.0111 | -0.0058 | -0.0001 | 0.0031 | -0.007 | -0.0014 | 0.0061 | -0.0012 | 0.0099 | 0.2681 | 0.0016 | 0.0027 | 0 | 0 | -0.0003 | 0.0005
IECS1 (Al 1000 Int.) | QC | 2024-08-12 10:04 | -0.0003 | 958.4498 | -0.0043 | 0.0099 | 0.0092 | 0 | 0.0046 | 0.0006 | 0.0001 | 0.0005 | 0.002 | 0.0079 | -0.0088 | -0.005 | 0.0043 | 0 | 0.0065 | 0.0002 | -0.0015 | 0.012 | 0.0343 | 0.0187 | 0.6129 K | -0.0078 | 0.009 | 0.0001 | 0.001 | 0.0004 | 0.0035
% Recovery |  |  | Al: 95.84%
Rinse1 | QC | 2024-08-12 10:07 | -0.0002 | 0.0073 | 0.0012 | 0.0062 | 0.0001 | 0 | -0.0068 | 0 | 0 | 0.0001 | 0.0001 | 0.0033 | -0.0022 | -0.0116 | -0.0035 | 0.0001 | 0.0026 | -0.0047 | -0.0012 | 0.0061 | -0.0026 | 0.0072 | 0.2523 | 0.0009 | 0.0014 | 0 | 0.0001 | 0.0002 | 0.0001
Rinse | QC | 2024-08-12 10:09 | 0.0002 | 0.0022 | 0.0027 | 0.0058 | 0 | -0.0001 | -0.0076 | 0.0002 | 0.0001 | 0.0002 | 0.0006 | 0.0004 | -0.0015 | -0.0028 | -0.0047 | 0 | 0.0024 | -0.0069 | -0.0006 | 0.0077 | -0.0024 | 0.0066 | 0.26 | 0.0083 | 0.0012 | 0 | 0.0001 | 0.0001 | 0.0005
IECS2 (Fe 1000 Int.) | QC | 2024-08-12 10:12 | 0.0001 | 0.0243 | 0.001 | -0.0059 | -0.0001 | 0 | -0.0204 | -0.0034 | 0.0059 | 0.0084 | -0.0126 | 1022.0713 | -0.0052 | -0.0033 | 0.9082 K | 0.0092 | 0.0014 | -0.0083 | 0.0071 | 0.0044 | -0.0032 | 0.0151 | 0.171 | -0.0513 | -0.0037 | 0.0005 | 0.0001 | -0.0046 | 0.0026
% Recovery |  |  | Fe: 102.2%
Rinse1 | QC | 2024-08-12 10:14 | -0.0004 | 0.0017 | 0.0026 | 0.0038 | 0 | 0.0001 | -0.0087 | 0.0001 | -0.0005 | 0.0001 | -0.0006 | 0.0184 | 0.0026 | -0.0048 | -0.0081 | -0.0001 | 0.0017 | -0.0071 | -0.0009 | 0.0044 | -0.0033 | 0.0115 | 0.2718 | 0.0082 | 0.0019 | 0 | 0 | 0.0002 | -0.0002
Rinse | QC | 2024-08-12 10:17 | -0.0001 | -0.0055 | 0 | 0.0035 | 0 | 0 | -0.0113 | 0.0001 | 0.0001 | -0.0004 | 0 | 0.0056 | -0.012 | -0.0022 | -0.0079 | 0 | 0.0016 | -0.0075 | -0.0012 | 0.0073 | -0.0041 | 0.0104 | 0.2552 | 0.0038 | 0.0013 | 0 | 0.0001 | -0.0002 | 0.0003
IECS3 (Mn 100 Int.) | QC | 2024-08-12 10:19 | -0.002 | -0.0005 | 0.0005 | -0.0012 | 0 | 0 | 0.7062 K | 0.0001 | -0.0013 | -0.0132 | 0.0008 | 0.006 | -0.0111 | -0.0053 | -0.0068 | 100.3548 | 0.0014 | -0.0086 | -0.0023 | 0.0049 | -0.0002 | -0.0428 | 0.2854 | 0.0328 | 0.006 | 0 | 0 | -0.0006 | 0.0002
% Recovery |  |  | Mn: 100.4%
IECS4 (Ca Mg 1000 Int.) | QC | 2024-08-12 10:22 | -0.0004 | 0.0523 | 0.005 | 0.0071 | 0.0019 | 0 | 985.224 | 0.0086 | 0.0003 | 0.0014 | 0.0019 | 0.0051 | -0.014 | -0.0005 | 997.0819 | 0.0095 | 0.0011 | 0.0063 | 0.0026 | 0.0068 | 0.0032 | 0.0556 | 0.3062 | 0.0032 | 0.0198 | 0.0077 | -0.0003 | 0.0015 | 0.0122
% Recovery |  |  | Ca: 98.52%; Mg: 99.71%
Rinse1 | QC | 2024-08-12 10:24 | -0.0003 | -0.0034 | 0.0042 | 0.0033 | 0 | 0 | 0.0002 | 0 | 0 | -0.0002 | -0.0007 | 0.0024 | 0.0029 | -0.0022 | 0.0025 | 0.0005 | 0.0019 | -0.0069 | -0.0011 | 0.0057 | -0.0026 | 0.0072 | 0.2733 | 0.0036 | 0.0016 | 0 | 0 | 0.0001 | -0.0003
Rinse | QC | 2024-08-12 10:27 | 0 | -0.0037 | 0.0029 | 0.0024 | 0.0001 | 0.0001 | -0.0044 | 0 | 0.0003 | 0 | 0.0001 | 0.0016 | -0.0009 | -0.0016 | -0.0045 | -0.0002 | 0.0023 | -0.0064 | -0.0014 | 0.0057 | -0.002 | 0.0077 | 0.2835 | 0.0065 | 0.0006 | 0 | 0.0001 | -0.0001 | 0.0005
IECS5 (Ti 20 Int.) | QC | 2024-08-12 10:29 | 0.0003 | 0.0011 | 0.0023 | 0.0017 | -0.0003 | -0.0004 | 0.0882 | 0.0001 | 0.0004 | -0.0001 | -0.0011 | 0.0016 | -0.0002 | -0.0041 | -0.0027 | 0.0005 | 0.0025 | -0.0067 | -0.0028 | 0.0037 | -0.0064 | 0.0113 | 0.2589 | 0.0073 | 0.0053 | 0 | 19.2902 K | -0.0016 | -0.0001
% Recovery
Rinse1 | QC | 2024-08-12 10:32 | -0.0003 | -0.0059 | 0.0032 | 0.0018 | -0.0001 | 0 | -0.0065 | -0.0001 | 0 | -0.0001 | 0.0001 | 0.0009 | 0.0015 | -0.0015 | -0.0055 | 0.0003 | 0.0024 | -0.0081 | -0.0022 | 0.0069 | -0.0018 | 0.0087 | 0.2606 | 0.0057 | 0.0021 | 0 | 0.0017 | 0.0001 | 0.0001
Rinse | QC | 2024-08-12 10:34 | -0.0003 | -0.0004 | 0.0017 | 0.002 | 0 | 0.0001 | -0.006 | 0 | 0.0002 | 0.0001 | 0.0006 | 0.0006 | 0.0043 | -0.0014 | -0.0047 | 0.0002 | 0.0012 | -0.0101 | -0.002 | 0.0037 | -0.0041 | 0.0079 | 0.2656 | 0.0084 | 0.0022 | 0 | 0.0004 | -0.0001 | 0.0003
ICB | QC | 2024-08-12 10:37 | -0.0002 | -0.0028 | 0.0017 | 0.0009 | 0 | 0 | -0.0099 | 0.0001 | 0 | 0.0001 | -0.0002 | 0.0005 | -0.0085 | -0.004 | -0.0076 | 0.0001 | 0.0011 | -0.0059 | -0.002 | 0.0056 | -0.0016 | 0.0091 | 0.2739 | 0.0069 | 0.0002 | 0 | 0.0002 | -0.0005 | -0.0003
ICV (SS 5ppm All) | QC | 2024-08-12 10:40 | 4.9161 | 4.9594 | 4.9733 | 4.8884 | 4.9278 | 4.907 | 4.9228 | 4.9016 | 4.8858 | 4.8772 | 4.8224 | 5.0993 | 4.8883 | 4.8949 | 4.9812 | 5.0631 | 5.0078 | 4.9085 | 4.8967 | 4.9619 | 4.9388 | 4.7556 | 3.9935 Q | 4.9407 | 4.7517 | 4.9519 | 4.8977 | 4.911 | 4.834
% Recovery |  |  | 98.32% | 99.19% | 99.47% | 97.77% | 98.56% | 98.14% | 98.46% | 98.03% | 97.72% | 97.54% | 96.45% | 102% | 97.77% | 97.9% | 99.62% | 101.3% | 100.2% | 98.17% | 97.93% | 99.24% | 98.78% | 95.11% |  | 98.81% | 95.03% | 99.04% | 97.95% | 98.22% | 96.68%
Rinse | QC | 2024-08-12 10:42 | -0.0005 | -0.0027 | 0.0156 | 0.0023 | 0.0003 | 0 | -0.0115 | 0.0001 | 0.0002 | 0 | 0.0002 | 0.0014 | 0.0025 | -0.0019 | -0.0072 | 0 | 0.0515 | -0.0078 | -0.0005 | 0.0074 | -0.0049 | 0.005 | 0.2104 | 0.0279 | 0.0019 | 0.0001 | 0.0014 | 0.0004 | 0.0005
CCB | QC | 2024-08-12 10:45 | 0.0001 | -0.0033 | 0.0069 | 0.0021 | 0 | -0.0001 | -0.0087 | 0.0001 | 0.0002 | 0.0003 | 0.0007 | 0.0016 | 0.0056 | -0.0024 | -0.0081 | 0 | 0.015 | -0.0087 | -0.0016 | 0.0068 | -0.0041 | 0.0136 | 0.2181 | 0.0149 | 0.0017 | 0 | 0.0004 | 0 | 0.0001
0.5ppm All (LLQC) | QC | 2024-08-12 10:48 | 0.4774 | 0.4891 | 0.4742 | 0.4982 | 0.488 | 0.4825 | 0.4795 | 0.4851 | 0.4875 | 0.4866 | 0.4735 | 0.5105 | 0.4291 Q | 0.4893 | 0.4733 | 0.5067 | 0.4554 | 0.5004 | 0.4854 | 0.4873 | 0.4878 | 0.4663 | 0.4515 | 0.4711 | 0.4772 | 0.4898 | 0.4755 | 0.4775 | 0.4772
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-12 10:50 | 0.4739 | 0.4874 | 0.474 | 0.4981 | 0.4867 | 0.4825 | 0.4804 | 0.484 | 0.4873 | 0.4862 | 0.4734 | 0.5087 | 0.4556 | 0.4888 | 0.4711 | 0.5065 | 0.4693 | 0.5019 | 0.4846 | 0.4855 | 0.4836 | 0.4692 | 0.4838 | 0.4833 | 0.4772 | 0.4887 | 0.4747 | 0.4771 | 0.476
RPD |  |  | 0.7358% | 0.3482% | 0.04219% | 0.02007% | 0.2667% | 0% | 0.1875% | 0.227% | 0.04103% | 0.08224% | 0.02112% | 0.3532% | #VALUE! | 1.626% | 0.4659% | 0.03948% | 3.006% | 0.2993% | 0.1649% | 0.3701% | 0.8647% | 0.62% | 6.907% | 2.557% | 0% | 0.2248% | 0.1684% | 0.0838% | 0.2518%
0.5ppm All (LLQC) | QC | 2024-08-12 10:48 | 0.4774 | 0.4891 | 0.4742 | 0.4982 | 0.488 | 0.4825 | 0.4795 | 0.4851 | 0.4875 | 0.4866 | 0.4735 | 0.5105 | 0.4291 Q | 0.4893 | 0.4733 | 0.5067 | 0.4554 | 0.5004 | 0.4854 | 0.4873 | 0.4878 | 0.4663 | 0.4515 | 0.4711 | 0.4772 | 0.4898 | 0.4755 | 0.4775 | 0.4772
% Recovery |  |  | 95.48% | 97.82% | 94.84% | 99.64% | 97.6% | 96.5% | 95.9% | 97.02% | 97.5% | 97.32% | 94.7% | 102.1% |  | 97.86% | 94.66% | 101.3% | 91.08% | 100.1% | 97.08% | 97.46% | 97.56% | 93.26% | 90.3% | 94.22% | 95.44% | 97.96% | 95.1% | 95.5% | 95.44%
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-12 10:50 | 0.4739 | 0.4874 | 0.474 | 0.4981 | 0.4867 | 0.4825 | 0.4804 | 0.484 | 0.4873 | 0.4862 | 0.4734 | 0.5087 | 0.4556 | 0.4888 | 0.4711 | 0.5065 | 0.4693 | 0.5019 | 0.4846 | 0.4855 | 0.4836 | 0.4692 | 0.4838 | 0.4833 | 0.4772 | 0.4887 | 0.4747 | 0.4771 | 0.476
% Recovery |  |  | 94.78% | 97.48% | 94.8% | 99.62% | 97.34% | 96.5% | 96.08% | 96.8% | 97.46% | 97.24% | 94.68% | 101.7% | 91.12% | 97.76% | 94.22% | 101.3% | 93.86% | 100.4% | 96.92% | 97.1% | 96.72% | 93.84% | 96.76% | 96.66% | 95.44% | 97.74% | 94.94% | 95.42% | 95.2%
CCV (5ppm All) | QC | 2024-08-12 10:53 | 4.8226 | 4.8687 | 4.8814 | 4.9065 | 4.8478 | 4.841 | 4.8818 | 4.8899 | 4.8927 | 4.8737 | 4.8237 | 5.0318 | 4.8233 | 4.9498 | 4.8899 | 4.9886 | 4.9505 | 4.8363 | 4.8853 | 4.9433 | 4.9105 | 4.7589 | 4.0935 | 4.9177 | 4.9706 | 4.858 | 4.8225 | 4.8501 | 4.855
% Recovery |  |  | 96.45% | 97.37% | 97.63% | 98.13% | 96.96% | 96.82% | 97.64% | 97.8% | 97.85% | 97.47% | 96.47% | 100.6% | 96.47% | 99% | 97.8% | 99.77% | 99.01% | 96.73% | 97.71% | 98.87% | 98.21% | 95.18% | 81.87% | 98.35% | 99.41% | 97.16% | 96.45% | 97% | 97.1%
CCB | QC | 2024-08-12 10:55 | -0.0001 | -0.0034 | 0.0186 | 0.0034 | 0.0001 | 0.0001 | -0.0121 | 0.0001 | 0.0003 | -0.0002 | 0.0001 | 0.0009 | -0.005 | -0.0026 | -0.0077 | 0 | 0.0522 | -0.0087 | -0.002 | 0.009 | -0.0014 | 0.0065 | 0.2233 | 0.0257 | 0.0025 | 0.0001 | 0.0016 | -0.0003 | -0.0002
Rinse | QC | 2024-08-12 13:08 | -0.0003 | -0.002 | 0.0012 | -0.0012 | -0.0001 | 0.0001 | -0.0105 | 0.0001 | 0.0004 | -0.0002 | -0.0005 | 0.0007 | -0.0009 | -0.0033 | -0.0105 | 0 | 0.0017 | -0.0074 | -0.0019 | 0.007 | -0.0033 | 0.0083 | 0.2475 | 0.0038 | -0.0003 | 0 | -0.0001 | -0.0002 | 0.0003
CCV (5ppm All) | QC | 2024-08-12 13:10 | 4.8177 | 4.8901 | 4.8594 | 4.8725 | 4.8541 | 4.8565 | 4.8712 | 4.8738 | 4.8726 | 4.8585 | 4.826 | 5.0357 | 4.8209 | 4.9536 | 4.862 | 4.9944 | 4.9168 | 4.8352 | 4.8786 | 5.0668 | 4.8838 | 4.7855 | 4.8472 | 4.9068 | 4.8931 | 4.8586 | 4.817 | 4.842 | 4.84
% Recovery |  |  | 96.35% | 97.8% | 97.19% | 97.45% | 97.08% | 97.13% | 97.42% | 97.48% | 97.45% | 97.17% | 96.52% | 100.7% | 96.42% | 99.07% | 97.24% | 99.89% | 98.34% | 96.7% | 97.57% | 101.3% | 97.68% | 95.71% | 96.94% | 98.14% | 97.86% | 97.17% | 96.34% | 96.84% | 96.8%
CCB | QC | 2024-08-12 13:13 | -0.0001 | -0.0027 | 0.0173 | 0.001 | 0.0002 | 0.0001 | -0.0143 | 0.0001 | 0 | 0.0001 | -0.0006 | 0.0002 | 0.0017 | -0.0024 | -0.0116 | 0.0001 | 0.0578 | -0.0102 | -0.0024 | 0.0009 | -0.0037 | 0.0022 | 0.226 | 0.0309 | 0.0017 | 0.0001 | 0.0018 | 0 | 0
08122024 LB | Sample | 2024-08-12 13:15 | -0.0003 | -0.0028 | 0.0055 | 0.0004 | 0 | 0.0001 | 0.0117 | 0.0001 | 0.0001 | 0 | 0 | 0.0036 | -0.0031 | -0.0121 | -0.0096 | 0.0002 | 0.0163 | 0.0001 | -0.0029 | 0.0057 | -0.0038 | 0.0131 | 0.2045 | 0.0121 | 0.0053 | 0 | 0.0005 | -0.0004 | 0.0024
10x 08122024 LCS | Sample | 2024-08-12 13:18 | 0.0091 | 0.4808 | 0.2381 | 0.049 | 0.2464 | 0.0484 | 0.4819 | 0.097 | 0.4828 | 0.2427 | 0.2364 | 0.5098 | 0.4665 | -0.0031 | 0.4748 | 0.1016 | 0.09 | 0.5027 | 0.2372 | 0.105 | 0.2364 | 0.2413 | 0.2395 | 0.2451 | 0.0953 | 0.099 | 0.096 | 0.24 | 0.0955
10x 08122024 LCS- Analytical Duplicate | Sample | 2024-08-12 13:20 | 0.0096 | 0.4706 | 0.2391 | 0.0478 | 0.2449 | 0.0478 | 0.4823 | 0.0967 | 0.4839 | 0.2431 | 0.2358 | 0.5085 | 0.4641 | -0.0035 | 0.4727 | 0.1003 | 0.0919 | 0.4959 | 0.238 | 0.1112 | 0.2349 | 0.2455 | 0.272 | 0.2403 | 0.0951 | 0.0981 | 0.0959 | 0.2402 | 0.096
RPD |  |  | 5.348% | 2.144% | 0.4191% | 2.479% | 0.6106% | 1.247% | 0.08297% | 0.3098% | 0.2276% | 0.1647% | 0.2541% | 0.2553% | 0.5158% | 1.626% | 0.4433% | 1.288% | 2.089% | 1.362% | 0.3367% | 5.735% | 0.6365% | 1.726% | 12.71% | 1.978% | 0.2101% | 0.9132% | 0.1042% | 0.0833% | 0.5222%
08122024 LB | Sample | 2024-08-12 13:15 | -0.0003 | -0.0028 | 0.0055 | 0.0004 | 0 | 0.0001 | 0.0117 | 0.0001 | 0.0001 | 0 | 0 | 0.0036 | -0.0031 | -0.0121 | -0.0096 | 0.0002 | 0.0163 | 0.0001 | -0.0029 | 0.0057 | -0.0038 | 0.0131 | 0.2045 | 0.0121 | 0.0053 | 0 | 0.0005 | -0.0004 | 0.0024
10x 08122024 LCS | Sample | 2024-08-12 13:18 | 0.0091 | 0.4808 | 0.2381 | 0.049 | 0.2464 | 0.0484 | 0.4819 | 0.097 | 0.4828 | 0.2427 | 0.2364 | 0.5098 | 0.4665 | -0.0031 | 0.4748 | 0.1016 | 0.09 | 0.5027 | 0.2372 | 0.105 | 0.2364 | 0.2413 | 0.2395 | 0.2451 | 0.0953 | 0.099 | 0.096 | 0.24 | 0.0955
R% |  |  | 94% | 96.72% | 93.04% | 97.2% | 98.56% | 96.6% | 94.04% | 96.9% | 96.54% | 97.08% | 94.56% | 101.2% | 93.92% |  | 96.88% | 101.4% | 73.7% | 100.5% | 96.04% | 99.3% | 96.08% | 91.28% | 14% | 93.2% | 90% | 99% | 95.5% | 96.16% | 93.1%
08122024 LB | Sample | 2024-08-12 13:15 | -0.0003 | -0.0028 | 0.0055 | 0.0004 | 0 | 0.0001 | 0.0117 | 0.0001 | 0.0001 | 0 | 0 | 0.0036 | -0.0031 | -0.0121 | -0.0096 | 0.0002 | 0.0163 | 0.0001 | -0.0029 | 0.0057 | -0.0038 | 0.0131 | 0.2045 | 0.0121 | 0.0053 | 0 | 0.0005 | -0.0004 | 0.0024
10x 08122024 LCS- Analytical Duplicate | Sample | 2024-08-12 13:20 | 0.0096 | 0.4706 | 0.2391 | 0.0478 | 0.2449 | 0.0478 | 0.4823 | 0.0967 | 0.4839 | 0.2431 | 0.2358 | 0.5085 | 0.4641 | -0.0035 | 0.4727 | 0.1003 | 0.0919 | 0.4959 | 0.238 | 0.1112 | 0.2349 | 0.2455 | 0.272 | 0.2403 | 0.0951 | 0.0981 | 0.0959 | 0.2402 | 0.096
R% |  |  | 99% | 94.68% | 93.44% | 94.8% | 97.96% | 95.4% | 94.12% | 96.6% | 96.76% | 97.24% | 94.32% | 101% | 93.44% |  | 96.46% | 100.1% | 75.6% | 99.16% | 96.36% | 105.5% | 95.48% | 92.96% | 27% | 91.28% | 89.8% | 98.1% | 95.4% | 96.24% | 93.6%
08122024 LB | Sample | 2024-08-12 13:15 | -0.0003 | -0.0028 | 0.0055 | 0.0004 | 0 | 0.0001 | 0.0117 | 0.0001 | 0.0001 | 0 | 0 | 0.0036 | -0.0031 | -0.0121 | -0.0096 | 0.0002 | 0.0163 | 0.0001 | -0.0029 | 0.0057 | -0.0038 | 0.0131 | 0.2045 | 0.0121 | 0.0053 | 0 | 0.0005 | -0.0004 | 0.0024
5x 08122024 LCS | Sample | 2024-08-12 13:23 | 0.0188 | 0.9633 | 0.4861 | 0.0982 | 0.4967 | 0.0972 | 0.9646 | 0.196 | 0.9805 | 0.4875 | 0.4754 | 1.0228 | 0.9475 | -0.0056 | 0.9728 | 0.2039 | 0.1772 | 1.0081 | 0.4832 | 0.2067 | 0.4828 | 0.4827 | 0.4015 | 0.481 | 0.1909 | 0.1984 | 0.1928 | 0.4842 | 0.1926
R% |  |  | 95.5% | 96.61% | 96.12% | 97.8% | 99.34% | 97.1% | 95.29% | 97.95% | 98.04% | 97.5% | 95.08% | 101.9% | 95.06% |  | 98.24% | 101.8% | 80.45% | 100.8% | 97.22% | 100.5% | 97.32% | 93.92% | 39.4% | 93.78% | 92.8% | 99.2% | 96.15% | 96.92% | 95.1%
08122024 LCS | Sample | 2024-08-12 13:25 | 0.0951 | 4.7787 | 2.4618 | 0.4929 | 2.4434 | 0.4802 | 4.8454 | 0.985 | 4.9104 | 2.4231 | 2.3844 | 5.0502 | 4.9365 | -0.0039 | 4.9593 | 1.0057 | 0.9308 | 4.9235 | 2.437 | 1.0304 | 2.4452 | 2.418 | 2.2314 | 2.4796 | 0.9748 | 0.9806 | 0.9708 | 2.435 | 0.9729
08122024 LCS - Analytical Duplicate | Sample | 2024-08-12 13:28 | 0.094 | 4.779 | 2.4543 | 0.4892 | 2.4386 | 0.4797 | 4.8193 | 0.9718 | 4.8742 | 2.4055 | 2.3732 | 5.0517 | 4.8974 | -0.0045 | 4.9228 | 1.0044 | 0.9468 | 4.9058 | 2.421 | 1.0196 | 2.4296 | 2.4032 | 1.8238 | 2.4647 | 0.9709 | 0.978 | 0.9649 | 2.4248 | 0.967
RPD |  |  | 1.163% | 0.006278% | 0.3051% | 0.7535% | 0.1966% | 0.1042% | 0.5401% | 1.349% | 0.7399% | 0.729% | 0.4708% | 0.0297% | 0.7952% | 1.626% | 0.7387% | 0.1293% | 1.704% | 0.3601% | 0.6587% | 1.054% | 0.64% | 0.614% | 20.1% | 0.6027% | 0.4009% | 0.2655% | 0.6096% | 0.4198% | 0.6083%
08122024 LB | Sample | 2024-08-12 13:15 | -0.0003 | -0.0028 | 0.0055 | 0.0004 | 0 | 0.0001 | 0.0117 | 0.0001 | 0.0001 | 0 | 0 | 0.0036 | -0.0031 | -0.0121 | -0.0096 | 0.0002 | 0.0163 | 0.0001 | -0.0029 | 0.0057 | -0.0038 | 0.0131 | 0.2045 | 0.0121 | 0.0053 | 0 | 0.0005 | -0.0004 | 0.0024
08122024 LCS | Sample | 2024-08-12 13:25 | 0.0951 | 4.7787 | 2.4618 | 0.4929 | 2.4434 | 0.4802 | 4.8454 | 0.985 | 4.9104 | 2.4231 | 2.3844 | 5.0502 | 4.9365 | -0.0039 | 4.9593 | 1.0057 | 0.9308 | 4.9235 | 2.437 | 1.0304 | 2.4452 | 2.418 | 2.2314 | 2.4796 | 0.9748 | 0.9806 | 0.9708 | 2.435 | 0.9729
R% |  |  | 95.4% | 95.63% | 98.25% | 98.5% | 97.74% | 96.02% | 96.67% | 98.49% | 98.21% | 96.92% | 95.38% | 100.9% | 98.79% |  | 99.38% | 100.6% | 91.45% | 98.47% | 97.6% | 103% | 97.96% | 96.2% | 81.08% | 98.7% | 97.48% | 98.06% | 97.03% | 97.42% | 97.05%
08122024 LB | Sample | 2024-08-12 13:15 | -0.0003 | -0.0028 | 0.0055 | 0.0004 | 0 | 0.0001 | 0.0117 | 0.0001 | 0.0001 | 0 | 0 | 0.0036 | -0.0031 | -0.0121 | -0.0096 | 0.0002 | 0.0163 | 0.0001 | -0.0029 | 0.0057 | -0.0038 | 0.0131 | 0.2045 | 0.0121 | 0.0053 | 0 | 0.0005 | -0.0004 | 0.0024
08122024 LCS - Analytical Duplicate | Sample | 2024-08-12 13:28 | 0.094 | 4.779 | 2.4543 | 0.4892 | 2.4386 | 0.4797 | 4.8193 | 0.9718 | 4.8742 | 2.4055 | 2.3732 | 5.0517 | 4.8974 | -0.0045 | 4.9228 | 1.0044 | 0.9468 | 4.9058 | 2.421 | 1.0196 | 2.4296 | 2.4032 | 1.8238 | 2.4647 | 0.9709 | 0.978 | 0.9649 | 2.4248 | 0.967
R% |  |  | 94.3% | 95.64% | 97.95% | 97.76% | 97.54% | 95.92% | 96.15% | 97.17% | 97.48% | 96.22% | 94.93% | 101% | 98.01% |  | 98.65% | 100.4% | 93.05% | 98.11% | 96.96% | 102% | 97.34% | 95.6% | 64.77% | 98.1% | 97.09% | 97.8% | 96.44% | 97.01% | 96.46%
Rinse | QC | 2024-08-12 13:30 | -0.0001 | -0.0041 | 0.0158 | 0.0009 | 0 | 0 | -0.0098 | 0.0001 | 0.0002 | 0.0003 | 0.0001 | 0.0001 | -0.0031 | -0.004 | -0.0105 | 0.0001 | 0.0333 | -0.0089 | -0.0012 | 0.0058 | -0.0052 | 0.0052 | 0.2152 | 0.0216 | 0.0022 | 0 | 0.0003 | 0.0002 | 0.0005
CCV (5ppm All) | QC | 2024-08-12 13:33 | 4.8204 | 4.9388 | 4.8716 | 4.885 | 4.9001 | 4.8524 | 4.8688 | 4.8823 | 4.8834 | 4.8633 | 4.8335 | 5.0788 | 4.8391 | 4.9926 | 4.8699 | 5.058 | 4.9522 | 4.8706 | 4.8827 | 5.0858 | 4.891 | 4.8042 | 4.8612 | 4.9266 | 4.9455 | 4.905 | 4.8105 | 4.8496 | 4.8826
% Recovery |  |  | 96.41% | 98.78% | 97.43% | 97.7% | 98% | 97.05% | 97.38% | 97.65% | 97.67% | 97.27% | 96.67% | 101.6% | 96.78% | 99.85% | 97.4% | 101.2% | 99.04% | 97.41% | 97.65% | 101.7% | 97.82% | 96.08% | 97.22% | 98.53% | 98.91% | 98.1% | 96.21% | 96.99% | 97.65%
CCB | QC | 2024-08-12 13:35 | -0.0004 | -0.004 | 0.0187 | 0.0013 | 0 | 0.0001 | -0.0133 | 0.0001 | 0.0002 | -0.0002 | -0.0004 | 0.0012 | 0.0013 | -0.0035 | -0.0112 | 0.0003 | 0.0526 | -0.0097 | -0.0013 | 0.007 | -0.003 | 0.0072 | 0.207 | 0.0239 | 0.0015 | 0.0001 | 0.0015 | -0.0001 | 0
0.5ppm All (LLQC) | QC | 2024-08-12 13:38 | 0.4711 | 0.4883 | 0.4708 | 0.4908 | 0.4877 | 0.4833 | 0.4748 | 0.48 | 0.483 | 0.4819 | 0.4705 | 0.5111 | 0.4444 Q | 0.4924 | 0.4601 | 0.5063 | 0.4647 | 0.5038 | 0.481 | 0.4997 | 0.4797 | 0.4654 | 0.4662 | 0.4755 | 0.4719 | 0.49 | 0.4709 | 0.4733 | 0.4716
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-12 13:40 | 0.4738 | 0.4834 | 0.4806 | 0.4953 | 0.4826 | 0.4786 | 0.4786 | 0.4839 | 0.4871 | 0.4854 | 0.4751 | 0.5065 | 0.4521 | 0.4892 | 0.468 | 0.502 | 0.4749 | 0.5006 | 0.484 | 0.5017 | 0.4836 | 0.469 | 0.4799 | 0.481 | 0.4757 | 0.4842 | 0.4745 | 0.4775 | 0.476
RPD |  |  | 0.5715% | 1.009% | 2.06% | 0.9127% | 1.051% | 0.9772% | 0.7971% | 0.8092% | 0.8453% | 0.7237% | 0.9729% | 0.9041% | #VALUE! | 1.626% | 1.702% | 0.8529% | 2.171% | 0.6372% | 0.6218% | 0.3994% | 0.8097% | 0.7705% | 2.896% | 1.15% | 0.802% | 1.191% | 0.7616% | 0.8835% | 0.9287%
0.5ppm All (LLQC) | QC | 2024-08-12 13:38 | 0.4711 | 0.4883 | 0.4708 | 0.4908 | 0.4877 | 0.4833 | 0.4748 | 0.48 | 0.483 | 0.4819 | 0.4705 | 0.5111 | 0.4444 Q | 0.4924 | 0.4601 | 0.5063 | 0.4647 | 0.5038 | 0.481 | 0.4997 | 0.4797 | 0.4654 | 0.4662 | 0.4755 | 0.4719 | 0.49 | 0.4709 | 0.4733 | 0.4716
% Recovery |  |  | 94.22% | 97.66% | 94.16% | 98.16% | 97.54% | 96.66% | 94.96% | 96% | 96.6% | 96.38% | 94.1% | 102.2% |  | 98.48% | 92.02% | 101.3% | 92.94% | 100.8% | 96.2% | 99.94% | 95.94% | 93.08% | 93.24% | 95.1% | 94.38% | 98% | 94.18% | 94.66% | 94.32%
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-12 13:40 | 0.4738 | 0.4834 | 0.4806 | 0.4953 | 0.4826 | 0.4786 | 0.4786 | 0.4839 | 0.4871 | 0.4854 | 0.4751 | 0.5065 | 0.4521 | 0.4892 | 0.468 | 0.502 | 0.4749 | 0.5006 | 0.484 | 0.5017 | 0.4836 | 0.469 | 0.4799 | 0.481 | 0.4757 | 0.4842 | 0.4745 | 0.4775 | 0.476
% Recovery |  |  | 94.76% | 96.68% | 96.12% | 99.06% | 96.52% | 95.72% | 95.72% | 96.78% | 97.42% | 97.08% | 95.02% | 101.3% | 90.42% | 97.84% | 93.6% | 100.4% | 94.98% | 100.1% | 96.8% | 100.3% | 96.72% | 93.8% | 95.98% | 96.2% | 95.14% | 96.84% | 94.9% | 95.5% | 95.2%
CCS1 (SS 5ppm) | QC | 2024-08-12 13:43 | 4.929 | 4.9938 | 5.0039 | 4.9145 | 4.9299 | 4.9051 | 4.9229 | 4.9142 | 4.9038 | 4.882 | 4.8475 | 5.0823 | 4.9027 | 4.9604 | 4.9653 | 5.0749 | 5.0853 | 4.9409 | 4.9008 | 5.1355 | 4.9505 | 4.8105 | 4.8714 | 4.9513 | 4.8047 | 4.9634 | 4.9054 | 4.92 | 4.8307
% Recovery |  |  | 98.58% | 99.88% | 100.1% | 98.29% | 98.6% | 98.1% | 98.46% | 98.28% | 98.08% | 97.64% | 96.95% | 101.6% | 98.05% | 99.21% | 99.31% | 101.5% | 101.7% | 98.82% | 98.02% | 102.7% | 99.01% | 96.21% | 97.43% | 99.03% | 96.09% | 99.27% | 98.11% | 98.4% | 96.61%
CCB | QC | 2024-08-12 13:45 | -0.0001 | -0.0018 | 0.0204 | 0.0015 | 0.0001 | 0.0001 | -0.0107 | 0.0002 | 0.0002 | -0.0001 | 0.0001 | 0.0007 | 0.0164 | -0.0039 | -0.0117 | 0 | 0.0507 | -0.0085 | -0.0023 | 0.0046 | -0.0017 | 0.0076 | 0.1914 | 0.0245 | 0.0024 | 0.0001 | 0.0014 | 0.0002 | 0
MQ-Blank-01 | Sample | 2024-08-12 13:48 | -0.0001 | -0.0036 | 0.0016 | 0.0005 | -0.0001 | -0.0001 | -0.0137 | 0.0002 | 0.0003 | 0 | -0.0002 | 0 | -0.0096 | -0.0026 | -0.0129 | 0.0002 | 0.0049 | -0.009 | -0.002 | 0.0052 | -0.0055 | 0.0049 | 0.2592 | 0.0073 | 0.003 | 0 | 0.0002 | -0.0001 | 0.0001
Rinse | QC | 2024-08-12 13:50 | -0.0002 | -0.0034 | 0.0082 | 0.0004 | -0.0001 | 0 | -0.0106 | 0.0001 | 0.0002 | 0 | -0.0003 | 0.0001 | 0.0067 | -0.0046 | -0.0112 | -0.0002 | 0.0117 | -0.0086 | -0.002 | 0.0057 | -0.0061 | 0.0096 | 0.2206 | 0.0087 | 0.0025 | 0 | 0.0003 | -0.0004 | 0.0004
CCV (5ppm All) | QC | 2024-08-12 13:53 | 4.7817 | 4.8565 | 4.8297 | 4.8389 | 4.8123 | 4.8026 | 4.8308 | 4.8387 | 4.8396 | 4.8179 | 4.7917 | 4.9808 | 4.7889 | 4.9131 | 4.8222 | 4.958 | 4.9081 | 4.7942 | 4.8346 | 5.0436 | 4.8544 | 4.7437 | 3.9964 | 4.8748 | 4.9077 | 4.8182 | 4.7734 | 4.8035 | 4.8094
% Recovery |  |  | 95.63% | 97.13% | 96.59% | 96.78% | 96.25% | 96.05% | 96.62% | 96.77% | 96.79% | 96.36% | 95.83% | 99.62% | 95.78% | 98.26% | 96.44% | 99.16% | 98.16% | 95.88% | 96.69% | 100.9% | 97.09% | 94.87% | 79.93% | 97.5% | 98.15% | 96.36% | 95.47% | 96.07% | 96.19%
CCB | QC | 2024-08-12 13:55 | -0.0001 | -0.0038 | 0.0159 | 0.0021 | 0.0001 | 0.0001 | -0.0118 | 0.0001 | -0.0001 | 0 | 0.0003 | -0.0001 | 0.0049 | -0.0053 | -0.0066 | 0.0002 | 0.0522 | -0.0075 | -0.0022 | 0.0076 | -0.0049 | 0.0097 | 0.2125 | 0.0284 | 0.0031 | 0.0001 | 0.0016 | 0 | 0.0001
0.5% NaCl LFP Battery-TM | Sample | 2024-08-12 13:58 | -0.0012 | 557.7988 | 0.0258 | 0.0046 | 0.0923 | -0.0001 | 2.6321 | 0.0019 | 0.0114 | 0.143 | 0.0216 | 1314.2552 | 0.6112 | 2.4354 | 1.5195 | 5.0189 | 0.037 | 1471.3982 | 17.2464 | 57.286 | 0.0063 | 2.5667 | 0.2012 | -0.0521 | 3.8912 | 0.0362 | 0.139 | 0.0832 | 0.2196
1% NaCl LFP Battery-TM | Sample | 2024-08-12 14:00 | -0.0017 | 467.9349 | 0.0215 | 0.011 | 0.0733 | -0.0001 | 4.6387 | 0.0032 | 0.0006 | 0.1437 | 0.0259 | 989.5368 | 1.0757 | 5.0366 | 1.2138 | 4.4929 | 0.0272 |  | 13.531 | 111.2263 | 0.0068 | 3.4047 | 0.2015 | -0.0365 | 5.4906 | 0.0503 | 0.1286 | 0.0682 | 0.1464
5% NaCl LFP Battery-TM | Sample | 2024-08-12 14:03 | -0.0016 | 814.1376 | 0.0251 | 0.0328 | 0.0692 | -0.0002 | 18.1943 | 0.003 | 0.0032 | 0.2392 | 0.0786 | 858.3883 | 5.3255 | 23.4352 | 1.4624 | 3.4276 | 0.0422 |  | 13.4173 | 196.4268 | 0.0057 | 15.3393 | 0.205 | -0.0417 | 9.7089 | 0.2306 | 0.2101 | 0.1067 | 0.0789
10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:05 | -0.0017 | 1293.5693 | 0.0355 | 0.0604 | 0.0904 | -0.0002 | 33.5257 | -0.0002 | 0.0034 | 0.3419 | 0.1374 | 1055.2288 | 13.6907 | 32.0557 | 2.2245 | 3.8145 | 0.0519 |  | 17.1599 | 159.2554 | 0.0041 | 29.3639 | 0.302 | -0.0609 | 10.0706 | 0.4585 | 0.318 | 0.1462 | 0.1659
15% NaCl LFP Battery-TM | Sample | 2024-08-12 14:08 | -0.0014 | 963.621 | 0.0262 | 0.1545 | 0.0046 | -0.0003 | 46.0159 | 0.0015 | 0.0008 | 0.3604 | 0.1651 | 753.6171 | 24.654 | 14.3553 | 4.7087 | 0.7849 | 0.0206 |  | 10.7765 | 16.5137 | -0.0045 | 41.614 | 0.2007 | -0.0331 | 0.7093 | 0.677 | 0.2669 | 0.0657 | 0.0409
0.5% NaCl MMC Battery-TM | Sample | 2024-08-12 14:10 | -0.0007 | 456.0864 | 0.0159 | 0.0858 | 0.0079 | -0.0002 | 2.7328 | 0.004 | -0.0061 | 0.2566 | 0.0556 | 1202.9261 | 0.7907 | 1.2715 | 3.2923 | 1.0584 | 0.0114 | 1666.8729 | 13.7941 | 18.9552 | 0.0044 | 1.7617 | 0.2311 | -0.055 | 0.5751 | 0.0324 | 0.1634 | 0.0449 | 0.0646
1% NaCl MMC Battery-TM | Sample | 2024-08-12 14:13 | -0.0004 | 656.2899 | 0.0243 | 0.1069 | 0.0083 | -0.0001 | 4.0848 | 0.0036 | 0.0092 | 0.3919 | 0.1357 | 1271.6013 | 1.1073 | 1.3916 | 4.9223 | 1.1126 | 0.0174 |  | 15.8358 | 22.503 | 0.0029 | 2.9429 | 0.2691 | -0.0593 | 0.6418 | 0.0542 | 0.2544 | 0.0666 | 0.0698
5% NaCl MMC Battery-TM | Sample | 2024-08-12 14:15 | -0.0015 | 793.9437 | 0.0196 | 0.1193 | 0.006 | -0.0002 | 15.8604 | 0.004 | 0.185 | 0.4084 | 0.1636 | 971.2853 | 7.1104 | 8.1921 | 4.9497 | 0.8952 | 0.0178 |  | 13.0225 | 23.9881 | -0.0011 | 13.5099 | 0.249 | -0.0527 | 0.6662 | 0.2582 | 0.2724 | 0.0683 | 0.0548
10% NaCl MMC Battery-TM | Sample | 2024-08-12 14:18 | -0.0014 | 1049.6128 | 0.0364 | 0.1497 | 0.0045 | -0.0002 | 37.1117 | 0.0025 | 0.0143 | 0.4244 | 0.1993 | 900.7167 | 17.0243 | 15.7913 | 5.1979 | 0.8891 | 0.0191 |  | 12.539 | 20.2247 | -0.0034 | 33.0031 | 0.2312 | -0.0555 | 0.7378 | 0.5309 | 0.3155 | 0.0769 | 0.0485
15% NaCl MMC Battery-TM | Sample | 2024-08-12 14:20 | -0.0016 | 958.1671 | 0.0176 | 0.0771 | 0.0719 | -0.0001 | 45.0598 | 0.005 | 0.0053 | 0.1875 | 0.0846 | 730.2844 | 24.2192 | 30.2284 | 1.8596 | 3.2229 | 0.0328 |  | 11.9989 | 102.5226 | -0.0074 | 40.7388 | 0.2175 | -0.0551 | 8.7413 | 0.6679 | 0.2083 | 0.0876 | 0.064
Rinse1 | QC | 2024-08-12 14:23 | -0.0004 | 0.0048 | 0.0019 | -0.0014 | -0.0001 | 0 | -0.0135 | 0 | -0.0002 | 0 | 0 | 0.0143 | 0.0044 | -0.0025 | -0.0154 | 0.0002 | 0.0027 | 1.1196 | -0.0013 | 0.0079 | -0.0067 | 0.0098 | 0.2488 | 0.0008 | 0.0282 | 0 | -0.0001 | -0.0001 | 0.0005
Rinse | QC | 2024-08-12 14:25 | -0.0004 | -0.0018 | 0.0008 | -0.0016 | -0.0001 | 0 | -0.009 | -0.0001 | -0.0001 | -0.0004 | -0.0007 | 0.0089 | -0.0028 | -0.0041 | -0.0154 | 0.0001 | 0.0007 | 0.6289 | -0.0013 | 0.0055 | -0.0056 | 0.0069 | 0.2227 | 0.0089 | 0.0173 | 0 | -0.0001 | -0.0005 | 0.0009
CCV (5ppm All) | QC | 2024-08-12 14:28 | 4.8117 | 4.9228 | 4.8845 | 4.8448 | 4.8877 | 4.8767 | 4.859 | 4.8639 | 4.8814 | 4.858 | 4.8329 | 5.0646 | 4.8122 | 4.9741 | 4.8483 | 5.0073 | 4.9355 | 5.2778 | 4.8723 | 5.1044 | 4.8902 | 4.8087 | 3.9893 | 4.8967 | 4.954 | 4.8882 | 4.8045 | 4.8395 | 4.8355
% Recovery |  |  | 96.23% | 98.46% | 97.69% | 96.9% | 97.75% | 97.53% | 97.18% | 97.28% | 97.63% | 97.16% | 96.66% | 101.3% | 96.24% | 99.48% | 96.97% | 100.1% | 98.71% | 105.6% | 97.45% | 102.1% | 97.8% | 96.17% | 79.79% | 97.93% | 99.08% | 97.76% | 96.09% | 96.79% | 96.71%
CCB | QC | 2024-08-12 14:30 | -0.0004 | 0.0003 | 0.0131 | 0.0003 | 0.0001 | -0.0001 | -0.0126 | 0.0002 | 0.0001 | -0.0001 | -0.0004 | 0.0043 | 0.0103 | -0.0045 | -0.0114 | 0.0003 | 0.0447 | 0.392 | -0.0022 | 0.0088 | -0.0023 | 0.0072 | 0.2017 | 0.0214 | 0.0105 | 0.0001 | 0.0013 | -0.0002 | 0.0003
08092024 MB | Sample | 2024-08-12 14:33 | -0.0005 | 0.0018 | 0.0079 | -0.0001 | 0 | 0 | -0.0091 | 0.0002 | 0.0001 | -0.0001 | 0.0005 | 0.0072 | -0.001 | -0.0037 | -0.0098 | 0.0002 | 0.014 | 0.4276 | -0.0023 | 0.0084 | -0.0027 | 0.007 | 0.1698 | 0.0075 | 0.0144 | 0 | 0.0005 | -0.0001 | 0.0011
1008092024 MFB | Sample | 2024-08-12 14:35 | 0.0955 | 4.8215 | 2.4388 | 0.4803 | 2.4493 | 0.4779 | 4.8377 | 0.9756 | 4.9165 | 2.4226 | 2.4159 | 5.0866 | 4.9341 | -0.0121 | 4.9412 | 1.0065 | 0.9871 | 5.5818 | 2.4371 | 1.0079 | 2.4329 | 2.3612 | 2.1057 | 2.3829 | 1.0051 | 0.9825 | 0.9772 | 2.4441 | 0.9651
08092024 MFB - Analyticsal Duplicate | Sample | 2024-08-12 14:38 | 0.0948 | 4.7914 | 2.4428 | 0.4807 | 2.4307 | 0.4767 | 4.8348 | 0.9699 | 4.8704 | 2.4239 | 2.4111 | 5.0631 | 4.8962 | -0.004 | 4.9375 | 1.0018 | 0.9995 | 5.5088 | 2.4317 | 1.0115 | 2.432 | 2.3437 | 2.1615 | 2.3907 | 1.0066 | 0.9764 | 0.9766 | 2.4348 | 0.9642
RPD |  |  | 0.7357% | 0.6262% | 0.1639% | 0.08325% | 0.7623% | 0.2514% | 0.05996% | 0.586% | 0.9421% | 0.05365% | 0.1989% | 0.4631% | 0.7711% | 1.626% | 0.07491% | 0.4681% | 1.248% | 1.316% | 0.2218% | 0.3565% | 0.037% | 0.7439% | 2.615% | 0.3268% | 0.1491% | 0.6228% | 0.06142% | 0.3812% | 0.0933%
08092024 MB | Sample | 2024-08-12 14:33 | -0.0005 | 0.0018 | 0.0079 | -0.0001 | 0 | 0 | -0.0091 | 0.0002 | 0.0001 | -0.0001 | 0.0005 | 0.0072 | -0.001 | -0.0037 | -0.0098 | 0.0002 | 0.014 | 0.4276 | -0.0023 | 0.0084 | -0.0027 | 0.007 | 0.1698 | 0.0075 | 0.0144 | 0 | 0.0005 | -0.0001 | 0.0011
1008092024 MFB | Sample | 2024-08-12 14:35 | 0.0955 | 4.8215 | 2.4388 | 0.4803 | 2.4493 | 0.4779 | 4.8377 | 0.9756 | 4.9165 | 2.4226 | 2.4159 | 5.0866 | 4.9341 | -0.0121 | 4.9412 | 1.0065 | 0.9871 | 5.5818 | 2.4371 | 1.0079 | 2.4329 | 2.3612 | 2.1057 | 2.3829 | 1.0051 | 0.9825 | 0.9772 | 2.4441 | 0.9651
R% |  |  | 96% | 96.39% | 97.24% | 96.08% | 97.97% | 95.58% | 96.94% | 97.54% | 98.33% | 96.91% | 96.62% | 101.6% | 98.7% |  | 99.02% | 100.6% | 97.31% | 103.1% | 97.58% | 100.8% | 97.42% | 94.17% | 77.44% | 95.02% | 100.5% | 98.25% | 97.67% | 97.77% | 96.4%
08092024 MB | Sample | 2024-08-12 14:33 | -0.0005 | 0.0018 | 0.0079 | -0.0001 | 0 | 0 | -0.0091 | 0.0002 | 0.0001 | -0.0001 | 0.0005 | 0.0072 | -0.001 | -0.0037 | -0.0098 | 0.0002 | 0.014 | 0.4276 | -0.0023 | 0.0084 | -0.0027 | 0.007 | 0.1698 | 0.0075 | 0.0144 | 0 | 0.0005 | -0.0001 | 0.0011
08092024 MFB - Analyticsal Duplicate | Sample | 2024-08-12 14:38 | 0.0948 | 4.7914 | 2.4428 | 0.4807 | 2.4307 | 0.4767 | 4.8348 | 0.9699 | 4.8704 | 2.4239 | 2.4111 | 5.0631 | 4.8962 | -0.004 | 4.9375 | 1.0018 | 0.9995 | 5.5088 | 2.4317 | 1.0115 | 2.432 | 2.3437 | 2.1615 | 2.3907 | 1.0066 | 0.9764 | 0.9766 | 2.4348 | 0.9642
R% |  |  | 95.3% | 95.79% | 97.4% | 96.16% | 97.23% | 95.34% | 96.88% | 96.97% | 97.41% | 96.96% | 96.42% | 101.1% | 97.94% |  | 98.95% | 100.2% | 98.55% | 101.6% | 97.36% | 101.2% | 97.39% | 93.47% | 79.67% | 95.33% | 100.7% | 97.64% | 97.61% | 97.4% | 96.31%
08092024 MB | Sample | 2024-08-12 14:33 | -0.0005 | 0.0018 | 0.0079 | -0.0001 | 0 | 0 | -0.0091 | 0.0002 | 0.0001 | -0.0001 | 0.0005 | 0.0072 | -0.001 | -0.0037 | -0.0098 | 0.0002 | 0.014 | 0.4276 | -0.0023 | 0.0084 | -0.0027 | 0.007 | 0.1698 | 0.0075 | 0.0144 | 0 | 0.0005 | -0.0001 | 0.0011
5x 08092024 MFB | Sample | 2024-08-12 14:40 | 0.0188 | 0.9539 | 0.4915 | 0.0947 | 0.4846 | 0.0947 | 0.9759 | 0.193 | 0.9692 | 0.4822 | 0.4712 | 1.0115 | 0.9427 | -0.007 | 0.9587 | 0.1998 | 0.1988 | 1.2827 | 0.4788 | 0.2066 | 0.4784 | 0.471 | 0.459 | 0.4883 | 0.2018 | 0.1935 | 0.1914 | 0.4791 | 0.1898
R% |  |  | 96.5% | 95.21% | 96.72% | 94.8% | 96.92% | 94.7% | 98.5% | 96.4% | 96.91% | 96.46% | 94.14% | 100.4% | 94.37% |  | 96.85% | 99.8% | 92.4% | 85.51% | 96.22% | 99.1% | 96.22% | 92.8% | 57.84% | 96.16% | 93.7% | 96.75% | 95.45% | 95.84% | 94.35%
5x 10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:43 | -0.0008 | 215.6443 | 0.0122 | 0.0078 | 0.0176 | 0 | 7.4977 | 0.0001 | 0.0015 | 0.079 | 0.0257 | 243.4776 | 1.5368 | 3.0259 | 0.4649 | 0.8884 | 0.0221 |  | 4.1748 | 35.9493 | -0.0002 | 6.7007 | 0.153 | -0.0125 | 1.1647 | 0.0894 | 0.0671 | 0.0324 | 0.0436
5x 10% NaCl LFP Battery-TM - MS | Sample | 2024-08-12 14:45 | 0.0176 | 189.7969 | 0.5183 | 0.106 | 0.4979 | 0.0908 | 7.6228 | 0.1866 | 0.8229 | 0.5387 | 0.572 | 215.9512 | 2.5462 | 2.6037 | 1.3208 | 0.971 | 0.2052 |  | 4.0047 | 32.264 | 0.4631 | 6.4754 | 0.5718 | 0.4763 | 1.3654 | 0.2731 | 0.2521 | 0.5194 | 0.2224
R% |  |  | 92% | -2585% | 101.2% | 98.2% | 96.06% | 90.8% | 12.51% | 93.25% | 82.14% | 91.94% | 109.3% | -2753% | 100.9% |  | 85.59% | 41.3% | 91.55% | 0% | -34.02% | -1843% | 92.66% | -45.06% | 83.76% | 97.76% | 100.3% | 91.85% | 92.5% | 97.4% | 89.4%
10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:05 | -0.0017 | 1293.5693 | 0.0355 | 0.0604 | 0.0904 | -0.0002 | 33.5257 | -0.0002 | 0.0034 | 0.3419 | 0.1374 | 1055.2288 | 13.6907 | 32.0557 | 2.2245 | 3.8145 | 0.0519 |  | 17.1599 | 159.2554 | 0.0041 | 29.3639 | 0.302 | -0.0609 | 10.0706 | 0.4585 | 0.318 | 0.1462 | 0.1659
10% NaCl LFP Battery-TM - Analytical Duplicate | Sample | 2024-08-12 14:48 | -0.0018 | 1291.7171 | 0.0276 | 0.0566 | 0.0902 | -0.0003 | 33.4245 | 0.0036 | 0.0035 | 0.3409 | 0.1377 | 1058.1629 | 13.5663 | 31.5973 | 2.2124 | 3.8145 | 0.052 |  | 17.0984 | 158.9981 | 0.011 | 29.428 | 0.2068 | -0.0515 | 9.1782 | 0.4576 | 0.3173 | 0.152 | 0.1655
RPD |  |  | 5.714% | 0.1433% | 25.04% | 6.496% | 0.2215% | 40% | 0.3023% | 223.5% | 2.899% | 0.2929% | 0.2181% | 0.2777% | 0.9128% | 1.626% | 0.5454% | 0% | 0.1925% | #DIV/0! | 0.359% | 0.1617% | 91.39% | 0.2181% | 37.42% | 16.73% | 9.272% | 0.1965% | 0.2204% | 3.89% | 0.2414%
10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:05 | -0.0017 | 1293.5693 | 0.0355 | 0.0604 | 0.0904 | -0.0002 | 33.5257 | -0.0002 | 0.0034 | 0.3419 | 0.1374 | 1055.2288 | 13.6907 | 32.0557 | 2.2245 | 3.8145 | 0.0519 |  | 17.1599 | 159.2554 | 0.0041 | 29.3639 | 0.302 | -0.0609 | 10.0706 | 0.4585 | 0.318 | 0.1462 | 0.1659
10% NaCl LFP Battery-TM - Lab Duplicate | Sample | 2024-08-12 14:50 | -0.0017 | 1307.6594 | 0.0302 | 0.056 | 0.0905 | -0.0003 | 33.3458 | 0.004 | -0.0001 | 0.3441 | 0.137 | 1063.5783 | 13.9791 | 31.9609 | 2.2103 | 3.8301 | 0.047 | 2435.7052 | 16.914 | 158.4421 | 0.0061 | 29.3356 | 0.2646 | -0.0649 | 8.6894 | 0.459 | 0.3177 | 0.1492 | 0.1639
RPD |  |  | 0% | 1.083% | 16.13% | 7.56% | 0.1106% | 40% | 0.538% | 221.1% | 212.1% | 0.6414% | 0.2915% | 0.7881% | 2.085% | 1.626% | 0.6404% | 0.4081% | 9.909% | 200% | 1.443% | 0.512% | 39.22% | 0.09642% | 13.2% | 6.359% | 14.72% | 0.109% | 0.09438% | 2.031% | 1.213%
10% NaCl LFP Battery-TM - MS | Sample | 2024-08-12 14:53 | 0.1095 | 1125.5413 | 2.3395 | 0.5103 | 2.5982 | 0.406 | 34.1408 | 0.7543 | 3.4841 | 2.354 | 3.0146 | 952.7033 | 20.8246 | 26.5031 | 5.5548 | 4.2462 | 0.926 |  | 16.6532 | 143.6606 | 1.8626 | 28.9446 | 2.3307 | 2.0901 | 9.5887 | 1.4172 | 1.2005 | 2.3862 | 0.8689
10% NaCl LFP Battery-TM - MS - Analytical Duplicate | Sample | 2024-08-12 14:55 | 0.11 | 1121.6721 | 2.3542 | 0.5118 | 2.5932 | 0.4086 | 34.1838 | 0.7539 | 3.4883 | 2.3576 | 3.0374 | 954.4947 | 20.6754 | 26.3993 | 5.5434 | 4.238 | 0.9317 |  | 16.7073 | 143.7686 | 1.8719 | 28.9841 | 2.3846 | 2.102 | 9.5116 | 1.4139 | 1.2022 | 2.3926 | 0.868
RPD |  |  | 0.4556% | 0.3444% | 0.6264% | 0.2935% | 0.1926% | 0.6384% | 0.1259% | 0.05304% | 0.1205% | 0.1528% | 0.7535% | 0.1879% | 0.719% | 1.626% | 0.2054% | 0.1933% | 0.6137% | #DIV/0! | 0.3243% | 0.07515% | 0.4981% | 0.1364% | 2.286% | 0.5677% | 0.8073% | 0.2331% | 0.1415% | 0.2678% | 0.1036%
10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:05 | -0.0017 | 1293.5693 | 0.0355 | 0.0604 | 0.0904 | -0.0002 | 33.5257 | -0.0002 | 0.0034 | 0.3419 | 0.1374 | 1055.2288 | 13.6907 | 32.0557 | 2.2245 | 3.8145 | 0.0519 |  | 17.1599 | 159.2554 | 0.0041 | 29.3639 | 0.302 | -0.0609 | 10.0706 | 0.4585 | 0.318 | 0.1462 | 0.1659
10% NaCl LFP Battery-TM - MS | Sample | 2024-08-12 14:53 | 0.1095 | 1125.5413 | 2.3395 | 0.5103 | 2.5982 | 0.406 | 34.1408 | 0.7543 | 3.4841 | 2.354 | 3.0146 | 952.7033 | 20.8246 | 26.5031 | 5.5548 | 4.2462 | 0.926 |  | 16.6532 | 143.6606 | 1.8626 | 28.9446 | 2.3307 | 2.0901 | 9.5887 | 1.4172 | 1.2005 | 2.3862 | 0.8689
R% |  |  | 111.2% | -3361% | 92.16% | 89.98% | 100.3% | 81.24% | 12.3% | 75.45% | 69.61% | 80.48% | 115.1% | -2051% | 142.7% |  | 66.61% | 43.17% | 87.41% | 0% | -20.27% | 14366% | 74.34% | -16.77% | 81.15% | 86.04% | 958.9% | 95.87% | 88.25% | 89.6% | 70.3%
10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:05 | -0.0017 | 1293.5693 | 0.0355 | 0.0604 | 0.0904 | -0.0002 | 33.5257 | -0.0002 | 0.0034 | 0.3419 | 0.1374 | 1055.2288 | 13.6907 | 32.0557 | 2.2245 | 3.8145 | 0.0519 |  | 17.1599 | 159.2554 | 0.0041 | 29.3639 | 0.302 | -0.0609 | 10.0706 | 0.4585 | 0.318 | 0.1462 | 0.1659
10% NaCl LFP Battery-TM - MS - Analytical Duplicate | Sample | 2024-08-12 14:55 | 0.11 | 1121.6721 | 2.3542 | 0.5118 | 2.5932 | 0.4086 | 34.1838 | 0.7539 | 3.4883 | 2.3576 | 3.0374 | 954.4947 | 20.6754 | 26.3993 | 5.5434 | 4.238 | 0.9317 |  | 16.7073 | 143.7686 | 1.8719 | 28.9841 | 2.3846 | 2.102 | 9.5116 | 1.4139 | 1.2022 | 2.3926 | 0.868
R% |  |  | 111.7% | -3438% | 92.75% | 90.28% | 100.1% | 81.76% | 13.16% | 75.41% | 69.7% | 80.63% | 116% | -2015% | 139.7% |  | 66.38% | 42.35% | 87.98% | 0% | -18.1% | 14377% | 74.71% | -15.19% | 83.3% | 86.52% | 951.2% | 95.54% | 88.42% | 89.86% | 70.21%
Rinse1 | QC | 2024-08-12 14:58 | -0.0007 | 0.0163 | 0.0108 | -0.0006 | 0.0001 | 0 | -0.01 | 0.0001 | -0.0002 | -0.0001 | -0.0001 | 0.0247 | 0.0179 | -0.0127 | -0.0143 | 0.0002 | 0.0064 | 1.6323 | -0.0008 | 0.0126 | -0.0029 | 0.0097 | 0.2656 | 0.0188 | 0.0519 | 0 | 0 | 0 | 0.0007
Rinse | QC | 2024-08-12 15:00 | -0.0004 | 0.0032 | 0.0015 | -0.0012 | 0 | 0.0001 | -0.0101 | 0 | 0.0006 | -0.0002 | -0.0001 | 0.0106 | -0.0013 | -0.0032 | -0.0124 | 0.0002 | 0.0033 | 0.8374 | -0.0007 | 0.0102 | -0.0038 | 0.0076 | 0.2516 | 0.0082 | 0.0295 | 0.0001 | 0 | -0.0002 | 0.0006
CCV (5ppm All) | QC | 2024-08-12 15:03 | 4.7764 | 4.8639 | 4.8429 | 4.8016 | 4.8164 | 4.778 | 4.8126 | 4.8249 | 4.835 | 4.8195 | 4.7844 | 5.0057 | 4.8052 | 4.9436 | 4.8093 | 4.9846 | 4.9115 | 5.3647 | 4.8336 | 5.0551 | 4.8444 | 4.7526 | 4.0302 | 4.8756 | 4.9362 | 4.8259 | 4.7712 | 4.8008 | 4.7814
% Recovery |  |  | 95.53% | 97.28% | 96.86% | 96.03% | 96.33% | 95.56% | 96.25% | 96.5% | 96.7% | 96.39% | 95.69% | 100.1% | 96.1% | 98.87% | 96.19% | 99.69% | 98.23% | 107.3% | 96.67% | 101.1% | 96.89% | 95.05% | 80.6% | 97.51% | 98.72% | 96.52% | 95.42% | 96.02% | 95.63%
CCB | QC | 2024-08-12 15:06 | -0.0003 | 0.0005 | 0.0137 | 0.0001 | 0 | 0.0001 | -0.0123 | 0.0001 | 0.0002 | -0.0001 | 0.0001 | 0.0065 | 0.0156 | -0.0056 | -0.0097 | 0.0001 | 0.0407 | 0.5054 Z | -0.0012 | 0.0099 | -0.0054 | 0.0078 | 0.2069 | 0.0263 | 0.0172 | 0.0001 | 0.0012 | 0 | 0.0003
Rinse | QC | 2024-08-12 15:08 | 0.0002 | 0 | 0.0048 | -0.0005 | 0 | 0.0001 | -0.0089 | 0.0001 | 0 | -0.0001 | -0.0005 | 0.0063 | -0.003 | -0.0054 | -0.0149 | -0.0001 | 0.012 | 0.4398 | -0.001 | 0.0069 | -0.003 | 0.0081 | 0.2166 | 0.0103 | 0.0099 | 0 | 0.0002 | 0.0002 | 0.0006
0.5ppm All (LLQC) | QC | 2024-08-12 15:11 | 0.4719 | 0.4848 | 0.4732 | 0.4877 | 0.485 | 0.4765 | 0.4757 | 0.4808 | 0.485 | 0.4827 | 0.4712 | 0.507 | 0.4563 | 0.4847 | 0.464 | 0.5034 | 0.4606 | 0.8458 Q | 0.4824 | 0.5039 | 0.4832 | 0.4639 | 0.4815 | 0.4752 | 0.4805 | 0.4871 | 0.4714 | 0.4749 | 0.4725
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-12 15:13 | 0.4729 | 0.4823 | 0.4717 | 0.4871 | 0.48 | 0.4783 | 0.4709 | 0.4789 | 0.482 | 0.481 | 0.4702 | 0.5046 | 0.4359 | 0.4844 | 0.4606 | 0.4982 | 0.4694 | 0.8014 | 0.4808 | 0.4982 | 0.4772 | 0.4598 | 0.4612 | 0.4828 | 0.4754 | 0.4825 | 0.4705 | 0.4728 | 0.4704
RPD |  |  | 0.2117% | 0.517% | 0.3175% | 0.1231% | 1.036% | 0.377% | 1.014% | 0.396% | 0.6205% | 0.3528% | 0.2124% | 0.4745% | 4.573% | 1.626% | 0.7355% | 1.038% | 1.892% | #VALUE! | 0.3322% | 1.138% | 1.249% | 0.8877% | 4.307% | 1.587% | 1.067% | 0.9488% | 0.1911% | 0.4432% | 0.4454%
0.5ppm All (LLQC) | QC | 2024-08-12 15:11 | 0.4719 | 0.4848 | 0.4732 | 0.4877 | 0.485 | 0.4765 | 0.4757 | 0.4808 | 0.485 | 0.4827 | 0.4712 | 0.507 | 0.4563 | 0.4847 | 0.464 | 0.5034 | 0.4606 | 0.8458 Q | 0.4824 | 0.5039 | 0.4832 | 0.4639 | 0.4815 | 0.4752 | 0.4805 | 0.4871 | 0.4714 | 0.4749 | 0.4725
% Recovery |  |  | 94.38% | 96.96% | 94.64% | 97.54% | 97% | 95.3% | 95.14% | 96.16% | 97% | 96.54% | 94.24% | 101.4% | 91.26% | 96.94% | 92.8% | 100.7% | 92.12% |  | 96.48% | 100.8% | 96.64% | 92.78% | 96.3% | 95.04% | 96.1% | 97.42% | 94.28% | 94.98% | 94.5%
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-12 15:13 | 0.4729 | 0.4823 | 0.4717 | 0.4871 | 0.48 | 0.4783 | 0.4709 | 0.4789 | 0.482 | 0.481 | 0.4702 | 0.5046 | 0.4359 | 0.4844 | 0.4606 | 0.4982 | 0.4694 | 0.8014 | 0.4808 | 0.4982 | 0.4772 | 0.4598 | 0.4612 | 0.4828 | 0.4754 | 0.4825 | 0.4705 | 0.4728 | 0.4704
% Recovery |  |  | 94.58% | 96.46% | 94.34% | 97.42% | 96% | 95.66% | 94.18% | 95.78% | 96.4% | 96.2% | 94.04% | 100.9% | 87.18% | 96.88% | 92.12% | 99.64% | 93.88% | 160.3% | 96.16% | 99.64% | 95.44% | 91.96% | 92.24% | 96.56% | 95.08% | 96.5% | 94.1% | 94.56% | 94.08%
CCS1 (SS 5ppm) | QC | 2024-08-12 15:16 | 4.8877 | 4.9531 | 4.9736 | 4.8275 | 4.8945 | 4.8447 | 4.8682 | 4.8703 | 4.8681 | 4.8498 | 4.8154 | 5.0542 | 4.8359 | 4.9059 | 4.9156 | 4.9842 | 5.0683 | 5.099 | 4.8619 | 5.1228 | 4.9146 | 4.7687 | 4.8104 | 4.9157 | 4.7513 | 4.9309 | 4.8611 | 4.8818 | 4.7833
% Recovery |  |  | 97.75% | 99.06% | 99.47% | 96.55% | 97.89% | 96.89% | 97.36% | 97.41% | 97.36% | 97% | 96.31% | 101.1% | 96.72% | 98.12% | 98.31% | 99.68% | 101.4% | 102% | 97.24% | 102.5% | 98.29% | 95.37% | 96.21% | 98.31% | 95.03% | 98.62% | 97.22% | 97.64% | 95.67%
CCB | QC | 2024-08-12 15:18 | -0.0001 | -0.0022 | 0.0167 | 0.0007 | 0.0001 | 0.0002 | -0.0111 | 0.0001 | 0 | 0 | 0.0004 | 0.0011 | 0.0047 | -0.008 | -0.0115 | 0.0001 | 0.0396 | 0.249 | -0.0018 | 0.0076 | -0.0029 | 0.0048 | 0.1845 | 0.0232 | 0.0038 | 0.0001 | 0.0011 | 0.0003 | 0.0003
Rinse1 | QC | 2024-08-12 15:21 | 0 | -0.0038 | 0.0028 | 0.0008 | 0 | 0.0001 | -0.0127 | 0 | 0.0004 | 0.0001 | -0.0001 | 0.002 | -0.0085 | -0.0053 | -0.0119 | 0.0003 | 0.0139 | 0.2272 | -0.0024 | 0.0069 | -0.0046 | 0.0056 | 0.184 | 0.0119 | 0.0047 | 0.0001 | 0.0003 | -0.0001 | 0.0003
Rinse | QC | 2024-08-12 15:23 | -0.0002 | -0.0031 | 0.0028 | 0 | 0 | 0 | -0.0144 | 0 | 0.0001 | -0.0004 | -0.0006 | 0.0011 | 0.0023 | -0.0064 | -0.011 | 0 | 0.0078 | 0.2003 | -0.0023 | 0.0057 | -0.0039 | 0.008 | 0.2215 | 0.0102 | 0.0019 | 0 | 0.0001 | -0.0004 | 0.0003
MQ-Blank-02 | Sample | 2024-08-12 15:26 | 0.0001 | -0.0042 | 0.0018 | -0.0012 | -0.0001 | 0 | -0.0131 | 0 | 0.0004 | -0.0002 | 0.0001 | 0.0028 | -0.0038 | -0.0042 | -0.0124 | 0 | 0.0019 | 0.1758 | -0.0024 | 0.0026 | -0.0044 | 0.0057 | 0.2462 | 0.0043 | 0.0035 | 0 | 0 | -0.0001 | 0.0011
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Lowest Cal Std |  |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Highest Cal Std |  |  | 20 | 200 | 20 | 20 | 50 | 20 | 2000 | 20 | 20 | 50 | 100 | 1000 | 100 | 100 | 2000 | 50 | 20 | 100 | 50 | 200 | 200 | 100 | 50 | 20 | 100 | 20 | 50 | 20 | 100
Correlation |  |  | 0.99997 | 0.99997 | 0.99997 | 1 | 1 | 0.99999 | 0.99987 | 1 | 0.99998 | 0.99999 | 0.99999 | 0.99998 | 0.99998 | 0.99999 | 0.99996 | 0.99988 | 0.99973 | 0.99998 | 0.99999 | 0.9999 | 0.99974 | 0.99964 | 0.9998 | 0.99992 | 0.99933 | 0.99999 | 0.99996 | 0.99999 | 0.99999
ICP-OES MDL |  |  | 0.0045598 | 0.0048715 | 0.0060725 | 0.0016216 | 0.00011879 | 0.000168 | 0.0038865 | 0.00024729 | 0.00024729 | 0.0007 | 0.00076374 | 0.0031683 | 0.032873 | 0.0015923 | 0.0080477 | 0.0008057 | 0.00050865 | 0.001351 | 0.0024937 | 0.0072662 | 0.0049325 | 0.0058515 | 0.0056274 | 0.0064431 | 0.002358 | 0.000168 | 0.00011879 | 0.00093287 | 0.00059397
ICP-OES MRL/PRDL |  |  | 0.045598 | 0.048715 | 0.060725 | 0.016216 | 0.0011879 | 0.00168 | 0.038865 | 0.0024729 | 0.0024729 | 0.0074 | 0.0076374 | 0.031683 | 0.32873 | 0.015923 | 0.080477 | 0.008057 | 0.0050865 | 0.01351 | 0.024937 | 0.072662 | 0.049325 | 0.058515 | 0.056274 | 0.064431 | 0.02358 | 0.00168 | 0.0011879 | 0.0093287 | 0.0059397
Where |  |  | Ba: QC Recoveries
| Calibration Standards |  | Ba: 100%; Be: 90<R%>110
| IEC Check Standards |  | Ba: 80%; Be: 70<R%>80
| CCV & Secondary source QC check standards |  | Ba: 70%; Be: 70%>R%>120%
| Low Level QC chek standards |  | Ba: 60%; Be: 65%>R%>135%
| Calculated/Estimated MDLs & MRLs |  | Ba: 100%; Be: 65<R%>135
| Dilution required |  | Ba: 20%; Be: RPD<20%
* | Sampling Error |  | Ba: 21%; Be: RPD>20%
** | Different concentration than the normal |  | Ba: 20; Be: RSD<20%
| Mislabeled |  | Ba: 21; Be: RSD>20%
[7 empty] | 20 | STDEV<20

## 08122024ICP-OES_TO178-444RD
Worksheet exported from: C:\Users\USEPA\Documents\Agilent\ICP Expert\My Results\08122024_ICP-OES_TO 058-16 + TO 178-444.esws
Solution Label | Type | Date Time | LIMS ID | Customer | Customer Ref. | Sample Date Time | Sample Site | Description | Ag (328.068 nm) ppm | Al (396.152 nm) ppm | As (188.980 nm) ppm | B (249.678 nm) ppm | Ba (455.403 nm) ppm | Be (313.107 nm) ppm | Ca (370.602 nm) ppm | Cd (214.439 nm) ppm | Co (238.892 nm) ppm | Cr (267.716 nm) ppm | Cu (324.754 nm) ppm | Fe (259.940 nm) ppm | K (766.491 nm) ppm | Li (670.783 nm) ppm | Mg (279.078 nm) ppm | Mn (257.610 nm) ppm | Mo (202.032 nm) ppm | Na (588.995 nm) ppm | Ni (231.604 nm) ppm | P (178.222 nm) ppm | Pb (220.353 nm) ppm | S (181.972 nm) ppm | Sb (217.582 nm) ppm | Se (196.026 nm) ppm | Si (251.611 nm) ppm | Sn (189.925 nm) ppm | Sr (421.552 nm) ppm | Ti (334.941 nm) ppm | V (292.401 nm) ppm | Zn (206.200 nm) ppm | Y (371.029 nm) Ratio | Y (371.029 nm) Ratio 1 | Y (488.368 nm) Ratio | Ag 328.068 C ppm | Al 396.152 C ppm | As 188.980 C ppm | B 249.678 C ppm | Ba 455.403 C ppm | Be 313.107 C ppm | Ca 370.602 C ppm | Cd 214.439 C ppm | Co 238.892 C ppm | Cr 267.716 C ppm | Cu 324.754 C ppm | Fe 259.940 C ppm | K 766.491 C ppm | Li 670.783 C ppm | Mg 279.078 C ppm | Mn 257.610 C ppm | Mo 202.032 C ppm | Na 588.995 C ppm | Ni 231.604 C ppm | P 178.222 C ppm | Pb 220.353 C ppm | S 181.972 C ppm | Sb 217.582 C ppm | Se 196.026 C ppm | Si 251.611 C ppm | Sr 421.552 C ppm | Ti 334.941 C ppm | V 292.401 C ppm | Zn 206.200 C ppm
Calibration Blank | Blank | 2024-08-12 09:19 | [6 empty] | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 1.00 ! | 1.00 ! | 1.00 !
Calibration Standard 0.5ppm | Standard | 2024-08-12 09:21 | [6 empty] | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 1.00 ! | 1.01 ! | 1.01 !
Calibration Standard 1ppm | Standard | 2024-08-12 09:24 | [6 empty] | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.01 ! | 1.01 ! | 1.01 !
Calibration Standard 5ppm | Standard | 2024-08-12 09:26 | [6 empty] | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 1.00 ! | 0.99 ! | 1.00 !
Calibration Standard 10ppm | Standard | 2024-08-12 09:29 | [6 empty] | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 1.00 ! | 0.99 ! | 0.99 !
Calibration Standard 20ppm | Standard | 2024-08-12 09:31 | [6 empty] | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 0.98 ! | 0.98 ! | 0.99 !
Calibration Standard 50ppm | Standard | 2024-08-12 09:34 | [7 empty] | 50.0000 ! |  |  | 50.0000 ! |  | 50.0000 ! |  |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! |  |  | 50.0000 ! |  | 50.0000 ! | 0.97 ! | 0.95 ! | 0.97 !
Calibration Standard-A 100ppm | Standard | 2024-08-12 09:36 | [7 empty] | 100.0000 ! |  |  |  |  | 100.0000 ! |  |  |  | 100.0000 ! | 100.0000 ! | 100.0000 ! | 100.0000 ! | 100.0000 ! |  |  | 100.0000 ! |  |  | 100.0000 ! |  |  |  | 100.0000 ! |  |  |  |  | 100.0000 ! | 0.93 ! | 0.93 ! | 0.94 !
Calibration Standard-B 100ppm | Standard | 2024-08-12 09:39 | [25 empty] | 100.0000 ! |  | 100.0000 ! | [8 empty] | 1.02 ! | 1.00 ! | 1.01 !
Calibration Standard Al+Fe+Ca+Mg200ppm | Standard | 2024-08-12 09:41 | [7 empty] | 200.0000 ! |  |  |  |  | 200.0000 ! |  |  |  |  | 200.0000 ! |  |  | 200.0000 ! | [15 empty] | 0.95 ! | 0.96 ! | 0.97 !
Calibration Standard Ca+Mg500ppm | Standard | 2024-08-12 09:44 | [12 empty] | 500.0000 ! | [7 empty] | 500.0000 ! | [15 empty] | 0.94 ! | 0.93 ! | 0.94 !
Calibration Standard Ca+Mg1000ppm | Standard | 2024-08-12 09:47 | [12 empty] | 1000.0000 ! | [7 empty] | 1000.0000 ! | [15 empty] | 0.91 ! | 0.90 ! | 0.90 !
Calibration Standard Ca+Mg2000ppm | Standard | 2024-08-12 09:49 | [12 empty] | 2000.0000 ! | [7 empty] | 2000.0000 ! | [15 empty] | 0.87 ! | 0.85 ! | 0.85 !
Calibration Standard P+S200ppm | Standard | 2024-08-12 09:52 | [25 empty] | 200.0000 ! |  | 200.0000 ! | [8 empty] | 1.02 ! | 1.03 ! | 1.03 !
Calibration Standard Pb200ppm | Standard | 2024-08-12 09:54 | [26 empty] | 200.0000 ! | [9 empty] | 1.01 ! | 1.01 ! | 1.02 !
Calibration Standard Fe 1000ppm | Standard | 2024-08-12 09:57 | [17 empty] | 1000.0000 ! | [18 empty] | 0.98 ! | 0.99 ! | 1.00 !
Rinse1 | QC | 2024-08-12 09:59 | [6 empty] | -0.0001 | -0.0031 | 0.0024 | 0.0095 | -0.0001 | 0 | -0.0063 | 0.0002 | -0.0003 | 0.0003 | -0.0003 | 0.0239 | -0.0058 | -0.0047 | -0.0035 | 0.0002 | 0.004 | -0.0083 | -0.0016 | 0.0057 | -0.003 | 0.013 | 0.2725 | 0.0064 | 0.0037 | 0.0009 | 0 | 0.0001 | -0.0002 | 0.0003 | 1.02 | 1 | 1.01 | -0.0001 | -0.0031 | 0.0024 | 0.0095 | -0.0001 | 0 | -0.0063 | 0.0002 | -0.0003 | 0.0003 | -0.0003 | 0.0239 | -0.0058 | -0.0047 | -0.0035 | 0.0002 | 0.004 | -0.0083 | -0.0016 | 0.0057 | -0.003 | 0.013 | 0.2725 | 0.0064 | 0.0037 | 0 | 0.0001 | -0.0002 | 0.0003
Rinse | QC | 2024-08-12 10:02 | [6 empty] | -0.0002 | -0.0039 | 0.0018 | 0.0084 | 0 | 0 | -0.0065 | 0.0002 | 0.0003 | -0.0002 | 0 | 0.005 | 0.0081 | -0.0111 | -0.0058 | -0.0001 | 0.0031 | -0.007 | -0.0014 | 0.0061 | -0.0012 | 0.0099 | 0.2681 | 0.0016 | 0.0027 | 0.0011 | 0 | 0 | -0.0003 | 0.0005 | 1.02 | 1 | 1.01 | -0.0002 | -0.0039 | 0.0018 | 0.0084 | 0 | 0 | -0.0065 | 0.0002 | 0.0003 | -0.0002 | 0 | 0.005 | 0.0081 | -0.0111 | -0.0058 | -0.0001 | 0.0031 | -0.007 | -0.0014 | 0.0061 | -0.0012 | 0.0099 | 0.2681 | 0.0016 | 0.0027 | 0 | 0 | -0.0003 | 0.0005
IECS1 (Al 1000 Int.) | QC | 2024-08-12 10:04 | [6 empty] | -0.0003 | 958.4498 o | -0.0043 | 0.0099 | 0.0092 | 0 | 0.0046 | 0.0006 | 0.0001 | 0.0005 | 0.002 | 0.0079 | -0.0088 | -0.005 | 0.0043 | 0 | 0.0065 | 0.0002 | -0.0015 | 0.012 | 0.0343 | 0.0187 | 0.6129 | 0.0018 | 0.009 | 0.0025 | 0.0001 | 0.001 | 0.0004 | 0.0035 | 0.94 | 1 | 1 | -0.0003 | 958.4498 | -0.0043 | 0.0099 | 0.0092 | 0 | 0.0046 | 0.0006 | 0.0001 | 0.0005 | 0.002 | 0.0079 | -0.0088 | -0.005 | 0.0043 | 0 | 0.0065 | 0.0002 | -0.0015 | 0.012 | 0.0343 | 0.0187 | 0.6129 K | -0.0078 | 0.009 | 0.0001 | 0.001 | 0.0004 | 0.0035
Rinse1 | QC | 2024-08-12 10:07 | [6 empty] | -0.0002 | 0.0073 | 0.0012 | 0.0062 | 0.0001 | 0 | -0.0068 | 0 | 0 | 0.0001 | 0.0001 | 0.0033 | -0.0022 | -0.0116 | -0.0035 | 0.0001 | 0.0026 | -0.0047 | -0.0012 | 0.0061 | -0.0026 | 0.0072 | 0.2523 | 0.0009 | 0.0014 | -0.0001 | 0 | 0.0001 | 0.0002 | 0.0001 | 1.02 | 1.01 | 1.01 | -0.0002 | 0.0073 | 0.0012 | 0.0062 | 0.0001 | 0 | -0.0068 | 0 | 0 | 0.0001 | 0.0001 | 0.0033 | -0.0022 | -0.0116 | -0.0035 | 0.0001 | 0.0026 | -0.0047 | -0.0012 | 0.0061 | -0.0026 | 0.0072 | 0.2523 | 0.0009 | 0.0014 | 0 | 0.0001 | 0.0002 | 0.0001
Rinse | QC | 2024-08-12 10:09 | [6 empty] | 0.0002 | 0.0022 | 0.0027 | 0.0058 | 0 | -0.0001 | -0.0076 | 0.0002 | 0.0001 | 0.0002 | 0.0006 | 0.0004 | -0.0015 | -0.0028 | -0.0047 | 0 | 0.0024 | -0.0069 | -0.0006 | 0.0077 | -0.0024 | 0.0066 | 0.26 | 0.0083 | 0.0012 | 0.0013 | 0 | 0.0001 | 0.0001 | 0.0005 | 1.02 | 1 | 1.01 | 0.0002 | 0.0022 | 0.0027 | 0.0058 | 0 | -0.0001 | -0.0076 | 0.0002 | 0.0001 | 0.0002 | 0.0006 | 0.0004 | -0.0015 | -0.0028 | -0.0047 | 0 | 0.0024 | -0.0069 | -0.0006 | 0.0077 | -0.0024 | 0.0066 | 0.26 | 0.0083 | 0.0012 | 0 | 0.0001 | 0.0001 | 0.0005
IECS2 (Fe 1000 Int.) | QC | 2024-08-12 10:12 | [6 empty] | 0.0001 | 0.0243 | 0.001 | -0.0059 | -0.0001 | 0 | -0.0204 | 0.0119 | 0.0212 | 0.0084 | -0.0238 | 1022.0713 | -0.0052 | -0.0033 | 1.0819 | 0.0092 | 0.0014 | -0.0083 | 0.0071 | 0.0198 | 0.0152 | 0.0151 | 0.1608 | -0.0329 | -0.0037 | 0.0054 | 0.0005 | 0.0001 | 0.0016 | 0.0026 | 0.99 | 0.98 | 1 | 0.0001 | 0.0243 | 0.001 | -0.0059 | -0.0001 | 0 | -0.0204 | -0.0034 | 0.0059 | 0.0084 | -0.0126 | 1022.0713 | -0.0052 | -0.0033 | 0.9082 K | 0.0092 | 0.0014 | -0.0083 | 0.0071 | 0.0044 | -0.0032 | 0.0151 | 0.171 | -0.0513 | -0.0037 | 0.0005 | 0.0001 | -0.0046 | 0.0026
Rinse1 | QC | 2024-08-12 10:14 | [6 empty] | -0.0004 | 0.0017 | 0.0026 | 0.0038 | 0 | 0.0001 | -0.0087 | 0.0001 | -0.0005 | 0.0001 | -0.0006 | 0.0184 | 0.0026 | -0.0048 | -0.0081 | -0.0001 | 0.0017 | -0.0071 | -0.0009 | 0.0044 | -0.0033 | 0.0115 | 0.2718 | 0.0082 | 0.0019 | 0.0016 | 0 | 0 | 0.0002 | -0.0002 | 1.02 | 1.01 | 1.02 | -0.0004 | 0.0017 | 0.0026 | 0.0038 | 0 | 0.0001 | -0.0087 | 0.0001 | -0.0005 | 0.0001 | -0.0006 | 0.0184 | 0.0026 | -0.0048 | -0.0081 | -0.0001 | 0.0017 | -0.0071 | -0.0009 | 0.0044 | -0.0033 | 0.0115 | 0.2718 | 0.0082 | 0.0019 | 0 | 0 | 0.0002 | -0.0002
Rinse | QC | 2024-08-12 10:17 | [6 empty] | -0.0001 | -0.0055 | 0 | 0.0035 | 0 | 0 | -0.0113 | 0.0001 | 0.0001 | -0.0004 | 0 | 0.0056 | -0.012 | -0.0022 | -0.0079 | 0 | 0.0016 | -0.0075 | -0.0012 | 0.0073 | -0.0041 | 0.0104 | 0.2552 | 0.0038 | 0.0013 | 0 | 0 | 0.0001 | -0.0002 | 0.0003 | 1.01 | 1 | 1.01 | -0.0001 | -0.0055 | 0 | 0.0035 | 0 | 0 | -0.0113 | 0.0001 | 0.0001 | -0.0004 | 0 | 0.0056 | -0.012 | -0.0022 | -0.0079 | 0 | 0.0016 | -0.0075 | -0.0012 | 0.0073 | -0.0041 | 0.0104 | 0.2552 | 0.0038 | 0.0013 | 0 | 0.0001 | -0.0002 | 0.0003
IECS3 (Mn 100 Int.) | QC | 2024-08-12 10:19 | [6 empty] | 0.008 | -0.0005 | 0.0005 | -0.0012 | 0 | 0 | 0.7062 | 0.0001 | 0.0037 | 0.0149 | 0.0008 | 0.006 | -0.0111 | -0.0053 | -0.0068 | 100.3548 o | 0.0014 | -0.0086 | -0.0023 | 0.0049 | 0.0128 | 0.4991 | 0.2854 | 0.09 | 0.0441 | 0.0009 | 0 | 0 | -0.0006 | 0.0002 | 1.02 | 1.01 | 1.02 | -0.002 | -0.0005 | 0.0005 | -0.0012 | 0 | 0 | 0.7062 K | 0.0001 | -0.0013 | -0.0132 | 0.0008 | 0.006 | -0.0111 | -0.0053 | -0.0068 | 100.3548 | 0.0014 | -0.0086 | -0.0023 | 0.0049 | -0.0002 | -0.0428 | 0.2854 | 0.0328 | 0.006 | 0 | 0 | -0.0006 | 0.0002
IECS4 (Ca Mg 1000 Int.) | QC | 2024-08-12 10:22 | [6 empty] | -0.0004 | 0.0523 | 0.005 | 0.0071 | 0.0019 | 0 | 985.224 | 0.0086 | 0.0003 | 0.0014 | 0.0019 | 0.0051 | -0.014 | -0.0005 | 997.0819 | 0.0095 | 0.0011 | 0.0063 | 0.0026 | 0.0068 | 0.0032 | 0.0556 | 0.3062 | 0.0032 | 0.0198 | 0.0009 | 0.0077 | -0.0003 | 0.0015 | 0.0122 | 0.92 | 0.9 | 0.91 | -0.0004 | 0.0523 | 0.005 | 0.0071 | 0.0019 | 0 | 985.224 | 0.0086 | 0.0003 | 0.0014 | 0.0019 | 0.0051 | -0.014 | -0.0005 | 997.0819 | 0.0095 | 0.0011 | 0.0063 | 0.0026 | 0.0068 | 0.0032 | 0.0556 | 0.3062 | 0.0032 | 0.0198 | 0.0077 | -0.0003 | 0.0015 | 0.0122
Rinse1 | QC | 2024-08-12 10:24 | [6 empty] | -0.0003 | -0.0034 | 0.0042 | 0.0033 | 0 | 0 | 0.0002 | 0 | 0 | -0.0002 | -0.0007 | 0.0024 | 0.0029 | -0.0022 | 0.0025 | 0.0005 | 0.0019 | -0.0069 | -0.0011 | 0.0057 | -0.0026 | 0.0072 | 0.2733 | 0.0036 | 0.0016 | 0.0002 | 0 | 0 | 0.0001 | -0.0003 | 1.02 | 1 | 1.01 | -0.0003 | -0.0034 | 0.0042 | 0.0033 | 0 | 0 | 0.0002 | 0 | 0 | -0.0002 | -0.0007 | 0.0024 | 0.0029 | -0.0022 | 0.0025 | 0.0005 | 0.0019 | -0.0069 | -0.0011 | 0.0057 | -0.0026 | 0.0072 | 0.2733 | 0.0036 | 0.0016 | 0 | 0 | 0.0001 | -0.0003
Rinse | QC | 2024-08-12 10:27 | [6 empty] | 0 | -0.0037 | 0.0029 | 0.0024 | 0.0001 | 0.0001 | -0.0044 | 0 | 0.0003 | 0 | 0.0001 | 0.0016 | -0.0009 | -0.0016 | -0.0045 | -0.0002 | 0.0023 | -0.0064 | -0.0014 | 0.0057 | -0.002 | 0.0077 | 0.2835 | 0.0065 | 0.0006 | 0.0003 | 0 | 0.0001 | -0.0001 | 0.0005 | 1.02 | 1.01 | 1.01 | 0 | -0.0037 | 0.0029 | 0.0024 | 0.0001 | 0.0001 | -0.0044 | 0 | 0.0003 | 0 | 0.0001 | 0.0016 | -0.0009 | -0.0016 | -0.0045 | -0.0002 | 0.0023 | -0.0064 | -0.0014 | 0.0057 | -0.002 | 0.0077 | 0.2835 | 0.0065 | 0.0006 | 0 | 0.0001 | -0.0001 | 0.0005
IECS5 (Ti 20 Int.) | QC | 2024-08-12 10:29 | [6 empty] | 0.0003 | 0.0011 | 0.0023 | 0.0017 | -0.0003 | -0.0004 | 0.0882 | 0.0001 | 0.0004 | -0.0001 | -0.0011 | 0.0016 | -0.0002 | -0.0041 | -0.0027 | 0.0005 | 0.0025 | -0.0067 | -0.0028 | 0.0037 | -0.0064 | 0.0113 | 0.2589 | 0.0073 | 0.0053 | -0.0015 | 0 | 19.2902 | 0.0158 | -0.0001 | 1.03 | 1.02 | 1.03 | 0.0003 | 0.0011 | 0.0023 | 0.0017 | -0.0003 | -0.0004 | 0.0882 | 0.0001 | 0.0004 | -0.0001 | -0.0011 | 0.0016 | -0.0002 | -0.0041 | -0.0027 | 0.0005 | 0.0025 | -0.0067 | -0.0028 | 0.0037 | -0.0064 | 0.0113 | 0.2589 | 0.0073 | 0.0053 | 0 | 19.2902 K | -0.0016 | -0.0001
Rinse1 | QC | 2024-08-12 10:32 | [6 empty] | -0.0003 | -0.0059 | 0.0032 | 0.0018 | -0.0001 | 0 | -0.0065 | -0.0001 | 0 | -0.0001 | 0.0001 | 0.0009 | 0.0015 | -0.0015 | -0.0055 | 0.0003 | 0.0024 | -0.0081 | -0.0022 | 0.0069 | -0.0018 | 0.0087 | 0.2606 | 0.0057 | 0.0021 | 0.0009 | 0 | 0.0017 | 0.0001 | 0.0001 | 1.02 | 1 | 1.01 | -0.0003 | -0.0059 | 0.0032 | 0.0018 | -0.0001 | 0 | -0.0065 | -0.0001 | 0 | -0.0001 | 0.0001 | 0.0009 | 0.0015 | -0.0015 | -0.0055 | 0.0003 | 0.0024 | -0.0081 | -0.0022 | 0.0069 | -0.0018 | 0.0087 | 0.2606 | 0.0057 | 0.0021 | 0 | 0.0017 | 0.0001 | 0.0001
Rinse | QC | 2024-08-12 10:34 | [6 empty] | -0.0003 | -0.0004 | 0.0017 | 0.002 | 0 | 0.0001 | -0.006 | 0 | 0.0002 | 0.0001 | 0.0006 | 0.0006 | 0.0043 | -0.0014 | -0.0047 | 0.0002 | 0.0012 | -0.0101 | -0.002 | 0.0037 | -0.0041 | 0.0079 | 0.2656 | 0.0084 | 0.0022 | 0.0008 | 0 | 0.0004 | -0.0001 | 0.0003 | 1.02 | 1 | 1.01 | -0.0003 | -0.0004 | 0.0017 | 0.002 | 0 | 0.0001 | -0.006 | 0 | 0.0002 | 0.0001 | 0.0006 | 0.0006 | 0.0043 | -0.0014 | -0.0047 | 0.0002 | 0.0012 | -0.0101 | -0.002 | 0.0037 | -0.0041 | 0.0079 | 0.2656 | 0.0084 | 0.0022 | 0 | 0.0004 | -0.0001 | 0.0003
ICB | QC | 2024-08-12 10:37 | [6 empty] | -0.0002 | -0.0028 | 0.0017 | 0.0009 | 0 | 0 | -0.0099 | 0.0001 | 0 | 0.0001 | -0.0002 | 0.0005 | -0.0085 | -0.004 | -0.0076 | 0.0001 | 0.0011 | -0.0059 | -0.002 | 0.0056 | -0.0016 | 0.0091 | 0.2739 | 0.0069 | 0.0002 | -0.0005 | 0 | 0.0002 | -0.0005 | -0.0003 | 1.03 | 1.01 | 1.01 | -0.0002 | -0.0028 | 0.0017 | 0.0009 | 0 | 0 | -0.0099 | 0.0001 | 0 | 0.0001 | -0.0002 | 0.0005 | -0.0085 | -0.004 | -0.0076 | 0.0001 | 0.0011 | -0.0059 | -0.002 | 0.0056 | -0.0016 | 0.0091 | 0.2739 | 0.0069 | 0.0002 | 0 | 0.0002 | -0.0005 | -0.0003
ICV (SS 5ppm All) | QC | 2024-08-12 10:40 | [6 empty] | 4.9166 | 4.9594 | 4.9733 | 4.8884 | 4.9278 | 4.907 | 4.9228 | 4.9017 | 4.8861 | 4.8786 | 4.8223 | 5.0993 | 4.8883 | 4.8949 | 4.9821 | 5.0631 | 5.0078 | 4.9085 | 4.8967 | 4.962 | 4.9396 | 4.7829 | 3.9934 | 4.9437 | 4.7537 | 4.9328 | 4.9519 | 4.8977 | 4.9154 | 4.834 | 1.01 | 1 | 1.01 | 4.9161 | 4.9594 | 4.9733 | 4.8884 | 4.9278 | 4.907 | 4.9228 | 4.9016 | 4.8858 | 4.8772 | 4.8224 | 5.0993 | 4.8883 | 4.8949 | 4.9812 | 5.0631 | 5.0078 | 4.9085 | 4.8967 | 4.9619 | 4.9388 | 4.7556 | 3.9935 Q | 4.9407 | 4.7517 | 4.9519 | 4.8977 | 4.911 | 4.834
Rinse | QC | 2024-08-12 10:42 | [6 empty] | -0.0005 | -0.0027 | 0.0156 | 0.0023 | 0.0003 | 0 | -0.0115 | 0.0001 | 0.0002 | 0 | 0.0002 | 0.0014 | 0.0025 | -0.0019 | -0.0072 | 0 | 0.0515 | -0.0078 | -0.0005 | 0.0074 | -0.0049 | 0.005 | 0.2104 | 0.0279 | 0.0019 | 0.0043 | 0.0001 | 0.0014 | 0.0004 | 0.0005 | 1.02 | 1.01 | 1.02 | -0.0005 | -0.0027 | 0.0156 | 0.0023 | 0.0003 | 0 | -0.0115 | 0.0001 | 0.0002 | 0 | 0.0002 | 0.0014 | 0.0025 | -0.0019 | -0.0072 | 0 | 0.0515 | -0.0078 | -0.0005 | 0.0074 | -0.0049 | 0.005 | 0.2104 | 0.0279 | 0.0019 | 0.0001 | 0.0014 | 0.0004 | 0.0005
CCB | QC | 2024-08-12 10:45 | [6 empty] | 0.0001 | -0.0033 | 0.0069 | 0.0021 | 0 | -0.0001 | -0.0087 | 0.0001 | 0.0002 | 0.0003 | 0.0007 | 0.0016 | 0.0056 | -0.0024 | -0.0081 | 0 | 0.015 | -0.0087 | -0.0016 | 0.0068 | -0.0041 | 0.0136 | 0.2181 | 0.0149 | 0.0017 | 0.0007 | 0 | 0.0004 | 0 | 0.0001 | 1.02 | 1 | 1 | 0.0001 | -0.0033 | 0.0069 | 0.0021 | 0 | -0.0001 | -0.0087 | 0.0001 | 0.0002 | 0.0003 | 0.0007 | 0.0016 | 0.0056 | -0.0024 | -0.0081 | 0 | 0.015 | -0.0087 | -0.0016 | 0.0068 | -0.0041 | 0.0136 | 0.2181 | 0.0149 | 0.0017 | 0 | 0.0004 | 0 | 0.0001
0.5ppm All (LLQC) | QC | 2024-08-12 10:48 | [6 empty] | 0.4775 | 0.4891 | 0.4742 | 0.4982 | 0.488 | 0.4825 | 0.4795 | 0.4851 | 0.4875 | 0.4868 | 0.4735 | 0.5105 | 0.4291 | 0.4893 | 0.4734 | 0.5067 | 0.4554 | 0.5004 | 0.4854 | 0.4873 | 0.4879 | 0.469 | 0.4515 | 0.4714 | 0.4774 | 0.4729 | 0.4898 | 0.4755 | 0.4779 | 0.4772 | 1.03 | 1.01 | 1.02 | 0.4774 | 0.4891 | 0.4742 | 0.4982 | 0.488 | 0.4825 | 0.4795 | 0.4851 | 0.4875 | 0.4866 | 0.4735 | 0.5105 | 0.4291 Q | 0.4893 | 0.4733 | 0.5067 | 0.4554 | 0.5004 | 0.4854 | 0.4873 | 0.4878 | 0.4663 | 0.4515 | 0.4711 | 0.4772 | 0.4898 | 0.4755 | 0.4775 | 0.4772
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-12 10:50 | [6 empty] | 0.474 | 0.4874 | 0.474 | 0.4981 | 0.4867 | 0.4825 | 0.4804 | 0.484 | 0.4873 | 0.4863 | 0.4734 | 0.5087 | 0.4556 | 0.4888 | 0.4712 | 0.5065 | 0.4693 | 0.5019 | 0.4846 | 0.4855 | 0.4836 | 0.4719 | 0.4838 | 0.4836 | 0.4774 | 0.4752 | 0.4887 | 0.4747 | 0.4776 | 0.476 | 1.03 | 1.01 | 1.02 | 0.4739 | 0.4874 | 0.474 | 0.4981 | 0.4867 | 0.4825 | 0.4804 | 0.484 | 0.4873 | 0.4862 | 0.4734 | 0.5087 | 0.4556 | 0.4888 | 0.4711 | 0.5065 | 0.4693 | 0.5019 | 0.4846 | 0.4855 | 0.4836 | 0.4692 | 0.4838 | 0.4833 | 0.4772 | 0.4887 | 0.4747 | 0.4771 | 0.476
CCV (5ppm All) | QC | 2024-08-12 10:53 | [6 empty] | 4.8231 | 4.8687 | 4.8814 | 4.9065 | 4.8478 | 4.841 | 4.8818 | 4.89 | 4.893 | 4.8751 | 4.8237 | 5.0318 | 4.8233 | 4.9498 | 4.8908 | 4.9886 | 4.9505 | 4.8363 | 4.8853 | 4.9434 | 4.9112 | 4.7859 | 4.0935 Q | 4.9207 | 4.9724 | 4.857 | 4.858 | 4.8225 | 4.8545 | 4.855 | 1.01 | 1 | 1.01 | 4.8226 | 4.8687 | 4.8814 | 4.9065 | 4.8478 | 4.841 | 4.8818 | 4.8899 | 4.8927 | 4.8737 | 4.8237 | 5.0318 | 4.8233 | 4.9498 | 4.8899 | 4.9886 | 4.9505 | 4.8363 | 4.8853 | 4.9433 | 4.9105 | 4.7589 | 4.0935 | 4.9177 | 4.9706 | 4.858 | 4.8225 | 4.8501 | 4.855
CCB | QC | 2024-08-12 10:55 | [6 empty] | -0.0001 | -0.0034 | 0.0186 | 0.0034 | 0.0001 | 0.0001 | -0.0121 | 0.0001 | 0.0003 | -0.0002 | 0.0001 | 0.0009 | -0.005 | -0.0026 | -0.0077 | 0 | 0.0522 | -0.0087 | -0.002 | 0.009 | -0.0014 | 0.0065 | 0.2233 | 0.0257 | 0.0025 | 0.0036 | 0.0001 | 0.0016 | -0.0003 | -0.0002 | 1.02 | 1.01 | 1.02 | -0.0001 | -0.0034 | 0.0186 | 0.0034 | 0.0001 | 0.0001 | -0.0121 | 0.0001 | 0.0003 | -0.0002 | 0.0001 | 0.0009 | -0.005 | -0.0026 | -0.0077 | 0 | 0.0522 | -0.0087 | -0.002 | 0.009 | -0.0014 | 0.0065 | 0.2233 | 0.0257 | 0.0025 | 0.0001 | 0.0016 | -0.0003 | -0.0002
Rinse | QC | 2024-08-12 13:08 | [6 empty] | -0.0003 | -0.002 | 0.0012 | -0.0012 | -0.0001 | 0.0001 | -0.0105 | 0.0001 | 0.0004 | -0.0002 | -0.0005 | 0.0007 | -0.0009 | -0.0033 | -0.0105 | 0 | 0.0017 | -0.0074 | -0.0019 | 0.007 | -0.0033 | 0.0083 | 0.2475 | 0.0038 | -0.0003 | -0.0001 | 0 | -0.0001 | -0.0002 | 0.0003 | 1.04 | 1.01 | 1.01 | -0.0003 | -0.002 | 0.0012 | -0.0012 | -0.0001 | 0.0001 | -0.0105 | 0.0001 | 0.0004 | -0.0002 | -0.0005 | 0.0007 | -0.0009 | -0.0033 | -0.0105 | 0 | 0.0017 | -0.0074 | -0.0019 | 0.007 | -0.0033 | 0.0083 | 0.2475 | 0.0038 | -0.0003 | 0 | -0.0001 | -0.0002 | 0.0003
CCV (5ppm All) | QC | 2024-08-12 13:10 | [6 empty] | 4.8182 | 4.8901 | 4.8594 | 4.8725 | 4.8541 | 4.8565 | 4.8712 | 4.8739 | 4.873 | 4.8599 | 4.8259 | 5.0357 | 4.8209 | 4.9536 | 4.8629 | 4.9944 | 4.9168 | 4.8352 | 4.8786 | 5.0669 | 4.8845 | 4.8124 | 4.8472 | 4.9097 | 4.895 | 4.8443 | 4.8586 | 4.817 | 4.8464 | 4.84 | 1.03 | 1.01 | 1.01 | 4.8177 | 4.8901 | 4.8594 | 4.8725 | 4.8541 | 4.8565 | 4.8712 | 4.8738 | 4.8726 | 4.8585 | 4.826 | 5.0357 | 4.8209 | 4.9536 | 4.862 | 4.9944 | 4.9168 | 4.8352 | 4.8786 | 5.0668 | 4.8838 | 4.7855 | 4.8472 | 4.9068 | 4.8931 | 4.8586 | 4.817 | 4.842 | 4.84
CCB | QC | 2024-08-12 13:13 | [6 empty] | -0.0001 | -0.0027 | 0.0173 | 0.001 | 0.0002 | 0.0001 | -0.0143 | 0.0001 | 0 | 0.0001 | -0.0006 | 0.0002 | 0.0017 | -0.0024 | -0.0116 | 0.0001 | 0.0578 | -0.0102 | -0.0024 | 0.0009 | -0.0037 | 0.0022 | 0.226 | 0.0309 | 0.0017 | 0.0035 | 0.0001 | 0.0018 | 0 | 0 | 1.03 | 1.01 | 1.01 | -0.0001 | -0.0027 | 0.0173 | 0.001 | 0.0002 | 0.0001 | -0.0143 | 0.0001 | 0 | 0.0001 | -0.0006 | 0.0002 | 0.0017 | -0.0024 | -0.0116 | 0.0001 | 0.0578 | -0.0102 | -0.0024 | 0.0009 | -0.0037 | 0.0022 | 0.226 | 0.0309 | 0.0017 | 0.0001 | 0.0018 | 0 | 0
08122024 LB | Sample | 2024-08-12 13:15 | [6 empty] | -0.0003 | -0.0028 | 0.0055 | 0.0004 | 0 | 0.0001 | 0.0117 | 0.0001 | 0.0001 | 0 | 0 | 0.0036 | -0.0031 | -0.0121 | -0.0096 | 0.0002 | 0.0163 | 0.0001 | -0.0029 | 0.0057 | -0.0038 | 0.0131 | 0.2045 | 0.0121 | 0.0053 | 0.0008 | 0 | 0.0005 | -0.0004 | 0.0024 | 1.04 | 1.01 | 1.02 | -0.0003 | -0.0028 | 0.0055 | 0.0004 | 0 | 0.0001 | 0.0117 | 0.0001 | 0.0001 | 0 | 0 | 0.0036 | -0.0031 | -0.0121 | -0.0096 | 0.0002 | 0.0163 | 0.0001 | -0.0029 | 0.0057 | -0.0038 | 0.0131 | 0.2045 | 0.0121 | 0.0053 | 0 | 0.0005 | -0.0004 | 0.0024
10x 08122024 LCS | Sample | 2024-08-12 13:18 | [6 empty] | 0.0091 | 0.4808 | 0.2381 | 0.049 | 0.2464 | 0.0484 | 0.4819 | 0.097 | 0.4828 | 0.2427 | 0.2363 | 0.5098 | 0.4665 | -0.0031 | 0.4749 | 0.1016 | 0.09 | 0.5027 | 0.2372 | 0.105 | 0.2365 | 0.2418 | 0.2395 | 0.2452 | 0.0953 | 0.0952 | 0.099 | 0.096 | 0.2401 | 0.0955 | 1.05 | 1.01 | 1.02 | 0.0091 | 0.4808 | 0.2381 | 0.049 | 0.2464 | 0.0484 | 0.4819 | 0.097 | 0.4828 | 0.2427 | 0.2364 | 0.5098 | 0.4665 | -0.0031 | 0.4748 | 0.1016 | 0.09 | 0.5027 | 0.2372 | 0.105 | 0.2364 | 0.2413 | 0.2395 | 0.2451 | 0.0953 | 0.099 | 0.096 | 0.24 | 0.0955
10x 08122024 LCS- Analytical Duplicate | Sample | 2024-08-12 13:20 | [6 empty] | 0.0096 | 0.4706 | 0.2391 | 0.0478 | 0.2449 | 0.0478 | 0.4823 | 0.0967 | 0.4839 | 0.2431 | 0.2358 | 0.5085 | 0.4641 | -0.0035 | 0.4728 | 0.1003 | 0.0919 | 0.4959 | 0.238 | 0.1112 | 0.235 | 0.246 | 0.272 | 0.2403 | 0.0952 | 0.0953 | 0.0981 | 0.0959 | 0.2402 | 0.096 | 1.05 | 1.01 | 1.02 | 0.0096 | 0.4706 | 0.2391 | 0.0478 | 0.2449 | 0.0478 | 0.4823 | 0.0967 | 0.4839 | 0.2431 | 0.2358 | 0.5085 | 0.4641 | -0.0035 | 0.4727 | 0.1003 | 0.0919 | 0.4959 | 0.238 | 0.1112 | 0.2349 | 0.2455 | 0.272 | 0.2403 | 0.0951 | 0.0981 | 0.0959 | 0.2402 | 0.096
5x 08122024 LCS | Sample | 2024-08-12 13:23 | [6 empty] | 0.0188 | 0.9633 | 0.4861 | 0.0982 | 0.4967 | 0.0972 | 0.9646 | 0.196 | 0.9806 | 0.4875 | 0.4754 | 1.0228 | 0.9475 | -0.0056 | 0.9729 | 0.2039 | 0.1772 | 1.0081 | 0.4832 | 0.2067 | 0.4829 | 0.4838 | 0.4015 | 0.4812 | 0.191 | 0.195 | 0.1984 | 0.1928 | 0.4843 | 0.1926 | 1.04 | 1.01 | 1.02 | 0.0188 | 0.9633 | 0.4861 | 0.0982 | 0.4967 | 0.0972 | 0.9646 | 0.196 | 0.9805 | 0.4875 | 0.4754 | 1.0228 | 0.9475 | -0.0056 | 0.9728 | 0.2039 | 0.1772 | 1.0081 | 0.4832 | 0.2067 | 0.4828 | 0.4827 | 0.4015 | 0.481 | 0.1909 | 0.1984 | 0.1928 | 0.4842 | 0.1926
08122024 LCS | Sample | 2024-08-12 13:25 | [6 empty] | 0.0952 | 4.7787 | 2.4618 | 0.4929 | 2.4434 | 0.4802 | 4.8454 | 0.9851 | 4.9105 | 2.4234 | 2.3844 | 5.0502 | 4.9365 | -0.0039 | 4.9601 | 1.0057 | 0.9308 | 4.9235 | 2.437 | 1.0305 | 2.4454 | 2.4234 | 2.2314 | 2.4803 | 0.9752 | 0.9776 | 0.9806 | 0.9708 | 2.4359 | 0.9729 | 1.04 | 1.01 | 1.02 | 0.0951 | 4.7787 | 2.4618 | 0.4929 | 2.4434 | 0.4802 | 4.8454 | 0.985 | 4.9104 | 2.4231 | 2.3844 | 5.0502 | 4.9365 | -0.0039 | 4.9593 | 1.0057 | 0.9308 | 4.9235 | 2.437 | 1.0304 | 2.4452 | 2.418 | 2.2314 | 2.4796 | 0.9748 | 0.9806 | 0.9708 | 2.435 | 0.9729
08122024 LCS - Analytical Duplicate | Sample | 2024-08-12 13:28 | [6 empty] | 0.0941 | 4.779 | 2.4543 | 0.4892 | 2.4386 | 0.4797 | 4.8193 | 0.9719 | 4.8743 | 2.4058 | 2.3731 | 5.0517 | 4.8974 | -0.0045 | 4.9237 | 1.0044 | 0.9468 | 4.9058 | 2.421 | 1.0197 | 2.4298 | 2.4086 | 1.8238 | 2.4654 | 0.9712 | 0.9741 | 0.978 | 0.9649 | 2.4257 | 0.967 | 1.03 | 1.01 | 1.02 | 0.094 | 4.779 | 2.4543 | 0.4892 | 2.4386 | 0.4797 | 4.8193 | 0.9718 | 4.8742 | 2.4055 | 2.3732 | 5.0517 | 4.8974 | -0.0045 | 4.9228 | 1.0044 | 0.9468 | 4.9058 | 2.421 | 1.0196 | 2.4296 | 2.4032 | 1.8238 | 2.4647 | 0.9709 | 0.978 | 0.9649 | 2.4248 | 0.967
Rinse | QC | 2024-08-12 13:30 | [6 empty] | -0.0001 | -0.0041 | 0.0158 | 0.0009 | 0 | 0 | -0.0098 | 0.0001 | 0.0002 | 0.0003 | 0.0001 | 0.0001 | -0.0031 | -0.004 | -0.0105 | 0.0001 | 0.0333 | -0.0089 | -0.0012 | 0.0058 | -0.0052 | 0.0052 | 0.2152 | 0.0216 | 0.0022 | 0.0029 | 0 | 0.0003 | 0.0002 | 0.0005 | 1.04 | 1.02 | 1.03 | -0.0001 | -0.0041 | 0.0158 | 0.0009 | 0 | 0 | -0.0098 | 0.0001 | 0.0002 | 0.0003 | 0.0001 | 0.0001 | -0.0031 | -0.004 | -0.0105 | 0.0001 | 0.0333 | -0.0089 | -0.0012 | 0.0058 | -0.0052 | 0.0052 | 0.2152 | 0.0216 | 0.0022 | 0 | 0.0003 | 0.0002 | 0.0005
CCV (5ppm All) | QC | 2024-08-12 13:33 | [6 empty] | 4.8209 | 4.9388 | 4.8716 | 4.885 | 4.9001 | 4.8524 | 4.8688 | 4.8824 | 4.8837 | 4.8647 | 4.8335 | 5.0788 | 4.8391 | 4.9926 | 4.8707 | 5.058 | 4.9522 | 4.8706 | 4.8827 | 5.0859 | 4.8917 | 4.8315 | 4.8612 | 4.9296 | 4.9474 | 4.8505 | 4.905 | 4.8105 | 4.854 | 4.8826 | 1.03 | 1 | 1.01 | 4.8204 | 4.9388 | 4.8716 | 4.885 | 4.9001 | 4.8524 | 4.8688 | 4.8823 | 4.8834 | 4.8633 | 4.8335 | 5.0788 | 4.8391 | 4.9926 | 4.8699 | 5.058 | 4.9522 | 4.8706 | 4.8827 | 5.0858 | 4.891 | 4.8042 | 4.8612 | 4.9266 | 4.9455 | 4.905 | 4.8105 | 4.8496 | 4.8826
CCB | QC | 2024-08-12 13:35 | [6 empty] | -0.0004 | -0.004 | 0.0187 | 0.0013 | 0 | 0.0001 | -0.0133 | 0.0001 | 0.0002 | -0.0002 | -0.0004 | 0.0012 | 0.0013 | -0.0035 | -0.0112 | 0.0003 | 0.0526 | -0.0097 | -0.0013 | 0.007 | -0.003 | 0.0072 | 0.207 | 0.0239 | 0.0015 | 0.0028 | 0.0001 | 0.0015 | -0.0001 | 0 | 1.04 | 1.01 | 1.02 | -0.0004 | -0.004 | 0.0187 | 0.0013 | 0 | 0.0001 | -0.0133 | 0.0001 | 0.0002 | -0.0002 | -0.0004 | 0.0012 | 0.0013 | -0.0035 | -0.0112 | 0.0003 | 0.0526 | -0.0097 | -0.0013 | 0.007 | -0.003 | 0.0072 | 0.207 | 0.0239 | 0.0015 | 0.0001 | 0.0015 | -0.0001 | 0
0.5ppm All (LLQC) | QC | 2024-08-12 13:38 | [6 empty] | 0.4711 | 0.4883 | 0.4708 | 0.4908 | 0.4877 | 0.4833 | 0.4748 | 0.48 | 0.4831 | 0.4821 | 0.4705 | 0.5111 | 0.4444 | 0.4924 | 0.4602 | 0.5063 | 0.4647 | 0.5038 | 0.481 | 0.4997 | 0.4798 | 0.4682 | 0.4662 | 0.4758 | 0.4721 | 0.4691 | 0.49 | 0.4709 | 0.4737 | 0.4716 | 1.05 | 1.01 | 1.02 | 0.4711 | 0.4883 | 0.4708 | 0.4908 | 0.4877 | 0.4833 | 0.4748 | 0.48 | 0.483 | 0.4819 | 0.4705 | 0.5111 | 0.4444 Q | 0.4924 | 0.4601 | 0.5063 | 0.4647 | 0.5038 | 0.481 | 0.4997 | 0.4797 | 0.4654 | 0.4662 | 0.4755 | 0.4719 | 0.49 | 0.4709 | 0.4733 | 0.4716
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-12 13:40 | [6 empty] | 0.4738 | 0.4834 | 0.4806 | 0.4953 | 0.4826 | 0.4786 | 0.4786 | 0.4839 | 0.4872 | 0.4855 | 0.4751 | 0.5065 | 0.4521 | 0.4892 | 0.468 | 0.502 | 0.4749 | 0.5006 | 0.484 | 0.5017 | 0.4837 | 0.4717 | 0.4798 | 0.4813 | 0.4758 | 0.4761 | 0.4842 | 0.4745 | 0.478 | 0.476 | 1.04 | 1.02 | 1.02 | 0.4738 | 0.4834 | 0.4806 | 0.4953 | 0.4826 | 0.4786 | 0.4786 | 0.4839 | 0.4871 | 0.4854 | 0.4751 | 0.5065 | 0.4521 | 0.4892 | 0.468 | 0.502 | 0.4749 | 0.5006 | 0.484 | 0.5017 | 0.4836 | 0.469 | 0.4799 | 0.481 | 0.4757 | 0.4842 | 0.4745 | 0.4775 | 0.476
CCS1 (SS 5ppm) | QC | 2024-08-12 13:43 | [6 empty] | 4.9295 | 4.9938 | 5.0039 | 4.9145 | 4.9299 | 4.9051 | 4.9229 | 4.9142 | 4.9041 | 4.8835 | 4.8475 | 5.0823 | 4.9027 | 4.9604 | 4.9662 | 5.0749 | 5.0853 | 4.9409 | 4.9008 | 5.1356 | 4.9513 | 4.8379 | 4.8713 | 4.9543 | 4.8066 | 4.9497 | 4.9634 | 4.9054 | 4.9244 | 4.8307 | 1.02 | 1 | 1.01 | 4.929 | 4.9938 | 5.0039 | 4.9145 | 4.9299 | 4.9051 | 4.9229 | 4.9142 | 4.9038 | 4.882 | 4.8475 | 5.0823 | 4.9027 | 4.9604 | 4.9653 | 5.0749 | 5.0853 | 4.9409 | 4.9008 | 5.1355 | 4.9505 | 4.8105 | 4.8714 | 4.9513 | 4.8047 | 4.9634 | 4.9054 | 4.92 | 4.8307
CCB | QC | 2024-08-12 13:45 | [6 empty] | -0.0001 | -0.0018 | 0.0204 | 0.0015 | 0.0001 | 0.0001 | -0.0107 | 0.0002 | 0.0002 | -0.0001 | 0.0001 | 0.0007 | 0.0164 | -0.0039 | -0.0117 | 0 | 0.0507 | -0.0085 | -0.0023 | 0.0046 | -0.0017 | 0.0076 | 0.1914 | 0.0245 | 0.0024 | 0.0054 | 0.0001 | 0.0014 | 0.0002 | 0 | 1.03 | 1.01 | 1.02 | -0.0001 | -0.0018 | 0.0204 | 0.0015 | 0.0001 | 0.0001 | -0.0107 | 0.0002 | 0.0002 | -0.0001 | 0.0001 | 0.0007 | 0.0164 | -0.0039 | -0.0117 | 0 | 0.0507 | -0.0085 | -0.0023 | 0.0046 | -0.0017 | 0.0076 | 0.1914 | 0.0245 | 0.0024 | 0.0001 | 0.0014 | 0.0002 | 0
MQ-Blank-01 | Sample | 2024-08-12 13:48 | [6 empty] | -0.0001 | -0.0036 | 0.0016 | 0.0005 | -0.0001 | -0.0001 | -0.0137 | 0.0002 | 0.0003 | 0 | -0.0002 | 0 | -0.0096 | -0.0026 | -0.0129 | 0.0002 | 0.0049 | -0.009 | -0.002 | 0.0052 | -0.0055 | 0.0049 | 0.2592 | 0.0073 | 0.003 | -0.0005 | 0 | 0.0002 | -0.0001 | 0.0001 | 1.07 | 1.03 | 1.04 | -0.0001 | -0.0036 | 0.0016 | 0.0005 | -0.0001 | -0.0001 | -0.0137 | 0.0002 | 0.0003 | 0 | -0.0002 | 0 | -0.0096 | -0.0026 | -0.0129 | 0.0002 | 0.0049 | -0.009 | -0.002 | 0.0052 | -0.0055 | 0.0049 | 0.2592 | 0.0073 | 0.003 | 0 | 0.0002 | -0.0001 | 0.0001
Rinse | QC | 2024-08-12 13:50 | [6 empty] | -0.0002 | -0.0034 | 0.0082 | 0.0004 | -0.0001 | 0 | -0.0106 | 0.0001 | 0.0002 | 0 | -0.0003 | 0.0001 | 0.0067 | -0.0046 | -0.0112 | -0.0002 | 0.0117 | -0.0086 | -0.002 | 0.0057 | -0.0061 | 0.0096 | 0.2206 | 0.0087 | 0.0025 | 0.0021 | 0 | 0.0003 | -0.0004 | 0.0004 | 1.04 | 1.02 | 1.03 | -0.0002 | -0.0034 | 0.0082 | 0.0004 | -0.0001 | 0 | -0.0106 | 0.0001 | 0.0002 | 0 | -0.0003 | 0.0001 | 0.0067 | -0.0046 | -0.0112 | -0.0002 | 0.0117 | -0.0086 | -0.002 | 0.0057 | -0.0061 | 0.0096 | 0.2206 | 0.0087 | 0.0025 | 0 | 0.0003 | -0.0004 | 0.0004
CCV (5ppm All) | QC | 2024-08-12 13:53 | [6 empty] | 4.7822 | 4.8565 | 4.8297 | 4.8389 | 4.8123 | 4.8026 | 4.8308 | 4.8388 | 4.8399 | 4.8193 | 4.7916 | 4.9808 | 4.7889 | 4.9131 | 4.8231 | 4.958 | 4.9081 | 4.7942 | 4.8346 | 5.0437 | 4.8551 | 4.7704 | 3.9963 Q | 4.8777 | 4.9095 | 4.8052 | 4.8182 | 4.7734 | 4.8078 | 4.8094 | 1.03 | 1.02 | 1.03 | 4.7817 | 4.8565 | 4.8297 | 4.8389 | 4.8123 | 4.8026 | 4.8308 | 4.8387 | 4.8396 | 4.8179 | 4.7917 | 4.9808 | 4.7889 | 4.9131 | 4.8222 | 4.958 | 4.9081 | 4.7942 | 4.8346 | 5.0436 | 4.8544 | 4.7437 | 3.9964 | 4.8748 | 4.9077 | 4.8182 | 4.7734 | 4.8035 | 4.8094
CCB | QC | 2024-08-12 13:55 | [6 empty] | -0.0001 | -0.0038 | 0.0159 | 0.0021 | 0.0001 | 0.0001 | -0.0118 | 0.0001 | -0.0001 | 0 | 0.0003 | -0.0001 | 0.0049 | -0.0053 | -0.0066 | 0.0002 | 0.0522 | -0.0075 | -0.0022 | 0.0076 | -0.0049 | 0.0097 | 0.2125 | 0.0284 | 0.0031 | 0.004 | 0.0001 | 0.0016 | 0 | 0.0001 | 1.04 | 1.02 | 1.02 | -0.0001 | -0.0038 | 0.0159 | 0.0021 | 0.0001 | 0.0001 | -0.0118 | 0.0001 | -0.0001 | 0 | 0.0003 | -0.0001 | 0.0049 | -0.0053 | -0.0066 | 0.0002 | 0.0522 | -0.0075 | -0.0022 | 0.0076 | -0.0049 | 0.0097 | 0.2125 | 0.0284 | 0.0031 | 0.0001 | 0.0016 | 0 | 0.0001
0.5% NaCl LFP Battery-TM | Sample | 2024-08-12 13:58 | [6 empty] | -0.0007 | 557.7988 o | 0.0258 | 0.0046 | 0.0923 | -0.0001 | 2.6321 | 0.0216 | 0.0314 | 0.1444 | 0.0072 | 1314.2552 o | 0.6112 | 2.4354 | 1.7429 | 5.0189 | 0.037 | 1471.3982 o | 17.2464 | 57.3057 | 0.0306 | 2.5938 | 0.1881 | -0.02 | 3.8931 | 0.0188 | 0.0362 | 0.139 | 0.0912 | 0.2196 | 0.91 | 0.92 | 0.93 | -0.0012 | 557.7988 | 0.0258 | 0.0046 | 0.0923 | -0.0001 | 2.6321 | 0.0019 | 0.0114 | 0.143 | 0.0216 | 1314.2552 | 0.6112 | 2.4354 | 1.5195 | 5.0189 | 0.037 | 1471.3982 | 17.2464 | 57.286 | 0.0063 | 2.5667 | 0.2012 | -0.0521 | 3.8912 | 0.0362 | 0.139 | 0.0832 | 0.2196
1% NaCl LFP Battery-TM | Sample | 2024-08-12 14:00 | [6 empty] | -0.0013 | 467.9349 o | 0.0215 | 0.011 | 0.0733 | -0.0001 | 4.6387 | 0.0181 | 0.0157 | 0.145 | 0.015 | 989.5368 | 1.0757 | 5.0366 | 1.382 | 4.4929 | 0.0272 | #### | 13.531 | 111.2412 | 0.0252 | 3.4289 | 0.1917 | -0.0115 | 5.4923 | 0.017 | 0.0503 | 0.1286 | 0.0742 | 0.1464 | 0.89 | 0.89 | 0.9 | -0.0017 | 467.9349 | 0.0215 | 0.011 | 0.0733 | -0.0001 | 4.6387 | 0.0032 | 0.0006 | 0.1437 | 0.0259 | 989.5368 | 1.0757 | 5.0366 | 1.2138 | 4.4929 | 0.0272 |  | 13.531 | 111.2263 | 0.0068 | 3.4047 | 0.2015 | -0.0365 | 5.4906 | 0.0503 | 0.1286 | 0.0682 | 0.1464
5% NaCl LFP Battery-TM | Sample | 2024-08-12 14:03 | [6 empty] | -0.0012 | 814.1376 o | 0.0251 | 0.0328 | 0.0692 | -0.0002 | 18.1943 | 0.0159 | 0.0163 | 0.2402 | 0.0691 | 858.3883 | 5.3255 | 23.4352 | 1.6083 | 3.4276 | 0.0422 | #### | 13.4173 | 196.4397 | 0.0215 | 15.3578 | 0.1964 | -0.0162 | 9.7102 | 0.0229 | 0.2306 | 0.2101 | 0.1121 | 0.0789 | 0.8 | 0.76 | 0.76 | -0.0016 | 814.1376 | 0.0251 | 0.0328 | 0.0692 | -0.0002 | 18.1943 | 0.003 | 0.0032 | 0.2392 | 0.0786 | 858.3883 | 5.3255 | 23.4352 | 1.4624 | 3.4276 | 0.0422 |  | 13.4173 | 196.4268 | 0.0057 | 15.3393 | 0.205 | -0.0417 | 9.7089 | 0.2306 | 0.2101 | 0.1067 | 0.0789
10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:05 | [6 empty] | -0.0013 | 1293.5693 o | 0.0355 | 0.0604 | 0.0904 | -0.0002 | 33.5257 | 0.0156 | 0.0194 | 0.343 | 0.1258 | 1055.2288 | 13.6907 | 32.0557 | 2.4039 | 3.8145 | 0.0519 | #### | 17.1599 | 159.2712 | 0.0236 | 29.3845 | 0.2915 | -0.0268 | 10.0721 | 0.023 | 0.4585 | 0.318 | 0.1528 | 0.1659 | 0.72 | 0.66 | 0.66 | -0.0017 | 1293.5693 | 0.0355 | 0.0604 | 0.0904 | -0.0002 | 33.5257 | -0.0002 | 0.0034 | 0.3419 | 0.1374 | 1055.2288 | 13.6907 | 32.0557 | 2.2245 | 3.8145 | 0.0519 |  | 17.1599 | 159.2554 | 0.0041 | 29.3639 | 0.302 | -0.0609 | 10.0706 | 0.4585 | 0.318 | 0.1462 | 0.1659
15% NaCl LFP Battery-TM | Sample | 2024-08-12 14:08 | [6 empty] | -0.0013 | 963.6210 o | 0.0262 | 0.1545 | 0.0046 | -0.0003 | 46.0159 | 0.0128 | 0.0121 | 0.3606 | 0.1568 | 753.6171 | 24.654 | 14.3553 | 4.8368 | 0.7849 | 0.0206 | #### | 10.7765 | 16.525 | 0.0092 | 41.6182 | 0.1931 | -0.0094 | 0.7096 | 0.0229 | 0.677 | 0.2669 | 0.0704 | 0.0409 | 0.68 | 0.58 | 0.59 | -0.0014 | 963.621 | 0.0262 | 0.1545 | 0.0046 | -0.0003 | 46.0159 | 0.0015 | 0.0008 | 0.3604 | 0.1651 | 753.6171 | 24.654 | 14.3553 | 4.7087 | 0.7849 | 0.0206 |  | 10.7765 | 16.5137 | -0.0045 | 41.614 | 0.2007 | -0.0331 | 0.7093 | 0.677 | 0.2669 | 0.0657 | 0.0409
0.5% NaCl MMC Battery-TM | Sample | 2024-08-12 14:10 | [6 empty] | -0.0006 | 456.0864 o | 0.0159 | 0.0858 | 0.0079 | -0.0002 | 2.7328 | 0.0221 | 0.012 | 0.2569 | 0.0424 | 1202.9261 o | 0.7907 | 1.2715 | 3.4968 | 1.0584 | 0.0114 | 1666.8729 o | 13.7941 | 18.9733 | 0.0262 | 1.7674 | 0.2191 | -0.0282 | 0.5755 | 0.0177 | 0.0324 | 0.1634 | 0.0523 | 0.0646 | 0.96 | 0.91 | 0.92 | -0.0007 | 456.0864 | 0.0159 | 0.0858 | 0.0079 | -0.0002 | 2.7328 | 0.004 | -0.0061 | 0.2566 | 0.0556 | 1202.9261 | 0.7907 | 1.2715 | 3.2923 | 1.0584 | 0.0114 | 1666.8729 | 13.7941 | 18.9552 | 0.0044 | 1.7617 | 0.2311 | -0.055 | 0.5751 | 0.0324 | 0.1634 | 0.0449 | 0.0646
1% NaCl MMC Battery-TM | Sample | 2024-08-12 14:13 | [6 empty] | -0.0003 | 656.2899 o | 0.0243 | 0.1069 | 0.0083 | -0.0001 | 4.0848 | 0.0227 | 0.0283 | 0.3922 | 0.1217 | 1271.6013 o | 1.1073 | 1.3916 | 5.1384 | 1.1126 | 0.0174 | #### | 15.8358 | 22.5221 | 0.0259 | 2.9489 | 0.2564 | -0.0292 | 0.6422 | 0.0252 | 0.0542 | 0.2544 | 0.0744 | 0.0698 | 0.93 | 0.88 | 0.88 | -0.0004 | 656.2899 | 0.0243 | 0.1069 | 0.0083 | -0.0001 | 4.0848 | 0.0036 | 0.0092 | 0.3919 | 0.1357 | 1271.6013 | 1.1073 | 1.3916 | 4.9223 | 1.1126 | 0.0174 |  | 15.8358 | 22.503 | 0.0029 | 2.9429 | 0.2691 | -0.0593 | 0.6418 | 0.0542 | 0.2544 | 0.0666 | 0.0698
5% NaCl MMC Battery-TM | Sample | 2024-08-12 14:15 | [6 empty] | -0.0014 | 793.9437 o | 0.0196 | 0.1193 | 0.006 | -0.0002 | 15.8604 | 0.0186 | 0.1996 | 0.4086 | 0.153 | 971.2853 | 7.1104 | 8.1921 | 5.1148 | 0.8952 | 0.0178 | #### | 13.0225 | 24.0027 | 0.0165 | 13.5148 | 0.2393 | -0.0268 | 0.6666 | 0.0236 | 0.2582 | 0.2724 | 0.0744 | 0.0548 | 0.82 | 0.74 | 0.75 | -0.0015 | 793.9437 | 0.0196 | 0.1193 | 0.006 | -0.0002 | 15.8604 | 0.004 | 0.185 | 0.4084 | 0.1636 | 971.2853 | 7.1104 | 8.1921 | 4.9497 | 0.8952 | 0.0178 |  | 13.0225 | 23.9881 | -0.0011 | 13.5099 | 0.249 | -0.0527 | 0.6662 | 0.2582 | 0.2724 | 0.0683 | 0.0548
10% NaCl MMC Battery-TM | Sample | 2024-08-12 14:18 | [6 empty] | -0.0013 | 1049.6128 o | 0.0364 | 0.1497 | 0.0045 | -0.0002 | 37.1117 | 0.016 | 0.0278 | 0.4247 | 0.1894 | 900.7167 | 17.0243 | 15.7913 | 5.3511 | 0.8891 | 0.0191 | #### | 12.539 | 20.2382 | 0.0129 | 33.0079 | 0.2222 | -0.0283 | 0.7381 | 0.0242 | 0.5309 | 0.3155 | 0.0826 | 0.0485 | 0.74 | 0.63 | 0.63 | -0.0014 | 1049.6128 | 0.0364 | 0.1497 | 0.0045 | -0.0002 | 37.1117 | 0.0025 | 0.0143 | 0.4244 | 0.1993 | 900.7167 | 17.0243 | 15.7913 | 5.1979 | 0.8891 | 0.0191 |  | 12.539 | 20.2247 | -0.0034 | 33.0031 | 0.2312 | -0.0555 | 0.7378 | 0.5309 | 0.3155 | 0.0769 | 0.0485
15% NaCl MMC Battery-TM | Sample | 2024-08-12 14:20 | [6 empty] | -0.0013 | 958.1671 o | 0.0176 | 0.0771 | 0.0719 | -0.0001 | 45.0598 | 0.0159 | 0.0164 | 0.1884 | 0.0766 | 730.2844 | 24.2192 | 30.2284 | 1.9837 | 3.2229 | 0.0328 | #### | 11.9989 | 102.5335 | 0.0061 | 40.7562 | 0.2102 | -0.0305 | 8.7425 | 0.0136 | 0.6679 | 0.2083 | 0.0922 | 0.064 | 0.69 | 0.58 | 0.58 | -0.0016 | 958.1671 | 0.0176 | 0.0771 | 0.0719 | -0.0001 | 45.0598 | 0.005 | 0.0053 | 0.1875 | 0.0846 | 730.2844 | 24.2192 | 30.2284 | 1.8596 | 3.2229 | 0.0328 |  | 11.9989 | 102.5226 | -0.0074 | 40.7388 | 0.2175 | -0.0551 | 8.7413 | 0.6679 | 0.2083 | 0.0876 | 0.064
Rinse1 | QC | 2024-08-12 14:23 | [6 empty] | -0.0004 | 0.0048 | 0.0019 | -0.0014 | -0.0001 | 0 | -0.0135 | 0 | -0.0002 | 0 | 0 | 0.0143 | 0.0044 | -0.0025 | -0.0154 | 0.0002 | 0.0027 | 1.1196 | -0.0013 | 0.0079 | -0.0067 | 0.0098 | 0.2488 | 0.0008 | 0.0282 | 0.0004 | 0 | -0.0001 | -0.0001 | 0.0005 | 1.1 | 1.02 | 1.02 | -0.0004 | 0.0048 | 0.0019 | -0.0014 | -0.0001 | 0 | -0.0135 | 0 | -0.0002 | 0 | 0 | 0.0143 | 0.0044 | -0.0025 | -0.0154 | 0.0002 | 0.0027 | 1.1196 | -0.0013 | 0.0079 | -0.0067 | 0.0098 | 0.2488 | 0.0008 | 0.0282 | 0 | -0.0001 | -0.0001 | 0.0005
Rinse | QC | 2024-08-12 14:25 | [6 empty] | -0.0004 | -0.0018 | 0.0008 | -0.0016 | -0.0001 | 0 | -0.009 | -0.0001 | -0.0001 | -0.0004 | -0.0007 | 0.0089 | -0.0028 | -0.0041 | -0.0154 | 0.0001 | 0.0007 | 0.6289 | -0.0013 | 0.0055 | -0.0056 | 0.0069 | 0.2227 | 0.0089 | 0.0173 | 0.0012 | 0 | -0.0001 | -0.0005 | 0.0009 | 1.09 | 1.01 | 1.02 | -0.0004 | -0.0018 | 0.0008 | -0.0016 | -0.0001 | 0 | -0.009 | -0.0001 | -0.0001 | -0.0004 | -0.0007 | 0.0089 | -0.0028 | -0.0041 | -0.0154 | 0.0001 | 0.0007 | 0.6289 | -0.0013 | 0.0055 | -0.0056 | 0.0069 | 0.2227 | 0.0089 | 0.0173 | 0 | -0.0001 | -0.0005 | 0.0009
CCV (5ppm All) | QC | 2024-08-12 14:28 | [6 empty] | 4.8122 | 4.9228 | 4.8845 | 4.8448 | 4.8877 | 4.8767 | 4.859 | 4.864 | 4.8817 | 4.8594 | 4.8328 | 5.0646 | 4.8122 | 4.9741 | 4.8491 | 5.0073 | 4.9355 | 5.2778 | 4.8723 | 5.1045 | 4.8909 | 4.8358 | 3.9892 Q | 4.8997 | 4.9559 | 4.845 | 4.8882 | 4.8045 | 4.8438 | 4.8355 | 1.06 | 1.01 | 1.02 | 4.8117 | 4.9228 | 4.8845 | 4.8448 | 4.8877 | 4.8767 | 4.859 | 4.8639 | 4.8814 | 4.858 | 4.8329 | 5.0646 | 4.8122 | 4.9741 | 4.8483 | 5.0073 | 4.9355 | 5.2778 | 4.8723 | 5.1044 | 4.8902 | 4.8087 | 3.9893 | 4.8967 | 4.954 | 4.8882 | 4.8045 | 4.8395 | 4.8355
CCB | QC | 2024-08-12 14:30 | [6 empty] | -0.0004 | 0.0003 | 0.0131 | 0.0003 | 0.0001 | -0.0001 | -0.0126 | 0.0002 | 0.0001 | -0.0001 | -0.0004 | 0.0043 | 0.0103 | -0.0045 | -0.0114 | 0.0003 | 0.0447 | 0.392 | -0.0022 | 0.0088 | -0.0023 | 0.0072 | 0.2017 | 0.0214 | 0.0105 | 0.0016 | 0.0001 | 0.0013 | -0.0002 | 0.0003 | 1.06 | 1.02 | 1.03 | -0.0004 | 0.0003 | 0.0131 | 0.0003 | 0.0001 | -0.0001 | -0.0126 | 0.0002 | 0.0001 | -0.0001 | -0.0004 | 0.0043 | 0.0103 | -0.0045 | -0.0114 | 0.0003 | 0.0447 | 0.392 | -0.0022 | 0.0088 | -0.0023 | 0.0072 | 0.2017 | 0.0214 | 0.0105 | 0.0001 | 0.0013 | -0.0002 | 0.0003
08092024 MB | Sample | 2024-08-12 14:33 | [6 empty] | -0.0005 | 0.0018 | 0.0079 | -0.0001 | 0 | 0 | -0.0091 | 0.0002 | 0.0001 | -0.0001 | 0.0005 | 0.0072 | -0.001 | -0.0037 | -0.0098 | 0.0002 | 0.014 | 0.4276 | -0.0023 | 0.0084 | -0.0027 | 0.007 | 0.1698 | 0.0075 | 0.0144 | 0.0018 | 0 | 0.0005 | -0.0001 | 0.0011 | 1.04 | 1.01 | 1.02 | -0.0005 | 0.0018 | 0.0079 | -0.0001 | 0 | 0 | -0.0091 | 0.0002 | 0.0001 | -0.0001 | 0.0005 | 0.0072 | -0.001 | -0.0037 | -0.0098 | 0.0002 | 0.014 | 0.4276 | -0.0023 | 0.0084 | -0.0027 | 0.007 | 0.1698 | 0.0075 | 0.0144 | 0 | 0.0005 | -0.0001 | 0.0011
1008092024 MFB | Sample | 2024-08-12 14:35 | [6 empty] | 0.0956 | 4.8215 | 2.4388 | 0.4803 | 2.4493 | 0.4779 | 4.8377 | 0.9757 | 4.9166 | 2.4229 | 2.4159 | 5.0866 | 4.9341 | -0.0121 | 4.9421 | 1.0065 | 0.9871 | 5.5818 | 2.4371 | 1.0079 | 2.4332 | 2.3666 | 2.1057 | 2.3836 | 1.0055 | 0.9853 | 0.9825 | 0.9772 | 2.445 | 0.9651 | 1.03 | 1.01 | 1.01 | 0.0955 | 4.8215 | 2.4388 | 0.4803 | 2.4493 | 0.4779 | 4.8377 | 0.9756 | 4.9165 | 2.4226 | 2.4159 | 5.0866 | 4.9341 | -0.0121 | 4.9412 | 1.0065 | 0.9871 | 5.5818 | 2.4371 | 1.0079 | 2.4329 | 2.3612 | 2.1057 | 2.3829 | 1.0051 | 0.9825 | 0.9772 | 2.4441 | 0.9651
08092024 MFB - Analyticsal Duplicate | Sample | 2024-08-12 14:38 | [6 empty] | 0.0949 | 4.7914 | 2.4428 | 0.4807 | 2.4307 | 0.4767 | 4.8348 | 0.97 | 4.8705 | 2.4242 | 2.411 | 5.0631 | 4.8962 | -0.004 | 4.9384 | 1.0018 | 0.9995 | 5.5088 | 2.4317 | 1.0116 | 2.4322 | 2.3491 | 2.1614 | 2.3914 | 1.0069 | 0.9828 | 0.9764 | 0.9766 | 2.4357 | 0.9642 | 1.02 | 1.01 | 1.02 | 0.0948 | 4.7914 | 2.4428 | 0.4807 | 2.4307 | 0.4767 | 4.8348 | 0.9699 | 4.8704 | 2.4239 | 2.4111 | 5.0631 | 4.8962 | -0.004 | 4.9375 | 1.0018 | 0.9995 | 5.5088 | 2.4317 | 1.0115 | 2.432 | 2.3437 | 2.1615 | 2.3907 | 1.0066 | 0.9764 | 0.9766 | 2.4348 | 0.9642
5x 08092024 MFB | Sample | 2024-08-12 14:40 | [6 empty] | 0.0188 | 0.9539 | 0.4915 | 0.0947 | 0.4846 | 0.0947 | 0.9759 | 0.193 | 0.9693 | 0.4823 | 0.4712 | 1.0115 | 0.9427 | -0.007 | 0.9589 | 0.1998 | 0.1988 | 1.2827 | 0.4788 | 0.2066 | 0.4785 | 0.4721 | 0.459 | 0.4884 | 0.2019 | 0.1922 | 0.1935 | 0.1914 | 0.4793 | 0.1898 | 1.05 | 1.02 | 1.03 | 0.0188 | 0.9539 | 0.4915 | 0.0947 | 0.4846 | 0.0947 | 0.9759 | 0.193 | 0.9692 | 0.4822 | 0.4712 | 1.0115 | 0.9427 | -0.007 | 0.9587 | 0.1998 | 0.1988 | 1.2827 | 0.4788 | 0.2066 | 0.4784 | 0.471 | 0.459 | 0.4883 | 0.2018 | 0.1935 | 0.1914 | 0.4791 | 0.1898
5x 10% NaCl LFP Battery-TM | Sample | 2024-08-12 14:43 | [6 empty] | -0.0007 | 215.6443 | 0.0122 | 0.0078 | 0.0176 | 0 | 7.4977 | 0.0038 | 0.0052 | 0.0793 | 0.023 | 243.4776 | 1.5368 | 3.0259 | 0.5063 | 0.8884 | 0.0221 | #### | 4.1748 | 35.953 | 0.0043 | 6.7055 | 0.1506 | -0.0055 | 1.165 | 0.0071 | 0.0894 | 0.0671 | 0.0339 | 0.0436 | 0.87 | 0.87 | 0.88 | -0.0008 | 215.6443 | 0.0122 | 0.0078 | 0.0176 | 0 | 7.4977 | 0.0001 | 0.0015 | 0.079 | 0.0257 | 243.4776 | 1.5368 | 3.0259 | 0.4649 | 0.8884 | 0.0221 |  | 4.1748 | 35.9493 | -0.0002 | 6.7007 | 0.153 | -0.0125 | 1.1647 | 0.0894 | 0.0671 | 0.0324 | 0.0436
5x 10% NaCl LFP Battery-TM - MS | Sample | 2024-08-12 14:45 | [6 empty] | 0.0177 | 189.7969 | 0.5183 | 0.106 | 0.4979 | 0.0908 | 7.6228 | 0.1898 | 0.8261 | 0.539 | 0.5697 | 215.9512 | 2.5462 | 2.6037 | 1.3575 | 0.971 | 0.2052 | #### | 4.0047 | 32.2673 | 0.4672 | 6.4807 | 0.5696 | 0.4827 | 1.3658 | 0.1921 | 0.2731 | 0.2521 | 0.5209 | 0.2224 | 0.87 | 0.89 | 0.89 | 0.0176 | 189.7969 | 0.5183 | 0.106 | 0.4979 | 0.0908 | 7.6228 | 0.1866 | 0.8229 | 0.5387 | 0.572 | 215.9512 | 2.5462 | 2.6037 | 1.3208 | 0.971 | 0.2052 |  | 4.0047 | 32.264 | 0.4631 | 6.4754 | 0.5718 | 0.4763 | 1.3654 | 0.2731 | 0.2521 | 0.5194 | 0.2224
10% NaCl LFP Battery-TM - Analytical Duplicate | Sample | 2024-08-12 14:48 | [6 empty] | -0.0015 | 1291.7171 o | 0.0276 | 0.0566 | 0.0902 | -0.0003 | 33.4245 | 0.0194 | 0.0195 | 0.342 | 0.1261 | 1058.1629 | 13.5663 | 31.5973 | 2.3923 | 3.8145 | 0.052 | #### | 17.0984 | 159.014 | 0.0305 | 29.4486 | 0.1963 | -0.0173 | 9.1797 | 0.0243 | 0.4576 | 0.3173 | 0.1586 | 0.1655 | 0.72 | 0.66 | 0.65 | -0.0018 | 1291.7171 | 0.0276 | 0.0566 | 0.0902 | -0.0003 | 33.4245 | 0.0036 | 0.0035 | 0.3409 | 0.1377 | 1058.1629 | 13.5663 | 31.5973 | 2.2124 | 3.8145 | 0.052 |  | 17.0984 | 158.9981 | 0.011 | 29.428 | 0.2068 | -0.0515 | 9.1782 | 0.4576 | 0.3173 | 0.152 | 0.1655
10% NaCl LFP Battery-TM - Lab Duplicate | Sample | 2024-08-12 14:50 | [6 empty] | -0.0013 | 1307.6594 o | 0.0302 | 0.056 | 0.0905 | -0.0003 | 33.3458 | 0.0199 | 0.0161 | 0.3451 | 0.1253 | 1063.5783 | 13.9791 | 31.9609 | 2.3911 | 3.8301 | 0.047 | 2435.7052 o | 16.914 | 158.4581 | 0.0258 | 29.3562 | 0.254 | -0.0305 | 8.6909 | 0.0239 | 0.459 | 0.3177 | 0.1559 | 0.1639 | 0.72 | 0.66 | 0.66 | -0.0017 | 1307.6594 | 0.0302 | 0.056 | 0.0905 | -0.0003 | 33.3458 | 0.004 | -0.0001 | 0.3441 | 0.137 | 1063.5783 | 13.9791 | 31.9609 | 2.2103 | 3.8301 | 0.047 | 2435.7052 | 16.914 | 158.4421 | 0.0061 | 29.3356 | 0.2646 | -0.0649 | 8.6894 | 0.459 | 0.3177 | 0.1492 | 0.1639
10% NaCl LFP Battery-TM - MS | Sample | 2024-08-12 14:53 | [6 empty] | 0.1099 | 1125.5413 o | 2.3395 | 0.5103 | 2.5982 | 0.406 | 34.1408 | 0.7686 | 3.4986 | 2.3552 | 3.0041 | 952.7033 | 20.8246 | 26.5031 | 5.7168 | 4.2462 | 0.926 | #### | 16.6532 | 143.6749 | 1.8803 | 28.9675 | 2.3211 | 2.1209 | 9.5903 | 0.7722 | 1.4172 | 1.2005 | 2.393 | 0.8689 | 0.74 | 0.68 | 0.68 | 0.1095 | 1125.5413 | 2.3395 | 0.5103 | 2.5982 | 0.406 | 34.1408 | 0.7543 | 3.4841 | 2.354 | 3.0146 | 952.7033 | 20.8246 | 26.5031 | 5.5548 | 4.2462 | 0.926 |  | 16.6532 | 143.6606 | 1.8626 | 28.9446 | 2.3307 | 2.0901 | 9.5887 | 1.4172 | 1.2005 | 2.3862 | 0.8689
10% NaCl LFP Battery-TM - MS - Analytical Duplicate | Sample | 2024-08-12 14:55 | [6 empty] | 0.1104 | 1121.6721 o | 2.3542 | 0.5118 | 2.5932 | 0.4086 | 34.1838 | 0.7682 | 3.5028 | 2.3588 | 3.0269 | 954.4947 | 20.6754 | 26.3993 | 5.7057 | 4.238 | 0.9317 | #### | 16.7073 | 143.7829 | 1.8896 | 29.007 | 2.3751 | 2.1328 | 9.5133 | 0.7706 | 1.4139 | 1.2022 | 2.3994 | 0.868 | 0.74 | 0.68 | 0.68 | 0.11 | 1121.6721 | 2.3542 | 0.5118 | 2.5932 | 0.4086 | 34.1838 | 0.7539 | 3.4883 | 2.3576 | 3.0374 | 954.4947 | 20.6754 | 26.3993 | 5.5434 | 4.238 | 0.9317 |  | 16.7073 | 143.7686 | 1.8719 | 28.9841 | 2.3846 | 2.102 | 9.5116 | 1.4139 | 1.2022 | 2.3926 | 0.868
Rinse1 | QC | 2024-08-12 14:58 | [6 empty] | -0.0007 | 0.0163 | 0.0108 | -0.0006 | 0.0001 | 0 | -0.01 | 0.0001 | -0.0002 | -0.0001 | -0.0001 | 0.0247 | 0.0179 | -0.0127 | -0.0143 | 0.0002 | 0.0064 | 1.6323 | -0.0008 | 0.0126 | -0.0029 | 0.0097 | 0.2656 | 0.0188 | 0.0519 | 0 | 0 | 0 | 0 | 0.0007 | 1.05 | 1.02 | 1.03 | -0.0007 | 0.0163 | 0.0108 | -0.0006 | 0.0001 | 0 | -0.01 | 0.0001 | -0.0002 | -0.0001 | -0.0001 | 0.0247 | 0.0179 | -0.0127 | -0.0143 | 0.0002 | 0.0064 | 1.6323 | -0.0008 | 0.0126 | -0.0029 | 0.0097 | 0.2656 | 0.0188 | 0.0519 | 0 | 0 | 0 | 0.0007
Rinse | QC | 2024-08-12 15:00 | [6 empty] | -0.0004 | 0.0032 | 0.0015 | -0.0012 | 0 | 0.0001 | -0.0101 | 0 | 0.0006 | -0.0002 | -0.0001 | 0.0106 | -0.0013 | -0.0032 | -0.0124 | 0.0002 | 0.0033 | 0.8374 | -0.0007 | 0.0102 | -0.0038 | 0.0076 | 0.2516 | 0.0082 | 0.0295 | 0.0006 | 0.0001 | 0 | -0.0002 | 0.0006 | 1.04 | 1.02 | 1.03 | -0.0004 | 0.0032 | 0.0015 | -0.0012 | 0 | 0.0001 | -0.0101 | 0 | 0.0006 | -0.0002 | -0.0001 | 0.0106 | -0.0013 | -0.0032 | -0.0124 | 0.0002 | 0.0033 | 0.8374 | -0.0007 | 0.0102 | -0.0038 | 0.0076 | 0.2516 | 0.0082 | 0.0295 | 0.0001 | 0 | -0.0002 | 0.0006
CCV (5ppm All) | QC | 2024-08-12 15:03 | [6 empty] | 4.7769 | 4.8639 | 4.8429 | 4.8016 | 4.8164 | 4.778 | 4.8126 | 4.825 | 4.8353 | 4.8209 | 4.7844 | 5.0057 | 4.8052 | 4.9436 | 4.8101 | 4.9846 | 4.9115 | 5.3647 | 4.8336 | 5.0551 | 4.8451 | 4.7796 | 4.0301 Q | 4.8785 | 4.9381 | 4.8098 | 4.8259 | 4.7712 | 4.8051 | 4.7814 | 1.03 | 1.02 | 1.02 | 4.7764 | 4.8639 | 4.8429 | 4.8016 | 4.8164 | 4.778 | 4.8126 | 4.8249 | 4.835 | 4.8195 | 4.7844 | 5.0057 | 4.8052 | 4.9436 | 4.8093 | 4.9846 | 4.9115 | 5.3647 | 4.8336 | 5.0551 | 4.8444 | 4.7526 | 4.0302 | 4.8756 | 4.9362 | 4.8259 | 4.7712 | 4.8008 | 4.7814
CCB | QC | 2024-08-12 15:06 | [6 empty] | -0.0003 | 0.0005 | 0.0137 | 0.0001 | 0 | 0.0001 | -0.0123 | 0.0001 | 0.0002 | -0.0001 | 0.0001 | 0.0065 | 0.0156 | -0.0056 | -0.0097 | 0.0001 | 0.0407 | 0.5054 | -0.0012 | 0.0099 | -0.0054 | 0.0078 | 0.2069 | 0.0263 | 0.0172 | 0.0015 | 0.0001 | 0.0012 | 0 | 0.0003 | 1.04 | 1.01 | 1.02 | -0.0003 | 0.0005 | 0.0137 | 0.0001 | 0 | 0.0001 | -0.0123 | 0.0001 | 0.0002 | -0.0001 | 0.0001 | 0.0065 | 0.0156 | -0.0056 | -0.0097 | 0.0001 | 0.0407 | 0.5054 Z | -0.0012 | 0.0099 | -0.0054 | 0.0078 | 0.2069 | 0.0263 | 0.0172 | 0.0001 | 0.0012 | 0 | 0.0003
Rinse | QC | 2024-08-12 15:08 | [6 empty] | 0.0002 | 0 | 0.0048 | -0.0005 | 0 | 0.0001 | -0.0089 | 0.0001 | 0 | -0.0001 | -0.0005 | 0.0063 | -0.003 | -0.0054 | -0.0149 | -0.0001 | 0.012 | 0.4398 | -0.001 | 0.0069 | -0.003 | 0.0081 | 0.2166 | 0.0103 | 0.0099 | 0.0003 | 0 | 0.0002 | 0.0002 | 0.0006 | 1.04 | 1.02 | 1.03 | 0.0002 | 0 | 0.0048 | -0.0005 | 0 | 0.0001 | -0.0089 | 0.0001 | 0 | -0.0001 | -0.0005 | 0.0063 | -0.003 | -0.0054 | -0.0149 | -0.0001 | 0.012 | 0.4398 | -0.001 | 0.0069 | -0.003 | 0.0081 | 0.2166 | 0.0103 | 0.0099 | 0 | 0.0002 | 0.0002 | 0.0006
0.5ppm All (LLQC) | QC | 2024-08-12 15:11 | [6 empty] | 0.4719 | 0.4848 | 0.4732 | 0.4877 | 0.485 | 0.4765 | 0.4757 | 0.4808 | 0.4851 | 0.4828 | 0.4712 | 0.507 | 0.4563 | 0.4847 | 0.4641 | 0.5034 | 0.4606 | 0.8458 | 0.4824 | 0.5039 | 0.4833 | 0.4666 | 0.4815 | 0.4755 | 0.4807 | 0.4686 | 0.4871 | 0.4714 | 0.4753 | 0.4725 | 1.05 | 1.02 | 1.04 | 0.4719 | 0.4848 | 0.4732 | 0.4877 | 0.485 | 0.4765 | 0.4757 | 0.4808 | 0.485 | 0.4827 | 0.4712 | 0.507 | 0.4563 | 0.4847 | 0.464 | 0.5034 | 0.4606 | 0.8458 Q | 0.4824 | 0.5039 | 0.4832 | 0.4639 | 0.4815 | 0.4752 | 0.4805 | 0.4871 | 0.4714 | 0.4749 | 0.4725
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-12 15:13 | [6 empty] | 0.473 | 0.4823 | 0.4717 | 0.4871 | 0.48 | 0.4783 | 0.4709 | 0.4789 | 0.4821 | 0.4812 | 0.4702 | 0.5046 | 0.4359 Q | 0.4844 | 0.4607 | 0.4982 | 0.4694 | 0.8014 Q | 0.4808 | 0.4982 | 0.4772 | 0.4624 | 0.4612 | 0.4831 | 0.4756 | 0.4724 | 0.4825 | 0.4705 | 0.4732 | 0.4704 | 1.04 | 1.02 | 1.03 | 0.4729 | 0.4823 | 0.4717 | 0.4871 | 0.48 | 0.4783 | 0.4709 | 0.4789 | 0.482 | 0.481 | 0.4702 | 0.5046 | 0.4359 | 0.4844 | 0.4606 | 0.4982 | 0.4694 | 0.8014 | 0.4808 | 0.4982 | 0.4772 | 0.4598 | 0.4612 | 0.4828 | 0.4754 | 0.4825 | 0.4705 | 0.4728 | 0.4704
CCS1 (SS 5ppm) | QC | 2024-08-12 15:16 | [6 empty] | 4.8882 | 4.9531 | 4.9736 | 4.8275 | 4.8945 | 4.8447 | 4.8682 | 4.8704 | 4.8684 | 4.8512 | 4.8154 | 5.0542 | 4.8359 | 4.9059 | 4.9164 | 4.9842 | 5.0683 | 5.099 | 4.8619 | 5.1228 | 4.9153 | 4.7956 | 4.8103 | 4.9187 | 4.7531 | 4.912 | 4.9309 | 4.8611 | 4.8862 | 4.7833 | 1.03 | 1.01 | 1.02 | 4.8877 | 4.9531 | 4.9736 | 4.8275 | 4.8945 | 4.8447 | 4.8682 | 4.8703 | 4.8681 | 4.8498 | 4.8154 | 5.0542 | 4.8359 | 4.9059 | 4.9156 | 4.9842 | 5.0683 | 5.099 | 4.8619 | 5.1228 | 4.9146 | 4.7687 | 4.8104 | 4.9157 | 4.7513 | 4.9309 | 4.8611 | 4.8818 | 4.7833
CCB | QC | 2024-08-12 15:18 | [6 empty] | -0.0001 | -0.0022 | 0.0167 | 0.0007 | 0.0001 | 0.0002 | -0.0111 | 0.0001 | 0 | 0 | 0.0004 | 0.0011 | 0.0047 | -0.008 | -0.0115 | 0.0001 | 0.0396 | 0.249 | -0.0018 | 0.0076 | -0.0029 | 0.0048 | 0.1845 | 0.0232 | 0.0038 | 0.003 | 0.0001 | 0.0011 | 0.0003 | 0.0003 | 1.04 | 1.02 | 1.02 | -0.0001 | -0.0022 | 0.0167 | 0.0007 | 0.0001 | 0.0002 | -0.0111 | 0.0001 | 0 | 0 | 0.0004 | 0.0011 | 0.0047 | -0.008 | -0.0115 | 0.0001 | 0.0396 | 0.249 | -0.0018 | 0.0076 | -0.0029 | 0.0048 | 0.1845 | 0.0232 | 0.0038 | 0.0001 | 0.0011 | 0.0003 | 0.0003
Rinse1 | QC | 2024-08-12 15:21 | [6 empty] | 0 | -0.0038 | 0.0028 | 0.0008 | 0 | 0.0001 | -0.0127 | 0 | 0.0004 | 0.0001 | -0.0001 | 0.002 | -0.0085 | -0.0053 | -0.0119 | 0.0003 | 0.0139 | 0.2272 | -0.0024 | 0.0069 | -0.0046 | 0.0056 | 0.184 | 0.0119 | 0.0047 | 0.0009 | 0.0001 | 0.0003 | -0.0001 | 0.0003 | 1.03 | 1.02 | 1.03 | 0 | -0.0038 | 0.0028 | 0.0008 | 0 | 0.0001 | -0.0127 | 0 | 0.0004 | 0.0001 | -0.0001 | 0.002 | -0.0085 | -0.0053 | -0.0119 | 0.0003 | 0.0139 | 0.2272 | -0.0024 | 0.0069 | -0.0046 | 0.0056 | 0.184 | 0.0119 | 0.0047 | 0.0001 | 0.0003 | -0.0001 | 0.0003
Rinse | QC | 2024-08-12 15:23 | [6 empty] | -0.0002 | -0.0031 | 0.0028 | 0 | 0 | 0 | -0.0144 | 0 | 0.0001 | -0.0004 | -0.0006 | 0.0011 | 0.0023 | -0.0064 | -0.011 | 0 | 0.0078 | 0.2003 | -0.0023 | 0.0057 | -0.0039 | 0.008 | 0.2215 | 0.0102 | 0.0019 | 0.0006 | 0 | 0.0001 | -0.0004 | 0.0003 | 1.04 | 1.02 | 1.04 | -0.0002 | -0.0031 | 0.0028 | 0 | 0 | 0 | -0.0144 | 0 | 0.0001 | -0.0004 | -0.0006 | 0.0011 | 0.0023 | -0.0064 | -0.011 | 0 | 0.0078 | 0.2003 | -0.0023 | 0.0057 | -0.0039 | 0.008 | 0.2215 | 0.0102 | 0.0019 | 0 | 0.0001 | -0.0004 | 0.0003
MQ-Blank-02 | Sample | 2024-08-12 15:26 | [6 empty] | 0.0001 | -0.0042 | 0.0018 | -0.0012 | -0.0001 | 0 | -0.0131 | 0 | 0.0004 | -0.0002 | 0.0001 | 0.0028 | -0.0038 | -0.0042 | -0.0124 | 0 | 0.0019 | 0.1758 | -0.0024 | 0.0026 | -0.0044 | 0.0057 | 0.2462 | 0.0043 | 0.0035 | 0.0003 | 0 | 0 | -0.0001 | 0.0011 | 1.06 | 1.03 | 1.04 | 0.0001 | -0.0042 | 0.0018 | -0.0012 | -0.0001 | 0 | -0.0131 | 0 | 0.0004 | -0.0002 | 0.0001 | 0.0028 | -0.0038 | -0.0042 | -0.0124 | 0 | 0.0019 | 0.1758 | -0.0024 | 0.0026 | -0.0044 | 0.0057 | 0.2462 | 0.0043 | 0.0035 | 0 | 0 | -0.0001 | 0.0011