## 08262024ICP_TO178-455SampleData
08262024_ICP-OES_TO 178-455 Sample Data
ANALYSIS TECHNIUE: |  | ICP-OES |  |  |  | Calibration date: |  | 2024-08-26
ANALYSIS METHOD: |  | 30M |  |  |  | Calibration Standards: |  | 2024-08-21
ANALYSIS BATCH: |  | TO 178-453 454 & 455 |  |  |  | Task Order# |  | 68HERC22F0178
Task Order: |  | TO 058 |  |  |  | Instrument : |  | Agilent 5900 SVDV ICP-OES (S/N: MY2138CP13) with Agilent SPS 4 Autosampler (S/N: AU213011544)
UNITS: |  | mg/L (ppm) |  |  |  | Analyst: |  | Mahendranath Arambewela
Comments: | [5 empty] | SOP# |  | K-LRTD-SOP-1185-2
Sample results are reported as mg/L (ppm)
Worksheet exported from: C:\Users\USEPA\Documents\Agilent\ICP Expert\My Results\08262024_ICP-OES_TO 178-453 454 & 455.esws
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Solution Label | Type | Date Time | Ag 328.068 C ppm | Al 396.152 C ppm | As 188.980 C ppm | B 249.678 C ppm | Ba 455.403 C ppm | Be 313.107 C ppm | Ca 370.602 C ppm | Cd 214.439 C ppm | Co 238.892 C ppm | Cr 267.716 C ppm | Cu 324.754 C ppm | Fe 259.940 C ppm | K 766.491 C ppm | Li 670.783 C ppm | Mg 279.078 C ppm | Mn 257.610 C ppm | Mo 202.032 C ppm | Na 588.995 C ppm | Ni 231.604 C ppm | P 178.222 C ppm | Pb 220.353 C ppm | S 181.972 C ppm | Sb 217.582 C ppm | Se 196.026 C ppm | Si 251.611 C ppm | Sr 421.552 C ppm | Ti 334.941 C ppm | V 292.401 C ppm | Zn 206.200 C ppm
08262024 LB-2 | Sample | 2024-08-26 15:17 | 0.0002 | 0.0001 | 0.0052 | -0.0019 | 0.0001 | -0.0001 | -0.0115 | 0 | 0.0003 | -0.0002 | 0.0005 | 0.0007 | 0.0277 | -0.0078 | -0.0066 | -0.0003 | 0.0145 | -0.0147 | -0.001 | 0.0039 | -0.0087 | -0.0036 | 0.116 | 0.0132 | 0.002 | 0 | 0.0004 | 0.0001 | 0.0005
10x 08265024 LCS-2 | Sample | 2024-08-26 15:19 | 0.0091 | 0.4747 | 0.245 | 0.0466 | 0.2424 | 0.0484 | 0.474 | 0.0979 | 0.4877 | 0.2465 | 0.2302 | 0.497 | 0.4924 | 0 | 0.4806 | 0.0995 | 0.0921 | 0.4818 | 0.241 | 0.1057 | 0.2418 | 0.2439 | 0.2421 | 0.2456 | 0.0954 | 0.1001 | 0.0977 | 0.2447 | 0.0958
10x 08262024 LCS-2- Analytical Duplicate | Sample | 2024-08-26 15:22 | 0.0098 | 0.4716 | 0.2447 | 0.047 | 0.2402 | 0.048 | 0.4719 | 0.0979 | 0.4881 | 0.2469 | 0.23 | 0.4937 | 0.4729 | -0.0089 | 0.4814 | 0.0985 | 0.0911 | 0.4819 | 0.2415 | 0.1077 | 0.2429 | 0.2408 | 0.2613 | 0.2502 | 0.0946 | 0.0992 | 0.0978 | 0.2443 | 0.0959
5x 08262024 LCS-2 | Sample | 2024-08-26 15:24 | 0.019 | 0.9546 | 0.4902 | 0.0948 | 0.484 | 0.0971 | 0.9666 | 0.1962 | 0.982 | 0.4911 | 0.4582 | 0.9936 | 0.9653 | -0.0003 | 0.9807 | 0.1978 | 0.1799 | 0.9785 | 0.4847 | 0.2071 | 0.4874 | 0.4869 | 0.473 | 0.4898 | 0.1861 | 0.199 | 0.1944 | 0.4873 | 0.1946
08262024 LCS-2 | Sample | 2024-08-26 15:27 | 0.0963 | 4.6699 | 2.4802 | 0.4939 | 2.3903 | 0.4793 | 4.8571 | 0.9906 | 4.9267 | 2.4482 | 2.2906 | 4.8599 | 4.8696 | -0.0012 | 4.993 | 0.9732 | 0.9373 | 4.828 | 2.4507 | 1.0387 | 2.4953 | 2.5074 | 2.5632 | 2.4876 | 0.9582 | 0.9815 | 0.9822 | 2.4491 | 0.977
08262024 LCS-2- Analytical Duplicate | Sample | 2024-08-26 15:29 | 0.0954 | 4.6689 | 2.4679 | 0.4892 | 2.395 | 0.477 | 4.8094 | 0.981 | 4.8862 | 2.4261 | 2.2895 | 4.8664 | 4.8059 | -0.0077 | 4.9477 | 0.9754 | 0.9521 | 4.7946 | 2.4232 | 1.0303 | 2.4739 | 2.4844 | 2.5859 | 2.4727 | 0.9547 | 0.9829 | 0.973 | 2.4409 | 0.9666
0.5% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 15:59 | -0.0004 | 0.1804 | 0.0154 | 0.0066 | 0.0418 | -0.0001 | 0.3158 | 0 | 0.0002 | 0.0005 | 0.0025 | 0.1616 | 6.6046 | 0.0095 | 0.0343 | 0.001 | 0.0174 | 1080.1282 | 1.455 | 0.1984 | -0.003 | 0.1711 | 0.1259 | 0.0175 | 0.0436 | 0.0013 | 0.0013 | 0.0004 | 0.0059
1% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:02 | -0.0002 | 0.1093 | 0.0126 | 0.0075 | 0.0466 | 0 | 0.5437 | 0.0001 | 0.0003 | 0.0001 | 0.003 | 0.0639 | 0.6065 | 0.001 | 0.0789 | 0.0004 | 0.0091 | 1782.5297 | 1.364 | 0.1745 | -0.0046 | 0.3071 | 0.1101 | 0.0136 | 0.0677 | 0.0021 | 0.0013 | 0.0003 | 0.0057
5% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:04 | -0.0008 | 0.1806 | 0.0084 | 0.0292 | 0.0881 | -0.0001 | 2.2565 | -0.0001 | 0.0003 | 0.0004 | 0.0022 | 0.1344 | 1.2881 | 0.0147 | 0.3968 | 0.0027 | 0.0082 |  | 0.8643 | 0.1541 | -0.0106 | 1.07 | 0.1085 | 0.008 | 0.3693 | 0.0079 | 0.0072 | 0.0015 | 0.0106
10% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:07 | -0.0012 | 0.1245 | 0.0069 | 0.0561 | 0.1333 | 0.0001 | 4.198 | -0.0001 | -0.0002 | 0.0005 | 0.0006 | 0.2474 | 2.5736 | 0.0322 | 0.7236 | 0.0088 | 0.0076 |  | 0.2948 | 0.2303 | -0.0143 | 1.9988 | 0.1094 | 0.0094 | 0.643 | 0.0154 | 0.0105 | 0.003 | 0.0114
15% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:09 | -0.0013 | 0.2724 | 0.0081 | 0.069 | 0.1331 | -0.0001 | 4.4436 | 0.0002 | -0.0003 | 0.0006 | 0 | 1.1278 | 5.9549 | 0.0369 | 0.7401 | 0.0075 | 0.0078 |  | 0.2598 | 0.4848 | -0.0148 | 2.921 | 0.1073 | 0.0067 | 0.9399 | 0.0171 | 0.0148 | 0.0046 | 0.014
0.5% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:12 | -0.0002 | 0.3046 | 0.0061 | 0.003 | 0.003 | -0.0001 | 0.2535 | 0 | 0.0011 | 0.0004 | 0.0024 | 0.1919 | 0.4832 | 0.0093 | 0.0299 | 0.0006 | 0.0037 | 1061.1289 | 4.6501 | 1.6312 | -0.0029 | 0.1404 | 0.12 | 0.0111 | 0.0661 | 0.0007 | 0.0006 | 0.0008 | 0.0056
1% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:14 | -0.0002 | 0.4131 | 0.0041 | 0.0076 | 0.004 | -0.0001 | 0.5031 | 0 | 0.0012 | 0.0004 | 0.0027 | 0.099 | 0.5255 | 0.0028 | 0.0744 | 0.0011 | 0.004 | 1752.4356 | 5.0444 | 2.0521 | -0.0019 | 0.2448 | 0.1071 | 0.015 | 0.1281 | 0.0016 | 0.0016 | 0.0005 | 0.0057
5% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:17 | -0.001 | 2.8795 | 0.0061 | 0.0277 | 0.03 | 0 | 2.4264 | 0 | 0.0019 | 0.0025 | 0.0025 | 4.3695 | 1.4801 | 0.0114 | 0.4514 | 0.0089 | 0.005 |  | 19.6945 | 4.4225 | -0.0107 | 1.1654 | 0.1183 | 0.0117 | 0.7978 | 0.0083 | 0.0172 | 0.0018 | 0.0124
10% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:19 | -0.0009 | 0.0738 | 0.0069 | 0.0524 | 0.0639 | -0.0001 | 3.8722 | -0.0003 | 0.0009 | 0.0029 | -0.0001 | 6.2456 | 3.0391 | 0.0282 | 0.6991 | 0.0087 | 0.0049 |  | 9.1367 | 3.6268 | -0.0167 | 2.4262 | 0.104 | 0.0087 | 1.3056 | 0.0153 | 0.0175 | 0.0039 | 0.0053
15% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:22 | -0.0012 | 0.0874 | 0.0059 | 0.0702 | 0.0723 | 0 | 4.1223 | -0.0002 | 0.0003 | 0.0022 | -0.001 | 6.0412 | 7.8595 | 0.038 | 0.7319 | 0.0097 | 0.0068 |  | 9.3892 | 3.4593 | -0.0156 | 3.5913 | 0.1042 | 0.0016 | 1.9866 | 0.0168 | 0.0242 | 0.0059 | 0.0044
08162024 MB | Sample | 2024-08-26 16:51 | 0.0001 | 0.0018 | 0.0046 | -0.0029 | 0.0001 | -0.0001 | -0.0144 | 0 | 0.0008 | -0.0002 | 0.0008 | 0.0011 | 0.0277 | 0.0013 | -0.0076 | -0.0002 | 0.0054 | 0.1339 | -0.0002 | 0.0024 | -0.0058 | -0.0035 | 0.1194 | 0.0148 | 0.0035 | 0 | 0.0002 | 0.0003 | 0.0013
08162024 MFB | Sample | 2024-08-26 16:54 | 0.0979 | 4.7972 | 2.4586 | 0.4941 | 2.4539 | 0.4817 | 4.9762 | 1.0022 | 5.0185 | 2.5104 | 2.387 | 4.9622 | 4.944 | 0.0001 | 5.0833 | 1.0047 | 0.9835 | 5.0818 | 2.4711 | 1.0582 | 2.5382 | 2.4606 | 2.623 | 2.4157 | 0.9733 | 1.0096 | 1.0026 | 2.518 | 0.9784
08162024 MFB - Analyticsal Duplicate | Sample | 2024-08-26 16:56 | 0.098 | 4.7714 | 2.4724 | 0.4945 | 2.4449 | 0.4804 | 4.965 | 1.0023 | 5.021 | 2.506 | 2.3818 | 4.9418 | 4.9284 | 0.0002 | 5.0728 | 1 | 1.0013 | 5.0693 | 2.4681 | 1.0578 | 2.5369 | 2.4574 | 2.6724 | 2.419 | 0.9814 | 1.0046 | 1.0016 | 2.5237 | 0.9774
5x 08162024 MFB | Sample | 2024-08-26 16:59 | 0.0194 | 0.9528 | 0.5033 | 0.0973 | 0.4873 | 0.0968 | 1.1014 | 0.1996 | 1.0015 | 0.5023 | 0.4701 | 0.991 | 0.9784 | 0.0077 | 0.9977 | 0.1998 | 0.1983 | 1.0942 | 0.4886 | 0.2111 | 0.5039 | 0.4887 | 0.5494 | 0.5 | 0.1922 | 0.2005 | 0.1973 | 0.4992 | 0.1944
5x 1% NaHCO3 NMC Battery-TM | Sample | 2024-08-26 17:01 | 0.0002 | 0.0804 | 0.0086 | -0.0006 | 0.001 | -0.0002 | 0.1845 | 0 | 0.0003 | 0.0001 | 0.0016 | 0.0196 | 0.0928 | -0.0064 | 0.0112 | 0.0001 | 0.0119 | 466.6491 | 1.0209 | 0.4067 | -0.0033 | 0.045 | 0.1394 | 0.0218 | 0.0262 | 0.0004 | 0.0006 | 0.0003 | 0.0019
5x 1% NaHCO3 NMC Battery-TM - MS | Sample | 2024-08-26 17:04 | 0.0199 | 0.9541 | 0.5124 | 0.1006 | 0.4883 | 0.096 | 1.1555 | 0.1992 | 1.0048 | 0.4966 | 0.5107 | 0.9758 | 1.0543 | -0.006 | 0.971 | 0.1921 | 0.1878 | 416.5691 | 0.7397 | 0.2539 | 0.4934 | 0.567 | 0.5127 | 0.5196 | 0.2116 | 0.1977 | 0.2015 | 0.5109 | 0.1978
5x 1% NaHCO3 NMC Battery-TM -MS -Analytical Duplicate | Sample | 2024-08-26 17:06 | 0.0187 | 0.6318 | 0.4854 | 0.0955 | 0.3204 | 0.0633 | 1.098 | 0.1886 | 0.9502 | 0.47 | 0.4832 | 0.6476 | 0.7085 | -0.0178 | 0.9119 | 0.1276 | 0.1825 | 278.896 | 0.6987 | 0.2401 | 0.4649 | 0.535 | 0.5213 | 0.4911 | 0.1989 | 0.13 | 0.1906 | 0.4833 | 0.1865
1% NaHCO3 NMC Battery-TM - Lab Duplicate | Sample | 2024-08-26 17:09 | -0.0004 | 0.4206 | 0.0072 | 0.0051 | 0.004 | -0.0001 | 0.4985 | 0 | 0.0015 | 0.0005 | 0.0024 | 0.0996 | 0.5282 | 0.0117 | 0.0729 | 0.0012 | 0.0112 | 1766.2352 | 5.0468 | 2.0369 | -0.0064 | 0.2381 | 0.138 | 0.019 | 0.1297 | 0.0016 | 0.0014 | 0.0007 | 0.0059
1% NaHCO3- NMC Battery-TM - MS | Sample | 2024-08-26 17:11 | -0.0001 | 0.1121 | 0.012 | 0.0053 | 0.0465 | -0.0002 | 0.541 | -0.0001 | 0 | 0.0003 | 0.0034 | 0.0679 | 0.6131 | 0.0019 | 0.0757 | 0.0002 | 0.0044 | 1760.8728 | 1.3672 | 0.1803 | -0.0042 | 0.3097 | 0.1151 | 0.0178 | 0.072 | 0.002 | 0.0009 | 0.0006 | 0.0051
1% NaHCO3- NMC Battery-TM - MS - Analytical Duplicate | Sample | 2024-08-26 17:14 | -0.0005 | 0.0097 | 0.0037 | -0.003 | 0.0049 | -0.0003 | 0.1786 | -0.0001 | -0.0004 | -0.0007 | -0.0042 | 0.0065 | 0.0788 | -0.0281 | 0.0129 | -0.0003 | 0.0024 | 249.0177 | 0.4664 | 0.0607 | -0.0134 | 0.0974 | 0.1198 | 0.0093 | 0.0214 | 0.0002 | 0.0003 | 0.0004 | 0.0017
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Lowest Cal Std |  |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Highest Cal Std |  |  | 20 | 200 | 20 | 20 | 50 | 20 | 2000 | 20 | 20 | 50 | 100 | 1000 | 100 | 100 | 2000 | 50 | 20 | 100 | 50 | 200 | 200 | 100 | 50 | 20 | 100 | 20 | 50 | 20 | 100
Correlation |  |  | 0.99998 | 0.99993 | 0.99997 | 1 | 0.99937 | 0.99996 | 0.99999 | 1 | 1 | 1 | 0.99996 | 0.99998 | 1 | 0.99999 | 0.99996 | 0.99992 | 0.99986 | 0.99999 | 1 | 0.99998 | 0.99978 | 0.99998 | 0.99936 | 0.99996 | 0.99927 | 0.99996 | 0.99999 | 0.99999 | 1
ICP-OES MDL |  |  | 0.0045598 | 0.0048715 | 0.0060725 | 0.0016216 | 0.00011879 | 0.000168 | 0.0038865 | 0.00024729 | 0.00024729 | 0.0007 | 0.00076374 | 0.0031683 | 0.032873 | 0.0015923 | 0.0080477 | 0.0008057 | 0.00050865 | 0.001351 | 0.0024937 | 0.0072662 | 0.0049325 | 0.0058515 | 0.0056274 | 0.0064431 | 0.002358 | 0.000168 | 0.00011879 | 0.00093287 | 0.00059397
ICP-OES MRL/PRDL |  |  | 0.045598 | 0.048715 | 0.060725 | 0.016216 | 0.0011879 | 0.00168 | 0.038865 | 0.0024729 | 0.0024729 | 0.0074 | 0.0076374 | 0.031683 | 0.32873 | 0.015923 | 0.080477 | 0.008057 | 0.0050865 | 0.01351 | 0.024937 | 0.072662 | 0.049325 | 0.058515 | 0.056274 | 0.064431 | 0.02358 | 0.00168 | 0.0011879 | 0.0093287 | 0.0059397
Where
Calculated/Estimated MDLs & MRLs
Dilution required
* | Sampling Error
** | Different concentration than the normal
Mislabeled

## 08262024ICP-OES_TO178-455Data
08262024_ICP-OES_TO 178-455 Data Report
ANALYSIS TECHNIUE: |  | ICP-OES |  |  |  | Calibration date: |  | 2024-08-26
ANALYSIS METHOD: |  | 30M |  |  |  | Calibration Standards: |  | 2024-08-21
ANALYSIS BATCH: |  | TO 178-453 454 & 455 |  |  |  | Task Order# |  | 68HERC22F0178
Task Order: |  | TO 058 |  |  |  | Instrument : |  | Agilent 5900 SVDV ICP-OES (S/N: MY2138CP13) with Agilent SPS 4 Autosampler (S/N: AU213011544)
UNITS: |  | mg/L (ppm) |  |  |  | Analyst: |  | Mahendranath Arambewela
Comments: | [5 empty] | SOP# |  | K-LRTD-SOP-1185-2
Sample results are reported as mg/L (ppm)
Worksheet exported from: C:\Users\USEPA\Documents\Agilent\ICP Expert\My Results\08262024_ICP-OES_TO 178-453 454 & 455.esws
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
|  |  | Ag (328.068 nm) ppm | Al (396.152 nm) ppm | As (188.980 nm) ppm | B (249.678 nm) ppm | Ba (455.403 nm) ppm | Be (313.107 nm) ppm | Ca (370.602 nm) ppm | Cd (214.439 nm) ppm | Co (238.892 nm) ppm | Cr (267.716 nm) ppm | Cu (324.754 nm) ppm | Fe (259.940 nm) ppm | K (766.491 nm) ppm | Li (670.783 nm) ppm | Mg (279.078 nm) ppm | Mn (257.610 nm) ppm | Mo (202.032 nm) ppm | Na (588.995 nm) ppm | Ni (231.604 nm) ppm | P (178.222 nm) ppm | Pb (220.353 nm) ppm | S (181.972 nm) ppm | Sb (217.582 nm) ppm | Se (196.026 nm) ppm | Si (251.611 nm) ppm | Sr (421.552 nm) ppm | Ti (334.941 nm) ppm | V (292.401 nm) ppm | Zn (206.200 nm) ppm
IECS1 (Al 1000 Int.) | IECS |  | Al: 1000
IECS2 (Fe 1000 Int.) | IECS |  | Fe: 1000
IECS3 (Mn 100 Int.) | IECS |  | Mn: 100
IECS4 (Ca Mg 1000 Int.) | IECS |  | Ca: 1000; Mg: 1000
IECS5 (Ti 20 Int.) | IECS |  | Ti: 20
ICV 0.5ppm All | QC |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
ICV 5ppm All | QC |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
CCS1 (SS 5ppm) | QC |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
CCV2 5ppm All | QC |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
ICV 1ppm All | QC | bar: Ag to Zn
|  | MDL | 0.0045598 | 0.0048715 | 0.0060725 | 0.0016216 | 0.00011879 | 0.000168 | 0.0038865 | 0.00024729 | 0.00024729 | 0.00038187 | 0.00076374 | 0.0031683 | 0.032873 | 0.0015923 | 0.0080477 | 0.0008057 | 0.00050865 | 0.001351 | 0.0024937 | 0.0072662 | 0.0049325 | 0.0058515 | 0.0056274 | 0.0064431 | 0.002358 | 0.000168 | 0.00011879 | 0.00093287 | 0.00059397
|  | MRL | 0.045598 | 0.048715 | 0.060725 | 0.016216 | 0.0011879 | 0.00168 | 0.038865 | 0.0024729 | 0.0024729 | 0.0038187 | 0.0076374 | 0.031683 | 0.32873 | 0.015923 | 0.080477 | 0.008057 | 0.0050865 | 0.01351 | 0.024937 | 0.072662 | 0.049325 | 0.058515 | 0.056274 | 0.064431 | 0.02358 | 0.00168 | 0.0011879 | 0.0093287 | 0.0059397
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Solution Label | Type | Date Time | Ag 328.068 C ppm | Al 396.152 C ppm | As 188.980 C ppm | B 249.678 C ppm | Ba 455.403 C ppm | Be 313.107 C ppm | Ca 370.602 C ppm | Cd 214.439 C ppm | Co 238.892 C ppm | Cr 267.716 C ppm | Cu 324.754 C ppm | Fe 259.940 C ppm | K 766.491 C ppm | Li 670.783 C ppm | Mg 279.078 C ppm | Mn 257.610 C ppm | Mo 202.032 C ppm | Na 588.995 C ppm | Ni 231.604 C ppm | P 178.222 C ppm | Pb 220.353 C ppm | S 181.972 C ppm | Sb 217.582 C ppm | Se 196.026 C ppm | Si 251.611 C ppm | Sr 421.552 C ppm | Ti 334.941 C ppm | V 292.401 C ppm | Zn 206.200 C ppm
Calibration Blank | Blank | 2024-08-26 10:07 | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 !
Calibration Standard 0.5ppm | Standard | 2024-08-26 10:10 | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 !
Calibration Standard 1ppm | Standard | 2024-08-26 10:12 | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 !
Calibration Standard 5ppm | Standard | 2024-08-26 10:15 | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 !
Calibration Standard 10ppm | Standard | 2024-08-26 10:17 | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 !
Calibration Standard 20ppm | Standard | 2024-08-26 10:20 | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 !
Calibration Standard 50ppm | Standard | 2024-08-26 10:22 |  | 50.0000 ! |  |  | 50.0000 ! |  | 50.0000 ! |  |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! |  | 50.0000 ! |  | 50.0000 !
Calibration Standard-A 100ppm | Standard | 2024-08-26 10:25 | Al: 100.0000 !; Ca: 100.0000 !; Cu: 100.0000 !; Fe: 100.0000 !; K: 100.0000 !; Li: 100.0000 !; Mg: 100.0000 !; Na: 100.0000 !; Pb: 100.0000 !; Si: 100.0000 !; Zn: 100.0000 !
Calibration Standard-B 100ppm | Standard | 2024-08-26 10:27 | P: 100.0000 !; S: 100.0000 !
Calibration Standard Al+Fe+Ca+Mg200ppm | Standard | 2024-08-26 10:30 | Al: 200.0000 !; Ca: 200.0000 !; Fe: 200.0000 !; Mg: 200.0000 !
Calibration Standard Ca+Mg500ppm | Standard | 2024-08-26 10:32 | Ca: 500.0000 !; Mg: 500.0000 !
Calibration Standard Ca+Mg1000ppm | Standard | 2024-08-26 10:35 | Ca: 1000.0000 !; Mg: 1000.0000 !
Calibration Standard Ca+Mg2000ppm | Standard | 2024-08-26 10:37 | Ca: 2000.0000 !; Mg: 2000.0000 !
Calibration Standard P+S200ppm | Standard | 2024-08-26 10:40 | P: 200.0000 !; S: 200.0000 !
Calibration Standard Pb200ppm | Standard | 2024-08-26 10:42 | Pb: 200.0000 !
Calibration Standard Fe 1000ppm | Standard | 2024-08-26 10:45 | Fe: 1000.0000 !
Rinse1 | QC | 2024-08-26 10:47 | 0.0002 | -0.0022 | 0.0031 | 0.0082 | 0 | -0.0001 | -0.005 | 0 | -0.0003 | -0.0001 | 0.0006 | 0.0191 | 0.023 | 0.0114 | 0.0028 | -0.0003 | 0.0034 | -0.0145 | -0.0006 | 0.0045 | -0.0036 | -0.0037 | 0.1102 | 0.0082 | 0.0046 | 0 | -0.0001 | 0.0002 | 0.0006
Rinse | QC | 2024-08-26 10:50 | 0.0001 | -0.0034 | 0.003 | 0.0056 | 0.0001 | -0.0002 | -0.0094 | 0 | 0.0007 | -0.0001 | 0.0003 | 0.0041 | 0.0266 | -0.0025 | -0.0014 | -0.0003 | 0.0028 | -0.0146 | -0.0005 | 0.0047 | -0.0062 | -0.0017 | 0.1072 | 0.0112 | 0.0033 | 0 | -0.0001 | -0.0001 | 0
IECS1 (Al 1000 Int.) | QC | 2024-08-26 10:52 | 0.0006 | 955.6486 | 0.0017 | 0.0091 | 0.0093 | 0.0001 | 0.0137 | 0.0007 | -0.0002 | 0.0009 | 0.0036 | 0.0468 | 0.0478 | 0.0014 | 0.0187 | 0 | 0.0068 | -0.002 | 0.0002 | 0.0129 | 0.0357 | 0.0091 | 0.1153 | 0.0067 | 0.0212 | 0.0004 | 0.0013 | 0.0003 | 0.0083
% Recovery |  |  | Al: 95.56%
Rinse1 | QC | 2024-08-26 10:55 | 0 | 0.0111 | 0 | 0.0034 | 0 | 0 | -0.0059 | -0.0001 | 0.0004 | -0.0002 | 0.0002 | 0.0023 | 0.038 | 0.0057 | -0.0023 | -0.0002 | 0.0019 | -0.0113 | -0.0005 | 0.0007 | -0.0043 | 0.0013 | 0.1098 | 0.0064 | 0.0015 | 0 | 0 | 0.0008 | 0.0004
Rinse | QC | 2024-08-26 10:57 | 0.0003 | 0.0046 | 0.0026 | 0.0033 | 0 | -0.0001 | -0.0148 | 0.0001 | 0.0001 | -0.0004 | 0.0005 | 0.0016 | 0.0273 | 0.0058 | -0.0022 | -0.0002 | 0.0017 | -0.0139 | -0.0015 | 0.0003 | -0.0069 | -0.0029 | 0.1124 | 0.0114 | 0.0031 | 0 | 0.0001 | 0.0007 | 0.0005
IECS2 (Fe 1000 Int.) | QC | 2024-08-26 11:00 | 0.0024 | 0.0635 | 0.0035 | -0.0111 | 0.0002 | 0.0001 | -0.0199 | -0.0036 | 0.0114 | 0.0089 | 0.0016 | 1045.4062 | 0.0391 | -0.0005 | 0.8779 K | 0.0097 | 0.0026 | -0.0123 | 0.0078 | -0.0009 | -0.0071 | 0.0025 | 0.1144 | -0.0225 | 0.0046 | 0.0009 | 0.0002 | -0.0001 | 0.0049
% Recovery |  |  | Fe: 104.5%
Rinse1 | QC | 2024-08-26 11:02 | -0.0003 | 0.0048 | 0.0006 | 0.0019 | 0 | -0.0001 | -0.0083 | -0.0001 | 0.0003 | -0.0003 | 0.0012 | 0.019 | 0.0344 | -0.003 | -0.0061 | -0.0001 | 0.0023 | -0.0139 | -0.0008 | -0.0017 | -0.0055 | -0.0025 | 0.1089 | 0.0139 | 0.0022 | 0 | -0.0001 | -0.0001 | 0.0006
Rinse | QC | 2024-08-26 11:05 | 0.0001 | -0.0016 | 0.0037 | 0.0009 | 0.0001 | -0.0002 | -0.0125 | 0 | 0.0004 | -0.0003 | 0.0001 | 0.0058 | 0.0394 | 0.0105 | -0.0013 | -0.0002 | 0.0007 | -0.0145 | -0.001 | 0.0037 | -0.0092 | -0.0029 | 0.1087 | 0.007 | 0.0001 | 0 | 0 | 0.0007 | 0.0002
IECS3 (Mn 100 Int.) | QC | 2024-08-26 11:08 | -0.0006 | 0.0194 | 0.0028 | -0.0033 | 0.0001 | -0.0001 | 0.7542 K | 0.0002 | -0.0017 | -0.0102 | 0.0004 | 0.0361 | 0.0279 | 0.0059 | 0.0069 | 99.3876 | 0.0016 | -0.0106 | -0.0023 | 0.0035 | -0.004 | -0.0174 | 0.1287 | 0.0352 | 0.0189 | 0.0002 | 0.0003 | 0.0004 | 0.0026
% Recovery |  |  | Mn: 99.39%
IECS4 (Ca Mg 1000 Int.) | QC | 2024-08-26 11:10 | 0.0005 | 0.0519 | 0.0033 | 0.0053 | 0.0022 | -0.0001 | 1010.9843 | 0.0086 | 0.0002 | 0.0016 | 0.003 | 0.0199 | 0.0303 | -0.0012 | 1026.6746 | 0.0267 | 0.0016 | 0.0038 | 0.0036 | 0.0085 | 0.0015 | 0.043 | 0.1095 | 0.009 | 0.0274 | 0.008 | -0.0003 | 0.0015 | 0.0141
% Recovery |  |  | Ca: 101.1%; Mg: 102.7%
Rinse1 | QC | 2024-08-26 11:13 | 0 | -0.0029 | 0.0007 | 0 | 0.0001 | -0.0002 | 0.0005 | 0 | 0.0003 | -0.0002 | 0 | 0.002 | 0.0371 | -0.0024 | 0.0033 | 0.0005 | 0.0013 | -0.0081 | 0.0003 | 0.0005 | -0.0075 | -0.0018 | 0.108 | 0.0092 | 0.0021 | 0 | -0.0001 | 0.0001 | 0.0004
Rinse | QC | 2024-08-26 11:15 | -0.0001 | -0.0004 | -0.0001 | -0.0003 | 0.0001 | 0 | -0.0078 | 0 | 0.0007 | -0.0002 | 0.0007 | 0.0009 | 0.0383 | -0.0032 | 0.0004 | 0.0001 | 0.001 | -0.0096 | -0.0001 | 0.0033 | -0.0081 | -0.0032 | 0.1103 | 0.0056 | 0.0017 | 0 | -0.0001 | 0 | 0.0005
IECS5 (Ti 20 Int.) | QC | 2024-08-26 11:18 | 0.0009 | 0.0088 | 0.0048 | 0.0004 | 0 | -0.0006 | 0.1258 | 0.0001 | 0.0001 | -0.0001 | -0.0003 | 0.0111 | 0.0172 | -0.002 | 0.0247 | 0.001 | 0.0014 | -0.0063 | -0.001 | 0.0028 | -0.011 | -0.0032 | 0.1089 | 0.0042 | 0.0131 | 0.0001 | 19.8735 | -0.0007 | 0.0015
% Recovery |  |  | Ti: 99.37%
Rinse1 | QC | 2024-08-26 11:20 | 0.0003 | 0.0003 | 0.0006 | -0.0008 | 0.0001 | -0.0002 | -0.0066 | 0 | -0.0001 | -0.0001 | 0.0008 | 0.0003 | 0.033 | -0.0018 | -0.0041 | -0.0002 | 0.0016 | -0.01 | -0.0005 | 0.0019 | -0.0077 | 0.0002 | 0.1111 | 0.0091 | 0.0015 | 0 | 0.0015 | 0.0001 | 0.0005
Rinse | QC | 2024-08-26 11:23 | 0.0001 | -0.003 | 0.0012 | -0.0012 | 0.0001 | -0.0001 | -0.0082 | -0.0001 | 0.0007 | -0.0001 | 0.0006 | -0.0002 | 0.0428 | 0.0055 | -0.0076 | -0.0002 | 0.0007 | -0.0103 | -0.0011 | 0.0033 | -0.0072 | 0.0015 | 0.1063 | 0.0092 | 0.0025 | 0 | 0.0004 | 0.0003 | 0.0002
ICB | QC | 2024-08-26 11:25 | -0.0001 | -0.0031 | -0.0011 | -0.0014 | 0 | -0.0002 | -0.0143 | -0.0001 | 0.0005 | -0.0001 | 0.0004 | 0.0007 | 0.0312 | -0.0041 | -0.0026 | -0.0002 | 0.0015 | -0.0127 | -0.0011 | 0.0037 | -0.0061 | -0.0048 | 0.1089 | 0.0136 | 0.0011 | 0 | 0.0001 | 0 | -0.0001
ICV (SS 5ppm All) | QC | 2024-08-26 11:28 | 4.9618 | 4.8891 | 4.9595 | 4.8833 | 4.8492 | 4.8377 | 4.9432 | 4.8909 | 4.8736 | 4.8898 | 4.6594 | 4.9181 | 4.8811 | 4.8129 | 4.9673 | 4.9005 | 4.9834 | 4.8674 | 4.8768 | 4.9181 | 5.0052 | 4.8413 | 5.1573 | 4.9204 | 4.67 | 4.9617 | 4.9407 | 4.934 | 4.8208
% Recovery |  |  | 99.24% | 97.78% | 99.19% | 97.67% | 96.98% | 96.75% | 98.86% | 97.82% | 97.47% | 97.8% | 93.19% | 98.36% | 97.62% | 96.26% | 99.35% | 98.01% | 99.67% | 97.35% | 97.54% | 98.36% | 100.1% | 96.83% | 103.1% | 98.41% | 93.4% | 99.23% | 98.81% | 98.68% | 96.42%
Rinse | QC | 2024-08-26 11:30 | 0.0003 | -0.0024 | 0.0148 | -0.0002 | 0.0002 | 0 | -0.0126 | 0.0002 | 0.0002 | -0.0001 | 0.0002 | -0.0008 | 0.034 | 0.0103 | -0.0038 | -0.0003 | 0.045 | -0.0112 | -0.0002 | 0.0008 | -0.0069 | -0.0039 | 0.1204 | 0.0255 | 0.0026 | 0.0001 | 0.0013 | 0.0005 | 0.0004
CCB | QC | 2024-08-26 11:33 | -0.0002 | -0.0022 | 0.0035 | 0.0002 | 0 | -0.0001 | -0.0187 | 0 | 0.0004 | -0.0003 | 0.0005 | 0.0005 | 0.0336 | 0.0054 | -0.0026 | 0 | 0.0132 | -0.0095 | 0 | 0.0062 | -0.0079 | -0.0024 | 0.1179 | 0.0164 | 0.0008 | 0 | 0.0004 | -0.0001 | 0
0.5ppm All (LLQC) | QC | 2024-08-26 11:35 | 0.4771 | 0.4765 | 0.4711 | 0.4746 | 0.468 | 0.4755 | 0.4793 | 0.4775 | 0.4796 | 0.4818 | 0.4528 | 0.4854 | 0.4717 | 0.4833 | 0.4657 | 0.4872 | 0.4495 | 0.4757 | 0.4797 | 0.476 | 0.477 | 0.4601 | 0.4674 | 0.4692 | 0.4569 | 0.4848 | 0.4756 | 0.4744 | 0.4706
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-26 11:38 | 0.4802 | 0.4764 | 0.4755 | 0.4779 | 0.4684 | 0.4771 | 0.4852 | 0.4811 | 0.4828 | 0.4853 | 0.4563 | 0.4859 | 0.4759 | 0.4847 | 0.4701 | 0.4869 | 0.4632 | 0.4802 | 0.4823 | 0.4773 | 0.4852 | 0.4655 | 0.5161 | 0.4808 | 0.4598 | 0.4856 | 0.4786 | 0.477 | 0.4737
RPD |  |  | 0.6477% | 0.02099% | 0.9296% | 0.6929% | 0.08543% | 0.3359% | 1.223% | 0.7511% | 0.665% | 0.7238% | 0.77% | 0.103% | 0.8864% | 1.626% | 0.9404% | 0.0616% | 3.002% | 0.9415% | 0.5405% | 0.2727% | 1.704% | 1.167% | 9.903% | 2.442% | 0.6327% | 0.1649% | 0.6288% | 0.5466% | 0.6566%
0.5ppm All (LLQC) | QC | 2024-08-26 11:35 | 0.4771 | 0.4765 | 0.4711 | 0.4746 | 0.468 | 0.4755 | 0.4793 | 0.4775 | 0.4796 | 0.4818 | 0.4528 | 0.4854 | 0.4717 | 0.4833 | 0.4657 | 0.4872 | 0.4495 | 0.4757 | 0.4797 | 0.476 | 0.477 | 0.4601 | 0.4674 | 0.4692 | 0.4569 | 0.4848 | 0.4756 | 0.4744 | 0.4706
% Recovery |  |  | 95.42% | 95.3% | 94.22% | 94.92% | 93.6% | 95.1% | 95.86% | 95.5% | 95.92% | 96.36% | 90.56% | 97.08% | 94.34% | 96.66% | 93.14% | 97.44% | 89.9% | 95.14% | 95.94% | 95.2% | 95.4% | 92.02% | 93.48% | 93.84% | 91.38% | 96.96% | 95.12% | 94.88% | 94.12%
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-26 11:38 | 0.4802 | 0.4764 | 0.4755 | 0.4779 | 0.4684 | 0.4771 | 0.4852 | 0.4811 | 0.4828 | 0.4853 | 0.4563 | 0.4859 | 0.4759 | 0.4847 | 0.4701 | 0.4869 | 0.4632 | 0.4802 | 0.4823 | 0.4773 | 0.4852 | 0.4655 | 0.5161 | 0.4808 | 0.4598 | 0.4856 | 0.4786 | 0.477 | 0.4737
% Recovery |  |  | 96.04% | 95.28% | 95.1% | 95.58% | 93.68% | 95.42% | 97.04% | 96.22% | 96.56% | 97.06% | 91.26% | 97.18% | 95.18% | 96.94% | 94.02% | 97.38% | 92.64% | 96.04% | 96.46% | 95.46% | 97.04% | 93.1% | 103.2% | 96.16% | 91.96% | 97.12% | 95.72% | 95.4% | 94.74%
CCV (5ppm All) | QC | 2024-08-26 11:40 | 4.8297 | 4.7689 | 4.8274 | 4.851 | 4.7504 | 4.7583 | 4.869 | 4.8471 | 4.839 | 4.8497 | 4.6283 | 4.8281 | 4.771 | 4.7916 | 4.8449 | 4.8287 | 4.8803 | 4.7391 | 4.8387 | 4.8406 | 4.931 | 4.8132 | 5.2404 | 4.8466 | 4.8671 | 4.8371 | 4.8284 | 4.8428 | 4.7968
% Recovery |  |  | 96.59% | 95.38% | 96.55% | 97.02% | 95.01% | 95.17% | 97.38% | 96.94% | 96.78% | 96.99% | 92.57% | 96.56% | 95.42% | 95.83% | 96.9% | 96.57% | 97.61% | 94.78% | 96.77% | 96.81% | 98.62% | 96.26% | 104.8% | 96.93% | 97.34% | 96.74% | 96.57% | 96.86% | 95.94%
CCB | QC | 2024-08-26 12:59 | 0 | -0.0034 | 0.0078 | -0.002 | 0.0003 | 0 | -0.0166 | 0 | 0.0006 | -0.0002 | 0.0012 | 0.0003 | 0.0228 | -0.0041 | -0.0049 | -0.0001 | 0.0209 | -0.0179 | -0.0008 | 0 | -0.0074 | -0.0061 | 0.163 | 0.0141 | 0.0025 | 0 | 0.0003 | 0.0003 | 0
Rinse | QC | 2024-08-26 15:09 | 0.0001 | -0.0018 | 0.0041 | -0.0032 | 0 | -0.0002 | -0.0149 | 0 | 0.0006 | -0.0002 | 0.0003 | 0.0007 | 0.024 | 0.0016 | -0.0053 | -0.0002 | 0.0027 | -0.0147 | 0.0001 | 0.0025 | -0.0067 | -0.0061 | 0.1114 | 0.0121 | 0.0011 | 0 | 0 | 0 | 0.0001
CCV (5ppm All) | QC | 2024-08-26 15:11 | 4.9388 | 4.8292 | 4.9747 | 4.9869 | 4.8694 | 4.8934 | 4.9624 | 4.9853 | 4.9768 | 4.9902 | 4.7354 | 4.9025 | 4.8388 | 4.8803 | 4.976 | 4.9054 | 5.0045 | 4.8047 | 4.9781 | 5.2294 | 5.0802 | 5.0456 | 5.321 | 5.0215 | 4.9475 | 4.9421 | 4.935 | 4.9634 | 4.9364
% Recovery |  |  | 98.78% | 96.58% | 99.49% | 99.74% | 97.39% | 97.87% | 99.25% | 99.71% | 99.54% | 99.8% | 94.71% | 98.05% | 96.78% | 97.61% | 99.52% | 98.11% | 100.1% | 96.09% | 99.56% | 104.6% | 101.6% | 100.9% | 106.4% | 100.4% | 98.95% | 98.84% | 98.7% | 99.27% | 98.73%
CCB | QC | 2024-08-26 15:14 | 0.0002 | -0.0002 | 0.0154 | -0.0014 | 0.0002 | 0 | -0.0143 | 0.0001 | 0.0005 | 0 | -0.0002 | 0.0017 | 0.0355 | -0.0098 | -0.0041 | 0 | 0.0536 | -0.0153 | -0.0007 | 0.0045 | -0.0076 | -0.0048 | 0.183 | 0.0334 | 0.0026 | 0.0001 | 0.0016 | 0.0001 | 0.0001
08262024 LB-2 | Sample | 2024-08-26 15:17 | 0.0002 | 0.0001 | 0.0052 | -0.0019 | 0.0001 | -0.0001 | -0.0115 | 0 | 0.0003 | -0.0002 | 0.0005 | 0.0007 | 0.0277 | -0.0078 | -0.0066 | -0.0003 | 0.0145 | -0.0147 | -0.001 | 0.0039 | -0.0087 | -0.0036 | 0.116 | 0.0132 | 0.002 | 0 | 0.0004 | 0.0001 | 0.0005
10x 08265024 LCS-2 | Sample | 2024-08-26 15:19 | 0.0091 | 0.4747 | 0.245 | 0.0466 | 0.2424 | 0.0484 | 0.474 | 0.0979 | 0.4877 | 0.2465 | 0.2302 | 0.497 | 0.4924 | 0 | 0.4806 | 0.0995 | 0.0921 | 0.4818 | 0.241 | 0.1057 | 0.2418 | 0.2439 | 0.2421 | 0.2456 | 0.0954 | 0.1001 | 0.0977 | 0.2447 | 0.0958
10x 08262024 LCS-2- Analytical Duplicate | Sample | 2024-08-26 15:22 | 0.0098 | 0.4716 | 0.2447 | 0.047 | 0.2402 | 0.048 | 0.4719 | 0.0979 | 0.4881 | 0.2469 | 0.23 | 0.4937 | 0.4729 | -0.0089 | 0.4814 | 0.0985 | 0.0911 | 0.4819 | 0.2415 | 0.1077 | 0.2429 | 0.2408 | 0.2613 | 0.2502 | 0.0946 | 0.0992 | 0.0978 | 0.2443 | 0.0959
RPD |  |  | 7.407% | 0.6552% | 0.1225% | 0.8547% | 0.9117% | 0.8299% | 0.444% | 0% | 0.08198% | 0.1621% | 0.08692% | 0.6662% | 4.04% | 1.626% | 0.1663% | 1.01% | 1.092% | 0.02075% | 0.2073% | 1.874% | 0.4539% | 1.279% | 7.628% | 1.856% | 0.8421% | 0.9032% | 0.1023% | 0.1636% | 0.1043%
08262024 LB-2 | Sample | 2024-08-26 15:17 | 0.0002 | 0.0001 | 0.0052 | -0.0019 | 0.0001 | -0.0001 | -0.0115 | 0 | 0.0003 | -0.0002 | 0.0005 | 0.0007 | 0.0277 | -0.0078 | -0.0066 | -0.0003 | 0.0145 | -0.0147 | -0.001 | 0.0039 | -0.0087 | -0.0036 | 0.116 | 0.0132 | 0.002 | 0 | 0.0004 | 0.0001 | 0.0005
10x 08265024 LCS-2 | Sample | 2024-08-26 15:19 | 0.0091 | 0.4747 | 0.245 | 0.0466 | 0.2424 | 0.0484 | 0.474 | 0.0979 | 0.4877 | 0.2465 | 0.2302 | 0.497 | 0.4924 | 0 | 0.4806 | 0.0995 | 0.0921 | 0.4818 | 0.241 | 0.1057 | 0.2418 | 0.2439 | 0.2421 | 0.2456 | 0.0954 | 0.1001 | 0.0977 | 0.2447 | 0.0958
R% |  |  | 89% | 94.92% | 95.92% | 97% | 96.92% | 97% | 97.1% | 97.9% | 97.48% | 98.68% | 91.88% | 99.26% | 92.94% |  | 97.44% | 99.8% | 77.6% | 99.3% | 96.8% | 101.8% | 100.2% | 99% | 50.44% | 92.96% | 93.4% | 100.1% | 97.3% | 97.84% | 95.3%
08262024 LB-2 | Sample | 2024-08-26 15:17 | 0.0002 | 0.0001 | 0.0052 | -0.0019 | 0.0001 | -0.0001 | -0.0115 | 0 | 0.0003 | -0.0002 | 0.0005 | 0.0007 | 0.0277 | -0.0078 | -0.0066 | -0.0003 | 0.0145 | -0.0147 | -0.001 | 0.0039 | -0.0087 | -0.0036 | 0.116 | 0.0132 | 0.002 | 0 | 0.0004 | 0.0001 | 0.0005
10x 08262024 LCS-2- Analytical Duplicate | Sample | 2024-08-26 15:22 | 0.0098 | 0.4716 | 0.2447 | 0.047 | 0.2402 | 0.048 | 0.4719 | 0.0979 | 0.4881 | 0.2469 | 0.23 | 0.4937 | 0.4729 | -0.0089 | 0.4814 | 0.0985 | 0.0911 | 0.4819 | 0.2415 | 0.1077 | 0.2429 | 0.2408 | 0.2613 | 0.2502 | 0.0946 | 0.0992 | 0.0978 | 0.2443 | 0.0959
R% |  |  | 96% | 94.3% | 95.8% | 97.8% | 96.04% | 96.2% | 96.68% | 97.9% | 97.56% | 98.84% | 91.8% | 98.6% | 89.04% |  | 97.6% | 98.8% | 76.6% | 99.32% | 97% | 103.8% | 100.6% | 97.76% | 58.12% | 94.8% | 92.6% | 99.2% | 97.4% | 97.68% | 95.4%
08262024 LB-2 | Sample | 2024-08-26 15:17 | 0.0002 | 0.0001 | 0.0052 | -0.0019 | 0.0001 | -0.0001 | -0.0115 | 0 | 0.0003 | -0.0002 | 0.0005 | 0.0007 | 0.0277 | -0.0078 | -0.0066 | -0.0003 | 0.0145 | -0.0147 | -0.001 | 0.0039 | -0.0087 | -0.0036 | 0.116 | 0.0132 | 0.002 | 0 | 0.0004 | 0.0001 | 0.0005
5x 08262024 LCS-2 | Sample | 2024-08-26 15:24 | 0.019 | 0.9546 | 0.4902 | 0.0948 | 0.484 | 0.0971 | 0.9666 | 0.1962 | 0.982 | 0.4911 | 0.4582 | 0.9936 | 0.9653 | -0.0003 | 0.9807 | 0.1978 | 0.1799 | 0.9785 | 0.4847 | 0.2071 | 0.4874 | 0.4869 | 0.473 | 0.4898 | 0.1861 | 0.199 | 0.1944 | 0.4873 | 0.1946
R% |  |  | 94% | 95.45% | 97% | 96.7% | 96.78% | 97.2% | 97.81% | 98.1% | 98.17% | 98.26% | 91.54% | 99.29% | 93.76% |  | 98.73% | 99.05% | 82.7% | 99.32% | 97.14% | 101.6% | 99.22% | 98.1% | 71.4% | 95.32% | 92.05% | 99.5% | 97% | 97.44% | 97.05%
08262024 LCS-2 | Sample | 2024-08-26 15:27 | 0.0963 | 4.6699 | 2.4802 | 0.4939 | 2.3903 | 0.4793 | 4.8571 | 0.9906 | 4.9267 | 2.4482 | 2.2906 | 4.8599 | 4.8696 | -0.0012 | 4.993 | 0.9732 | 0.9373 | 4.828 | 2.4507 | 1.0387 | 2.4953 | 2.5074 | 2.5632 | 2.4876 | 0.9582 | 0.9815 | 0.9822 | 2.4491 | 0.977
08262024 LCS-2- Analytical Duplicate | Sample | 2024-08-26 15:29 | 0.0954 | 4.6689 | 2.4679 | 0.4892 | 2.395 | 0.477 | 4.8094 | 0.981 | 4.8862 | 2.4261 | 2.2895 | 4.8664 | 4.8059 | -0.0077 | 4.9477 | 0.9754 | 0.9521 | 4.7946 | 2.4232 | 1.0303 | 2.4739 | 2.4844 | 2.5859 | 2.4727 | 0.9547 | 0.9829 | 0.973 | 2.4409 | 0.9666
RPD |  |  | 0.939% | 0.02142% | 0.4972% | 0.9562% | 0.1964% | 0.481% | 0.9869% | 0.9738% | 0.8254% | 0.9068% | 0.04803% | 0.1337% | 1.317% | 1.626% | 0.9114% | 0.2258% | 1.567% | 0.6942% | 1.128% | 0.812% | 0.8613% | 0.9215% | 0.8817% | 0.6008% | 0.3659% | 0.1425% | 0.9411% | 0.3354% | 1.07%
08262024 LB-2 | Sample | 2024-08-26 15:17 | 0.0002 | 0.0001 | 0.0052 | -0.0019 | 0.0001 | -0.0001 | -0.0115 | 0 | 0.0003 | -0.0002 | 0.0005 | 0.0007 | 0.0277 | -0.0078 | -0.0066 | -0.0003 | 0.0145 | -0.0147 | -0.001 | 0.0039 | -0.0087 | -0.0036 | 0.116 | 0.0132 | 0.002 | 0 | 0.0004 | 0.0001 | 0.0005
08262024 LCS-2 | Sample | 2024-08-26 15:27 | 0.0963 | 4.6699 | 2.4802 | 0.4939 | 2.3903 | 0.4793 | 4.8571 | 0.9906 | 4.9267 | 2.4482 | 2.2906 | 4.8599 | 4.8696 | -0.0012 | 4.993 | 0.9732 | 0.9373 | 4.828 | 2.4507 | 1.0387 | 2.4953 | 2.5074 | 2.5632 | 2.4876 | 0.9582 | 0.9815 | 0.9822 | 2.4491 | 0.977
R% |  |  | 96.1% | 93.4% | 99% | 99.16% | 95.61% | 95.88% | 97.37% | 99.06% | 98.53% | 97.94% | 91.6% | 97.18% | 96.84% |  | 99.99% | 97.35% | 92.28% | 96.85% | 98.07% | 103.9% | 100.2% | 100.4% | 97.89% | 98.98% | 95.82% | 98.15% | 98.18% | 97.96% | 97.65%
08262024 LB-2 | Sample | 2024-08-26 15:17 | 0.0002 | 0.0001 | 0.0052 | -0.0019 | 0.0001 | -0.0001 | -0.0115 | 0 | 0.0003 | -0.0002 | 0.0005 | 0.0007 | 0.0277 | -0.0078 | -0.0066 | -0.0003 | 0.0145 | -0.0147 | -0.001 | 0.0039 | -0.0087 | -0.0036 | 0.116 | 0.0132 | 0.002 | 0 | 0.0004 | 0.0001 | 0.0005
08262024 LCS-2- Analytical Duplicate | Sample | 2024-08-26 15:29 | 0.0954 | 4.6689 | 2.4679 | 0.4892 | 2.395 | 0.477 | 4.8094 | 0.981 | 4.8862 | 2.4261 | 2.2895 | 4.8664 | 4.8059 | -0.0077 | 4.9477 | 0.9754 | 0.9521 | 4.7946 | 2.4232 | 1.0303 | 2.4739 | 2.4844 | 2.5859 | 2.4727 | 0.9547 | 0.9829 | 0.973 | 2.4409 | 0.9666
R% |  |  | 95.2% | 93.38% | 98.51% | 98.22% | 95.8% | 95.42% | 96.42% | 98.1% | 97.72% | 97.05% | 91.56% | 97.31% | 95.56% |  | 99.09% | 97.57% | 93.76% | 96.19% | 96.97% | 103% | 99.3% | 99.52% | 98.8% | 98.38% | 95.47% | 98.29% | 97.26% | 97.63% | 96.61%
Rinse | QC | 2024-08-26 15:32 | -0.0003 | -0.0025 | 0.0141 | -0.0021 | 0.0002 | -0.0001 | -0.0112 | 0 | 0.0005 | -0.0003 | 0.0004 | 0.0004 | 0.0199 | -0.0006 | -0.005 | 0 | 0.0305 | -0.0155 | -0.0014 | 0.0008 | -0.0083 | -0.001 | 0.1824 | 0.0173 | 0.0027 | 0 | 0.0003 | 0.0001 | -0.0001
CCV (5ppm All) | QC | 2024-08-26 15:34 | 4.9389 | 4.8512 | 4.9794 | 4.9897 | 4.8841 | 4.9022 | 4.9564 | 4.9838 | 4.9764 | 4.9904 | 4.7376 | 4.9213 | 4.8156 | 4.8705 | 4.9696 | 4.9116 | 5.0241 | 4.8023 | 4.9756 | 5.2477 | 5.0848 | 5.0369 | 5.3825 | 5.027 | 4.9736 | 4.9615 | 4.9335 | 4.9644 | 4.9204
% Recovery |  |  | 98.78% | 97.02% | 99.59% | 99.79% | 97.68% | 98.04% | 99.13% | 99.68% | 99.53% | 99.81% | 94.75% | 98.43% | 96.31% | 97.41% | 99.39% | 98.23% | 100.5% | 96.05% | 99.51% | 105% | 101.7% | 100.7% | 107.7% | 100.5% | 99.47% | 99.23% | 98.67% | 99.29% | 98.41%
CCB | QC | 2024-08-26 15:37 | 0.0003 | -0.0053 | 0.0156 | -0.0012 | 0.0002 | -0.0001 | -0.0152 | 0.0001 | 0.0004 | 0 | 0.0003 | -0.0002 | 0.031 | -0.0007 | -0.0084 | -0.0001 | 0.0461 | -0.0125 | -0.001 | 0.006 | -0.0074 | -0.0093 | 0.1541 | 0.0243 | 0.0033 | 0.0001 | 0.0014 | -0.0001 | 0
0.5ppm All (LLQC) | QC | 2024-08-26 15:39 | 0.4906 | 0.4836 | 0.4869 | 0.4883 | 0.4772 | 0.4868 | 0.491 | 0.4917 | 0.4949 | 0.4969 | 0.465 | 0.4887 | 0.4833 | 0.4872 | 0.479 | 0.4956 | 0.4776 | 0.4797 | 0.4944 | 0.511 | 0.4949 | 0.4864 | 0.5089 | 0.4933 | 0.4689 | 0.4931 | 0.4883 | 0.488 | 0.4838
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-26 15:42 | 0.4882 | 0.4883 | 0.4903 | 0.4868 | 0.4806 | 0.4876 | 0.4951 | 0.4909 | 0.4935 | 0.4953 | 0.465 | 0.4942 | 0.4764 | 0.4839 | 0.4787 | 0.497 | 0.4781 | 0.4753 | 0.4924 | 0.5118 | 0.4962 | 0.4818 | 0.5352 | 0.4921 | 0.4672 | 0.497 | 0.4871 | 0.4878 | 0.4827
RPD |  |  | 0.4904% | 0.9672% | 0.6959% | 0.3077% | 0.71% | 0.1642% | 0.8316% | 0.1628% | 0.2833% | 0.3225% | 0% | 1.119% | 1.438% | 1.626% | 0.06265% | 0.2821% | 0.1046% | 0.9215% | 0.4054% | 0.1564% | 0.2623% | 0.9502% | 5.038% | 0.2436% | 0.3632% | 0.7878% | 0.2461% | 0.04099% | 0.2276%
0.5ppm All (LLQC) | QC | 2024-08-26 15:39 | 0.4906 | 0.4836 | 0.4869 | 0.4883 | 0.4772 | 0.4868 | 0.491 | 0.4917 | 0.4949 | 0.4969 | 0.465 | 0.4887 | 0.4833 | 0.4872 | 0.479 | 0.4956 | 0.4776 | 0.4797 | 0.4944 | 0.511 | 0.4949 | 0.4864 | 0.5089 | 0.4933 | 0.4689 | 0.4931 | 0.4883 | 0.488 | 0.4838
% Recovery |  |  | 98.12% | 96.72% | 97.38% | 97.66% | 95.44% | 97.36% | 98.2% | 98.34% | 98.98% | 99.38% | 93% | 97.74% | 96.66% | 97.44% | 95.8% | 99.12% | 95.52% | 95.94% | 98.88% | 102.2% | 98.98% | 97.28% | 101.8% | 98.66% | 93.78% | 98.62% | 97.66% | 97.6% | 96.76%
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-26 15:42 | 0.4882 | 0.4883 | 0.4903 | 0.4868 | 0.4806 | 0.4876 | 0.4951 | 0.4909 | 0.4935 | 0.4953 | 0.465 | 0.4942 | 0.4764 | 0.4839 | 0.4787 | 0.497 | 0.4781 | 0.4753 | 0.4924 | 0.5118 | 0.4962 | 0.4818 | 0.5352 | 0.4921 | 0.4672 | 0.497 | 0.4871 | 0.4878 | 0.4827
% Recovery |  |  | 97.64% | 97.66% | 98.06% | 97.36% | 96.12% | 97.52% | 99.02% | 98.18% | 98.7% | 99.06% | 93% | 98.84% | 95.28% | 96.78% | 95.74% | 99.4% | 95.62% | 95.06% | 98.48% | 102.4% | 99.24% | 96.36% | 107% | 98.42% | 93.44% | 99.4% | 97.42% | 97.56% | 96.54%
CCS1 (SS 5ppm) | QC | 2024-08-26 15:44 | 5.0765 | 4.9933 | 5.1119 | 5.0207 | 5.0011 | 5.019 | 5.0408 | 5.0254 | 5.0331 | 5.0271 | 4.7777 | 5.0138 | 4.9132 | 4.8744 | 5.083 | 5.0055 | 5.1896 | 4.9194 | 5.0073 | 5.3275 | 5.1391 | 5.0734 | 5.4158 | 5.0701 | 4.8141 | 5.1066 | 5.0508 | 5.0636 | 4.9732
% Recovery |  |  | 101.5% | 99.87% | 102.2% | 100.4% | 100% | 100.4% | 100.8% | 100.5% | 100.7% | 100.5% | 95.55% | 100.3% | 98.26% | 97.49% | 101.7% | 100.1% | 103.8% | 98.39% | 100.1% | 106.5% | 102.8% | 101.5% | 108.3% | 101.4% | 96.28% | 102.1% | 101% | 101.3% | 99.46%
CCB | QC | 2024-08-26 15:47 | -0.0001 | -0.0025 | 0.0143 | -0.0005 | 0.0002 | 0 | -0.0142 | 0.0001 | 0.0005 | -0.0004 | 0.0005 | 0.0008 | 0.0391 | -0.008 | -0.007 | -0.0001 | 0.0439 | -0.0143 | -0.0012 | 0.0031 | -0.0094 | -0.0061 | 0.1272 | 0.0237 | 0.0028 | 0.0001 | 0.0012 | 0.0002 | 0.0003
MQ-Blank-02 | Sample | 2024-08-26 15:49 | 0.0001 | 0.0012 | 0 | -0.0026 | 0 | 0.0001 | -0.0038 | 0.0001 | 0.0003 | -0.0002 | 0.0004 | -0.0002 | 0.0338 | -0.0068 | -0.0082 | 0 | 0.0043 | -0.0137 | 0.0002 | 0.0037 | -0.0084 | -0.0051 | 0.1059 | 0.0116 | 0.0031 | 0.0001 | 0.0001 | 0.0003 | 0.0004
Rinse | QC | 2024-08-26 15:52 | 0.0001 | 0.0004 | 0.0044 | -0.0017 | 0 | -0.0002 | -0.0177 | 0.0001 | 0.0004 | -0.0002 | 0.0008 | -0.0003 | 0.0301 | -0.0084 | -0.0031 | -0.0003 | 0.0099 | -0.0148 | -0.0001 | -0.0013 | -0.0084 | -0.0053 | 0.1127 | 0.0122 | 0.0005 | 0 | 0.0002 | 0.0002 | 0.0004
CCV (5ppm All) | QC | 2024-08-26 15:54 | 4.9369 | 4.874 | 4.9746 | 4.9946 | 4.8944 | 4.8892 | 4.9557 | 4.976 | 4.9722 | 4.9812 | 4.7368 | 4.9438 | 4.8666 | 4.8894 | 4.9599 | 4.9434 | 5.0179 | 4.8117 | 4.9632 | 5.2589 | 5.0722 | 5.0212 | 5.3428 | 4.9922 | 4.9746 | 4.9731 | 4.9332 | 4.9609 | 4.9532
% Recovery |  |  | 98.74% | 97.48% | 99.49% | 99.89% | 97.89% | 97.78% | 99.11% | 99.52% | 99.44% | 99.62% | 94.74% | 98.88% | 97.33% | 97.79% | 99.2% | 98.87% | 100.4% | 96.23% | 99.26% | 105.2% | 101.4% | 100.4% | 106.9% | 99.84% | 99.49% | 99.46% | 98.66% | 99.22% | 99.06%
CCB | QC | 2024-08-26 15:57 | -0.0001 | -0.0009 | 0.0147 | 0.0009 | 0.0004 | 0.0001 | -0.0142 | 0.0001 | 0.0001 | 0 | 0.0004 | 0.0005 | 0.0169 | -0.0067 | -0.0079 | 0.0002 | 0.0507 | -0.0154 | -0.0007 | 0.0033 | -0.0082 | -0.0047 | 0.1762 | 0.0281 | 0.0024 | 0.0003 | 0.0016 | 0.0006 | 0.0005
0.5% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 15:59 | -0.0004 | 0.1804 | 0.0154 | 0.0066 | 0.0418 | -0.0001 | 0.3158 | 0 | 0.0002 | 0.0005 | 0.0025 | 0.1616 | 6.6046 | 0.0095 | 0.0343 | 0.001 | 0.0174 | 1080.1282 | 1.455 | 0.1984 | -0.003 | 0.1711 | 0.1259 | 0.0175 | 0.0436 | 0.0013 | 0.0013 | 0.0004 | 0.0059
1% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:02 | -0.0002 | 0.1093 | 0.0126 | 0.0075 | 0.0466 | 0 | 0.5437 | 0.0001 | 0.0003 | 0.0001 | 0.003 | 0.0639 | 0.6065 | 0.001 | 0.0789 | 0.0004 | 0.0091 | 1782.5297 | 1.364 | 0.1745 | -0.0046 | 0.3071 | 0.1101 | 0.0136 | 0.0677 | 0.0021 | 0.0013 | 0.0003 | 0.0057
5% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:04 | -0.0008 | 0.1806 | 0.0084 | 0.0292 | 0.0881 | -0.0001 | 2.2565 | -0.0001 | 0.0003 | 0.0004 | 0.0022 | 0.1344 | 1.2881 | 0.0147 | 0.3968 | 0.0027 | 0.0082 |  | 0.8643 | 0.1541 | -0.0106 | 1.07 | 0.1085 | 0.008 | 0.3693 | 0.0079 | 0.0072 | 0.0015 | 0.0106
10% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:07 | -0.0012 | 0.1245 | 0.0069 | 0.0561 | 0.1333 | 0.0001 | 4.198 | -0.0001 | -0.0002 | 0.0005 | 0.0006 | 0.2474 | 2.5736 | 0.0322 | 0.7236 | 0.0088 | 0.0076 |  | 0.2948 | 0.2303 | -0.0143 | 1.9988 | 0.1094 | 0.0094 | 0.643 | 0.0154 | 0.0105 | 0.003 | 0.0114
15% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:09 | -0.0013 | 0.2724 | 0.0081 | 0.069 | 0.1331 | -0.0001 | 4.4436 | 0.0002 | -0.0003 | 0.0006 | 0 | 1.1278 | 5.9549 | 0.0369 | 0.7401 | 0.0075 | 0.0078 |  | 0.2598 | 0.4848 | -0.0148 | 2.921 | 0.1073 | 0.0067 | 0.9399 | 0.0171 | 0.0148 | 0.0046 | 0.014
0.5% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:12 | -0.0002 | 0.3046 | 0.0061 | 0.003 | 0.003 | -0.0001 | 0.2535 | 0 | 0.0011 | 0.0004 | 0.0024 | 0.1919 | 0.4832 | 0.0093 | 0.0299 | 0.0006 | 0.0037 | 1061.1289 | 4.6501 | 1.6312 | -0.0029 | 0.1404 | 0.12 | 0.0111 | 0.0661 | 0.0007 | 0.0006 | 0.0008 | 0.0056
1% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:14 | -0.0002 | 0.4131 | 0.0041 | 0.0076 | 0.004 | -0.0001 | 0.5031 | 0 | 0.0012 | 0.0004 | 0.0027 | 0.099 | 0.5255 | 0.0028 | 0.0744 | 0.0011 | 0.004 | 1752.4356 | 5.0444 | 2.0521 | -0.0019 | 0.2448 | 0.1071 | 0.015 | 0.1281 | 0.0016 | 0.0016 | 0.0005 | 0.0057
5% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:17 | -0.001 | 2.8795 | 0.0061 | 0.0277 | 0.03 | 0 | 2.4264 | 0 | 0.0019 | 0.0025 | 0.0025 | 4.3695 | 1.4801 | 0.0114 | 0.4514 | 0.0089 | 0.005 |  | 19.6945 | 4.4225 | -0.0107 | 1.1654 | 0.1183 | 0.0117 | 0.7978 | 0.0083 | 0.0172 | 0.0018 | 0.0124
10% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:19 | -0.0009 | 0.0738 | 0.0069 | 0.0524 | 0.0639 | -0.0001 | 3.8722 | -0.0003 | 0.0009 | 0.0029 | -0.0001 | 6.2456 | 3.0391 | 0.0282 | 0.6991 | 0.0087 | 0.0049 |  | 9.1367 | 3.6268 | -0.0167 | 2.4262 | 0.104 | 0.0087 | 1.3056 | 0.0153 | 0.0175 | 0.0039 | 0.0053
15% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:22 | -0.0012 | 0.0874 | 0.0059 | 0.0702 | 0.0723 | 0 | 4.1223 | -0.0002 | 0.0003 | 0.0022 | -0.001 | 6.0412 | 7.8595 | 0.038 | 0.7319 | 0.0097 | 0.0068 |  | 9.3892 | 3.4593 | -0.0156 | 3.5913 | 0.1042 | 0.0016 | 1.9866 | 0.0168 | 0.0242 | 0.0059 | 0.0044
Rinse1 | QC | 2024-08-26 16:24 | -0.0003 | -0.0019 | 0.0033 | -0.0032 | -0.0001 | -0.0001 | -0.0093 | -0.0001 | 0.0006 | -0.0002 | 0.0008 | 0.0014 | 0.0378 | -0.0084 | -0.008 | -0.0002 | 0.0022 | 0.7828 | -0.001 | 0.0029 | -0.0074 | -0.0042 | 0.12 | 0.0085 | 0.0015 | 0 | -0.0001 | -0.0002 | 0.0021
Rinse1 | QC | 2024-08-26 16:27 | -0.0003 | -0.0035 | -0.0003 | -0.0034 | 0.0002 | -0.0002 | -0.0085 | 0 | 0.0003 | 0.0001 | 0.0001 | 0.0009 | 0.0298 | -0.0072 | -0.0049 | -0.0002 | 0.0022 | 0.4574 | -0.0012 | 0.0035 | -0.0101 | -0.0039 | 0.1124 | 0.0059 | 0.0012 | 0 | -0.0001 | 0.0005 | 0.0023
Rinse | QC | 2024-08-26 16:29 | 0.0001 | -0.0025 | -0.0013 | -0.0035 | 0.0001 | -0.0001 | -0.0172 | 0 | 0.0005 | -0.0001 | 0.001 | 0.0001 | 0.0354 | -0.0075 | -0.0067 | -0.0003 | 0.0022 | 0.3793 | -0.0009 | -0.0007 | -0.0059 | -0.0025 | 0.1147 | 0.0111 | 0.0013 | 0 | 0 | 0.0002 | 0.0019
CCV (5ppm All) | QC | 2024-08-26 16:32 | 4.9243 | 4.8786 | 4.9225 | 4.9468 | 4.8842 | 4.9004 | 4.9412 | 4.9472 | 4.9429 | 4.9657 | 4.7225 | 4.9091 | 4.8315 | 4.8826 | 4.9254 | 4.9456 | 4.9739 | 5.032 | 4.9424 | 5.2544 | 5.042 | 4.9785 | 5.2856 | 4.962 | 4.8957 | 4.9329 | 4.9221 | 4.9492 | 4.9082
% Recovery |  |  | 98.49% | 97.57% | 98.45% | 98.94% | 97.68% | 98.01% | 98.82% | 98.94% | 98.86% | 99.31% | 94.45% | 98.18% | 96.63% | 97.65% | 98.51% | 98.91% | 99.48% | 100.6% | 98.85% | 105.1% | 100.8% | 99.57% | 105.7% | 99.24% | 97.91% | 98.66% | 98.44% | 98.98% | 98.16%
CCB | QC | 2024-08-26 16:34 | -0.0002 | -0.0029 | 0.0166 | -0.0014 | 0.0001 | 0 | -0.0141 | 0.0001 | 0.0005 | 0.0001 | 0.0006 | 0.001 | 0.0409 | 0.0017 | -0.0053 | -0.0003 | 0.0491 | 0.1999 | -0.0002 | 0.0027 | -0.0087 | -0.0064 | 0.1625 | 0.028 | 0.0027 | 0.0001 | 0.0015 | 0 | 0.0009
08162024 MB | Sample | 2024-08-26 16:51 | 0.0001 | 0.0018 | 0.0046 | -0.0029 | 0.0001 | -0.0001 | -0.0144 | 0 | 0.0008 | -0.0002 | 0.0008 | 0.0011 | 0.0277 | 0.0013 | -0.0076 | -0.0002 | 0.0054 | 0.1339 | -0.0002 | 0.0024 | -0.0058 | -0.0035 | 0.1194 | 0.0148 | 0.0035 | 0 | 0.0002 | 0.0003 | 0.0013
08162024 MFB | Sample | 2024-08-26 16:54 | 0.0979 | 4.7972 | 2.4586 | 0.4941 | 2.4539 | 0.4817 | 4.9762 | 1.0022 | 5.0185 | 2.5104 | 2.387 | 4.9622 | 4.944 | 0.0001 | 5.0833 | 1.0047 | 0.9835 | 5.0818 | 2.4711 | 1.0582 | 2.5382 | 2.4606 | 2.623 | 2.4157 | 0.9733 | 1.0096 | 1.0026 | 2.518 | 0.9784
08162024 MFB - Analyticsal Duplicate | Sample | 2024-08-26 16:56 | 0.098 | 4.7714 | 2.4724 | 0.4945 | 2.4449 | 0.4804 | 4.965 | 1.0023 | 5.021 | 2.506 | 2.3818 | 4.9418 | 4.9284 | 0.0002 | 5.0728 | 1 | 1.0013 | 5.0693 | 2.4681 | 1.0578 | 2.5369 | 2.4574 | 2.6724 | 2.419 | 0.9814 | 1.0046 | 1.0016 | 2.5237 | 0.9774
RPD |  |  | 0.1021% | 0.5393% | 0.5597% | 0.08092% | 0.3674% | 0.2702% | 0.2253% | 0.009978% | 0.0498% | 0.1754% | 0.2181% | 0.412% | 0.316% | 1.626% | 0.2068% | 0.4689% | 1.794% | 0.2463% | 0.1215% | 0.03781% | 0.05123% | 0.1301% | 1.866% | 0.1365% | 0.8288% | 0.4965% | 0.09979% | 0.2261% | 0.1023%
08162024 MB | Sample | 2024-08-26 16:51 | 0.0001 | 0.0018 | 0.0046 | -0.0029 | 0.0001 | -0.0001 | -0.0144 | 0 | 0.0008 | -0.0002 | 0.0008 | 0.0011 | 0.0277 | 0.0013 | -0.0076 | -0.0002 | 0.0054 | 0.1339 | -0.0002 | 0.0024 | -0.0058 | -0.0035 | 0.1194 | 0.0148 | 0.0035 | 0 | 0.0002 | 0.0003 | 0.0013
08162024 MFB | Sample | 2024-08-26 16:54 | 0.0979 | 4.7972 | 2.4586 | 0.4941 | 2.4539 | 0.4817 | 4.9762 | 1.0022 | 5.0185 | 2.5104 | 2.387 | 4.9622 | 4.944 | 0.0001 | 5.0833 | 1.0047 | 0.9835 | 5.0818 | 2.4711 | 1.0582 | 2.5382 | 2.4606 | 2.623 | 2.4157 | 0.9733 | 1.0096 | 1.0026 | 2.518 | 0.9784
R% |  |  | 97.8% | 95.91% | 98.16% | 99.4% | 98.15% | 96.36% | 99.81% | 100.2% | 100.4% | 100.4% | 95.45% | 99.22% | 98.33% |  | 101.8% | 100.5% | 97.81% | 98.96% | 98.85% | 105.8% | 101.8% | 98.56% | 100.1% | 96.04% | 97.33% | 101% | 100.2% | 100.7% | 97.71%
08162024 MB | Sample | 2024-08-26 16:51 | 0.0001 | 0.0018 | 0.0046 | -0.0029 | 0.0001 | -0.0001 | -0.0144 | 0 | 0.0008 | -0.0002 | 0.0008 | 0.0011 | 0.0277 | 0.0013 | -0.0076 | -0.0002 | 0.0054 | 0.1339 | -0.0002 | 0.0024 | -0.0058 | -0.0035 | 0.1194 | 0.0148 | 0.0035 | 0 | 0.0002 | 0.0003 | 0.0013
08162024 MFB - Analyticsal Duplicate | Sample | 2024-08-26 16:56 | 0.098 | 4.7714 | 2.4724 | 0.4945 | 2.4449 | 0.4804 | 4.965 | 1.0023 | 5.021 | 2.506 | 2.3818 | 4.9418 | 4.9284 | 0.0002 | 5.0728 | 1 | 1.0013 | 5.0693 | 2.4681 | 1.0578 | 2.5369 | 2.4574 | 2.6724 | 2.419 | 0.9814 | 1.0046 | 1.0016 | 2.5237 | 0.9774
R% |  |  | 97.9% | 95.39% | 98.71% | 99.48% | 97.79% | 96.1% | 99.59% | 100.2% | 100.4% | 100.2% | 95.24% | 98.81% | 98.01% |  | 101.6% | 100% | 99.59% | 98.71% | 98.73% | 105.8% | 101.7% | 98.44% | 102.1% | 96.17% | 98.14% | 100.5% | 100.1% | 100.9% | 97.61%
08162024 MB | Sample | 2024-08-26 16:51 | 0.0001 | 0.0018 | 0.0046 | -0.0029 | 0.0001 | -0.0001 | -0.0144 | 0 | 0.0008 | -0.0002 | 0.0008 | 0.0011 | 0.0277 | 0.0013 | -0.0076 | -0.0002 | 0.0054 | 0.1339 | -0.0002 | 0.0024 | -0.0058 | -0.0035 | 0.1194 | 0.0148 | 0.0035 | 0 | 0.0002 | 0.0003 | 0.0013
5x 08162024 MFB | Sample | 2024-08-26 16:59 | 0.0194 | 0.9528 | 0.5033 | 0.0973 | 0.4873 | 0.0968 | 1.1014 | 0.1996 | 1.0015 | 0.5023 | 0.4701 | 0.991 | 0.9784 | 0.0077 | 0.9977 | 0.1998 | 0.1983 | 1.0942 | 0.4886 | 0.2111 | 0.5039 | 0.4887 | 0.5494 | 0.5 | 0.1922 | 0.2005 | 0.1973 | 0.4992 | 0.1944
R% |  |  | 96.5% | 95.1% | 99.74% | 100.2% | 97.44% | 96.9% | 111.6% | 99.8% | 100.1% | 100.5% | 93.86% | 98.99% | 95.07% |  | 100.5% | 100% | 96.45% | 96.03% | 97.76% | 104.3% | 101.9% | 98.44% | 86% | 97.04% | 94.35% | 100.2% | 98.55% | 99.78% | 96.55%
5x 1% NaHCO3 NMC Battery-TM | Sample | 2024-08-26 17:01 | 0.0002 | 0.0804 | 0.0086 | -0.0006 | 0.001 | -0.0002 | 0.1845 | 0 | 0.0003 | 0.0001 | 0.0016 | 0.0196 | 0.0928 | -0.0064 | 0.0112 | 0.0001 | 0.0119 | 466.6491 | 1.0209 | 0.4067 | -0.0033 | 0.045 | 0.1394 | 0.0218 | 0.0262 | 0.0004 | 0.0006 | 0.0003 | 0.0019
5x 1% NaHCO3 NMC Battery-TM - MS | Sample | 2024-08-26 17:04 | 0.0199 | 0.9541 | 0.5124 | 0.1006 | 0.4883 | 0.096 | 1.1555 | 0.1992 | 1.0048 | 0.4966 | 0.5107 | 0.9758 | 1.0543 | -0.006 | 0.971 | 0.1921 | 0.1878 | 416.5691 | 0.7397 | 0.2539 | 0.4934 | 0.567 | 0.5127 | 0.5196 | 0.2116 | 0.1977 | 0.2015 | 0.5109 | 0.1978
5x 1% NaHCO3 NMC Battery-TM -MS -Analytical Duplicate | Sample | 2024-08-26 17:06 | 0.0187 | 0.6318 | 0.4854 | 0.0955 | 0.3204 | 0.0633 | 1.098 | 0.1886 | 0.9502 | 0.47 | 0.4832 | 0.6476 | 0.7085 | -0.0178 | 0.9119 | 0.1276 | 0.1825 | 278.896 | 0.6987 | 0.2401 | 0.4649 | 0.535 | 0.5213 | 0.4911 | 0.1989 | 0.13 | 0.1906 | 0.4833 | 0.1865
RPD |  |  | 6.218% | 40.65% | 5.412% | 5.201% | 41.52% | 41.05% | 5.103% | 5.467% | 5.586% | 5.504% | 5.534% | 40.43% | 39.23% | 1.626% | 6.278% | 40.35% | 2.863% | 39.59% | 5.701% | 5.587% | 5.948% | 5.808% | 1.663% | 5.64% | 6.188% | 41.32% | 5.56% | 5.552% | 5.881%
5x 1% NaHCO3 NMC Battery-TM | Sample | 2024-08-26 17:01 | 0.0002 | 0.0804 | 0.0086 | -0.0006 | 0.001 | -0.0002 | 0.1845 | 0 | 0.0003 | 0.0001 | 0.0016 | 0.0196 | 0.0928 | -0.0064 | 0.0112 | 0.0001 | 0.0119 | 466.6491 | 1.0209 | 0.4067 | -0.0033 | 0.045 | 0.1394 | 0.0218 | 0.0262 | 0.0004 | 0.0006 | 0.0003 | 0.0019
5x 1% NaHCO3 NMC Battery-TM - MS | Sample | 2024-08-26 17:04 | 0.0199 | 0.9541 | 0.5124 | 0.1006 | 0.4883 | 0.096 | 1.1555 | 0.1992 | 1.0048 | 0.4966 | 0.5107 | 0.9758 | 1.0543 | -0.006 | 0.971 | 0.1921 | 0.1878 | 416.5691 | 0.7397 | 0.2539 | 0.4934 | 0.567 | 0.5127 | 0.5196 | 0.2116 | 0.1977 | 0.2015 | 0.5109 | 0.1978
R% |  |  | 98.5% | 87.37% | 100.8% | 101.2% | 97.46% | 96.2% | 97.1% | 99.6% | 100.4% | 99.3% | 101.8% | 95.62% | 96.15% |  | 95.98% | 96% | 87.95% | -5008% | -56.24% | -76.4% | 99.34% | 104.4% | 74.66% | 99.56% | 92.7% | 98.65% | 100.4% | 102.1% | 97.95%
5x 1% NaHCO3 NMC Battery-TM | Sample | 2024-08-26 17:01 | 0.0002 | 0.0804 | 0.0086 | -0.0006 | 0.001 | -0.0002 | 0.1845 | 0 | 0.0003 | 0.0001 | 0.0016 | 0.0196 | 0.0928 | -0.0064 | 0.0112 | 0.0001 | 0.0119 | 466.6491 | 1.0209 | 0.4067 | -0.0033 | 0.045 | 0.1394 | 0.0218 | 0.0262 | 0.0004 | 0.0006 | 0.0003 | 0.0019
5x 1% NaHCO3 NMC Battery-TM -MS -Analytical Duplicate | Sample | 2024-08-26 17:06 | 0.0187 | 0.6318 | 0.4854 | 0.0955 | 0.3204 | 0.0633 | 1.098 | 0.1886 | 0.9502 | 0.47 | 0.4832 | 0.6476 | 0.7085 | -0.0178 | 0.9119 | 0.1276 | 0.1825 | 278.896 | 0.6987 | 0.2401 | 0.4649 | 0.535 | 0.5213 | 0.4911 | 0.1989 | 0.13 | 0.1906 | 0.4833 | 0.1865
R% |  |  | 92.5% | 55.14% | 95.36% | 96.1% | 63.88% | 63.5% | 91.35% | 94.3% | 94.99% | 93.98% | 96.32% | 62.8% | 61.57% |  | 90.07% | 63.75% | 85.3% | -18775% | -64.44% | -83.3% | 93.64% | 98% | 76.38% | 93.86% | 86.35% | 64.8% | 95% | 96.6% | 92.3%
1% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:14 | -0.0002 | 0.4131 | 0.0041 | 0.0076 | 0.004 | -0.0001 | 0.5031 | 0 | 0.0012 | 0.0004 | 0.0027 | 0.099 | 0.5255 | 0.0028 | 0.0744 | 0.0011 | 0.004 | 1752.4356 | 5.0444 | 2.0521 | -0.0019 | 0.2448 | 0.1071 | 0.015 | 0.1281 | 0.0016 | 0.0016 | 0.0005 | 0.0057
1% NaHCO3 NMC Battery-TM - Lab Duplicate | Sample | 2024-08-26 17:09 | -0.0004 | 0.4206 | 0.0072 | 0.0051 | 0.004 | -0.0001 | 0.4985 | 0 | 0.0015 | 0.0005 | 0.0024 | 0.0996 | 0.5282 | 0.0117 | 0.0729 | 0.0012 | 0.0112 | 1766.2352 | 5.0468 | 2.0369 | -0.0064 | 0.2381 | 0.138 | 0.019 | 0.1297 | 0.0016 | 0.0014 | 0.0007 | 0.0059
RPD |  |  | 66.67% | 1.799% | 54.87% | 39.37% | 0% | 0% | 0.9185% | #DIV/0! | 22.22% | 22.22% | 11.76% | 0.6042% | 0.5125% | 1.626% | 2.037% | 8.696% | 94.74% | 0.7844% | 0.04757% | 0.7435% | 108.4% | 2.775% | 25.21% | 23.53% | 1.241% | 0% | 13.33% | 33.33% | 3.448%
1% NaHCO3- NMC Battery-TM - MS | Sample | 2024-08-26 17:11 | -0.0001 | 0.1121 | 0.012 | 0.0053 | 0.0465 | -0.0002 | 0.541 | -0.0001 | 0 | 0.0003 | 0.0034 | 0.0679 | 0.6131 | 0.0019 | 0.0757 | 0.0002 | 0.0044 | 1760.8728 | 1.3672 | 0.1803 | -0.0042 | 0.3097 | 0.1151 | 0.0178 | 0.072 | 0.002 | 0.0009 | 0.0006 | 0.0051
1% NaHCO3- NMC Battery-TM - MS - Analytical Duplicate | Sample | 2024-08-26 17:14 | -0.0005 | 0.0097 | 0.0037 | -0.003 | 0.0049 | -0.0003 | 0.1786 | -0.0001 | -0.0004 | -0.0007 | -0.0042 | 0.0065 | 0.0788 | -0.0281 | 0.0129 | -0.0003 | 0.0024 | 249.0177 | 0.4664 | 0.0607 | -0.0134 | 0.0974 | 0.1198 | 0.0093 | 0.0214 | 0.0002 | 0.0003 | 0.0004 | 0.0017
Rinse1 | QC | 2024-08-26 17:16 | -0.0001 | -0.0029 | 0.0005 | -0.0033 | 0 | -0.0001 | -0.0089 | -0.0001 | 0.0002 | -0.0002 | 0.0001 | 0.0005 | 0.0267 | -0.006 | -0.0054 | -0.0004 | 0.0022 | 0.1691 | -0.0013 | -0.0006 | -0.0061 | -0.0029 | 0.1095 | 0.0094 | 0.0017 | 0 | 0 | 0.0003 | 0.0011
Rinse1 | QC | 2024-08-26 17:19 | -0.0004 | -0.0018 | 0.0006 | -0.0032 | 0.0001 | -0.0002 | -0.0074 | 0 | 0.0002 | -0.0002 | -0.0007 | -0.0006 | 0.0355 | -0.0073 | -0.0053 | 0 | 0.0016 | 0.1128 | -0.0014 | 0.005 | -0.0064 | -0.0098 | 0.1175 | 0.0042 | 0.0022 | 0 | -0.0002 | -0.0001 | 0.001
Rinse | QC | 2024-08-26 17:21 | -0.0001 | -0.0019 | 0.001 | -0.0038 | 0.0001 | -0.0001 | -0.0132 | -0.0002 | 0.0003 | -0.0003 | 0.0003 | -0.0003 | 0.0406 | -0.0082 | -0.0057 | -0.0003 | 0.001 | 0.0934 | -0.0014 | 0.0004 | -0.0047 | -0.0077 | 0.121 | 0.0032 | 0.0017 | 0 | -0.0002 | 0.0004 | 0.0012
0.5ppm All (LLQC) | QC | 2024-08-26 17:24 | 0.491 | 0.4834 | 0.4767 | 0.4824 | 0.4808 | 0.4846 | 0.4909 | 0.49 | 0.4929 | 0.4952 | 0.4634 | 0.4955 | 0.4607 | 0.4929 | 0.4759 | 0.4989 | 0.4481 | 0.5765 | 0.4917 | 0.5164 | 0.4977 | 0.4786 | 0.4485 | 0.4812 | 0.4633 | 0.4968 | 0.4869 | 0.4868 | 0.4826
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-26 17:26 | 0.4851 | 0.4867 | 0.4785 | 0.4803 | 0.481 | 0.4871 | 0.4897 | 0.4897 | 0.4909 | 0.4944 | 0.463 | 0.4933 | 0.492 | 0.4892 | 0.4744 | 0.4982 | 0.4702 | 0.5705 | 0.4914 | 0.5191 | 0.4974 | 0.4825 | 0.5184 | 0.4939 | 0.4618 | 0.4976 | 0.4856 | 0.4858 | 0.4813
RPD |  |  | 1.209% | 0.6803% | 0.3769% | 0.4363% | 0.04159% | 0.5146% | 0.2447% | 0.06124% | 0.4066% | 0.1617% | 0.08636% | 0.445% | 6.571% | 1.626% | 0.3157% | 0.1404% | 4.813% | 1.046% | 0.06103% | 0.5215% | 0.0603% | 0.8116% | 14.46% | 2.605% | 0.3243% | 0.1609% | 0.2674% | 0.2056% | 0.2697%
0.5ppm All (LLQC) | QC | 2024-08-26 17:24 | 0.491 | 0.4834 | 0.4767 | 0.4824 | 0.4808 | 0.4846 | 0.4909 | 0.49 | 0.4929 | 0.4952 | 0.4634 | 0.4955 | 0.4607 | 0.4929 | 0.4759 | 0.4989 | 0.4481 | 0.5765 | 0.4917 | 0.5164 | 0.4977 | 0.4786 | 0.4485 | 0.4812 | 0.4633 | 0.4968 | 0.4869 | 0.4868 | 0.4826
% Recovery |  |  | 98.2% | 96.68% | 95.34% | 96.48% | 96.16% | 96.92% | 98.18% | 98% | 98.58% | 99.04% | 92.68% | 99.1% | 92.14% | 98.58% | 95.18% | 99.78% | 89.62% | 115.3% | 98.34% | 103.3% | 99.54% | 95.72% | 89.7% | 96.24% | 92.66% | 99.36% | 97.38% | 97.36% | 96.52%
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-26 17:26 | 0.4851 | 0.4867 | 0.4785 | 0.4803 | 0.481 | 0.4871 | 0.4897 | 0.4897 | 0.4909 | 0.4944 | 0.463 | 0.4933 | 0.492 | 0.4892 | 0.4744 | 0.4982 | 0.4702 | 0.5705 | 0.4914 | 0.5191 | 0.4974 | 0.4825 | 0.5184 | 0.4939 | 0.4618 | 0.4976 | 0.4856 | 0.4858 | 0.4813
% Recovery |  |  | 97.02% | 97.34% | 95.7% | 96.06% | 96.2% | 97.42% | 97.94% | 97.94% | 98.18% | 98.88% | 92.6% | 98.66% | 98.4% | 97.84% | 94.88% | 99.64% | 94.04% | 114.1% | 98.28% | 103.8% | 99.48% | 96.5% | 103.7% | 98.78% | 92.36% | 99.52% | 97.12% | 97.16% | 96.26%
CCS1 (SS 5ppm) | QC | 2024-08-26 17:29 | 5.0454 | 4.9491 | 5.0689 | 4.9867 | 4.9221 | 4.9418 | 4.9939 | 4.9892 | 4.9974 | 4.9959 | 4.737 | 4.9595 | 4.89 | 4.8573 | 5.0343 | 4.9756 | 5.1416 | 4.9607 | 4.9696 | 5.3477 | 5.102 | 5.004 | 5.3579 | 5.0291 | 4.7235 | 5.0348 | 5.0173 | 5.028 | 4.9308
% Recovery |  |  | 100.9% | 98.98% | 101.4% | 99.73% | 98.44% | 98.84% | 99.88% | 99.78% | 99.95% | 99.92% | 94.74% | 99.19% | 97.8% | 97.15% | 100.7% | 99.51% | 102.8% | 99.21% | 99.39% | 107% | 102% | 100.1% | 107.2% | 100.6% | 94.47% | 100.7% | 100.3% | 100.6% | 98.62%
Rinse | QC | 2024-08-26 17:31 | 0.0001 | 0.001 | 0.0157 | -0.0024 | 0.0001 | -0.0001 | -0.0156 | 0 | 0.0005 | 0 | 0.0002 | 0.0005 | 0.0216 | 0.0019 | -0.0061 | -0.0002 | 0.0461 | 0.0694 | -0.0003 | 0.0017 | -0.0073 | -0.0076 | 0.1235 | 0.0254 | 0.0033 | 0.0001 | 0.0012 | -0.0002 | 0.0006
CCB | QC | 2024-08-26 17:34 | 0 | -0.0039 | 0.0074 | -0.0017 | 0 | -0.0001 | -0.0154 | -0.0001 | 0.0002 | -0.0004 | 0.0008 | -0.0002 | 0.0216 | -0.0077 | -0.0051 | -0.0002 | 0.015 | 0.0644 | 0.0002 | 0.0055 | -0.0043 | -0.0029 | 0.1151 | 0.0124 | 0.0022 | 0 | 0.0002 | 0.0004 | 0.0004
5% HNO3 | Sample | 2024-08-26 17:38 | 0.1026 | 4.8192 | 2.6053 | 0.5215 | 2.4429 | 0.4671 | 5.3022 | 0.9867 | 4.9342 | 2.4656 | 2.6398 | 4.7432 | 5.6634 | 0.0008 | 4.8247 | 0.9449 | 0.9928 | 1641.1815 | 3.6549 | 1.2738 | 2.4629 | 2.9375 | 2.5508 | 2.59 | 1.0941 | 0.9884 | 1.0167 | 2.5702 | 0.9743
MQ-Blank-03 | Sample | 2024-08-26 17:41 | 0 | -0.004 | 0.0015 | -0.0036 | 0.0001 | 0 | -0.0106 | 0.0003 | 0.0003 | 0 | 0.0004 | 0.0003 | 0.0488 | 0.0025 | -0.0065 | -0.0001 | 0.0044 | 0.1077 | -0.0005 | 0.0047 | -0.005 | -0.0038 | 0.1133 | 0.0135 | 0.003 | 0.0001 | 0.0001 | 0.0004 | 0.0014
Solution Label | Type | Date Time | Ag | Al | As | B | Ba | Be | Ca | Cd | Co | Cr | Cu | Fe | K | Li | Mg | Mn | Mo | Na | Ni | P | Pb | S | Sb | Se | Si | Sr | Ti | V | Zn
Lowest Cal Std |  |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Highest Cal Std |  |  | 20 | 200 | 20 | 20 | 50 | 20 | 2000 | 20 | 20 | 50 | 100 | 1000 | 100 | 100 | 2000 | 50 | 20 | 100 | 50 | 200 | 200 | 100 | 50 | 20 | 100 | 20 | 50 | 20 | 100
Correlation |  |  | 0.99998 | 0.99993 | 0.99997 | 1 | 0.99937 | 0.99996 | 0.99999 | 1 | 1 | 1 | 0.99996 | 0.99998 | 1 | 0.99999 | 0.99996 | 0.99992 | 0.99986 | 0.99999 | 1 | 0.99998 | 0.99978 | 0.99998 | 0.99936 | 0.99996 | 0.99927 | 0.99996 | 0.99999 | 0.99999 | 1
ICP-OES MDL |  |  | 0.0045598 | 0.0048715 | 0.0060725 | 0.0016216 | 0.00011879 | 0.000168 | 0.0038865 | 0.00024729 | 0.00024729 | 0.0007 | 0.00076374 | 0.0031683 | 0.032873 | 0.0015923 | 0.0080477 | 0.0008057 | 0.00050865 | 0.001351 | 0.0024937 | 0.0072662 | 0.0049325 | 0.0058515 | 0.0056274 | 0.0064431 | 0.002358 | 0.000168 | 0.00011879 | 0.00093287 | 0.00059397
ICP-OES MRL/PRDL |  |  | 0.045598 | 0.048715 | 0.060725 | 0.016216 | 0.0011879 | 0.00168 | 0.038865 | 0.0024729 | 0.0024729 | 0.0074 | 0.0076374 | 0.031683 | 0.32873 | 0.015923 | 0.080477 | 0.008057 | 0.0050865 | 0.01351 | 0.024937 | 0.072662 | 0.049325 | 0.058515 | 0.056274 | 0.064431 | 0.02358 | 0.00168 | 0.0011879 | 0.0093287 | 0.0059397
Where |  |  | Ba: QC Recoveries
| Calibration Standards |  | Ba: 100%; Be: 90<R%>110
| IEC Check Standards |  | Ba: 80%; Be: 70<R%>80
| CCV & Secondary source QC check standards |  | Ba: 70%; Be: 70%>R%>120%
| Low Level QC chek standards |  | Ba: 60%; Be: 65%>R%>135%
| Calculated/Estimated MDLs & MRLs |  | Ba: 100%; Be: 65<R%>135
| Dilution required |  | Ba: 20%; Be: RPD<20%
* | Sampling Error |  | Ba: 21%; Be: RPD>20%
** | Different concentration than the normal |  | Ba: 20; Be: RSD<20%
| Mislabeled |  | Ba: 21; Be: RSD>20%
[7 empty] | 20 | STDEV<20

## 08262024ICP-OES_TO178-455RD
Worksheet exported from: C:\Users\USEPA\Documents\Agilent\ICP Expert\My Results\08262024_ICP-OES_TO 178-453 454 & 455.esws
Solution Label | Type | Date Time | LIMS ID | Customer | Customer Ref. | Sample Date Time | Sample Site | Description | Ag (328.068 nm) ppm | Al (396.152 nm) ppm | As (188.980 nm) ppm | B (249.678 nm) ppm | Ba (455.403 nm) ppm | Be (313.107 nm) ppm | Ca (370.602 nm) ppm | Cd (214.439 nm) ppm | Co (238.892 nm) ppm | Cr (267.716 nm) ppm | Cu (324.754 nm) ppm | Fe (259.940 nm) ppm | K (766.491 nm) ppm | Li (670.783 nm) ppm | Mg (279.078 nm) ppm | Mn (257.610 nm) ppm | Mo (202.032 nm) ppm | Na (588.995 nm) ppm | Ni (231.604 nm) ppm | P (178.222 nm) ppm | Pb (220.353 nm) ppm | S (181.972 nm) ppm | Sb (217.582 nm) ppm | Se (196.026 nm) ppm | Si (251.611 nm) ppm | Sn (189.925 nm) ppm | Sr (421.552 nm) ppm | Ti (334.941 nm) ppm | V (292.401 nm) ppm | Zn (206.200 nm) ppm | Y (371.029 nm) Ratio | Y (371.029 nm) Ratio 1 | Y (488.368 nm) Ratio | Ag 328.068 C ppm | Al 396.152 C ppm | As 188.980 C ppm | B 249.678 C ppm | Ba 455.403 C ppm | Be 313.107 C ppm | Ca 370.602 C ppm | Cd 214.439 C ppm | Co 238.892 C ppm | Cr 267.716 C ppm | Cu 324.754 C ppm | Fe 259.940 C ppm | K 766.491 C ppm | Li 670.783 C ppm | Mg 279.078 C ppm | Mn 257.610 C ppm | Mo 202.032 C ppm | Na 588.995 C ppm | Ni 231.604 C ppm | P 178.222 C ppm | Pb 220.353 C ppm | S 181.972 C ppm | Sb 217.582 C ppm | Se 196.026 C ppm | Si 251.611 C ppm | Sr 421.552 C ppm | Ti 334.941 C ppm | V 292.401 C ppm | Zn 206.200 C ppm
Calibration Blank | Blank | 2024-08-26 10:07 | [6 empty] | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 0.0000 ! | 1.00 ! | 1.00 ! | 1.00 !
Calibration Standard 0.5ppm | Standard | 2024-08-26 10:10 | [6 empty] | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 0.5000 ! | 1.01 ! | 1.00 ! | 1.00 !
Calibration Standard 1ppm | Standard | 2024-08-26 10:12 | [6 empty] | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.0000 ! | 1.01 ! | 1.00 ! | 0.99 !
Calibration Standard 5ppm | Standard | 2024-08-26 10:15 | [6 empty] | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 5.0000 ! | 0.99 ! | 0.99 ! | 0.99 !
Calibration Standard 10ppm | Standard | 2024-08-26 10:17 | [6 empty] | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 10.0000 ! | 0.98 ! | 0.98 ! | 0.98 !
Calibration Standard 20ppm | Standard | 2024-08-26 10:20 | [6 empty] | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 20.0000 ! | 0.96 ! | 0.98 ! | 0.97 !
Calibration Standard 50ppm | Standard | 2024-08-26 10:22 | [7 empty] | 50.0000 ! |  |  | 50.0000 ! |  | 50.0000 ! |  |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! | 50.0000 ! |  | 50.0000 ! |  |  | 50.0000 ! |  | 50.0000 ! | 0.94 ! | 0.95 ! | 0.95 !
Calibration Standard-A 100ppm | Standard | 2024-08-26 10:25 | [7 empty] | 100.0000 ! |  |  |  |  | 100.0000 ! |  |  |  | 100.0000 ! | 100.0000 ! | 100.0000 ! | 100.0000 ! | 100.0000 ! |  |  | 100.0000 ! |  |  | 100.0000 ! |  |  |  | 100.0000 ! |  |  |  |  | 100.0000 ! | 0.92 ! | 0.94 ! | 0.93 !
Calibration Standard-B 100ppm | Standard | 2024-08-26 10:27 | [25 empty] | 100.0000 ! |  | 100.0000 ! | [8 empty] | 1.01 ! | 1.00 ! | 1.00 !
Calibration Standard Al+Fe+Ca+Mg200ppm | Standard | 2024-08-26 10:30 | [7 empty] | 200.0000 ! |  |  |  |  | 200.0000 ! |  |  |  |  | 200.0000 ! |  |  | 200.0000 ! | [15 empty] | 0.94 ! | 0.96 ! | 0.96 !
Calibration Standard Ca+Mg500ppm | Standard | 2024-08-26 10:32 | [12 empty] | 500.0000 ! | [7 empty] | 500.0000 ! | [15 empty] | 0.94 ! | 0.92 ! | 0.92 !
Calibration Standard Ca+Mg1000ppm | Standard | 2024-08-26 10:35 | [12 empty] | 1000.0000 ! | [7 empty] | 1000.0000 ! | [15 empty] | 0.91 ! | 0.89 ! | 0.88 !
Calibration Standard Ca+Mg2000ppm | Standard | 2024-08-26 10:37 | [12 empty] | 2000.0000 ! | [7 empty] | 2000.0000 ! | [15 empty] | 0.87 ! | 0.84 ! | 0.84 !
Calibration Standard P+S200ppm | Standard | 2024-08-26 10:40 | [25 empty] | 200.0000 ! |  | 200.0000 ! | [8 empty] | 1.02 ! | 1.00 ! | 1.00 !
Calibration Standard Pb200ppm | Standard | 2024-08-26 10:42 | [26 empty] | 200.0000 ! | [9 empty] | 1.01 ! | 1.01 ! | 1.00 !
Calibration Standard Fe 1000ppm | Standard | 2024-08-26 10:45 | [17 empty] | 1000.0000 ! | [18 empty] | 0.98 ! | 0.97 ! | 0.98 !
Rinse1 | QC | 2024-08-26 10:47 | [6 empty] | 0.0002 | -0.0022 | 0.0031 | 0.0082 | 0 | -0.0001 | -0.005 | 0 | -0.0003 | -0.0001 | 0.0006 | 0.0191 | 0.023 | 0.0114 | 0.0028 | -0.0003 | 0.0034 | -0.0145 | -0.0006 | 0.0045 | -0.0036 | -0.0037 | 0.1102 | 0.0082 | 0.0046 | -0.0016 | 0 | -0.0001 | 0.0002 | 0.0006 | 1.01 | 0.99 | 0.99 | 0.0002 | -0.0022 | 0.0031 | 0.0082 | 0 | -0.0001 | -0.005 | 0 | -0.0003 | -0.0001 | 0.0006 | 0.0191 | 0.023 | 0.0114 | 0.0028 | -0.0003 | 0.0034 | -0.0145 | -0.0006 | 0.0045 | -0.0036 | -0.0037 | 0.1102 | 0.0082 | 0.0046 | 0 | -0.0001 | 0.0002 | 0.0006
Rinse | QC | 2024-08-26 10:50 | [6 empty] | 0.0001 | -0.0034 | 0.003 | 0.0056 | 0.0001 | -0.0002 | -0.0094 | 0 | 0.0007 | -0.0001 | 0.0003 | 0.0041 | 0.0266 | -0.0025 | -0.0014 | -0.0003 | 0.0028 | -0.0146 | -0.0005 | 0.0047 | -0.0062 | -0.0017 | 0.1072 | 0.0112 | 0.0033 | -0.0009 | 0 | -0.0001 | -0.0001 | 0 | 1.02 | 0.99 | 0.99 | 0.0001 | -0.0034 | 0.003 | 0.0056 | 0.0001 | -0.0002 | -0.0094 | 0 | 0.0007 | -0.0001 | 0.0003 | 0.0041 | 0.0266 | -0.0025 | -0.0014 | -0.0003 | 0.0028 | -0.0146 | -0.0005 | 0.0047 | -0.0062 | -0.0017 | 0.1072 | 0.0112 | 0.0033 | 0 | -0.0001 | -0.0001 | 0
IECS1 (Al 1000 Int.) | QC | 2024-08-26 10:52 | [6 empty] | 0.0006 | 955.6486 o | 0.0017 | 0.0091 | 0.0093 | 0.0001 | 0.0137 | 0.0007 | -0.0002 | 0.0009 | 0.0036 | 0.0468 | 0.0478 | 0.0014 | 0.0187 | 0 | 0.0068 | -0.002 | 0.0002 | 0.0129 | 0.0357 | 0.0091 | 0.1152 | 0.0162 | 0.0212 | -0.0018 | 0.0004 | 0.0013 | 0.0003 | 0.0083 | 0.93 | 0.99 | 0.99 | 0.0006 | 955.6486 | 0.0017 | 0.0091 | 0.0093 | 0.0001 | 0.0137 | 0.0007 | -0.0002 | 0.0009 | 0.0036 | 0.0468 | 0.0478 | 0.0014 | 0.0187 | 0 | 0.0068 | -0.002 | 0.0002 | 0.0129 | 0.0357 | 0.0091 | 0.1153 | 0.0067 | 0.0212 | 0.0004 | 0.0013 | 0.0003 | 0.0083
Rinse1 | QC | 2024-08-26 10:55 | [6 empty] | 0 | 0.0111 | 0 | 0.0034 | 0 | 0 | -0.0059 | -0.0001 | 0.0004 | -0.0002 | 0.0002 | 0.0023 | 0.038 | 0.0057 | -0.0023 | -0.0002 | 0.0019 | -0.0113 | -0.0005 | 0.0007 | -0.0043 | 0.0013 | 0.1098 | 0.0064 | 0.0015 | -0.002 | 0 | 0 | 0.0008 | 0.0004 | 1.02 | 1 | 0.99 | 0 | 0.0111 | 0 | 0.0034 | 0 | 0 | -0.0059 | -0.0001 | 0.0004 | -0.0002 | 0.0002 | 0.0023 | 0.038 | 0.0057 | -0.0023 | -0.0002 | 0.0019 | -0.0113 | -0.0005 | 0.0007 | -0.0043 | 0.0013 | 0.1098 | 0.0064 | 0.0015 | 0 | 0 | 0.0008 | 0.0004
Rinse | QC | 2024-08-26 10:57 | [6 empty] | 0.0003 | 0.0046 | 0.0026 | 0.0033 | 0 | -0.0001 | -0.0148 | 0.0001 | 0.0001 | -0.0004 | 0.0005 | 0.0016 | 0.0273 | 0.0058 | -0.0022 | -0.0002 | 0.0017 | -0.0139 | -0.0015 | 0.0003 | -0.0069 | -0.0029 | 0.1124 | 0.0114 | 0.0031 | -0.0012 | 0 | 0.0001 | 0.0007 | 0.0005 | 1.01 | 1 | 0.99 | 0.0003 | 0.0046 | 0.0026 | 0.0033 | 0 | -0.0001 | -0.0148 | 0.0001 | 0.0001 | -0.0004 | 0.0005 | 0.0016 | 0.0273 | 0.0058 | -0.0022 | -0.0002 | 0.0017 | -0.0139 | -0.0015 | 0.0003 | -0.0069 | -0.0029 | 0.1124 | 0.0114 | 0.0031 | 0 | 0.0001 | 0.0007 | 0.0005
IECS2 (Fe 1000 Int.) | QC | 2024-08-26 11:00 | [6 empty] | 0.0024 | 0.0635 | 0.0035 | -0.0111 | 0.0002 | 0.0001 | -0.0199 | 0.0121 | 0.027 | 0.0089 | -0.0099 | 1045.4062 | 0.0391 | -0.0005 | 1.0556 | 0.0097 | 0.0026 | -0.0123 | 0.0078 | 0.0148 | 0.0117 | 0.0025 | 0.104 | -0.0037 | 0.0046 | 0.0038 | 0.0009 | 0.0002 | 0.0062 | 0.0049 | 0.98 | 0.98 | 0.98 | 0.0024 | 0.0635 | 0.0035 | -0.0111 | 0.0002 | 0.0001 | -0.0199 | -0.0036 | 0.0114 | 0.0089 | 0.0016 | 1045.4062 | 0.0391 | -0.0005 | 0.8779 K | 0.0097 | 0.0026 | -0.0123 | 0.0078 | -0.0009 | -0.0071 | 0.0025 | 0.1144 | -0.0225 | 0.0046 | 0.0009 | 0.0002 | -0.0001 | 0.0049
Rinse1 | QC | 2024-08-26 11:02 | [6 empty] | -0.0003 | 0.0048 | 0.0006 | 0.0019 | 0 | -0.0001 | -0.0083 | -0.0001 | 0.0003 | -0.0003 | 0.0012 | 0.019 | 0.0344 | -0.003 | -0.0061 | -0.0001 | 0.0023 | -0.0139 | -0.0008 | -0.0017 | -0.0055 | -0.0025 | 0.1089 | 0.0139 | 0.0022 | -0.0012 | 0 | -0.0001 | -0.0001 | 0.0006 | 1.01 | 1 | 1 | -0.0003 | 0.0048 | 0.0006 | 0.0019 | 0 | -0.0001 | -0.0083 | -0.0001 | 0.0003 | -0.0003 | 0.0012 | 0.019 | 0.0344 | -0.003 | -0.0061 | -0.0001 | 0.0023 | -0.0139 | -0.0008 | -0.0017 | -0.0055 | -0.0025 | 0.1089 | 0.0139 | 0.0022 | 0 | -0.0001 | -0.0001 | 0.0006
Rinse | QC | 2024-08-26 11:05 | [6 empty] | 0.0001 | -0.0016 | 0.0037 | 0.0009 | 0.0001 | -0.0002 | -0.0125 | 0 | 0.0004 | -0.0003 | 0.0001 | 0.0058 | 0.0394 | 0.0105 | -0.0013 | -0.0002 | 0.0007 | -0.0145 | -0.001 | 0.0037 | -0.0092 | -0.0029 | 0.1087 | 0.007 | 0.0001 | -0.0009 | 0 | 0 | 0.0007 | 0.0002 | 1.01 | 1 | 1 | 0.0001 | -0.0016 | 0.0037 | 0.0009 | 0.0001 | -0.0002 | -0.0125 | 0 | 0.0004 | -0.0003 | 0.0001 | 0.0058 | 0.0394 | 0.0105 | -0.0013 | -0.0002 | 0.0007 | -0.0145 | -0.001 | 0.0037 | -0.0092 | -0.0029 | 0.1087 | 0.007 | 0.0001 | 0 | 0 | 0.0007 | 0.0002
IECS3 (Mn 100 Int.) | QC | 2024-08-26 11:08 | [6 empty] | 0.0094 | 0.0194 | 0.0028 | -0.0033 | 0.0001 | -0.0001 | 0.7542 | 0.0002 | 0.0033 | 0.0177 | 0.0004 | 0.0361 | 0.0279 | 0.0059 | 0.0069 | 99.3876 o | 0.0016 | -0.0106 | -0.0023 | 0.0035 | 0.009 | 0.5193 | 0.1287 | 0.0919 | 0.0567 | -0.0016 | 0.0002 | 0.0003 | 0.0004 | 0.0026 | 1.01 | 1 | 0.99 | -0.0006 | 0.0194 | 0.0028 | -0.0033 | 0.0001 | -0.0001 | 0.7542 K | 0.0002 | -0.0017 | -0.0102 | 0.0004 | 0.0361 | 0.0279 | 0.0059 | 0.0069 | 99.3876 | 0.0016 | -0.0106 | -0.0023 | 0.0035 | -0.004 | -0.0174 | 0.1287 | 0.0352 | 0.0189 | 0.0002 | 0.0003 | 0.0004 | 0.0026
IECS4 (Ca Mg 1000 Int.) | QC | 2024-08-26 11:10 | [6 empty] | 0.0005 | 0.0519 | 0.0033 | 0.0053 | 0.0022 | -0.0001 | 1010.9843 | 0.0086 | 0.0002 | 0.0016 | 0.003 | 0.0199 | 0.0303 | -0.0012 | 1026.6746 | 0.0267 | 0.0016 | 0.0038 | 0.0036 | 0.0085 | 0.0015 | 0.0432 | 0.1095 | 0.009 | 0.0275 | -0.0011 | 0.008 | -0.0003 | 0.0015 | 0.0141 | 0.91 | 0.9 | 0.9 | 0.0005 | 0.0519 | 0.0033 | 0.0053 | 0.0022 | -0.0001 | 1010.9843 | 0.0086 | 0.0002 | 0.0016 | 0.003 | 0.0199 | 0.0303 | -0.0012 | 1026.6746 | 0.0267 | 0.0016 | 0.0038 | 0.0036 | 0.0085 | 0.0015 | 0.043 | 0.1095 | 0.009 | 0.0274 | 0.008 | -0.0003 | 0.0015 | 0.0141
Rinse1 | QC | 2024-08-26 11:13 | [6 empty] | 0 | -0.0029 | 0.0007 | 0 | 0.0001 | -0.0002 | 0.0005 | 0 | 0.0003 | -0.0002 | 0 | 0.002 | 0.0371 | -0.0024 | 0.0033 | 0.0005 | 0.0013 | -0.0081 | 0.0003 | 0.0005 | -0.0075 | -0.0018 | 0.108 | 0.0092 | 0.0021 | -0.001 | 0 | -0.0001 | 0.0001 | 0.0004 | 1.01 | 1 | 0.99 | 0 | -0.0029 | 0.0007 | 0 | 0.0001 | -0.0002 | 0.0005 | 0 | 0.0003 | -0.0002 | 0 | 0.002 | 0.0371 | -0.0024 | 0.0033 | 0.0005 | 0.0013 | -0.0081 | 0.0003 | 0.0005 | -0.0075 | -0.0018 | 0.108 | 0.0092 | 0.0021 | 0 | -0.0001 | 0.0001 | 0.0004
Rinse | QC | 2024-08-26 11:15 | [6 empty] | -0.0001 | -0.0004 | -0.0001 | -0.0003 | 0.0001 | 0 | -0.0078 | 0 | 0.0007 | -0.0002 | 0.0007 | 0.0009 | 0.0383 | -0.0032 | 0.0004 | 0.0001 | 0.001 | -0.0096 | -0.0001 | 0.0033 | -0.0081 | -0.0032 | 0.1103 | 0.0056 | 0.0017 | -0.0011 | 0 | -0.0001 | 0 | 0.0005 | 1.01 | 1 | 1 | -0.0001 | -0.0004 | -0.0001 | -0.0003 | 0.0001 | 0 | -0.0078 | 0 | 0.0007 | -0.0002 | 0.0007 | 0.0009 | 0.0383 | -0.0032 | 0.0004 | 0.0001 | 0.001 | -0.0096 | -0.0001 | 0.0033 | -0.0081 | -0.0032 | 0.1103 | 0.0056 | 0.0017 | 0 | -0.0001 | 0 | 0.0005
IECS5 (Ti 20 Int.) | QC | 2024-08-26 11:18 | [6 empty] | 0.0009 | 0.0088 | 0.0048 | 0.0004 | 0 | -0.0006 | 0.1258 | 0.0001 | 0.0001 | -0.0001 | -0.0003 | 0.0111 | 0.0172 | -0.002 | 0.0247 | 0.001 | 0.0014 | -0.0063 | -0.001 | 0.0028 | -0.011 | -0.0032 | 0.1089 | 0.0042 | 0.0131 | -0.0004 | 0.0001 | 19.8735 | 0.0172 | 0.0015 | 1.03 | 1 | 1 | 0.0009 | 0.0088 | 0.0048 | 0.0004 | 0 | -0.0006 | 0.1258 | 0.0001 | 0.0001 | -0.0001 | -0.0003 | 0.0111 | 0.0172 | -0.002 | 0.0247 | 0.001 | 0.0014 | -0.0063 | -0.001 | 0.0028 | -0.011 | -0.0032 | 0.1089 | 0.0042 | 0.0131 | 0.0001 | 19.8735 K | -0.0007 | 0.0015
Rinse1 | QC | 2024-08-26 11:20 | [6 empty] | 0.0003 | 0.0003 | 0.0006 | -0.0008 | 0.0001 | -0.0002 | -0.0066 | 0 | -0.0001 | -0.0001 | 0.0008 | 0.0003 | 0.033 | -0.0018 | -0.0041 | -0.0002 | 0.0016 | -0.01 | -0.0005 | 0.0019 | -0.0077 | 0.0002 | 0.1111 | 0.0091 | 0.0015 | -0.0015 | 0 | 0.0015 | 0.0001 | 0.0005 | 1.01 | 1 | 1.01 | 0.0003 | 0.0003 | 0.0006 | -0.0008 | 0.0001 | -0.0002 | -0.0066 | 0 | -0.0001 | -0.0001 | 0.0008 | 0.0003 | 0.033 | -0.0018 | -0.0041 | -0.0002 | 0.0016 | -0.01 | -0.0005 | 0.0019 | -0.0077 | 0.0002 | 0.1111 | 0.0091 | 0.0015 | 0 | 0.0015 | 0.0001 | 0.0005
Rinse | QC | 2024-08-26 11:23 | [6 empty] | 0.0001 | -0.003 | 0.0012 | -0.0012 | 0.0001 | -0.0001 | -0.0082 | -0.0001 | 0.0007 | -0.0001 | 0.0006 | -0.0002 | 0.0428 | 0.0055 | -0.0076 | -0.0002 | 0.0007 | -0.0103 | -0.0011 | 0.0033 | -0.0072 | 0.0015 | 0.1063 | 0.0092 | 0.0025 | -0.0005 | 0 | 0.0004 | 0.0003 | 0.0002 | 1.01 | 1 | 0.99 | 0.0001 | -0.003 | 0.0012 | -0.0012 | 0.0001 | -0.0001 | -0.0082 | -0.0001 | 0.0007 | -0.0001 | 0.0006 | -0.0002 | 0.0428 | 0.0055 | -0.0076 | -0.0002 | 0.0007 | -0.0103 | -0.0011 | 0.0033 | -0.0072 | 0.0015 | 0.1063 | 0.0092 | 0.0025 | 0 | 0.0004 | 0.0003 | 0.0002
ICB | QC | 2024-08-26 11:25 | [6 empty] | -0.0001 | -0.0031 | -0.0011 | -0.0014 | 0 | -0.0002 | -0.0143 | -0.0001 | 0.0005 | -0.0001 | 0.0004 | 0.0007 | 0.0312 | -0.0041 | -0.0026 | -0.0002 | 0.0015 | -0.0127 | -0.0011 | 0.0037 | -0.0061 | -0.0048 | 0.1089 | 0.0136 | 0.0011 | 0.0002 | 0 | 0.0001 | 0 | -0.0001 | 1.01 | 1 | 1 | -0.0001 | -0.0031 | -0.0011 | -0.0014 | 0 | -0.0002 | -0.0143 | -0.0001 | 0.0005 | -0.0001 | 0.0004 | 0.0007 | 0.0312 | -0.0041 | -0.0026 | -0.0002 | 0.0015 | -0.0127 | -0.0011 | 0.0037 | -0.0061 | -0.0048 | 0.1089 | 0.0136 | 0.0011 | 0 | 0.0001 | 0 | -0.0001
ICV (SS 5ppm All) | QC | 2024-08-26 11:28 | [6 empty] | 4.9623 | 4.8891 | 4.9595 | 4.8833 | 4.8492 | 4.8377 | 4.9432 | 4.8909 | 4.8739 | 4.8912 | 4.6594 | 4.9181 | 4.8811 | 4.8129 | 4.9681 | 4.9005 | 4.9834 | 4.8674 | 4.8768 | 4.9182 | 5.0059 | 4.8677 | 5.1573 | 4.9233 | 4.6719 | 4.9552 | 4.9617 | 4.9407 | 4.9385 | 4.8208 | 1.01 | 1 | 0.99 | 4.9618 | 4.8891 | 4.9595 | 4.8833 | 4.8492 | 4.8377 | 4.9432 | 4.8909 | 4.8736 | 4.8898 | 4.6594 | 4.9181 | 4.8811 | 4.8129 | 4.9673 | 4.9005 | 4.9834 | 4.8674 | 4.8768 | 4.9181 | 5.0052 | 4.8413 | 5.1573 | 4.9204 | 4.67 | 4.9617 | 4.9407 | 4.934 | 4.8208
Rinse | QC | 2024-08-26 11:30 | [6 empty] | 0.0003 | -0.0024 | 0.0148 | -0.0002 | 0.0002 | 0 | -0.0126 | 0.0002 | 0.0002 | -0.0001 | 0.0002 | -0.0008 | 0.034 | 0.0103 | -0.0038 | -0.0003 | 0.045 | -0.0112 | -0.0002 | 0.0008 | -0.0069 | -0.0039 | 0.1204 | 0.0255 | 0.0026 | 0.0019 | 0.0001 | 0.0013 | 0.0005 | 0.0004 | 1.02 | 1.01 | 1 | 0.0003 | -0.0024 | 0.0148 | -0.0002 | 0.0002 | 0 | -0.0126 | 0.0002 | 0.0002 | -0.0001 | 0.0002 | -0.0008 | 0.034 | 0.0103 | -0.0038 | -0.0003 | 0.045 | -0.0112 | -0.0002 | 0.0008 | -0.0069 | -0.0039 | 0.1204 | 0.0255 | 0.0026 | 0.0001 | 0.0013 | 0.0005 | 0.0004
CCB | QC | 2024-08-26 11:33 | [6 empty] | -0.0002 | -0.0022 | 0.0035 | 0.0002 | 0 | -0.0001 | -0.0187 | 0 | 0.0004 | -0.0003 | 0.0005 | 0.0005 | 0.0336 | 0.0054 | -0.0026 | 0 | 0.0132 | -0.0095 | 0 | 0.0062 | -0.0079 | -0.0024 | 0.1179 | 0.0164 | 0.0008 | -0.0004 | 0 | 0.0004 | -0.0001 | 0 | 1.02 | 0.99 | 0.99 | -0.0002 | -0.0022 | 0.0035 | 0.0002 | 0 | -0.0001 | -0.0187 | 0 | 0.0004 | -0.0003 | 0.0005 | 0.0005 | 0.0336 | 0.0054 | -0.0026 | 0 | 0.0132 | -0.0095 | 0 | 0.0062 | -0.0079 | -0.0024 | 0.1179 | 0.0164 | 0.0008 | 0 | 0.0004 | -0.0001 | 0
0.5ppm All (LLQC) | QC | 2024-08-26 11:35 | [6 empty] | 0.4772 | 0.4765 | 0.4711 | 0.4746 | 0.468 | 0.4755 | 0.4793 | 0.4775 | 0.4796 | 0.4819 | 0.4528 | 0.4854 | 0.4717 | 0.4833 | 0.4658 | 0.4872 | 0.4495 | 0.4757 | 0.4797 | 0.476 | 0.4771 | 0.4627 | 0.4674 | 0.4695 | 0.4571 | 0.4709 | 0.4848 | 0.4756 | 0.4748 | 0.4706 | 1.02 | 1 | 1 | 0.4771 | 0.4765 | 0.4711 | 0.4746 | 0.468 | 0.4755 | 0.4793 | 0.4775 | 0.4796 | 0.4818 | 0.4528 | 0.4854 | 0.4717 | 0.4833 | 0.4657 | 0.4872 | 0.4495 Q | 0.4757 | 0.4797 | 0.476 | 0.477 | 0.4601 | 0.4674 | 0.4692 | 0.4569 | 0.4848 | 0.4756 | 0.4744 | 0.4706
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-26 11:38 | [6 empty] | 0.4802 | 0.4764 | 0.4755 | 0.4779 | 0.4684 | 0.4771 | 0.4852 | 0.4811 | 0.4829 | 0.4854 | 0.4563 | 0.4859 | 0.4759 | 0.4847 | 0.4702 | 0.4869 | 0.4632 | 0.4802 | 0.4823 | 0.4773 | 0.4853 | 0.4681 | 0.5161 | 0.4811 | 0.46 | 0.476 | 0.4856 | 0.4786 | 0.4774 | 0.4737 | 1.02 | 1 | 1 | 0.4802 | 0.4764 | 0.4755 | 0.4779 | 0.4684 | 0.4771 | 0.4852 | 0.4811 | 0.4828 | 0.4853 | 0.4563 | 0.4859 | 0.4759 | 0.4847 | 0.4701 | 0.4869 | 0.4632 | 0.4802 | 0.4823 | 0.4773 | 0.4852 | 0.4655 | 0.5161 | 0.4808 | 0.4598 | 0.4856 | 0.4786 | 0.477 | 0.4737
CCV (5ppm All) | QC | 2024-08-26 11:40 | [6 empty] | 4.8302 | 4.7689 | 4.8274 | 4.851 | 4.7504 | 4.7583 | 4.869 | 4.8472 | 4.8393 | 4.8511 | 4.6283 | 4.8281 | 4.771 | 4.7916 | 4.8457 | 4.8287 | 4.8803 | 4.7391 | 4.8387 | 4.8407 | 4.9318 | 4.8393 | 5.2404 | 4.8495 | 4.869 | 4.8317 | 4.8371 | 4.8284 | 4.8472 | 4.7968 | 1 | 1 | 0.99 | 4.8297 | 4.7689 | 4.8274 | 4.851 | 4.7504 | 4.7583 | 4.869 | 4.8471 | 4.839 | 4.8497 | 4.6283 | 4.8281 | 4.771 | 4.7916 | 4.8449 | 4.8287 | 4.8803 | 4.7391 | 4.8387 | 4.8406 | 4.931 | 4.8132 | 5.2404 | 4.8466 | 4.8671 | 4.8371 | 4.8284 | 4.8428 | 4.7968
CCB | QC | 2024-08-26 12:59 | [6 empty] | 0.0000 ! | -0.0034 ! | 0.0078 ! | -0.0020 ! | 0.0003 ! | 0.0000 ! | -0.0166 ! | 0.0000 ! | 0.0006 ! | -0.0002 ! | 0.0012 ! | 0.0003 ! | 0.0228 ! | -0.0041 ! | -0.0049 ! | -0.0001 ! | 0.0209 ! | -0.0179 ! | -0.0008 ! | 0.0000 ! | -0.0074 ! | -0.0061 ! | 0.1630 ! | 0.0141 ! | 0.0025 ! | 0.0011 ! | 0.0000 ! | 0.0003 ! | 0.0003 ! | 0.0000 ! | 1.00 ! | 1.00 ! | 1.01 ! | 0 | -0.0034 | 0.0078 | -0.002 | 0.0003 | 0 | -0.0166 | 0 | 0.0006 | -0.0002 | 0.0012 | 0.0003 | 0.0228 | -0.0041 | -0.0049 | -0.0001 | 0.0209 | -0.0179 | -0.0008 | 0 | -0.0074 | -0.0061 | 0.163 | 0.0141 | 0.0025 | 0 | 0.0003 | 0.0003 | 0
Rinse | QC | 2024-08-26 15:09 | [6 empty] | 0.0001 | -0.0018 | 0.0041 | -0.0032 | 0 | -0.0002 | -0.0149 | 0 | 0.0006 | -0.0002 | 0.0003 | 0.0007 | 0.024 | 0.0016 | -0.0053 | -0.0002 | 0.0027 | -0.0147 | 0.0001 | 0.0025 | -0.0067 | -0.0061 | 0.1114 | 0.0121 | 0.0011 | -0.0015 | 0 | 0 | 0 | 0.0001 | 1.03 | 1 | 1.01 | 0.0001 | -0.0018 | 0.0041 | -0.0032 | 0 | -0.0002 | -0.0149 | 0 | 0.0006 | -0.0002 | 0.0003 | 0.0007 | 0.024 | 0.0016 | -0.0053 | -0.0002 | 0.0027 | -0.0147 | 0.0001 | 0.0025 | -0.0067 | -0.0061 | 0.1114 | 0.0121 | 0.0011 | 0 | 0 | 0 | 0.0001
CCV (5ppm All) | QC | 2024-08-26 15:11 | [6 empty] | 4.9393 | 4.8292 | 4.9747 | 4.9869 | 4.8694 | 4.8934 | 4.9624 | 4.9853 | 4.9771 | 4.9916 | 4.7353 | 4.9025 | 4.8388 | 4.8803 | 4.9769 | 4.9054 | 5.0045 | 4.8047 | 4.9781 | 5.2294 | 5.0809 | 5.0721 | 5.3209 | 5.0244 | 4.9494 | 4.9716 | 4.9421 | 4.935 | 4.9679 | 4.9364 | 1.01 | 1 | 1 | 4.9388 | 4.8292 | 4.9747 | 4.9869 | 4.8694 | 4.8934 | 4.9624 | 4.9853 | 4.9768 | 4.9902 | 4.7354 | 4.9025 | 4.8388 | 4.8803 | 4.976 | 4.9054 | 5.0045 | 4.8047 | 4.9781 | 5.2294 | 5.0802 | 5.0456 | 5.321 | 5.0215 | 4.9475 | 4.9421 | 4.935 | 4.9634 | 4.9364
CCB | QC | 2024-08-26 15:14 | [6 empty] | 0.0002 | -0.0002 | 0.0154 | -0.0014 | 0.0002 | 0 | -0.0143 | 0.0001 | 0.0005 | 0 | -0.0002 | 0.0017 | 0.0355 | -0.0098 | -0.0041 | 0 | 0.0536 | -0.0153 | -0.0007 | 0.0045 | -0.0076 | -0.0048 | 0.183 | 0.0334 | 0.0026 | 0.0011 | 0.0001 | 0.0016 | 0.0001 | 0.0001 | 1.02 | 1 | 1.01 | 0.0002 | -0.0002 | 0.0154 | -0.0014 | 0.0002 | 0 | -0.0143 | 0.0001 | 0.0005 | 0 | -0.0002 | 0.0017 | 0.0355 | -0.0098 | -0.0041 | 0 | 0.0536 | -0.0153 | -0.0007 | 0.0045 | -0.0076 | -0.0048 | 0.183 | 0.0334 | 0.0026 | 0.0001 | 0.0016 | 0.0001 | 0.0001
08262024 LB-2 | Sample | 2024-08-26 15:17 | [6 empty] | 0.0002 | 0.0001 | 0.0052 | -0.0019 | 0.0001 | -0.0001 | -0.0115 | 0 | 0.0003 | -0.0002 | 0.0005 | 0.0007 | 0.0277 | -0.0078 | -0.0066 | -0.0003 | 0.0145 | -0.0147 | -0.001 | 0.0039 | -0.0087 | -0.0036 | 0.116 | 0.0132 | 0.002 | -0.0001 | 0 | 0.0004 | 0.0001 | 0.0005 | 1.03 | 0.99 | 1 | 0.0002 | 0.0001 | 0.0052 | -0.0019 | 0.0001 | -0.0001 | -0.0115 | 0 | 0.0003 | -0.0002 | 0.0005 | 0.0007 | 0.0277 | -0.0078 | -0.0066 | -0.0003 | 0.0145 | -0.0147 | -0.001 | 0.0039 | -0.0087 | -0.0036 | 0.116 | 0.0132 | 0.002 | 0 | 0.0004 | 0.0001 | 0.0005
10x 08265024 LCS-2 | Sample | 2024-08-26 15:19 | [6 empty] | 0.0091 | 0.4747 | 0.245 | 0.0466 | 0.2424 | 0.0484 | 0.474 | 0.0979 | 0.4878 | 0.2465 | 0.2302 | 0.497 | 0.4924 | 0 | 0.4807 | 0.0995 | 0.0921 | 0.4818 | 0.241 | 0.1057 | 0.2418 | 0.2445 | 0.242 | 0.2457 | 0.0954 | 0.0958 | 0.1001 | 0.0977 | 0.2448 | 0.0958 | 1.04 | 1 | 1.01 | 0.0091 | 0.4747 | 0.245 | 0.0466 | 0.2424 | 0.0484 | 0.474 | 0.0979 | 0.4877 | 0.2465 | 0.2302 | 0.497 | 0.4924 | 0 | 0.4806 | 0.0995 | 0.0921 | 0.4818 | 0.241 | 0.1057 | 0.2418 | 0.2439 | 0.2421 | 0.2456 | 0.0954 | 0.1001 | 0.0977 | 0.2447 | 0.0958
10x 08262024 LCS-2- Analytical Duplicate | Sample | 2024-08-26 15:22 | [6 empty] | 0.0098 | 0.4716 | 0.2447 | 0.047 | 0.2402 | 0.048 | 0.4719 | 0.0979 | 0.4881 | 0.247 | 0.23 | 0.4937 | 0.4729 | -0.0089 | 0.4815 | 0.0985 | 0.0911 | 0.4819 | 0.2415 | 0.1077 | 0.2429 | 0.2413 | 0.2613 | 0.2503 | 0.0946 | 0.0971 | 0.0992 | 0.0978 | 0.2444 | 0.0959 | 1.03 | 0.99 | 1 | 0.0098 | 0.4716 | 0.2447 | 0.047 | 0.2402 | 0.048 | 0.4719 | 0.0979 | 0.4881 | 0.2469 | 0.23 | 0.4937 | 0.4729 | -0.0089 | 0.4814 | 0.0985 | 0.0911 | 0.4819 | 0.2415 | 0.1077 | 0.2429 | 0.2408 | 0.2613 | 0.2502 | 0.0946 | 0.0992 | 0.0978 | 0.2443 | 0.0959
5x 08262024 LCS-2 | Sample | 2024-08-26 15:24 | [6 empty] | 0.0191 | 0.9546 | 0.4902 | 0.0948 | 0.484 | 0.0971 | 0.9666 | 0.1962 | 0.982 | 0.4911 | 0.4582 | 0.9936 | 0.9653 | -0.0003 | 0.9809 | 0.1978 | 0.1799 | 0.9785 | 0.4847 | 0.2071 | 0.4875 | 0.488 | 0.473 | 0.49 | 0.1862 | 0.1948 | 0.199 | 0.1944 | 0.4875 | 0.1946 | 1.03 | 0.99 | 1 | 0.019 | 0.9546 | 0.4902 | 0.0948 | 0.484 | 0.0971 | 0.9666 | 0.1962 | 0.982 | 0.4911 | 0.4582 | 0.9936 | 0.9653 | -0.0003 | 0.9807 | 0.1978 | 0.1799 | 0.9785 | 0.4847 | 0.2071 | 0.4874 | 0.4869 | 0.473 | 0.4898 | 0.1861 | 0.199 | 0.1944 | 0.4873 | 0.1946
08262024 LCS-2 | Sample | 2024-08-26 15:27 | [6 empty] | 0.0964 | 4.6699 | 2.4802 | 0.4939 | 2.3903 | 0.4793 | 4.8571 | 0.9906 | 4.9269 | 2.4485 | 2.2906 | 4.8599 | 4.8696 | -0.0012 | 4.9938 | 0.9732 | 0.9373 | 4.828 | 2.4507 | 1.0388 | 2.4955 | 2.5127 | 2.5631 | 2.4883 | 0.9586 | 0.9909 | 0.9815 | 0.9822 | 2.45 | 0.977 | 1.01 | 1 | 1 | 0.0963 | 4.6699 | 2.4802 | 0.4939 | 2.3903 | 0.4793 | 4.8571 | 0.9906 | 4.9267 | 2.4482 | 2.2906 | 4.8599 | 4.8696 | -0.0012 | 4.993 | 0.9732 | 0.9373 | 4.828 | 2.4507 | 1.0387 | 2.4953 | 2.5074 | 2.5632 | 2.4876 | 0.9582 | 0.9815 | 0.9822 | 2.4491 | 0.977
08262024 LCS-2- Analytical Duplicate | Sample | 2024-08-26 15:29 | [6 empty] | 0.0955 | 4.6689 | 2.4679 | 0.4892 | 2.395 | 0.477 | 4.8094 | 0.9811 | 4.8864 | 2.4264 | 2.2894 | 4.8664 | 4.8059 | -0.0077 | 4.9485 | 0.9754 | 0.9521 | 4.7946 | 2.4232 | 1.0304 | 2.4741 | 2.4897 | 2.5859 | 2.4734 | 0.955 | 0.9818 | 0.9829 | 0.973 | 2.4418 | 0.9666 | 1.01 | 0.99 | 1 | 0.0954 | 4.6689 | 2.4679 | 0.4892 | 2.395 | 0.477 | 4.8094 | 0.981 | 4.8862 | 2.4261 | 2.2895 | 4.8664 | 4.8059 | -0.0077 | 4.9477 | 0.9754 | 0.9521 | 4.7946 | 2.4232 | 1.0303 | 2.4739 | 2.4844 | 2.5859 | 2.4727 | 0.9547 | 0.9829 | 0.973 | 2.4409 | 0.9666
Rinse | QC | 2024-08-26 15:32 | [6 empty] | -0.0003 | -0.0025 | 0.0141 | -0.0021 | 0.0002 | -0.0001 | -0.0112 | 0 | 0.0005 | -0.0003 | 0.0004 | 0.0004 | 0.0199 | -0.0006 | -0.005 | 0 | 0.0305 | -0.0155 | -0.0014 | 0.0008 | -0.0083 | -0.001 | 0.1824 | 0.0173 | 0.0027 | 0.0013 | 0 | 0.0003 | 0.0001 | -0.0001 | 1.02 | 1 | 1.01 | -0.0003 | -0.0025 | 0.0141 | -0.0021 | 0.0002 | -0.0001 | -0.0112 | 0 | 0.0005 | -0.0003 | 0.0004 | 0.0004 | 0.0199 | -0.0006 | -0.005 | 0 | 0.0305 | -0.0155 | -0.0014 | 0.0008 | -0.0083 | -0.001 | 0.1824 | 0.0173 | 0.0027 | 0 | 0.0003 | 0.0001 | -0.0001
CCV (5ppm All) | QC | 2024-08-26 15:34 | [6 empty] | 4.9394 | 4.8512 | 4.9794 | 4.9897 | 4.8841 | 4.9022 | 4.9564 | 4.9839 | 4.9768 | 4.9918 | 4.7376 | 4.9213 | 4.8156 | 4.8705 | 4.9705 | 4.9116 | 5.0241 | 4.8023 | 4.9756 | 5.2478 | 5.0856 | 5.0635 | 5.3825 | 5.0299 | 4.9755 | 4.9661 | 4.9615 | 4.9335 | 4.9689 | 4.9204 | 1 | 0.99 | 1 | 4.9389 | 4.8512 | 4.9794 | 4.9897 | 4.8841 | 4.9022 | 4.9564 | 4.9838 | 4.9764 | 4.9904 | 4.7376 | 4.9213 | 4.8156 | 4.8705 | 4.9696 | 4.9116 | 5.0241 | 4.8023 | 4.9756 | 5.2477 | 5.0848 | 5.0369 | 5.3825 | 5.027 | 4.9736 | 4.9615 | 4.9335 | 4.9644 | 4.9204
CCB | QC | 2024-08-26 15:37 | [6 empty] | 0.0003 | -0.0053 | 0.0156 | -0.0012 | 0.0002 | -0.0001 | -0.0152 | 0.0001 | 0.0004 | 0 | 0.0003 | -0.0002 | 0.031 | -0.0007 | -0.0084 | -0.0001 | 0.0461 | -0.0125 | -0.001 | 0.006 | -0.0074 | -0.0093 | 0.1541 | 0.0243 | 0.0033 | 0.0025 | 0.0001 | 0.0014 | -0.0001 | 0 | 1.02 | 0.99 | 1 | 0.0003 | -0.0053 | 0.0156 | -0.0012 | 0.0002 | -0.0001 | -0.0152 | 0.0001 | 0.0004 | 0 | 0.0003 | -0.0002 | 0.031 | -0.0007 | -0.0084 | -0.0001 | 0.0461 | -0.0125 | -0.001 | 0.006 | -0.0074 | -0.0093 | 0.1541 | 0.0243 | 0.0033 | 0.0001 | 0.0014 | -0.0001 | 0
0.5ppm All (LLQC) | QC | 2024-08-26 15:39 | [6 empty] | 0.4906 | 0.4836 | 0.4869 | 0.4883 | 0.4772 | 0.4868 | 0.491 | 0.4917 | 0.4949 | 0.4971 | 0.465 | 0.4887 | 0.4833 | 0.4872 | 0.4791 | 0.4956 | 0.4776 | 0.4797 | 0.4944 | 0.511 | 0.495 | 0.489 | 0.5089 | 0.4936 | 0.4691 | 0.4861 | 0.4931 | 0.4883 | 0.4884 | 0.4838 | 1.03 | 1 | 1 | 0.4906 | 0.4836 | 0.4869 | 0.4883 | 0.4772 | 0.4868 | 0.491 | 0.4917 | 0.4949 | 0.4969 | 0.465 | 0.4887 | 0.4833 | 0.4872 | 0.479 | 0.4956 | 0.4776 | 0.4797 | 0.4944 | 0.511 | 0.4949 | 0.4864 | 0.5089 | 0.4933 | 0.4689 | 0.4931 | 0.4883 | 0.488 | 0.4838
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-26 15:42 | [6 empty] | 0.4882 | 0.4883 | 0.4903 | 0.4868 | 0.4806 | 0.4876 | 0.4951 | 0.4909 | 0.4935 | 0.4955 | 0.465 | 0.4942 | 0.4764 | 0.4839 | 0.4788 | 0.497 | 0.4781 | 0.4753 | 0.4924 | 0.5118 | 0.4962 | 0.4845 | 0.5351 | 0.4924 | 0.4674 | 0.4839 | 0.497 | 0.4871 | 0.4882 | 0.4827 | 1.03 | 0.99 | 1.01 | 0.4882 | 0.4883 | 0.4903 | 0.4868 | 0.4806 | 0.4876 | 0.4951 | 0.4909 | 0.4935 | 0.4953 | 0.465 | 0.4942 | 0.4764 | 0.4839 | 0.4787 | 0.497 | 0.4781 | 0.4753 | 0.4924 | 0.5118 | 0.4962 | 0.4818 | 0.5352 | 0.4921 | 0.4672 | 0.497 | 0.4871 | 0.4878 | 0.4827
CCS1 (SS 5ppm) | QC | 2024-08-26 15:44 | [6 empty] | 5.077 | 4.9933 | 5.1119 | 5.0207 | 5.0011 | 5.019 | 5.0408 | 5.0255 | 5.0334 | 5.0285 | 4.7776 | 5.0138 | 4.9132 | 4.8744 | 5.0838 | 5.0055 | 5.1896 | 4.9194 | 5.0073 | 5.3276 | 5.1398 | 5.1004 | 5.4157 | 5.0731 | 4.816 | 5.0871 | 5.1066 | 5.0508 | 5.0682 | 4.9732 | 1.01 | 0.99 | 1 | 5.0765 | 4.9933 | 5.1119 | 5.0207 | 5.0011 | 5.019 | 5.0408 | 5.0254 | 5.0331 | 5.0271 | 4.7777 | 5.0138 | 4.9132 | 4.8744 | 5.083 | 5.0055 | 5.1896 | 4.9194 | 5.0073 | 5.3275 | 5.1391 | 5.0734 | 5.4158 | 5.0701 | 4.8141 | 5.1066 | 5.0508 | 5.0636 | 4.9732
CCB | QC | 2024-08-26 15:47 | [6 empty] | -0.0001 | -0.0025 | 0.0143 | -0.0005 | 0.0002 | 0 | -0.0142 | 0.0001 | 0.0005 | -0.0004 | 0.0005 | 0.0008 | 0.0391 | -0.008 | -0.007 | -0.0001 | 0.0439 | -0.0143 | -0.0012 | 0.0031 | -0.0094 | -0.0061 | 0.1272 | 0.0237 | 0.0028 | 0.0039 | 0.0001 | 0.0012 | 0.0002 | 0.0003 | 1.03 | 1 | 1 | -0.0001 | -0.0025 | 0.0143 | -0.0005 | 0.0002 | 0 | -0.0142 | 0.0001 | 0.0005 | -0.0004 | 0.0005 | 0.0008 | 0.0391 | -0.008 | -0.007 | -0.0001 | 0.0439 | -0.0143 | -0.0012 | 0.0031 | -0.0094 | -0.0061 | 0.1272 | 0.0237 | 0.0028 | 0.0001 | 0.0012 | 0.0002 | 0.0003
MQ-Blank-02 | Sample | 2024-08-26 15:49 | [6 empty] | 0.0001 | 0.0012 | 0 | -0.0026 | 0 | 0.0001 | -0.0038 | 0.0001 | 0.0003 | -0.0002 | 0.0004 | -0.0002 | 0.0338 | -0.0068 | -0.0082 | 0 | 0.0043 | -0.0137 | 0.0002 | 0.0037 | -0.0084 | -0.0051 | 0.1059 | 0.0116 | 0.0031 | -0.0002 | 0.0001 | 0.0001 | 0.0003 | 0.0004 | 1.06 | 1 | 1 | 0.0001 | 0.0012 | 0 | -0.0026 | 0 | 0.0001 | -0.0038 | 0.0001 | 0.0003 | -0.0002 | 0.0004 | -0.0002 | 0.0338 | -0.0068 | -0.0082 | 0 | 0.0043 | -0.0137 | 0.0002 | 0.0037 | -0.0084 | -0.0051 | 0.1059 | 0.0116 | 0.0031 | 0.0001 | 0.0001 | 0.0003 | 0.0004
Rinse | QC | 2024-08-26 15:52 | [6 empty] | 0.0001 | 0.0004 | 0.0044 | -0.0017 | 0 | -0.0002 | -0.0177 | 0.0001 | 0.0004 | -0.0002 | 0.0008 | -0.0003 | 0.0301 | -0.0084 | -0.0031 | -0.0003 | 0.0099 | -0.0148 | -0.0001 | -0.0013 | -0.0084 | -0.0053 | 0.1127 | 0.0122 | 0.0005 | 0.0005 | 0 | 0.0002 | 0.0002 | 0.0004 | 1.02 | 1 | 1.01 | 0.0001 | 0.0004 | 0.0044 | -0.0017 | 0 | -0.0002 | -0.0177 | 0.0001 | 0.0004 | -0.0002 | 0.0008 | -0.0003 | 0.0301 | -0.0084 | -0.0031 | -0.0003 | 0.0099 | -0.0148 | -0.0001 | -0.0013 | -0.0084 | -0.0053 | 0.1127 | 0.0122 | 0.0005 | 0 | 0.0002 | 0.0002 | 0.0004
CCV (5ppm All) | QC | 2024-08-26 15:54 | [6 empty] | 4.9374 | 4.874 | 4.9746 | 4.9946 | 4.8944 | 4.8892 | 4.9557 | 4.9761 | 4.9726 | 4.9826 | 4.7368 | 4.9438 | 4.8666 | 4.8894 | 4.9608 | 4.9434 | 5.0179 | 4.8117 | 4.9632 | 5.259 | 5.0729 | 5.0479 | 5.3427 | 4.9952 | 4.9765 | 4.9578 | 4.9731 | 4.9332 | 4.9653 | 4.9532 | 1 | 0.98 | 0.99 | 4.9369 | 4.874 | 4.9746 | 4.9946 | 4.8944 | 4.8892 | 4.9557 | 4.976 | 4.9722 | 4.9812 | 4.7368 | 4.9438 | 4.8666 | 4.8894 | 4.9599 | 4.9434 | 5.0179 | 4.8117 | 4.9632 | 5.2589 | 5.0722 | 5.0212 | 5.3428 | 4.9922 | 4.9746 | 4.9731 | 4.9332 | 4.9609 | 4.9532
CCB | QC | 2024-08-26 15:57 | [6 empty] | -0.0001 | -0.0009 | 0.0147 | 0.0009 | 0.0004 | 0.0001 | -0.0142 | 0.0001 | 0.0001 | 0 | 0.0004 | 0.0005 | 0.0169 | -0.0067 | -0.0079 | 0.0002 | 0.0507 | -0.0154 | -0.0007 | 0.0033 | -0.0082 | -0.0047 | 0.1762 | 0.0281 | 0.0024 | 0.0021 | 0.0003 | 0.0016 | 0.0006 | 0.0005 | 1.01 | 1 | 1 | -0.0001 | -0.0009 | 0.0147 | 0.0009 | 0.0004 | 0.0001 | -0.0142 | 0.0001 | 0.0001 | 0 | 0.0004 | 0.0005 | 0.0169 | -0.0067 | -0.0079 | 0.0002 | 0.0507 | -0.0154 | -0.0007 | 0.0033 | -0.0082 | -0.0047 | 0.1762 | 0.0281 | 0.0024 | 0.0003 | 0.0016 | 0.0006 | 0.0005
0.5% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 15:59 | [6 empty] | -0.0004 | 0.1804 | 0.0154 | 0.0066 | 0.0418 | -0.0001 | 0.3158 | 0.0001 | 0.0002 | 0.0005 | 0.0025 | 0.1616 | 6.6046 | 0.0095 | 0.0343 | 0.001 | 0.0174 | 1080.1282 o | 1.455 | 0.1984 | -0.003 | 0.1711 | 0.1259 | 0.0175 | 0.0436 | 0.0036 | 0.0013 | 0.0013 | 0.0004 | 0.0059 | 0.92 | 0.93 | 0.93 | -0.0004 | 0.1804 | 0.0154 | 0.0066 | 0.0418 | -0.0001 | 0.3158 | 0 | 0.0002 | 0.0005 | 0.0025 | 0.1616 | 6.6046 | 0.0095 | 0.0343 | 0.001 | 0.0174 | 1080.1282 | 1.455 | 0.1984 | -0.003 | 0.1711 | 0.1259 | 0.0175 | 0.0436 | 0.0013 | 0.0013 | 0.0004 | 0.0059
1% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:02 | [6 empty] | -0.0002 | 0.1093 | 0.0126 | 0.0075 | 0.0466 | 0 | 0.5437 | 0.0001 | 0.0003 | 0.0001 | 0.003 | 0.0639 | 0.6065 | 0.001 | 0.0789 | 0.0004 | 0.0091 | 1782.5297 o | 1.364 | 0.1745 | -0.0046 | 0.3071 | 0.1101 | 0.0136 | 0.0677 | 0.0034 | 0.0021 | 0.0013 | 0.0003 | 0.0057 | 0.9 | 0.9 | 0.91 | -0.0002 | 0.1093 | 0.0126 | 0.0075 | 0.0466 | 0 | 0.5437 | 0.0001 | 0.0003 | 0.0001 | 0.003 | 0.0639 | 0.6065 | 0.001 | 0.0789 | 0.0004 | 0.0091 | 1782.5297 | 1.364 | 0.1745 | -0.0046 | 0.3071 | 0.1101 | 0.0136 | 0.0677 | 0.0021 | 0.0013 | 0.0003 | 0.0057
5% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:04 | [6 empty] | -0.0008 | 0.1806 | 0.0084 | 0.0292 | 0.0881 | -0.0001 | 2.2565 | 0 | 0.0003 | 0.0004 | 0.0022 | 0.1344 | 1.2881 | 0.0147 | 0.3968 | 0.0027 | 0.0082 | #### | 0.8643 | 0.1541 | -0.0106 | 1.07 | 0.1085 | 0.008 | 0.3693 | 0.0019 | 0.0079 | 0.0072 | 0.0016 | 0.0106 | 0.82 | 0.78 | 0.79 | -0.0008 | 0.1806 | 0.0084 | 0.0292 | 0.0881 | -0.0001 | 2.2565 | -0.0001 | 0.0003 | 0.0004 | 0.0022 | 0.1344 | 1.2881 | 0.0147 | 0.3968 | 0.0027 | 0.0082 |  | 0.8643 | 0.1541 | -0.0106 | 1.07 | 0.1085 | 0.008 | 0.3693 | 0.0079 | 0.0072 | 0.0015 | 0.0106
10% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:07 | [6 empty] | -0.0012 | 0.1245 | 0.0069 | 0.0561 | 0.1333 | 0.0001 | 4.198 | -0.0001 | -0.0002 | 0.0005 | 0.0006 | 0.2474 | 2.5736 | 0.0322 | 0.7236 | 0.0088 | 0.0076 | #### | 0.2948 | 0.2303 | -0.0143 | 1.9989 | 0.1094 | 0.0095 | 0.643 | 0.002 | 0.0154 | 0.0105 | 0.003 | 0.0114 | 0.74 | 0.67 | 0.67 | -0.0012 | 0.1245 | 0.0069 | 0.0561 | 0.1333 | 0.0001 | 4.198 | -0.0001 | -0.0002 | 0.0005 | 0.0006 | 0.2474 | 2.5736 | 0.0322 | 0.7236 | 0.0088 | 0.0076 |  | 0.2948 | 0.2303 | -0.0143 | 1.9988 | 0.1094 | 0.0094 | 0.643 | 0.0154 | 0.0105 | 0.003 | 0.0114
15% NaHCO3 LFP Battery-TM | Sample | 2024-08-26 16:09 | [6 empty] | -0.0013 | 0.2724 | 0.0081 | 0.069 | 0.1331 | -0.0001 | 4.4436 | 0.0002 | -0.0003 | 0.0006 | 0 | 1.1278 | 5.9549 | 0.0369 | 0.7403 | 0.0075 | 0.0078 | #### | 0.2598 | 0.4848 | -0.0148 | 2.9211 | 0.1073 | 0.0067 | 0.9399 | 0.0002 | 0.0171 | 0.0148 | 0.0046 | 0.014 | 0.74 | 0.67 | 0.67 | -0.0013 | 0.2724 | 0.0081 | 0.069 | 0.1331 | -0.0001 | 4.4436 | 0.0002 | -0.0003 | 0.0006 | 0 | 1.1278 | 5.9549 | 0.0369 | 0.7401 | 0.0075 | 0.0078 |  | 0.2598 | 0.4848 | -0.0148 | 2.921 | 0.1073 | 0.0067 | 0.9399 | 0.0171 | 0.0148 | 0.0046 | 0.014
0.5% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:12 | [6 empty] | -0.0002 | 0.3046 | 0.0061 | 0.003 | 0.003 | -0.0001 | 0.2535 | 0 | 0.0011 | 0.0004 | 0.0024 | 0.1919 | 0.4832 | 0.0093 | 0.0299 | 0.0006 | 0.0037 | 1061.1289 o | 4.6501 | 1.6312 | -0.0029 | 0.1404 | 0.12 | 0.0111 | 0.0661 | 0.0027 | 0.0007 | 0.0006 | 0.0008 | 0.0056 | 0.9 | 0.92 | 0.92 | -0.0002 | 0.3046 | 0.0061 | 0.003 | 0.003 | -0.0001 | 0.2535 | 0 | 0.0011 | 0.0004 | 0.0024 | 0.1919 | 0.4832 | 0.0093 | 0.0299 | 0.0006 | 0.0037 | 1061.1289 | 4.6501 | 1.6312 | -0.0029 | 0.1404 | 0.12 | 0.0111 | 0.0661 | 0.0007 | 0.0006 | 0.0008 | 0.0056
1% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:14 | [6 empty] | -0.0002 | 0.4131 | 0.0041 | 0.0076 | 0.004 | -0.0001 | 0.5031 | 0 | 0.0012 | 0.0004 | 0.0027 | 0.099 | 0.5255 | 0.0028 | 0.0744 | 0.0011 | 0.004 | 1752.4356 o | 5.0444 | 2.0521 | -0.0019 | 0.2448 | 0.1071 | 0.015 | 0.1281 | 0.0015 | 0.0016 | 0.0016 | 0.0005 | 0.0057 | 0.87 | 0.89 | 0.9 | -0.0002 | 0.4131 | 0.0041 | 0.0076 | 0.004 | -0.0001 | 0.5031 | 0 | 0.0012 | 0.0004 | 0.0027 | 0.099 | 0.5255 | 0.0028 | 0.0744 | 0.0011 | 0.004 | 1752.4356 | 5.0444 | 2.0521 | -0.0019 | 0.2448 | 0.1071 | 0.015 | 0.1281 | 0.0016 | 0.0016 | 0.0005 | 0.0057
5% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:17 | [6 empty] | -0.001 | 2.8795 | 0.0061 | 0.0277 | 0.03 | 0 | 2.4264 | 0.0001 | 0.002 | 0.0025 | 0.0024 | 4.3695 | 1.4801 | 0.0114 | 0.4522 | 0.0089 | 0.005 | #### | 19.6945 | 4.4226 | -0.0106 | 1.1655 | 0.1182 | 0.0118 | 0.7978 | 0.0018 | 0.0083 | 0.0172 | 0.0019 | 0.0124 | 0.81 | 0.77 | 0.78 | -0.001 | 2.8795 | 0.0061 | 0.0277 | 0.03 | 0 | 2.4264 | 0 | 0.0019 | 0.0025 | 0.0025 | 4.3695 | 1.4801 | 0.0114 | 0.4514 | 0.0089 | 0.005 |  | 19.6945 | 4.4225 | -0.0107 | 1.1654 | 0.1183 | 0.0117 | 0.7978 | 0.0083 | 0.0172 | 0.0018 | 0.0124
10% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:19 | [6 empty] | -0.0009 | 0.0738 | 0.0069 | 0.0524 | 0.0639 | -0.0001 | 3.8722 | -0.0002 | 0.001 | 0.0029 | -0.0002 | 6.2456 | 3.0391 | 0.0282 | 0.7002 | 0.0087 | 0.0049 | #### | 9.1367 | 3.6269 | -0.0166 | 2.4263 | 0.1039 | 0.0088 | 1.3056 | 0.0016 | 0.0153 | 0.0175 | 0.0039 | 0.0053 | 0.74 | 0.68 | 0.68 | -0.0009 | 0.0738 | 0.0069 | 0.0524 | 0.0639 | -0.0001 | 3.8722 | -0.0003 | 0.0009 | 0.0029 | -0.0001 | 6.2456 | 3.0391 | 0.0282 | 0.6991 | 0.0087 | 0.0049 |  | 9.1367 | 3.6268 | -0.0167 | 2.4262 | 0.104 | 0.0087 | 1.3056 | 0.0153 | 0.0175 | 0.0039 | 0.0053
15% NaHCO3 MMC Battery-TM | Sample | 2024-08-26 16:22 | [6 empty] | -0.0012 | 0.0874 | 0.0059 | 0.0702 | 0.0723 | 0 | 4.1223 | -0.0001 | 0.0004 | 0.0022 | -0.001 | 6.0412 | 7.8595 | 0.038 | 0.7329 | 0.0097 | 0.0068 | #### | 9.3892 | 3.4593 | -0.0154 | 3.5913 | 0.1041 | 0.0017 | 1.9866 | 0.0007 | 0.0168 | 0.0242 | 0.006 | 0.0044 | 0.74 | 0.67 | 0.67 | -0.0012 | 0.0874 | 0.0059 | 0.0702 | 0.0723 | 0 | 4.1223 | -0.0002 | 0.0003 | 0.0022 | -0.001 | 6.0412 | 7.8595 | 0.038 | 0.7319 | 0.0097 | 0.0068 |  | 9.3892 | 3.4593 | -0.0156 | 3.5913 | 0.1042 | 0.0016 | 1.9866 | 0.0168 | 0.0242 | 0.0059 | 0.0044
Rinse1 | QC | 2024-08-26 16:24 | [6 empty] | -0.0003 | -0.0019 | 0.0033 | -0.0032 | -0.0001 | -0.0001 | -0.0093 | -0.0001 | 0.0006 | -0.0002 | 0.0008 | 0.0014 | 0.0378 | -0.0084 | -0.008 | -0.0002 | 0.0022 | 0.7828 | -0.001 | 0.0029 | -0.0074 | -0.0042 | 0.12 | 0.0085 | 0.0015 | -0.0004 | 0 | -0.0001 | -0.0002 | 0.0021 | 0.99 | 0.99 | 1 | -0.0003 | -0.0019 | 0.0033 | -0.0032 | -0.0001 | -0.0001 | -0.0093 | -0.0001 | 0.0006 | -0.0002 | 0.0008 | 0.0014 | 0.0378 | -0.0084 | -0.008 | -0.0002 | 0.0022 | 0.7828 | -0.001 | 0.0029 | -0.0074 | -0.0042 | 0.12 | 0.0085 | 0.0015 | 0 | -0.0001 | -0.0002 | 0.0021
Rinse1 | QC | 2024-08-26 16:27 | [6 empty] | -0.0003 | -0.0035 | -0.0003 | -0.0034 | 0.0002 | -0.0002 | -0.0085 | 0 | 0.0003 | 0.0001 | 0.0001 | 0.0009 | 0.0298 | -0.0072 | -0.0049 | -0.0002 | 0.0022 | 0.4574 | -0.0012 | 0.0035 | -0.0101 | -0.0039 | 0.1124 | 0.0059 | 0.0012 | -0.0001 | 0 | -0.0001 | 0.0005 | 0.0023 | 1 | 0.98 | 0.99 | -0.0003 | -0.0035 | -0.0003 | -0.0034 | 0.0002 | -0.0002 | -0.0085 | 0 | 0.0003 | 0.0001 | 0.0001 | 0.0009 | 0.0298 | -0.0072 | -0.0049 | -0.0002 | 0.0022 | 0.4574 | -0.0012 | 0.0035 | -0.0101 | -0.0039 | 0.1124 | 0.0059 | 0.0012 | 0 | -0.0001 | 0.0005 | 0.0023
Rinse | QC | 2024-08-26 16:29 | [6 empty] | 0.0001 | -0.0025 | -0.0013 | -0.0035 | 0.0001 | -0.0001 | -0.0172 | 0 | 0.0005 | -0.0001 | 0.001 | 0.0001 | 0.0354 | -0.0075 | -0.0067 | -0.0003 | 0.0022 | 0.3793 | -0.0009 | -0.0007 | -0.0059 | -0.0025 | 0.1147 | 0.0111 | 0.0013 | -0.0003 | 0 | 0 | 0.0002 | 0.0019 | 0.99 | 0.98 | 0.99 | 0.0001 | -0.0025 | -0.0013 | -0.0035 | 0.0001 | -0.0001 | -0.0172 | 0 | 0.0005 | -0.0001 | 0.001 | 0.0001 | 0.0354 | -0.0075 | -0.0067 | -0.0003 | 0.0022 | 0.3793 | -0.0009 | -0.0007 | -0.0059 | -0.0025 | 0.1147 | 0.0111 | 0.0013 | 0 | 0 | 0.0002 | 0.0019
CCV (5ppm All) | QC | 2024-08-26 16:32 | [6 empty] | 4.9248 | 4.8786 | 4.9225 | 4.9468 | 4.8842 | 4.9004 | 4.9412 | 4.9473 | 4.9433 | 4.9671 | 4.7225 | 4.9091 | 4.8315 | 4.8826 | 4.9262 | 4.9456 | 4.9739 | 5.032 | 4.9424 | 5.2544 | 5.0427 | 5.0052 | 5.2856 | 4.965 | 4.8975 | 4.9396 | 4.9329 | 4.9221 | 4.9537 | 4.9082 | 0.98 | 0.98 | 0.98 | 4.9243 | 4.8786 | 4.9225 | 4.9468 | 4.8842 | 4.9004 | 4.9412 | 4.9472 | 4.9429 | 4.9657 | 4.7225 | 4.9091 | 4.8315 | 4.8826 | 4.9254 | 4.9456 | 4.9739 | 5.032 | 4.9424 | 5.2544 | 5.042 | 4.9785 | 5.2856 | 4.962 | 4.8957 | 4.9329 | 4.9221 | 4.9492 | 4.9082
CCB | QC | 2024-08-26 16:34 | [6 empty] | -0.0002 | -0.0029 | 0.0166 | -0.0014 | 0.0001 | 0 | -0.0141 | 0.0001 | 0.0005 | 0.0001 | 0.0006 | 0.001 | 0.0409 | 0.0017 | -0.0053 | -0.0003 | 0.0491 | 0.1999 | -0.0002 | 0.0027 | -0.0087 | -0.0064 | 0.1625 | 0.028 | 0.0027 | 0.0005 | 0.0001 | 0.0015 | 0 | 0.0009 | 0.99 | 0.98 | 0.99 | -0.0002 | -0.0029 | 0.0166 | -0.0014 | 0.0001 | 0 | -0.0141 | 0.0001 | 0.0005 | 0.0001 | 0.0006 | 0.001 | 0.0409 | 0.0017 | -0.0053 | -0.0003 | 0.0491 | 0.1999 | -0.0002 | 0.0027 | -0.0087 | -0.0064 | 0.1625 | 0.028 | 0.0027 | 0.0001 | 0.0015 | 0 | 0.0009
08162024 MB | Sample | 2024-08-26 16:51 | [6 empty] | 0.0001 ! | 0.0018 ! | 0.0046 ! | -0.0029 ! | 0.0001 ! | -0.0001 ! | -0.0144 ! | 0.0000 ! | 0.0008 ! | -0.0002 ! | 0.0008 ! | 0.0011 ! | 0.0277 ! | 0.0013 ! | -0.0076 ! | -0.0002 ! | 0.0054 ! | 0.1339 ! | -0.0002 ! | 0.0024 ! | -0.0058 ! | -0.0035 ! | 0.1194 ! | 0.0148 ! | 0.0035 ! | -0.0002 ! | 0.0000 ! | 0.0002 ! | 0.0003 ! | 0.0013 ! | 0.97 ! | 0.97 ! | 0.99 ! | 0.0001 | 0.0018 | 0.0046 | -0.0029 | 0.0001 | -0.0001 | -0.0144 | 0 | 0.0008 | -0.0002 | 0.0008 | 0.0011 | 0.0277 | 0.0013 | -0.0076 | -0.0002 | 0.0054 | 0.1339 | -0.0002 | 0.0024 | -0.0058 | -0.0035 | 0.1194 | 0.0148 | 0.0035 | 0 | 0.0002 | 0.0003 | 0.0013
08162024 MFB | Sample | 2024-08-26 16:54 | [6 empty] | 0.0980 ! | 4.7972 ! | 2.4586 ! | 0.4941 ! | 2.4539 ! | 0.4817 ! | 4.9762 ! | 1.0023 ! | 5.0186 ! | 2.5107 ! | 2.3870 ! | 4.9622 ! | 4.9440 ! | 0.0001 ! | 5.0841 ! | 1.0047 ! | 0.9835 ! | 5.0818 ! | 2.4711 ! | 1.0582 ! | 2.5385 ! | 2.4661 ! | 2.6229 ! | 2.4164 ! | 0.9737 ! | 1.0036 ! | 1.0096 ! | 1.0026 ! | 2.5189 ! | 0.9784 ! | 0.96 ! | 0.97 ! | 0.98 ! | 0.0979 | 4.7972 | 2.4586 | 0.4941 | 2.4539 | 0.4817 | 4.9762 | 1.0022 | 5.0185 | 2.5104 | 2.387 | 4.9622 | 4.944 | 0.0001 | 5.0833 | 1.0047 | 0.9835 | 5.0818 | 2.4711 | 1.0582 | 2.5382 | 2.4606 | 2.623 | 2.4157 | 0.9733 | 1.0096 | 1.0026 | 2.518 | 0.9784
08162024 MFB - Analyticsal Duplicate | Sample | 2024-08-26 16:56 | [6 empty] | 0.0981 ! | 4.7714 ! | 2.4724 ! | 0.4945 ! | 2.4449 ! | 0.4804 ! | 4.9650 ! | 1.0024 ! | 5.0211 ! | 2.5063 ! | 2.3817 ! | 4.9418 ! | 4.9284 ! | 0.0002 ! | 5.0737 ! | 1.0000 ! | 1.0013 ! | 5.0693 ! | 2.4681 ! | 1.0578 ! | 2.5371 ! | 2.4628 ! | 2.6724 ! | 2.4197 ! | 0.9818 ! | 1.0042 ! | 1.0046 ! | 1.0016 ! | 2.5246 ! | 0.9774 ! | 0.97 ! | 0.97 ! | 0.98 ! | 0.098 | 4.7714 | 2.4724 | 0.4945 | 2.4449 | 0.4804 | 4.965 | 1.0023 | 5.021 | 2.506 | 2.3818 | 4.9418 | 4.9284 | 0.0002 | 5.0728 | 1 | 1.0013 | 5.0693 | 2.4681 | 1.0578 | 2.5369 | 2.4574 | 2.6724 | 2.419 | 0.9814 | 1.0046 | 1.0016 | 2.5237 | 0.9774
5x 08162024 MFB | Sample | 2024-08-26 16:59 | [6 empty] | 0.0194 ! | 0.9528 ! | 0.5033 ! | 0.0973 ! | 0.4873 ! | 0.0968 ! | 1.1014 ! | 0.1996 ! | 1.0016 ! | 0.5024 ! | 0.4701 ! | 0.9910 ! | 0.9784 ! | 0.0077 ! | 0.9979 ! | 0.1998 ! | 0.1983 ! | 1.0942 ! | 0.4886 ! | 0.2111 ! | 0.5040 ! | 0.4898 ! | 0.5494 ! | 0.5002 ! | 0.1922 ! | 0.1970 ! | 0.2005 ! | 0.1973 ! | 0.4994 ! | 0.1944 ! | 0.99 ! | 0.98 ! | 0.99 ! | 0.0194 | 0.9528 | 0.5033 | 0.0973 | 0.4873 | 0.0968 | 1.1014 | 0.1996 | 1.0015 | 0.5023 | 0.4701 | 0.991 | 0.9784 | 0.0077 | 0.9977 | 0.1998 | 0.1983 | 1.0942 | 0.4886 | 0.2111 | 0.5039 | 0.4887 | 0.5494 | 0.5 | 0.1922 | 0.2005 | 0.1973 | 0.4992 | 0.1944
5x 1% NaHCO3 NMC Battery-TM | Sample | 2024-08-26 17:01 | [6 empty] | 0.0002 ! | 0.0804 ! | 0.0086 ! | -0.0006 ! | 0.0010 ! | -0.0002 ! | 0.1845 ! | 0.0000 ! | 0.0003 ! | 0.0001 ! | 0.0016 ! | 0.0196 ! | 0.0928 ! | -0.0064 ! | 0.0112 ! | 0.0001 ! | 0.0119 ! | 466.6491 !o | 1.0209 ! | 0.4067 ! | -0.0033 ! | 0.0450 ! | 0.1394 ! | 0.0218 ! | 0.0262 ! | 0.0001 ! | 0.0004 ! | 0.0006 ! | 0.0003 ! | 0.0019 ! | 0.94 ! | 0.95 ! | 0.96 ! | 0.0002 | 0.0804 | 0.0086 | -0.0006 | 0.001 | -0.0002 | 0.1845 | 0 | 0.0003 | 0.0001 | 0.0016 | 0.0196 | 0.0928 | -0.0064 | 0.0112 | 0.0001 | 0.0119 | 466.6491 | 1.0209 | 0.4067 | -0.0033 | 0.045 | 0.1394 | 0.0218 | 0.0262 | 0.0004 | 0.0006 | 0.0003 | 0.0019
5x 1% NaHCO3 NMC Battery-TM - MS | Sample | 2024-08-26 17:04 | [6 empty] | 0.0199 ! | 0.9541 ! | 0.5124 ! | 0.1006 ! | 0.4883 ! | 0.0960 ! | 1.1555 ! | 0.1993 ! | 1.0049 ! | 0.4966 ! | 0.5107 ! | 0.9758 ! | 1.0543 ! | -0.0060 ! | 0.9711 ! | 0.1921 ! | 0.1878 ! | 416.5691 !o | 0.7397 ! | 0.2540 ! | 0.4935 ! | 0.5680 ! | 0.5127 ! | 0.5198 ! | 0.2117 ! | 0.1985 ! | 0.1977 ! | 0.2015 ! | 0.5111 ! | 0.1978 ! | 0.93 ! | 0.95 ! | 0.96 ! | 0.0199 | 0.9541 | 0.5124 | 0.1006 | 0.4883 | 0.096 | 1.1555 | 0.1992 | 1.0048 | 0.4966 | 0.5107 | 0.9758 | 1.0543 | -0.006 | 0.971 | 0.1921 | 0.1878 | 416.5691 | 0.7397 | 0.2539 | 0.4934 | 0.567 | 0.5127 | 0.5196 | 0.2116 | 0.1977 | 0.2015 | 0.5109 | 0.1978
5x 1% NaHCO3 NMC Battery-TM -MS -Analytical Duplicate | Sample | 2024-08-26 17:06 | [6 empty] | 0.0187 | 0.6318 | 0.4854 | 0.0955 | 0.3204 | 0.0633 | 1.098 | 0.1886 | 0.9502 | 0.47 | 0.4832 | 0.6476 | 0.7085 | -0.0178 | 0.912 | 0.1276 | 0.1825 | 278.8960 o | 0.6987 | 0.2401 | 0.4649 | 0.5357 | 0.5213 | 0.4912 | 0.199 | 0.1875 | 0.13 | 0.1906 | 0.4835 | 0.1865 | 0.95 | 1.23 | 1.24 | 0.0187 | 0.6318 | 0.4854 | 0.0955 | 0.3204 | 0.0633 | 1.098 | 0.1886 | 0.9502 | 0.47 | 0.4832 | 0.6476 | 0.7085 | -0.0178 | 0.9119 | 0.1276 | 0.1825 | 278.896 | 0.6987 | 0.2401 | 0.4649 | 0.535 | 0.5213 | 0.4911 | 0.1989 | 0.13 | 0.1906 | 0.4833 | 0.1865
1% NaHCO3 NMC Battery-TM - Lab Duplicate | Sample | 2024-08-26 17:09 | [6 empty] | -0.0004 | 0.4206 | 0.0072 | 0.0051 | 0.004 | -0.0001 | 0.4985 | 0 | 0.0015 | 0.0005 | 0.0024 | 0.0996 | 0.5282 | 0.0117 | 0.0729 | 0.0012 | 0.0112 | 1766.2352 o | 5.0468 | 2.0369 | -0.0064 | 0.2381 | 0.138 | 0.019 | 0.1297 | 0.0014 | 0.0016 | 0.0014 | 0.0007 | 0.0059 | 0.87 | 0.89 | 0.89 | -0.0004 | 0.4206 | 0.0072 | 0.0051 | 0.004 | -0.0001 | 0.4985 | 0 | 0.0015 | 0.0005 | 0.0024 | 0.0996 | 0.5282 | 0.0117 | 0.0729 | 0.0012 | 0.0112 | 1766.2352 | 5.0468 | 2.0369 | -0.0064 | 0.2381 | 0.138 | 0.019 | 0.1297 | 0.0016 | 0.0014 | 0.0007 | 0.0059
1% NaHCO3- NMC Battery-TM - MS | Sample | 2024-08-26 17:11 | [6 empty] | -0.0001 | 0.1121 | 0.012 | 0.0053 | 0.0465 | -0.0002 | 0.541 | -0.0001 | 0 | 0.0003 | 0.0034 | 0.0679 | 0.6131 | 0.0019 | 0.0758 | 0.0002 | 0.0044 | 1760.8728 o | 1.3672 | 0.1803 | -0.0042 | 0.3097 | 0.1151 | 0.0178 | 0.072 | 0.0031 | 0.002 | 0.0009 | 0.0006 | 0.0051 | 0.87 | 0.89 | 0.9 | -0.0001 | 0.1121 | 0.012 | 0.0053 | 0.0465 | -0.0002 | 0.541 | -0.0001 | 0 | 0.0003 | 0.0034 | 0.0679 | 0.6131 | 0.0019 | 0.0757 | 0.0002 | 0.0044 | 1760.8728 | 1.3672 | 0.1803 | -0.0042 | 0.3097 | 0.1151 | 0.0178 | 0.072 | 0.002 | 0.0009 | 0.0006 | 0.0051
1% NaHCO3- NMC Battery-TM - MS - Analytical Duplicate | Sample | 2024-08-26 17:14 | [6 empty] | -0.0005 | 0.0097 | 0.0037 | -0.003 | 0.0049 | -0.0003 | 0.1786 | -0.0001 | -0.0004 | -0.0007 | -0.0042 | 0.0065 | 0.0788 | -0.0281 | 0.0129 | -0.0003 | 0.0024 | 249.0177 o | 0.4664 | 0.0607 | -0.0134 | 0.0974 | 0.1198 | 0.0093 | 0.0214 | 0.0013 | 0.0002 | 0.0003 | 0.0004 | 0.0017 | 1.13 | 2.07 | 2.08 | -0.0005 | 0.0097 | 0.0037 | -0.003 | 0.0049 | -0.0003 | 0.1786 | -0.0001 | -0.0004 | -0.0007 | -0.0042 | 0.0065 | 0.0788 | -0.0281 | 0.0129 | -0.0003 | 0.0024 | 249.0177 | 0.4664 | 0.0607 | -0.0134 | 0.0974 | 0.1198 | 0.0093 | 0.0214 | 0.0002 | 0.0003 | 0.0004 | 0.0017
Rinse1 | QC | 2024-08-26 17:16 | [6 empty] | -0.0001 | -0.0029 | 0.0005 | -0.0033 | 0 | -0.0001 | -0.0089 | -0.0001 | 0.0002 | -0.0002 | 0.0001 | 0.0005 | 0.0267 | -0.006 | -0.0054 | -0.0004 | 0.0022 | 0.1691 | -0.0013 | -0.0006 | -0.0061 | -0.0029 | 0.1095 | 0.0094 | 0.0017 | -0.0003 | 0 | 0 | 0.0003 | 0.0011 | 0.99 | 0.98 | 0.98 | -0.0001 | -0.0029 | 0.0005 | -0.0033 | 0 | -0.0001 | -0.0089 | -0.0001 | 0.0002 | -0.0002 | 0.0001 | 0.0005 | 0.0267 | -0.006 | -0.0054 | -0.0004 | 0.0022 | 0.1691 | -0.0013 | -0.0006 | -0.0061 | -0.0029 | 0.1095 | 0.0094 | 0.0017 | 0 | 0 | 0.0003 | 0.0011
Rinse1 | QC | 2024-08-26 17:19 | [6 empty] | -0.0004 | -0.0018 | 0.0006 | -0.0032 | 0.0001 | -0.0002 | -0.0074 | 0 | 0.0002 | -0.0002 | -0.0007 | -0.0006 | 0.0355 | -0.0073 | -0.0053 | 0 | 0.0016 | 0.1128 | -0.0014 | 0.005 | -0.0064 | -0.0098 | 0.1175 | 0.0042 | 0.0022 | 0.0001 | 0 | -0.0002 | -0.0001 | 0.001 | 0.99 | 0.98 | 0.99 | -0.0004 | -0.0018 | 0.0006 | -0.0032 | 0.0001 | -0.0002 | -0.0074 | 0 | 0.0002 | -0.0002 | -0.0007 | -0.0006 | 0.0355 | -0.0073 | -0.0053 | 0 | 0.0016 | 0.1128 | -0.0014 | 0.005 | -0.0064 | -0.0098 | 0.1175 | 0.0042 | 0.0022 | 0 | -0.0002 | -0.0001 | 0.001
Rinse | QC | 2024-08-26 17:21 | [6 empty] | -0.0001 | -0.0019 | 0.001 | -0.0038 | 0.0001 | -0.0001 | -0.0132 | -0.0002 | 0.0003 | -0.0003 | 0.0003 | -0.0003 | 0.0406 | -0.0082 | -0.0057 | -0.0003 | 0.001 | 0.0934 | -0.0014 | 0.0004 | -0.0047 | -0.0077 | 0.121 | 0.0032 | 0.0017 | -0.0006 | 0 | -0.0002 | 0.0004 | 0.0012 | 0.99 | 0.98 | 1 | -0.0001 | -0.0019 | 0.001 | -0.0038 | 0.0001 | -0.0001 | -0.0132 | -0.0002 | 0.0003 | -0.0003 | 0.0003 | -0.0003 | 0.0406 | -0.0082 | -0.0057 | -0.0003 | 0.001 | 0.0934 | -0.0014 | 0.0004 | -0.0047 | -0.0077 | 0.121 | 0.0032 | 0.0017 | 0 | -0.0002 | 0.0004 | 0.0012
0.5ppm All (LLQC) | QC | 2024-08-26 17:24 | [6 empty] | 0.4911 | 0.4834 | 0.4767 | 0.4824 | 0.4808 | 0.4846 | 0.4909 | 0.49 | 0.493 | 0.4954 | 0.4634 | 0.4955 | 0.4607 | 0.4929 | 0.476 | 0.4989 | 0.4481 | 0.5765 | 0.4917 | 0.5164 | 0.4977 | 0.4813 | 0.4485 | 0.4815 | 0.4635 | 0.4825 | 0.4968 | 0.4869 | 0.4872 | 0.4826 | 0.99 | 0.98 | 0.99 | 0.491 | 0.4834 | 0.4767 | 0.4824 | 0.4808 | 0.4846 | 0.4909 | 0.49 | 0.4929 | 0.4952 | 0.4634 | 0.4955 | 0.4607 | 0.4929 | 0.4759 | 0.4989 | 0.4481 Q | 0.5765 Q | 0.4917 | 0.5164 | 0.4977 | 0.4786 | 0.4485 Q | 0.4812 | 0.4633 | 0.4968 | 0.4869 | 0.4868 | 0.4826
0.5ppm All (LLQC)-Analytical Duplicate | QC | 2024-08-26 17:26 | [6 empty] | 0.4851 | 0.4867 | 0.4785 | 0.4803 | 0.481 | 0.4871 | 0.4897 | 0.4897 | 0.4909 | 0.4946 | 0.463 | 0.4933 | 0.492 | 0.4892 | 0.4745 | 0.4982 | 0.4702 | 0.5705 Q | 0.4914 | 0.5191 | 0.4975 | 0.4852 | 0.5184 | 0.4942 | 0.462 | 0.4833 | 0.4976 | 0.4856 | 0.4862 | 0.4813 | 1 | 0.98 | 0.99 | 0.4851 | 0.4867 | 0.4785 | 0.4803 | 0.481 | 0.4871 | 0.4897 | 0.4897 | 0.4909 | 0.4944 | 0.463 | 0.4933 | 0.492 | 0.4892 | 0.4744 | 0.4982 | 0.4702 | 0.5705 | 0.4914 | 0.5191 | 0.4974 | 0.4825 | 0.5184 | 0.4939 | 0.4618 | 0.4976 | 0.4856 | 0.4858 | 0.4813
CCS1 (SS 5ppm) | QC | 2024-08-26 17:29 | [6 empty] | 5.0459 | 4.9491 | 5.0689 | 4.9867 | 4.9221 | 4.9418 | 4.9939 | 4.9893 | 4.9977 | 4.9973 | 4.7369 | 4.9595 | 4.89 | 4.8573 | 5.0351 | 4.9756 | 5.1416 | 4.9607 | 4.9696 | 5.3477 | 5.1028 | 5.0309 | 5.3579 | 5.0321 | 4.7254 | 5.0517 | 5.0348 | 5.0173 | 5.0325 | 4.9308 | 0.97 | 0.98 | 0.99 | 5.0454 | 4.9491 | 5.0689 | 4.9867 | 4.9221 | 4.9418 | 4.9939 | 4.9892 | 4.9974 | 4.9959 | 4.737 | 4.9595 | 4.89 | 4.8573 | 5.0343 | 4.9756 | 5.1416 | 4.9607 | 4.9696 | 5.3477 | 5.102 | 5.004 | 5.3579 | 5.0291 | 4.7235 | 5.0348 | 5.0173 | 5.028 | 4.9308
Rinse | QC | 2024-08-26 17:31 | [6 empty] | 0.0001 | 0.001 | 0.0157 | -0.0024 | 0.0001 | -0.0001 | -0.0156 | 0 | 0.0005 | 0 | 0.0002 | 0.0005 | 0.0216 | 0.0019 | -0.0061 | -0.0002 | 0.0461 | 0.0694 | -0.0003 | 0.0017 | -0.0073 | -0.0076 | 0.1235 | 0.0254 | 0.0033 | 0.0023 | 0.0001 | 0.0012 | -0.0002 | 0.0006 | 0.99 | 0.98 | 0.99 | 0.0001 | 0.001 | 0.0157 | -0.0024 | 0.0001 | -0.0001 | -0.0156 | 0 | 0.0005 | 0 | 0.0002 | 0.0005 | 0.0216 | 0.0019 | -0.0061 | -0.0002 | 0.0461 | 0.0694 | -0.0003 | 0.0017 | -0.0073 | -0.0076 | 0.1235 | 0.0254 | 0.0033 | 0.0001 | 0.0012 | -0.0002 | 0.0006
CCB | QC | 2024-08-26 17:34 | [6 empty] | 0 | -0.0039 | 0.0074 | -0.0017 | 0 | -0.0001 | -0.0154 | -0.0001 | 0.0002 | -0.0004 | 0.0008 | -0.0002 | 0.0216 | -0.0077 | -0.0051 | -0.0002 | 0.015 | 0.0644 | 0.0002 | 0.0055 | -0.0043 | -0.0029 | 0.1151 | 0.0124 | 0.0022 | 0.0005 | 0 | 0.0002 | 0.0004 | 0.0004 | 0.99 | 0.98 | 0.99 | 0 | -0.0039 | 0.0074 | -0.0017 | 0 | -0.0001 | -0.0154 | -0.0001 | 0.0002 | -0.0004 | 0.0008 | -0.0002 | 0.0216 | -0.0077 | -0.0051 | -0.0002 | 0.015 | 0.0644 | 0.0002 | 0.0055 | -0.0043 | -0.0029 | 0.1151 | 0.0124 | 0.0022 | 0 | 0.0002 | 0.0004 | 0.0004
5% HNO3 | Sample | 2024-08-26 17:38 | [6 empty] | 0.1027 ! | 4.8192 ! | 2.6053 ! | 0.5215 ! | 2.4429 ! | 0.4671 ! | 5.3022 ! | 0.9868 ! | 4.9343 ! | 2.4658 ! | 2.6397 ! | 4.7432 ! | 5.6634 ! | 0.0008 ! | 4.8255 ! | 0.9449 ! | 0.9928 ! | 1641.1815 !o | 3.6549 ! | 1.2739 ! | 2.4631 ! | 2.9426 ! | 2.5508 ! | 2.5907 ! | 1.0944 ! | 0.9920 ! | 0.9884 ! | 1.0167 ! | 2.5712 ! | 0.9743 ! | 0.87 ! | 0.90 ! | 0.91 ! | 0.1026 | 4.8192 | 2.6053 | 0.5215 | 2.4429 | 0.4671 | 5.3022 | 0.9867 | 4.9342 | 2.4656 | 2.6398 | 4.7432 | 5.6634 | 0.0008 | 4.8247 | 0.9449 | 0.9928 | 1641.1815 | 3.6549 | 1.2738 | 2.4629 | 2.9375 | 2.5508 | 2.59 | 1.0941 | 0.9884 | 1.0167 | 2.5702 | 0.9743
MQ-Blank-03 | Sample | 2024-08-26 17:41 | [6 empty] | 0 | -0.004 | 0.0015 | -0.0036 | 0.0001 | 0 | -0.0106 | 0.0003 | 0.0003 | 0 | 0.0004 | 0.0003 | 0.0488 | 0.0025 | -0.0065 | -0.0001 | 0.0044 | 0.1077 | -0.0005 | 0.0047 | -0.005 | -0.0038 | 0.1133 | 0.0135 | 0.003 | -0.0004 | 0.0001 | 0.0001 | 0.0004 | 0.0014 | 1.01 | 0.98 | 1 | 0 | -0.004 | 0.0015 | -0.0036 | 0.0001 | 0 | -0.0106 | 0.0003 | 0.0003 | 0 | 0.0004 | 0.0003 | 0.0488 | 0.0025 | -0.0065 | -0.0001 | 0.0044 | 0.1077 | -0.0005 | 0.0047 | -0.005 | -0.0038 | 0.1133 | 0.0135 | 0.003 | 0.0001 | 0.0001 | 0.0004 | 0.0014