## README_Logsheets
This Excel Book contains all QAQC log sheets related to Roof Rainwater Harvesting water quality monitoring in ORD Athens, Georgia location.
Note that we used the data from February 28 through August 1, 2024.
Sheet Name | Description
README_Logsheets | This is the current Sheet that provides the description of all the log sheets.
DEPTH Data Logsheet | This is the manual depth measurement log sheet.
Operational Check Data Logsheet | This is the operational check log sheet.
Cleaning equipment Logsheet | This is the equipment cleaning log sheet.
Data Sonde Cali Form | This is the Aqua TROLL 600 (the QAQC) sonde calibration, pre-calibration, post-calibration log sheet.
Reformatted Sonde Cali Form | This sheet contains the same information as the Data Sonde Cali Form sheet, but formatted differently.
QC Acceptance criteria | This sheet contains the QAQC Acceptance criteria.

## DEPTH Data Logsheet
Appendix H |  | Data collected by: | SRG; KW; JM |  |  | SRG = SANTOSH RAJ GHIMIRE; KW = KURT WOLFE; JM = JON MCMAHAN (UNTIL 4/24/2024).
Depth measurement log sheet:
Table 3. Water depth log sheet to be completed on site. |  |  | 1 in = 25.4 mm
Date | Time AM/PM (EST) | Water depth using metal ruler (inch) | Water depth using metal ruler (mm) | Data collection by (Initial) | Note
2024-02-28 | 10:00 | Not available | Not available |  | Performed Equipment calibration only
2024-04-08 | 10:00 | Not available | Not available |  | Performed Equipment calibration only
2024-04-24 | 11:10 | 27 | 685.8
2024-04-25 | 11:54 | 23 | 584.2
2024-04-25 | 12:00 | 19 | 482.6
2024-04-29 | 11:36 | 18 | 457.2
2024-04-29 | 11:42 | 14 | 355.6
2024-05-01 | 12:20 | 17 | 431.8
2024-05-01 | 12:26 | 13 | 330.2
2024-05-08 | 10:30 | 27.5 | 698.5
2024-05-22 | 9:06 | 27 | 685.8
2024-05-22 | 9:55 | 3.5 | 88.9
2024-06-05 | 10:15 | 29.25 | 742.95 |  | Full barrel
2024-06-20 | 9:45 | 3.25 | 82.55
2024-07-02 | 11:15 | 3.25 | 82.55
2024-07-18 | 10:34 | 0 | 0 |  | New barrel installed
2024-08-01 | 15:37 | 33 | 838.2
2024-08-15 | 11:35 | 33 | 838.2
2024-08-29 | 11:14 | 33 | 838.2
2024-09-12 | 11:54 | 30.5 | 774.7
2024-09-25 | 11:11 | 32.75 | 831.85
2024-10-09 | 11:48 | 33 | 838.2
Formulas:
Appendix H = [Appendix H, row above]+14

## Operational Check Data Logsheet
|  | Data collected by: | SRG; KW; JM |  |  | SRG = SANTOSH RAJ GHIMIRE; KW = KURT WOLFE; JM = JON MCMAHON (UNTIL 4/24/2024).
Appendix I
Operational checks log sheet:
Table 4. Operational check log sheet to be completed on site.
Date | Time AM/PM (EST) | Operation check parameter | Result within ± 10% of the standard (Yes/No) | Corrective Actions | Note
2024-02-28 | 10:00 | InSitu Sonde Conductivity and Temperature | Yes | No | This is the first calibration performed at EPA Region 4 laboratory.
2024-04-08 | 10:00 | InSitu Sonde Conductivity and Temperature | Yes | No | This is the calibration performed on site at EPA ORD in Athens, GA.
2024-04-24 | 11:00 | InSitu Sonde Conductivity and Temperature | Yes | No
2024-05-08 | 10:15 | InSitu Sonde Conductivity and Temperature | Yes | No | Operational check for conductivity only. Thermistor checked against a NIST-traceable thermometer on an annual basis.
2024-05-22 | 9:06 | InSitu Sonde Conductivity and Temperature | Yes | No | Only for conductivity. Thermistor checked against a NIST-traceable thermometer on an annual basis.
2024-06-05 | 10:30 | InSitu Sonde Conductivity and Temperature | Yes | No | Only for conductivity. Thermistor checked against a NIST-traceable thermometer on an annual basis.
2024-06-20 | 9:45 | Only clean up of equipment | NA | NA | Only for conductivity. Thermistor checked against a NIST-traceable thermometer on an annual basis.
2024-07-02 | 11:33 | InSitu Sonde Conductivity and Temperature | Yes | No | Only for conductivity. Thermistor checked against a NIST-traceable thermometer on an annual basis.
2024-07-18 | 10:34 | InSitu Sonde Conductivity and Temperature | Yes | No | Only for conductivity. Thermistor checked against a NIST-traceable thermometer on an annual basis.
2024-08-01 | 15:37 | InSitu Sonde Conductivity and Temperature | Yes | No | Only for conductivity. Thermistor checked against a NIST-traceable thermometer on an annual basis.
2024-08-15 | 11:36 | InSitu Sonde Conductivity and Temperature | Yes | No | Only for conductivity. Thermistor checked against a NIST-traceable thermometer on an annual basis.
2024-08-29 | 11:14 | InSitu Sonde Conductivity and Temperature | Yes | Yes
2024-09-12 | 11:54 | InSitu Sonde Conductivity and Temperature | Yes | Yes
2024-09-25 | 11:12 | InSitu Sonde Conductivity and Temperature | Yes | Yes
2024-10-09 | 11:56 | InSitu Sonde Conductivity and Temperature | Yes | Yes
Formulas:
Date = [Date, row above]+14

## Cleaning equipment Logsheet
|  | Data collected by: | SRG; KW; JM | SRG = SANTOSH RAJ GHIMIRE; KW = KURT WOLFE; JM = JON MCMAHON (UNTIL 4/24/2024).
Appendix J
Table 5. Monthly cleaning log sheet.
Date | Time AM/PM (EST) | Item cleaned | Note
2024-02-28 | 10:00 | InSitu Sonde | This is the first calibration performed at EPA Region 4 laboratory.
2024-04-08 | 10:00 | InSitu Sonde | This is the calibration performed on site at EPA ORD in Athens, GA.
2024-04-24 | 11:00 | InSitu Sonde | Order cleaning items as neded
2024-05-08 | 10:15 | InSitu Sonde
2024-05-22 | 9:06 | InSitu Sonde | Also drain the rain barrel to minimum level at 3.5 inch
2024-06-05 | 10:30 | InSitu Sonde
2024-06-20 | 10:30 | InSitu Sonde | Only clean up performed
2024-07-02 | 11:15 | InSitu Sonde | None
2024-08-01 | 15:37 | InSitu Sonde | Cleaned
2024-08-15 | 11:36 | InSitu Sonde | Cleaned
2024-08-29 | 11:24 | InSitu Sonde | Cleaned
2024-09-12 | 12:00 | InSitu Sonde | Cleaned
2024-09-25 | 11:25 | InSitu Sonde | Cleaned
2024-10-09 | 11:48 | InSitu Sonde | Cleaned
Formulas:
Date = [Date, row above]+14

## Data Sonde Cali Form
OPERATIONAL CHECK/CALIBRATION OF AQUA TROLL 600 SONDE
Note: On 6/20/2024 10:30:00 AM The operational checks and clean up performed but no Sonde data recorded on site.
| Data Sonde Calibration Form | Data collected by: | JM; SRG; KW | JM; SRG; KW | JM; SRG; KW | SRG; KW | SRG; KW | SRG; KW | SRG; KW | SRG; KW | SRG; KW | SRG; KW | SRG; KW | SRG; KW | SRG; KW | SRG; KW
| Date and Time | Begin | 2024-02-28 10:00 | 2024-04-08 10:00 | 2024-04-24 11:00 | 2024-05-08 10:10 | 2024-05-22 09:06 | 2024-06-05 10:33 | 2024-07-02 11:33 | 2024-07-18 10:34 | 2024-08-01 15:45 | 2024-08-15 11:48 | 2024-08-29 11:24 | 2024-09-12 12:03 | 2024-09-25 11:26 | 2024-10-09 11:57
|  | End | 2024-04-08 10:00 | 2024-04-24 11:00 | 2024-05-08 10:10 | 2024-05-22 09:06 | 2024-06-05 10:33 | 2024-07-02 11:33 | 2024-07-18 10:34 | 2024-08-01 15:45 | 2024-08-15 11:48 | 2024-08-29 11:24 | 2024-09-12 12:03 | 2024-09-25 11:26 | 2024-10-09 11:57
| Specific Conductance 1417 (µS/cm) | PreCal | 1422 | 1425 | 1428 | 1424.4 | 1438.1 | 1455.2 | 1439.8 | 1452.1 | 1437.4 | 1441.3 | 1364.2 | 1437.9 | 1421.7 | 1445.6
|  | PostCal | 1417 | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%) | Not needed (within ±10%)
|  | End Check | 1425 | 1428 | 1424.4 | 1438.1 | 1455.2 | 1439.8 | 1452.1 | 1437.4 | 1441.3 | 1364.2 | 1437.9 | 1421.7 | 1445.6
| Temperature (°C) verification NIST/Thermistor | Begin (NIST) | 21 | 21.2 | 21 | NA | NA | NA | NA | NA | NA | NA | 23 | 21 | 24 | 22
|  | Begin (Hydros 21 Thermistor) | 21.55 | 21.14 | 21.47 | 23.11 | 22.69 | 24.77 | 23.82 | 24.29 | 24.76 | 23.5 | 23.92 | 21.72 | 24.01 | 23
|  | End (NIST) | 21.2 | 21 | NA | NA | NA | NA | NA | NA | NA | 23 | 21 | 24 | 22
|  | End (Hydros 21 Thermistor) | 21.14 | 21.47 | 23.11 | 22.69 | 24.77 | 23.82 | 24.29 | 24.76 | 23.5 | 23.92 | 21.72 | 24.01 | 23
|  | 0.85 | 0.9 | 0.95 | 1 | 1.05 | 1.1 | 1.15
|  | 1204.45 | 1275.3 | 1346.15 | 1417 | 1487.85 | 1558.7 | 1629.55
| NOTES: | SRG = SANTOSH RAJ GHIMIRE; KW = KURT WOLFE; JM = JON MCMAHAN (JM, UNTIL 4/24/2024).
The "Begin" is the current date and time of calibration or operational check.
For all dates, the "End" Date and Time is the "Begin" Date and Time of succeeding calibration/operation check.
For all dates, the "End Check" Specific Conductance value is the "PreCal" Specific Conductance value of the succeeding "Begin" calibration/operation check value.
Calibration is needed only if operational check fails.
NaCl | Specific conductivity (µS/cm)
Standard | 1417
10% of Standard | 141.7
Standard + 10% | 1558.7
Standard - 10% | 1275.3
Formulas:
JM; SRG; KW = [JM; SRG; KW, row above]
SRG; KW = [SRG; KW, row above]

## Reformatted Sonde Cali Form
Reformatted data sonde calibration form
Date and Time |  | Specific Conductance 1417 (µS/cm) |  |  | Temperature (°C) verification NIST/Thermistor
Begin | End | PreCal | PostCal | End Check | Begin (NIST) | Begin (Hydros 21) | End (NIST) | End (Hydros 21)
2024-02-28 10:00 | 2024-04-08 10:00 | 1422 | 1417 | 1425 | 21 | 21.55 | 21.2 | 21.14
2024-04-08 10:00 | 2024-04-24 11:00 | 1425 | Not needed (within ±10%) | 1428 | 21.2 | 21.14 | 21 | 21.47
2024-04-24 11:00 | 2024-05-08 10:10 | 1428 | Not needed (within ±10%) | 1424.4 | 21 | 21.47 | NA | 23.11
2024-05-08 10:10 | 2024-05-22 09:06 | 1424.4 | Not needed (within ±10%) | 1438.1 | NA | 23.11 | NA | 22.69
2024-05-22 09:06 | 2024-06-05 10:33 | 1438.1 | Not needed (within ±10%) | 1455.2 | NA | 22.69 | NA | 24.77
2024-06-05 10:33 | 2024-07-02 11:33 | 1455.2 | Not needed (within ±10%) | 1439.8 | NA | 24.77 | NA | 23.82
2024-07-02 11:33 | 2024-07-18 10:34 | 1439.8 | Not needed (within ±10%) | 1452.1 | NA | 23.82 | NA | 24.29
2024-07-18 10:34 | 2024-08-01 15:45 | 1452.1 | Not needed (within ±10%) | 1437.4 | NA | 24.29 | NA | 24.76
2024-08-01 15:45 | 2024-08-15 11:48 | 1437.4 | Not needed (within ±10%) | 1441.3 | NA | 24.76 | NA | 23.5
2024-08-15 11:48 | 2024-08-29 11:24 | 1441.3 | Not needed (within ±10%) | 1364.2 | NA | 23.5 | NA | 23.92
2024-08-29 11:24 | 2024-09-12 12:03 | 1364.2 | Not needed (within ±10%) | 1437.9 | 23 | 23.92 | 21 | 21.72
2024-09-12 12:03 | 2024-09-25 11:26 | 1437.9 | Not needed (within ±10%) |  | 21 | 21.72
2024-09-25 11:26 | 2024-10-09 11:57 | 1421.7 | Not needed (within ±10%) | 1445.6 | 24 | 24.01 | 22 | 23
2024-10-09 11:57 |  | 1445.6 | Not needed (within ±10%) |  | 22 | 23
Formulas:
End = [Begin, row below]

## QC Acceptance criteria
Table 2. Quality control acceptance criteria.
Data Quality Step | Frequency | Reported Value | Acceptance | Corrective
|  |  | Criteria | Actions
Conductivity | Pre-Deployment | µs/cm | Pass in software | Do not deploy for conductivity if unable to pass calibration, purchase a new conductivity probe
Conductivity Calibration and Operational Check | Minimum of two weeks in addition to Post Deployment | µs/cm | ±10% (1275-1559 microsiemens/cm) | Flag data
Temperature | Pre and post deployments verification checks (two weeks) | °C | ±4°C | Factory Service
Water Depth | Pre and Post Deployment as needed | Inches or mm | ±0.50 inches measured within the 15-minutes time steps due to hand measurement error/irregular Barrel bottom or consistent sensor data offset | Flag sensor data or apply offset correction in software