Electronic Reference for NMR

(10-2-2024)

Bruker Manual TopSpin ERETIC 2 (Version 001)

(http://www.nmragenda-leiden.nl/static/Bruker_manual_4.1.4/topspin_pdf/eretic2.pdf)

http://www2.chem.uic.edu/nmr/downloads/BASHCD10/pdf/z31725.pdf

BR600

- Preparation of DSS reference standard: prepare 1 mL of 20 µM DSS/100 mM Phosphate buffer (pH=7.4) in 99.9% D2O.

- Add 0.6 mL DSS reference standard into a 5 mm NMR tube.

- Create a date file: loading a cpmg1d date file and then type new: Use current parameters.

- Type epa_setup to tune (atma exact), shim (topshim ordmax=6) and calibrate p1 and O1.

- Type new:  calibration for Name, CMC_PROTON for Experiment, D2O for Set solvent. Click OK.

- Edit acqupar (ACQUPARS). Change pulse program to cpmgpr1d and SET PULPROG.

- getprosol 1H <p1 value> -7.45 and set O1.

- Set sw=15 ppm, aq=2s, d1=2s, d20=1ms, L4=128, rg=203, ns=64 and ds=8.

- zg

- Copy all parameters to a new date file (new) and rerun.

- Copy all parameters to a new date file (new) and ns=32. Run again.

- Copy all parameters to a new date file (new) and ns=16. Run again.

- Data processing: lb=0.3, efp, .ph, abs 3 for all 4 data files.

- Measure linewidth of DSS for the spectra with ns=64: zoom the DSS peak and type peakw. Record the line width.

- Integrate DSS peak (from 0.03 to -0.03 ppm).

- Right click on DSS peak, then Eretic  Define as Eretic Reference.

- Enter 0.02 for Concentration [mmol/l], 0.6 for Sample volume [ml], 9 for Number of atoms, 0.03 to - 0.03 for Region [ppm], DSS for Molecule name. Click OK. Exit integrate menu.

- Edit processing parameters (PROCPARS). Click submenu E button.

- Enter, 0.0 for Position of signal, linewidth (2.241 Hz), 203 for rg.

- To test the correct reference integral, click Q

- Select eretic file (calibration\1\pdata\1\eretic) for RORIGIN, enter linewidth, 0.5 for position. Click A.

- If the peak at 0.5 ppm does not have same height as at 0 ppm, reprocess (efp abs 3) and adjust reference integral: PROCPAR, change value for EINT (integral of Eretic signal).

- Once both peaks have same intensity (7.2e+06), change REINT in eretic file to the value (7.2e+6): integrate menu, then do steps 15, 16 and 17.

- Load the 2nd DSS spectrum (ns=64). Then do steps 15-23.

- Test the eretic references using the other two spectra (ns=32 and ns=16).

- Load the spectrum to be calibrated.

- Click PROCPARS, SPECTYP=PROTON. Click submenu E.

- Select the correct path for the eretic reference file:

D:\nmr data17 sept2023-dec2024\calibration\1\pdata\1\eretic

- EINT=7.2e+6, ELW=2.241Hz, EPOS=0. Click submenu A.