## ReadMe
Figure S2. I-CIMS calibration curves for 4:2 FTOH (red), 6:2 FTOH (orange), 8:2 FTOH (light teal), and 10:2 FTOH (dark teal). Open circles represent the mean of the triplicate standard injections at each level and error bars represent one standard deviation of the mean. Solid lines represent an unweighted linear fit to the mean calibration data.
Table 3. Comparison of I-CIMS sensitivities to FTOHs reported in the literature to this study.
Calibration injections of combined FTOH standards at 0.2 and 2 ng/µL in ethyl acetate were performed at volumes ranging from 1-10 µL
The mass of each FTOH that was introduced by each calibration injection was converted to a mixing ratio in ppbv (see "MR Calculations")
The mean I-CIMS signal resulting from each calibration injection was plotted as a function of the mixing ratios of the injections to generate calibration curves.
The standard deviation of the I-CIMS signal at each calibration level was used to add error bars to the plots.
A linear fit was applied to the data, the slope of which represented the instrument's sensitivity to the analyte in terms of Hz/ppbv
The sensitivities were divided by 1000 to convert them to Hz/pptv and were summarized in Table 3.
Abbreviations |  | Data Dictionary
avg | average | final_MR_4_2 | 4:2 FTOH calibration mixing ratios (ppbv)
cal | calibration | final_signal_avg_4_2 | 4:2 FTOH calibration average I-CIMS signal
FTOH | fluorotelomer alcohol | final_signal_sdev_4_2 | 4:2 FTOH calibration I-CIMS signal standard deviation
g | grams | final_MR_6_2 | 6:2 FTOH calibration mixing ratios (ppbv)
Hz | Hertz | final_signal_avg_6_2 | 6:2 FTOH calibration average I-CIMS signal
I-CIMS | iodide chemical ionization mass spectrometer | final_signal_sdev_6_2 | 6:2 FTOH calibration I-CIMS signal standard deviation
L | liters | final_MR_8_2 | 8:2 FTOH calibration mixing ratios (ppbv)
min | minutes | final_signal_avg_8_2 | 8:2 FTOH calibration average I-CIMS signal
mL | milliliter | final_signal_sdev_8_2 | 8:2 FTOH calibration I-CIMS signal standard deviation
mol | mole | final_MR_10_2 | 10:2 FTOH calibration mixing ratios (ppbv)
MR | mixing ratio | final_signal_avg_10_2 | 10:2 FTOH calibration average I-CIMS signal
µg | microgram | final_signal_sdev_10_2 | 10:2 FTOH calibration I-CIMS signal standard deviation
µL | microliter
ng | nanogram
ppbv | parts per billion by volume
pptv | parts per trillion by volume
s | second(s)
sdev | standard deviation

## Raw Data
final_MR_4_2 | final_signal_avg_4_2 | final_signal_sdev_4_2 | final_MR_6_2 | final_signal_avg_6_2 | final_signal_sdev_6_2 | final_MR_8_2 | final_signal_avg_8_2 | final_signal_sdev_8_2 | final_MR_10_2 | final_signal_avg_10_2 | final_signal_sdev_10_2
0.6179 | 2295.77 | 137.477 | 0.4482 | 4333.7998 | 767.18 | 0.3516 | 3227.4199 | 437.059 | 0.2893 | 2941.02 | 396.252
1.5448 | 5431.9102 | 900.457 | 1.1205 | 10683 | 154.092 | 0.879 | 8067.9399 | 615.712 | 0.72325 | 7425.3501 | 833.324
3.0895 | 10349.4 | 902.287 | 2.241 | 19464 | 832.976 | 1.758 | 14895.9 | 1259.28 | 1.4465 | 14128 | 918.081
6.1792 | 21027.199 | 2226.6699 | 4.4819 | 40616.102 | 1677.4 | 3.5161 | 30079.801 | 3508.22 | 2.8927 | 26388.301 | 3064.6001
15.4481 | 50832.898 | 2513.4199 | 11.2049 | 98835.797 | 2462.6299 | 8.7902 | 74136.203 | 4231.9399 | 7.2318 | 66545.797 | 4393.1299
30.8962 | 100653 | 6878.3101 | 22.4097 | 196309 | 5700.9102 | 17.5803 | 146834 | 5920.1201 | 14.4636 | 133524 | 3829.1299
61.7924 | 197989 | 16457.199 | 44.8194 | 391672 | 4593.9502 | 35.1606 | 289275 | 19369.301 | 28.9273 | 260779 | 15419.9

## MR Calculations
Compound | 4:2 FTOH 0.2 ng/µL | 4:2 FTOH 2 ng/µL | 6:2 FTOH 0.2 ng/µL | 6:2 FTOH 2 ng/µL | 8:2 FTOH 0.2 ng/µL | 8:2 FTOH 2 ng/µL | 10:2 FTOH 0.2 ng/µL | 10:2 FTOH 2 ng/µL
Molecular mass (g/mol) | 264.09 | 264.09 | 364.1 | 364.1 | 464.12 | 464.12 | 564.13 | 564.13
INDIVIDUAL CONCENTRATED STOCK SOLNS
Actual Concentration (ng/µL) | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | Note: Wellington standards at 50 µg/mL (ng/µL)
CALIBRATION STANDARD
Aliquot of stock soln volume (µL) | 4 | 40 | 4 | 40 | 4 | 40 | 4 | 40
Aliquot of first dilution added to 1 mL solvent (ng/µL) | 0.2 | 2 | 0.2 | 2 | 0.2 | 2 | 0.2 | 2
INJECTION CONCENTRATIONS
Mass sampled per µL injection (ng) | 0.2 | 2 | 0.2 | 2 | 0.2 | 2 | 0.2 | 2
Mass sampled per µL injection (g) | 2e-10 | 2e-09 | 2e-10 | 2e-09 | 2e-10 | 2e-09 | 2e-10 | 2e-09
Mol sampled per µL injection | 7.5732e-13 | 7.5732e-12 | 5.493e-13 | 5.493e-12 | 4.3092e-13 | 4.3092e-12 | 3.5453e-13 | 3.5453e-12
Vol sampled per µL injection (L) | 1.8538e-11 | 1.8538e-10 | 1.3446e-11 | 1.3446e-10 | 1.0548e-11 | 1.0548e-10 | 8.6782e-12 | 8.6782e-11
IMR flow rate (L/min) | 1.8 | 1.8 | 1.8 | 1.8 | 1.8 | 1.8 | 1.8 | 1.8
IMR flow rate (L/s) | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
Summed 1 Hz mixing ratio per µL injection (ppbv) | 0.61792 | 6.1792 | 0.44819 | 4.4819 | 0.35161 | 3.5161 | 0.28927 | 2.8927
Summed mixing ratio (ppbv) - 1 µL injection | 0.61792 | 6.1792 | 0.44819 | 4.4819 | 0.35161 | 3.5161 | 0.28927 | 2.8927
Summed mixing ratio (ppbv) - 2.5 µL injection | 1.5448 | 15.4481 | 1.1205 | 11.2048 | 0.87902 | 8.7902 | 0.72318 | 7.2318
Summed mixing ratio (ppbv) - 5 µL injection | 3.0896 | 30.8962 | 2.241 | 22.4097 | 1.758 | 17.5803 | 1.4464 | 14.4636
Summed mixing ratio (ppbv) - 10 µL injection |  | 61.7924 |  | 44.8194 |  | 35.1606 |  | 28.9273

## Cal Curve Fits
fit_final_signal_avg_4_2= W_coef[0]+W_coef[1]*x | fit_final_signal_avg_6_2= W_coef[0]+W_coef[1]*x | fit_final_signal_avg_8_2= W_coef[0]+W_coef[1]*x | fit_final_signal_avg_10_2= W_coef[0]+W_coef[1]*x
W_coef={847.52,3200.2} | W_coef={743.95,8725.3} | W_coef={1022.2,8221.6} | W_coef={1000.8,9017.1}
V_chisq= 2.25747e+06;V_npnts= 7;V_numNaNs= 0;V_numINFs= 0; | V_chisq= 1.53599e+06;V_npnts= 7;V_numNaNs= 0;V_numINFs= 0; | V_chisq= 3.87565e+06;V_npnts= 7;V_numNaNs= 0;V_numINFs= 0; | V_chisq= 6.61147e+06;V_npnts= 7;V_numNaNs= 0;V_numINFs= 0;
V_startRow= 0;V_endRow= 6;V_q= 1;V_Rab= -0.635191;V_Pr= 0.999964; | V_startRow= 0;V_endRow= 6;V_q= 1;V_Rab= -0.635193;V_Pr= 0.999994; | V_startRow= 0;V_endRow= 6;V_q= 1;V_Rab= -0.635192;V_Pr= 0.999971; | V_startRow= 0;V_endRow= 6;V_q= 1;V_Rab= -0.635195;V_Pr= 0.999939;
V_r2= 0.999927; | V_r2= 0.999987; | V_r2= 0.999941; | V_r2= 0.999877;
W_sigma={329,12.2} | W_sigma={271,13.9} | W_sigma={431,28.2} | W_sigma={563,44.7}
Coefficient values ± one standard deviation | Coefficient values ± one standard deviation | Coefficient values ± one standard deviation | Coefficient values ± one standard deviation
a =847.52 ± 329 | a =743.95 ± 271 | a =1022.2 ± 431 | a =1000.8 ± 563
b =3200.2 ± 12.2 | b =8725.3 ± 13.9 | b =8221.6 ± 28.2 | b =9017.1 ± 44.7