## ReadMe
Table 1. TD-GC/MS MDLs and RLs and I-CIMS LODs and LOQs for the four target FTOHs.
Abbreviations |  | TD-GC/MS MDLs and RLs
Acq | acquisition | Eight PFAS TD tubes were spiked with 1 µL of the 0.25 ng/µL FTOH standards and four were spiked with I µL of the 0.75 ng/µL FTOH standard.
Avg | average | The spiked TD tubes were analyzed on the TD-GC/MS using the same method as used for the calibration standards and quantified using the calibration curves for the target FTOHs in Figure S1.
BG | background | Bromochloromethane and 1,4-dichlorobenzene were used as internal standards to monitor retention times and injection quality.
Cal(s) | calibration(s) | The four 0.75 ng standards were used as QC checks to ensure the calibration remained valid (within 30% of the original calibration curve) for the MDL test.
Calc | calculated, calculation | The first MDL injection was excluded from analysis due to over 400% difference in response for 8:2 FTOH compared to the calibration curve.
CFR | code of federal regulations | The remaining seven MDL injections were used to calculate the MDL according to the Code of Federal Regulations (40 CFR 136 Appendix B): MDL = t*s where t is the Student's t value for a 99% confidence level and n-1 degrees of freedom (3.143 for n=7) and s is the standard deviation of the calculated amounts in the seven replicates.
Cond | conditioned
cps | counts per second | The RLs were defined as 3x the MDLs.
Diff | difference | MDL and RL values in ng were converted to units of µg/m3 and pptv (see "Unit Conversions")
ext vol | extended volume
FTOH | fluorotelomer alcohol | I-CIMS LODs and LOQs
g | grams | The I-CIMS LODs and LOQs were calculated according to Bertram et al. 2011 (doi.org/10.5194/amt-4-1471-2011) using the instrument sensitivities determined by the calibrations in Figure S2 and the instrument background signal while sampling UZA prior to calibration injections.
GC | gas chromatograph/gas chromatography
I-CIMS | iodide chemical ionization mass spectrometer | See "LOD+LOQ Calc Explanation" for further details on the calculations.
Info | information | See "Raw I-CIMS BG Data" for the raw I-CIMS data used to calculate the LOD and LOQ.
L | liter | LOD and LOQ values in pptv were converted to units of µg/m3 (see "Unit Conversions")
LOD | limit of detection
LOQ | limit of quantitation
m | meter
MDL | method detection limit
Min | minute(s)
Misc | miscellaneous
mol | mole
mol wt | molecular weight
MS | mass spectrometer/mass spectrometry
µg | microgram
µL | microliter
ng | nanograms
PFAS | per- and polyfluorinated alkyl substances
pptv | parts per trillion by volume
QC | quality control
Ret Time | retention time
RL | reporting limit
sec | second(s)
SIM | selected ion monitoring
Std Dev | standard deviation
STP | standard temperature and pressure (25 °C and 1 atm)
TD | thermal desorption
UZA | ultra-zero air
Vol | volume

## TD Internal Standards
Compound | [6 empty] | Bromochloromethane | Bromochloromethane | Bromochloromethane | Bromochloromethane | 1,4 Difluorobenzene | 1,4 Difluorobenzene | 1,4 Difluorobenzene | 1,4 Difluorobenzene
Labels | Date Acquired | Data File Name | Acq. Method File | Sample Name | Misc Info | Vial Number | (# 1) Name | (# 1) Ret Time | (# 1) Amount | (# 1) Target Response | (# 2) Name | (# 2) Ret Time | (# 2) Amount | (# 2) Target Response
| 2025-04-16 11:29 | 20250415037.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902607 | PFAS Blank Cond. 4/15/25 | 31 | Bromochloromethane | 3.34 | 0 | 109824 | 1,4 Difluorobenzene | 4.69 | 0 | 445496
| 2025-04-15 17:25 | 20250415007.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | Bromochloromethane | 3.34 | 0 | 112751 | 1,4 Difluorobenzene | 4.69 | 0 | 458102
| 2025-04-15 18:01 | 20250415008.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902601 | 0.75 ng FTOHs, spiked 4/15/25 | 12 | Bromochloromethane | 3.34 | 0 | 111876 | 1,4 Difluorobenzene | 4.69 |  | 452467
| 2025-04-15 18:40 | 20250415009.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902608 | 0.75 ng FTOHs, spiked 4/15/25 | 13 | Bromochloromethane | 3.33 | 0 | 110093 | 1,4 Difluorobenzene | 4.69 |  | 445529
| 2025-04-15 19:18 | 20250415010.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | Bromochloromethane | 3.34 | 0 | 109404 | 1,4 Difluorobenzene | 4.69 |  | 440732
| 2025-04-15 19:52 | 20250415011.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | Bromochloromethane | 3.34 | 0 | 109546 | 1,4 Difluorobenzene | 4.69 |  | 441204
| 2025-04-15 20:27 | 20250415012.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | Bromochloromethane | 3.34 | 0 | 112903 | 1,4 Difluorobenzene | 4.69 |  | 455821
| 2025-04-15 21:04 | 20250415013.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902630 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 14 | Bromochloromethane | 3.34 | 0 | 113113 | 1,4 Difluorobenzene | 4.69 |  | 456237
| 2025-04-15 21:41 | 20250415014.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902633 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 15 | Bromochloromethane | 3.34 | 0 | 109961 | 1,4 Difluorobenzene | 4.69 |  | 444638
| 2025-04-15 22:18 | 20250415015.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902612 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 16 | Bromochloromethane | 3.34 | 0 | 109563 | 1,4 Difluorobenzene | 4.69 |  | 442516
| 2025-04-15 22:55 | 20250415016.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902637 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 17 | Bromochloromethane | 3.34 | 0 | 109738 | 1,4 Difluorobenzene | 4.69 |  | 443016
| 2025-04-15 23:31 | 20250415017.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902617 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 18 | Bromochloromethane | 3.34 | 0 | 108483 | 1,4 Difluorobenzene | 4.69 |  | 438819
| 2025-04-16 00:08 | 20250415018.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902614 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 19 | Bromochloromethane | 3.34 | 0 | 108701 | 1,4 Difluorobenzene | 4.69 |  | 440825
| 2025-04-16 00:44 | 20250415019.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902604 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 20 | Bromochloromethane | 3.34 | 0 | 109505 | 1,4 Difluorobenzene | 4.69 |  | 441553
| 2025-04-16 01:20 | 20250415020.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902606 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 21 | Bromochloromethane | 3.34 | 0 | 108707 | 1,4 Difluorobenzene | 4.69 |  | 439290
| 2025-04-16 01:56 | 20250415021.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | Bromochloromethane | 3.34 | 0 | 108676 | 1,4 Difluorobenzene | 4.69 |  | 439163
| 2025-04-16 02:31 | 20250415022.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902611 | 0.75 ng FTOHs, spiked 4/15/25 | 22 | Bromochloromethane | 3.34 | 0 | 110322 | 1,4 Difluorobenzene | 4.69 |  | 443483
| 2025-04-16 03:08 | 20250415023.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902615 | 0.75 ng FTOHs, spiked 4/15/25 | 23 | Bromochloromethane | 3.34 | 0 | 108876 | 1,4 Difluorobenzene | 4.69 |  | 442070
| 2025-04-16 03:44 | 20250415024.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902640 | PFAS Blank Cond. 4/15/25 | 24 | Bromochloromethane | 3.34 | 0 | 109147 | 1,4 Difluorobenzene | 4.69 |  | 442336
| 2025-04-16 04:20 | 20250415025.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902640 | PFAS Blank Cond. 4/15/25 | 24 | Bromochloromethane | 3.34 | 0 | 108503 | 1,4 Difluorobenzene | 4.69 |  | 439765
| 2025-04-16 04:56 | 20250415026.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902640 | PFAS Blank Cond. 4/15/25 | 24 | Bromochloromethane | 3.34 | 0 | 108782 | 1,4 Difluorobenzene | 4.69 |  | 440467

## TD MDLs+RLs
Compound Name | [6 empty] | 4:2 FTOH | 4:2 FTOH | 4:2 FTOH | 4:2 FTOH |  |  | 6:2 FTOH | 6:2 FTOH | 6:2 FTOH | 6:2 FTOH |  |  | 8:2 FTOH | 8:2 FTOH | 8:2 FTOH | 8:2 FTOH |  |  | 10:2 FTOH | 10:2 FTOH | 10:2 FTOH | 10:2 FTOH
Labels | Date Acquired | Data File Name | Acq. Method File | Sample Name | Misc Info | Vial Number | (# 3) Name | (# 3) Ret Time | (# 3) Amount | (# 3) Target Response | Calc Amount (ng) | % Diff | (# 4) Name | (# 4) Ret Time | (# 4) Amount | (# 4) Target Response | Calc Amount (ng) | % Diff | (# 5) Name | (# 5) Ret Time | (# 5) Amount | (# 5) Target Response | Calc Amount (ng) | % Diff | (# 6) Name | (# 6) Ret Time | (# 6) Amount | (# 6) Target Response | Calc Amount (ng) | % Diff
| 2025-04-16 11:29 | 20250415037.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902607 | PFAS Blank Cond. 4/15/25 | 31 | 4:2 FTOH | 5.89 | 0 | 7 | 0.034855 |  | 6:2 FTOH | 7.58 | 0 | 21 | 0.035684 |  | 8:2 FTOH | 8.98 | 0 | 68 | 0.10121 |  | 10:2 FTOH | 10.28 | 0 | 142 | 0.25473
| 2025-04-15 17:25 | 20250415007.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | 4:2 FTOH | 0 | 0 | 0 |  |  | 6:2 FTOH | 0 | 0 | 0 |  |  | 8:2 FTOH | 0 | 0 | 0 |  |  | 10:2 FTOH | 0 | 0 | 0
| 2025-04-15 18:01 | 20250415008.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902601 | 0.75 ng FTOHs, spiked 4/15/25 | 12 | 4:2 FTOH | 5.86 | 0.75 | 438 | 0.84075 | 12.1% | 6:2 FTOH | 7.55 | 0.75 | 843 | 0.7935 | 5.8% | 8:2 FTOH | 9 | 0.75 | 814 | 0.72329 | -3.561% | 10:2 FTOH | 10.27 | 0.75 | 717 | 0.80613 | 7.484%
| 2025-04-15 18:40 | 20250415009.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902608 | 0.75 ng FTOHs, spiked 4/15/25 | 13 | 4:2 FTOH | 5.86 | 0.75 | 439 | 0.84262 | 12.35% | 6:2 FTOH | 7.54 | 0.75 | 856 | 0.80548 | 7.398% | 8:2 FTOH | 9 | 0.75 | 827 | 0.73413 | -2.116% | 10:2 FTOH | 10.27 | 0.75 | 588 | 0.68242 | -9.01%
| 2025-04-15 19:18 | 20250415010.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | 4:2 FTOH | 0 | 0 | 0 |  |  | 6:2 FTOH | 7.55 | 0 | 54 | 0.066107 |  | 8:2 FTOH | 9 | 0 | 78 | 0.10955 |  | 10:2 FTOH | 10.28 | 0 | 169 | 0.28062
| 2025-04-15 19:52 | 20250415011.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | 4:2 FTOH | 0 | 0 | 0 |  |  | 6:2 FTOH | 7.54 | 0 | 26 | 0.040293 |  | 8:2 FTOH | 0 | 0 | 0 |  |  | 10:2 FTOH | 10.3 | 0 | 78 | 0.19335
| 2025-04-15 20:27 | 20250415012.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | 4:2 FTOH | 0 | 0 | 0 |  |  | 6:2 FTOH | 0 | 0 | 0 |  |  | 8:2 FTOH | 0 | 0 | 0 |  |  | 10:2 FTOH | 10.31 | 0 | 34 | 0.15116
| 2025-04-15 21:04 | 20250415013.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902630 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 14 | 4:2 FTOH | 5.86 | 0.25 | 118 | 0.24241 | -3.038% | 6:2 FTOH | 7.55 | 0.25 | 269 | 0.26432 | 5.727% | 8:2 FTOH | 9 | 0.25 | 1565 | 1.3495 | 439.8% | 10:2 FTOH | 10.27 | 0.25 | 480 | 0.57886 | 131.5%
| 2025-04-15 21:41 | 20250415014.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902633 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 15 | 4:2 FTOH | 5.86 | 0.25 | 109 | 0.22558 | -9.769% | 6:2 FTOH | 7.54 | 0.25 | 255 | 0.25141 | 0.5646% | 8:2 FTOH | 9 | 0.25 | 431 | 0.40391 | 61.56% | 10:2 FTOH | 10.27 | 0.25 | 266 | 0.37364 | 49.46%
| 2025-04-15 22:18 | 20250415015.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902612 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 16 | 4:2 FTOH | 5.86 | 0.25 | 105 | 0.2181 | -12.76% | 6:2 FTOH | 7.55 | 0.25 | 253 | 0.24957 | -0.173% | 8:2 FTOH | 9 | 0.25 | 338 | 0.32636 | 30.54% | 10:2 FTOH | 10.27 | 0.25 | 247 | 0.35542 | 42.17%
| 2025-04-15 22:55 | 20250415016.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902637 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 17 | 4:2 FTOH | 5.86 | 0.25 | 86 | 0.18257 | -26.97% | 6:2 FTOH | 7.55 | 0.25 | 207 | 0.20716 | -17.14% | 8:2 FTOH | 9 | 0.25 | 266 | 0.26632 | 6.528% | 10:2 FTOH | 10.27 | 0.25 | 242 | 0.35062 | 40.25%
| 2025-04-15 23:31 | 20250415017.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902617 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 18 | 4:2 FTOH | 5.86 | 0.25 | 116 | 0.23867 | -4.534% | 6:2 FTOH | 7.55 | 0.25 | 280 | 0.27446 | 9.784% | 8:2 FTOH | 9 | 0.25 | 323 | 0.31385 | 25.54% | 10:2 FTOH | 10.27 | 0.25 | 293 | 0.39953 | 59.81%
| 2025-04-16 00:08 | 20250415018.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902614 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 19 | 4:2 FTOH | 5.86 | 0.25 | 104 | 0.21623 | -13.51% | 6:2 FTOH | 7.54 | 0.25 | 237 | 0.23482 | -6.073% | 8:2 FTOH | 9 | 0.25 | 309 | 0.30218 | 20.87% | 10:2 FTOH | 10.27 | 0.25 | 280 | 0.38706 | 54.83%
| 2025-04-16 00:44 | 20250415019.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902604 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 20 | 4:2 FTOH | 5.86 | 0.25 | 135 | 0.27419 | 9.677% | 6:2 FTOH | 7.54 | 0.25 | 266 | 0.26155 | 4.621% | 8:2 FTOH | 8.99 | 0.25 | 380 | 0.36138 | 44.55% | 10:2 FTOH | 10.27 | 0.25 | 469 | 0.56831 | 127.3%
| 2025-04-16 01:20 | 20250415020.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902606 | 0.25 ng FTOHs MDL, spiked 4/15/25 | 21 | 4:2 FTOH | 5.86 | 0.25 | 127 | 0.25923 | 3.694% | 6:2 FTOH | 7.55 | 0.25 | 302 | 0.29474 | 17.9% | 8:2 FTOH | 9 | 0.25 | 363 | 0.34721 | 38.88% | 10:2 FTOH | 10.27 | 0.25 | 407 | 0.50885 | 103.5%
| 2025-04-16 01:56 | 20250415021.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902603 | PFAS Blank Cond. 4/15/25 | 11 | 4:2 FTOH | 0 | 0 | 0 |  |  | 6:2 FTOH | 7.54 | 0 | 28 | 0.042137 |  | 8:2 FTOH | 9 | 0 | 61 | 0.095372 |  | 10:2 FTOH | 10.29 | 0 | 75 | 0.19048
| 2025-04-16 02:31 | 20250415022.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902611 | 0.75 ng FTOHs, spiked 4/15/25 | 22 | 4:2 FTOH | 5.86 | 0.75 | 318 | 0.61637 | -17.82% | 6:2 FTOH | 7.54 | 0.75 | 647 | 0.6128 | -18.29% | 8:2 FTOH | 9 | 0.75 | 729 | 0.65241 | -13.01% | 10:2 FTOH | 10.27 | 0.75 | 509 | 0.60666 | -19.11%
| 2025-04-16 03:08 | 20250415023.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902615 | 0.75 ng FTOHs, spiked 4/15/25 | 23 | 4:2 FTOH | 5.86 | 0.75 | 314 | 0.60889 | -18.81% | 6:2 FTOH | 7.54 | 0.75 | 718 | 0.67826 | -9.566% | 8:2 FTOH | 9 | 0.75 | 815 | 0.72412 | -3.45% | 10:2 FTOH | 10.27 | 0.75 | 576 | 0.67091 | -10.54%
| 2025-04-16 03:44 | 20250415024.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902640 | PFAS Blank Cond. 4/15/25 | 24 | 4:2 FTOH | 0 | 0 | 0 |  |  | 6:2 FTOH | 7.54 | 0 | 34 | 0.047668 |  | 8:2 FTOH | 9 | 0 | 114 | 0.13957 |  | 10:2 FTOH | 10.28 | 0 | 189 | 0.2998
| 2025-04-16 04:20 | 20250415025.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902640 | PFAS Blank Cond. 4/15/25 | 24 | 4:2 FTOH | 0 | 0 | 0 |  |  | 6:2 FTOH | 7.54 | 0 | 18 | 0.032918 |  | 8:2 FTOH | 9 | 0 | 25 | 0.065352 |  | 10:2 FTOH | 10.29 | 0 | 84 | 0.19911
| 2025-04-16 04:56 | 20250415026.D | VF-200ms SIM Scan Target 20250218_Etune.M | 902640 | PFAS Blank Cond. 4/15/25 | 24 | 4:2 FTOH | 0 | 0 | 0 |  |  | 6:2 FTOH | 0 | 0 | 0 |  |  | 8:2 FTOH | 0 | 0 | 0 |  |  | 10:2 FTOH | 0 | 0 | 0
[10 empty] | Std Dev | 0.030222 | [5 empty] | 0.028051 | [5 empty] | 0.044402 | [5 empty] | 0.084187
[10 empty] | MDL | 0.094987 | [5 empty] | 0.088165 | [5 empty] | 0.13956 | [5 empty] | 0.2646
[10 empty] | RL | 0.28496 | [5 empty] | 0.26449 | [5 empty] | 0.41867 | [5 empty] | 0.7938

## LOD+LOQ Calc Explanation
LOD and LOQ calculations were performed according to Bertram, et al. 2011 using the following equation:
Where B is the background count rate (counts per second), Cf is the calibration factor (a.k.a. slope of the calibration curve; counts per second per pptv), [X] is the mixing ratio (pptv), and t is the integration time (seconds)
To use this equation to calculate limits of detection (LODs) and limits of quantitation (LOQs), the S/N ratio is set to 3 and 10, respectively, and the equation is solved for [X]. The resulting mixing ratio [X] in pptv is the LOD or LOQ, respectively.
Example:
TFA
Cf | 12.543 | cps/pptv
[X] LOD | 6.0051 | pptv
[X] LOQ | 23.1876 | pptv
t | 1 | sec
B | 277.523 | cps
S/N LOD | 3 | Set to 3
S/N LOQ | 10 | Set to 10
Highlighted values are experiment-dependent values that are user updated; the Cf values are taken from calibration curves created and linearly fit in Igor and the B values are taken from background signal data processed in Tofware
For our instrument, t = 1 sec always
The cells labeled LOD and LOQ contain equations equivalent to the equation above
To solve for [X] LOD, the Solver function is used by selecting the S/N LOD cell as the objective, setting the target value to a value of 3, and selecting the [X] LOD cell for the variable cell to be changed (see screenshot below)
The same thing is done to solve for [X] LOQ using the [X] LOQ and S/N LOQ cells with a target value of 10

## Raw I-CIMS BG Data
Local Time | 4:2 FTOH | 6:2 FTOH | 8:2 FTOH | 10:2 FTOH
2025-04-01 11:02 | 0 | 59.4276 | 6.207 | 24.9318 | These are the first 2000 data points collected on the day of I-CIMS calibration while sampling UZA prior to any calibration injections.
2025-04-01 11:02 | 5.6302 | 71.6977 | 8.1807 | 10.8635 | The means of these signals were used for "B" in the calculation of the LODs and LOQs.
2025-04-01 11:02 | 9.8766 | 45.1403 | 5.2428 | 9.139
2025-04-01 11:02 | 6.6505 | 48.1844 | 2.2052 | 18.6779
2025-04-01 11:02 | 9.0532 | 47.3053 | 5.7856 | 15.709
2025-04-01 11:02 | 10.8981 | 48.1049 | 6.6055 | 24.4632
2025-04-01 11:02 | 0 | 41.1312 | 4.2747 | 13.0515
2025-04-01 11:02 | 8.3716 | 34.2584 | 3.5487 | 13.0317
2025-04-01 11:02 | 0 | 36.9473 | 3.3566 | 20.8056
2025-04-01 11:02 | 6.4299 | 29.5884 | 5.6016 | 7.5523
2025-04-01 11:02 | 16.7044 | 49.7264 | 13.0621 | 5.557
2025-04-01 11:02 | 0 | 59.7505 | 0.87098 | 11.2006
2025-04-01 11:02 | 8.4413 | 36.0212 | 8.0482 | 14.7621
2025-04-01 11:02 | 0 | 60.7442 | 7.5319 | 20.5066
2025-04-01 11:02 | 12.3074 | 53.492 | 3.244 | 19.3144
2025-04-01 11:02 | 9.7142 | 50.4675 | 4.0523 | 11.7345
2025-04-01 11:02 | 6.9647 | 45.706 | 6.4033 | 8.3047
2025-04-01 11:02 | 2.4887 | 32.8326 | 9.688 | 10.0854
2025-04-01 11:02 | 4.2534 | 41.7223 | 4.8055 | 14.3766
2025-04-01 11:02 | 13.7567 | 42.0397 | 4.7767 | 3.5325
2025-04-01 11:02 | 0 | 43.0625 | 9.1824 | 22.2897
2025-04-01 11:02 | 0 | 41.5094 | 3.0218 | 16.4994
2025-04-01 11:02 | 8.4489 | 54.5075 | 1.6293 | 11.9411
2025-04-01 11:02 | 0 | 40.8534 | 2.5244 | 14.5601
2025-04-01 11:02 | 0 | 45.2598 | 8.6902 | 10.0646
2025-04-01 11:02 | 6.9587 | 64.7392 | 3.4175 | 12.3421
2025-04-01 11:02 | 0 | 37.7306 | 1.8013 | 13.2255
2025-04-01 11:02 | 17.9568 | 35.9622 | 3.0826 | 14.8186
2025-04-01 11:02 | 0 | 41.4902 | 3.1739 | 13.8986
2025-04-01 11:02 | 4.8141 | 63.4528 | 2.6239 | 12.0741
2025-04-01 11:03 | 0 | 55.0247 | 6.0505 | 13.5124
2025-04-01 11:03 | 11.961 | 37.7635 | 3.7189 | 6.138
2025-04-01 11:03 | 7.767 | 34.2911 | 12.8267 | 33.3503
2025-04-01 11:03 | 0 | 33.1872 | 4.3019 | 19.4529
2025-04-01 11:03 | 16.6334 | 57.4973 | 3.3627 | 23.0844
2025-04-01 11:03 | 2.1566 | 54.2715 | 3.9317 | 5.7581
2025-04-01 11:03 | 8.6318 | 36.1418 | 1.8719 | 9.7273
2025-04-01 11:03 | 3.9615 | 33.9245 | 7.1568 | 10.8881
2025-04-01 11:03 | 0 | 23.8874 | 1.9708 | 9.8018
2025-04-01 11:03 | 2.666 | 45.4931 | 4.5066 | 16.5467
2025-04-01 11:03 | 9.3436 | 57.7248 | 1.4423 | 15.7987
2025-04-01 11:03 | 8.0789 | 35.0431 | 7.0096 | 19.7895
2025-04-01 11:03 | 0 | 67.3978 | 2.4542 | 14.4951
2025-04-01 11:03 | 5.7919 | 46.9135 | 1.5884 | 14.2051
2025-04-01 11:03 | 0 | 37.4521 | 2.7076 | 5.078
2025-04-01 11:03 | 0 | 40.8463 | 4.1762 | 33.2834
2025-04-01 11:03 | 1.2461 | 44.0182 | 3.9953 | 7.6305
2025-04-01 11:03 | 5.6517 | 45.0705 | 6.6012 | 14.6282
2025-04-01 11:03 | 4.017 | 39.911 | 6.5435 | 19.1418
2025-04-01 11:03 | 15.826 | 39.5482 | 2.6269 | 4.9795
2025-04-01 11:03 | 0.036693 | 56.8112 | 1.7789 | 17.3101
2025-04-01 11:03 | 9.7197 | 36.0421 | 4.207 | 8.1265
2025-04-01 11:03 | 2.1122 | 33.0801 | 8.4675 | 19.5101
2025-04-01 11:03 | 0 | 36.4217 | 4.3287 | 10.7058
2025-04-01 11:03 | 0 | 27.9047 | 1.2969 | 18.1236
2025-04-01 11:03 | 0 | 27.3282 | 4.7816 | 25.8983
2025-04-01 11:03 | 5.6905 | 54.2382 | 2.0904 | 21.6318
2025-04-01 11:03 | 0.60287 | 44.4925 | 4.0049 | 11.1904
2025-04-01 11:03 | 4.5477 | 67.4198 | 5.4293 | 12.6557
2025-04-01 11:03 | 3.4346 | 60.7565 | 3.3787 | 7.9709
2025-04-01 11:03 | 12.5246 | 38.0513 | 2.5365 | 24.5651
2025-04-01 11:03 | 3.4942 | 36.9559 | 4.1081 | 19.3785
2025-04-01 11:03 | 7.8982 | 48.8658 | 3.1339 | 15.6526
2025-04-01 11:03 | 10.6572 | 34.6557 | 5.9619 | 12.0634
2025-04-01 11:03 | 0 | 60.3141 | 3.0175 | 19.1943
2025-04-01 11:03 | 5.7406 | 41.5944 | 4.0234 | 8.7039
2025-04-01 11:03 | 1.7177 | 43.2794 | 4.3511 | 20.9431
2025-04-01 11:03 | 4.869 | 39.7212 | 1.9811 | 14.1677
2025-04-01 11:03 | 9.6683 | 31.1597 | 5.0478 | 10.4372
2025-04-01 11:03 | 0 | 44.2433 | 3.2981 | 8.1121
2025-04-01 11:03 | 0 | 62.4539 | 2.3753 | 15.3451
2025-04-01 11:03 | 4.7747 | 47.1398 | 1.9875 | 10.1622
2025-04-01 11:03 | 1.9071 | 41.8549 | 2.1008 | 17.6401
2025-04-01 11:03 | 6.4915 | 31.0569 | 7.1502 | 13.2095
2025-04-01 11:03 | 0 | 48.2058 | 4.128 | 5.2591
2025-04-01 11:03 | 6.728 | 58.5854 | 5.987 | 27.0054
2025-04-01 11:03 | 0 | 60.9353 | 2.5824 | 4.9631
2025-04-01 11:03 | 0 | 41.2883 | 1.786 | 32.3201
2025-04-01 11:03 | 0 | 45.2539 | 3.4773 | 14.2744
2025-04-01 11:03 | 10.3523 | 43.9032 | 3.9736 | 9.2443
2025-04-01 11:03 | 5.3859 | 37.9673 | 8.8646 | 8.7156
2025-04-01 11:03 | 8.4625 | 38.0102 | 3.5817 | 11.3705
2025-04-01 11:03 | 5.9254 | 40.3804 | 2.9465 | 14.4053
2025-04-01 11:03 | 9.4684 | 50.8984 | 3.2542 | 14.3353
2025-04-01 11:03 | 16.3269 | 50.8987 | 4.1355 | 9.2705
2025-04-01 11:03 | 5.1188 | 61.4734 | 0.20516 | 10.6161
2025-04-01 11:03 | 1.4386 | 29.511 | 5.1826 | 14.8095
2025-04-01 11:03 | 11.4108 | 62.5894 | 4.0852 | 17.5722
2025-04-01 11:03 | 10.3362 | 38.77 | 8.1713 | 15.6268
2025-04-01 11:03 | 5.8633 | 35.644 | 2.4099 | 16.65
2025-04-01 11:04 | 0 | 29.3899 | 3.9053 | 14.3671
2025-04-01 11:04 | 2.5909 | 47.5865 | 5.1708 | 13.7454
2025-04-01 11:04 | 20.5909 | 49.4939 | 2.9993 | 11.0548
2025-04-01 11:04 | 0 | 59.4608 | 1.043 | 5.1908
2025-04-01 11:04 | 9.4425 | 46.0021 | 2.9219 | 21.636
2025-04-01 11:04 | 1.7956 | 53.2113 | 0.46325 | 7.6744
2025-04-01 11:04 | 7.3018 | 42.6134 | 5.9241 | 3.1811
2025-04-01 11:04 | 0 | 42.6822 | 4.8087 | 10.1274
2025-04-01 11:04 | 2.9279 | 36.5315 | 2.9683 | 9.6583
2025-04-01 11:04 | 21.8231 | 41.2996 | 3.9672 | 18.7541
2025-04-01 11:04 | 8.81 | 36.8249 | 2.7243 | 34.8109
2025-04-01 11:04 | 3.7572 | 49.1799 | 5.2162 | 20.2527
2025-04-01 11:04 | 8.4573 | 59.178 | 2.4601 | 16.5428
2025-04-01 11:04 | 5.2171 | 30.6342 | 8 | 28.4053
2025-04-01 11:04 | 5.3484 | 57.9919 | 6.2849 | 9.9963
2025-04-01 11:04 | 17.2456 | 60.3265 | 9.4329 | 14.6295
2025-04-01 11:04 | 11.3138 | 35.2112 | 2.186 | 12.2313
2025-04-01 11:04 | 17.0151 | 34.4781 | 2.1159 | 21.4769
2025-04-01 11:04 | 1.783 | 44.7227 | 2.6203 | 13.6656
2025-04-01 11:04 | 10.9374 | 35.5429 | 2.6793 | 13.0894
2025-04-01 11:04 | 6.6912 | 50.5719 | 8.1092 | 16.4367
2025-04-01 11:04 | 6.5356 | 53.3739 | 3.3634 | 19.7173
2025-04-01 11:04 | 0 | 36.424 | 8.3809 | 5.1541
2025-04-01 11:04 | 9.3197 | 49.6297 | 7.3064 | 20.097
2025-04-01 11:04 | 0 | 50.5318 | 2.8866 | 11.324
2025-04-01 11:04 | 5.8681 | 41.6411 | 7.5777 | 24.4142
2025-04-01 11:04 | 4.8897 | 49.5681 | 5.417 | 13.6637
2025-04-01 11:04 | 4.3891 | 48.0278 | 5.4223 | 11.8564
2025-04-01 11:04 | 13.2115 | 35.4685 | 3.3609 | 15.3929
2025-04-01 11:04 | 2.5443 | 47.7378 | 3.9664 | 8.1361
2025-04-01 11:04 | 12.7192 | 29.9296 | 3.3955 | 5.2988
2025-04-01 11:04 | 1.3605 | 40.0792 | 2.2156 | 18.5705
2025-04-01 11:04 | 15.9598 | 39.5276 | 2.5819 | 14.5683
2025-04-01 11:04 | 5.176 | 60.9988 | 2.1706 | 18.5079
2025-04-01 11:04 | 3.6787 | 37.7093 | 4.0284 | 11.8049
2025-04-01 11:04 | 7.0601 | 55.3596 | 1.7709 | 7.368
2025-04-01 11:04 | 9.266 | 42.7359 | 3.7776 | 9.401
2025-04-01 11:04 | 0 | 48.3194 | 4.4048 | 9.2203
2025-04-01 11:04 | 1.3393 | 41.2054 | 4.0107 | 18.2692
2025-04-01 11:04 | 8.6848 | 34.4103 | 2.5798 | 13.1931
2025-04-01 11:04 | 2.8895 | 50.8264 | 2.0853 | 13.8583
2025-04-01 11:04 | 0 | 36.4493 | 1.0887 | 5.0869
2025-04-01 11:04 | 10.2778 | 40.5583 | 5.5376 | 22.2048
2025-04-01 11:04 | 1.4599 | 47.8551 | 4.3832 | 9.758
2025-04-01 11:04 | 0 | 33.0777 | 4.3088 | 17.4056
2025-04-01 11:04 | 5.3044 | 32.7711 | 2.3945 | 30.6269
2025-04-01 11:04 | 0 | 65.1344 | 9.1135 | 10.9434
2025-04-01 11:04 | 8.9359 | 33.4016 | 3.5679 | 31.4316
2025-04-01 11:04 | 1.5813 | 38.0432 | 4.0258 | 21.2569
2025-04-01 11:04 | 2.038 | 41.6395 | 2.9071 | 19.5664
2025-04-01 11:04 | 6.967 | 55.1975 | 1.6247 | 16.5522
2025-04-01 11:04 | 0.62301 | 27.9891 | 3.9785 | 23.9207
2025-04-01 11:04 | 4.9682 | 49.5663 | 2.0097 | 7.3675
2025-04-01 11:04 | 0 | 38.5585 | 4.1422 | 6.206
2025-04-01 11:04 | 2.5594 | 38.2488 | 7.1157 | 13.3917
2025-04-01 11:04 | 3.2618 | 23.2215 | 5.9134 | 9.2211
2025-04-01 11:04 | 8.0153 | 38.1922 | 2.4879 | 13.1338
2025-04-01 11:04 | 16.133 | 58.6377 | 3.4724 | 16.3901
2025-04-01 11:04 | 1.9398 | 35.9609 | 1.8716 | 22.936
2025-04-01 11:04 | 0 | 33.0004 | 4.1081 | 2.292
2025-04-01 11:05 | 0 | 44.6042 | 8.494 | 9.4305
2025-04-01 11:05 | 3.8203 | 30.1171 | 5.1024 | 16.2463
2025-04-01 11:05 | 8.8207 | 46.0337 | 5.2806 | 19.2283
2025-04-01 11:05 | 0 | 29.5059 | 7.8792 | 11.5745
2025-04-01 11:05 | 0 | 59.6455 | 1.9644 | 17.1771
2025-04-01 11:05 | 0 | 33.1244 | 2.7729 | 22.6779
2025-04-01 11:05 | 0 | 40.1454 | 6.2034 | 21.0063
2025-04-01 11:05 | 12.5705 | 29.0867 | 1.3648 | 19.8129
2025-04-01 11:05 | 3.0261 | 27.7634 | 6.1473 | 16.8069
2025-04-01 11:05 | 1.7032 | 35.9794 | 7.7274 | 10.3359
2025-04-01 11:05 | 5.9676 | 61.3067 | 3.7739 | 8.1772
2025-04-01 11:05 | 0 | 57.1841 | 3.2802 | 9.491
2025-04-01 11:05 | 8.0664 | 38.0154 | 4.981 | 18.0431
2025-04-01 11:05 | 15.872 | 38.9679 | 3.9464 | 13.1143
2025-04-01 11:05 | 3.9804 | 44.3139 | 3.1993 | 10.6773
2025-04-01 11:05 | 18.0358 | 59.3322 | 7.3981 | 11.4053
2025-04-01 11:05 | 0 | 31.1806 | 2.8147 | 12.1378
2025-04-01 11:05 | 0 | 32.133 | 8.7361 | 13.1233
2025-04-01 11:05 | 6.0333 | 43.8015 | 7.9988 | 11.9157
2025-04-01 11:05 | 0.7844 | 45.5695 | 3.3491 | 10.7684
2025-04-01 11:05 | 4.9267 | 36.5745 | 2.7063 | 13.9498
2025-04-01 11:05 | 4.4671 | 54.9921 | 8.0956 | 13.3867
2025-04-01 11:05 | 0 | 35.0799 | 4.9929 | 19.1913
2025-04-01 11:05 | 3.54 | 43.5527 | 4.7637 | 11.6456
2025-04-01 11:05 | 1.5099 | 46.0137 | 5.121 | 16.1852
2025-04-01 11:05 | 0.98611 | 45.0795 | 4.7424 | 16.1374
2025-04-01 11:05 | 0 | 44.097 | 5.7634 | 9.9831
2025-04-01 11:05 | 0 | 85.992 | 3.4144 | 29.128
2025-04-01 11:05 | 1.3472 | 28.5869 | 3.802 | 10.7107
2025-04-01 11:05 | 10.5695 | 39.6638 | 5.4636 | 22.2464
2025-04-01 11:05 | 3.1056 | 37.684 | 4.8947 | 27.1082
2025-04-01 11:05 | 3.0531 | 49.5299 | 4.0895 | 23.9896
2025-04-01 11:05 | 2.6663 | 64.952 | 5.4445 | 15.708
2025-04-01 11:05 | 0 | 42.4302 | 5.6811 | 10.1004
2025-04-01 11:05 | 15.9436 | 63.1145 | 2.8388 | 12.9433
2025-04-01 11:05 | 2.4199 | 43.3186 | 4.3104 | 6.4025
2025-04-01 11:05 | 20.0983 | 24.504 | 6.2865 | 5.7206
2025-04-01 11:05 | 0 | 54.3407 | 1.7843 | 18.3224
2025-04-01 11:05 | 14.3508 | 50.3483 | 5.9941 | 16.4561
2025-04-01 11:05 | 3.3053 | 39.2769 | 3.7742 | 11.9229
2025-04-01 11:05 | 8.1903 | 44.2643 | 1.741 | 19.1714
2025-04-01 11:05 | 12.1135 | 42.8023 | 1.2619 | 21.7673
2025-04-01 11:05 | 6.3427 | 53.2177 | 3.759 | 9.9628
2025-04-01 11:05 | 5.092 | 16.9577 | 2.872 | 19.2815
2025-04-01 11:05 | 6.7872 | 28.0679 | 5.3904 | 7.5815
2025-04-01 11:05 | 1.2224 | 49.5277 | 0.34674 | 18.4057
2025-04-01 11:05 | 0.31922 | 66.9982 | 3.7086 | 9.621
2025-04-01 11:05 | 1.6403 | 35.3386 | 2.6291 | 18.2789
2025-04-01 11:05 | 8.7786 | 31.3291 | 5.3713 | 18.0669
2025-04-01 11:05 | 16.3396 | 30.0501 | 4.0303 | 14.6108
2025-04-01 11:05 | 1.3897 | 50.9962 | 3.6989 | 11.3871
2025-04-01 11:05 | 4.0127 | 25.3644 | 3.6733 | 21.3495
2025-04-01 11:05 | 5.9738 | 56.7341 | 7.4343 | 28.0438
2025-04-01 11:05 | 12.764 | 30.4627 | 4.4945 | 15.1595
2025-04-01 11:05 | 1.1965 | 50.8763 | 6.3983 | 25.023
2025-04-01 11:05 | 9.0057 | 31.71 | 5.4395 | 5.9727
2025-04-01 11:05 | 3.4831 | 55.4099 | 4.696 | 7.3785
2025-04-01 11:05 | 17.3947 | 62.1445 | 2.7463 | 6.812
2025-04-01 11:05 | 0 | 51.4153 | 2.561 | 6.2354
2025-04-01 11:05 | 11.119 | 41.4548 | 1.9993 | 1.5493
2025-04-01 11:06 | 6.5479 | 41.8937 | 6.101 | 11.7594
2025-04-01 11:06 | 0 | 52.3968 | 6.6826 | 14.1377
2025-04-01 11:06 | 6.3885 | 32.0405 | 1.8232 | 20.7642
2025-04-01 11:06 | 2.4214 | 62.2971 | 2.1356 | 15.7959
2025-04-01 11:06 | 12.6471 | 58.5398 | 8.1013 | 25.4898
2025-04-01 11:06 | 0.57853 | 36.9057 | 3.9564 | 6.5734
2025-04-01 11:06 | 0.61745 | 35.2938 | 3.8147 | 13.4383
2025-04-01 11:06 | 4.5627 | 54.8638 | 2.3617 | 9.9687
2025-04-01 11:06 | 0 | 32.9808 | 1.8453 | 15.6552
2025-04-01 11:06 | 12.8446 | 35.7187 | 1.1169 | 18.8196
2025-04-01 11:06 | 4.0643 | 37.3237 | 5.4363 | 20.4555
2025-04-01 11:06 | 0 | 37.8857 | 1.5602 | 18.945
2025-04-01 11:06 | 4.1041 | 50.208 | 2.2912 | 11.6619
2025-04-01 11:06 | 4.1495 | 54.2891 | 5.2748 | 14.7622
2025-04-01 11:06 | 3.4063 | 59.7191 | 3.7379 | 8.1215
2025-04-01 11:06 | 6.6859 | 55.7612 | 1.4918 | 15.6335
2025-04-01 11:06 | 0 | 47.2231 | 3.8234 | 9.3929
2025-04-01 11:06 | 0 | 33.8321 | 4.1253 | 14.7293
2025-04-01 11:06 | 8.6292 | 23.9965 | 6.4748 | 17.7219
2025-04-01 11:06 | 13.652 | 33.1408 | 0.31032 | 16.0992
2025-04-01 11:06 | 5.5637 | 53.8052 | 3.1861 | 2.971
2025-04-01 11:06 | 6.0546 | 44.8367 | 1.692 | 13.0075
2025-04-01 11:06 | 7.8943 | 43.2598 | 6.1829 | 23.4292
2025-04-01 11:06 | 4.5261 | 63.0939 | 5.8639 | 18.6885
2025-04-01 11:06 | 16.1886 | 46.1814 | 5.7482 | 4.9155
2025-04-01 11:06 | 0 | 36.1283 | 2.2555 | 8.9099
2025-04-01 11:06 | 7.5233 | 64.64 | 5.6225 | 15.895
2025-04-01 11:06 | 8.5517 | 58.2041 | 1.5699 | 14.5955
2025-04-01 11:06 | 11.5726 | 37.7746 | 1.8039 | 25.3575
2025-04-01 11:06 | 3.0811 | 48.2669 | 4.0889 | 9.9105
2025-04-01 11:06 | 5.1602 | 54.2008 | 0.84332 | 12.6885
2025-04-01 11:06 | 0 | 34.0142 | 0.60976 | 10.0116
2025-04-01 11:06 | 0 | 49.0783 | 4.5327 | 3.0258
2025-04-01 11:06 | 4.6373 | 37.8309 | 3.4682 | 9.0561
2025-04-01 11:06 | 0 | 41.0316 | 5.2659 | 8.3351
2025-04-01 11:06 | 0 | 20.1336 | 3.6084 | 14.0874
2025-04-01 11:06 | 6.4191 | 34.1725 | 4.883 | 16.0651
2025-04-01 11:06 | 7.3665 | 47.0919 | 3.8513 | 7.9994
2025-04-01 11:06 | 0 | 32.7052 | 2.0799 | 7.3937
2025-04-01 11:06 | 6.0609 | 52.9053 | 1.7131 | 18.9181
2025-04-01 11:06 | 9.0912 | 45.3721 | 4.2984 | 8.4049
2025-04-01 11:06 | 3.1017 | 58.3568 | 4.2289 | 10.7185
2025-04-01 11:06 | 1.8428 | 37.928 | 3.4408 | 8.952
2025-04-01 11:06 | 0 | 35.918 | 3.5347 | 10.584
2025-04-01 11:06 | 0 | 45.1018 | 1.2084 | 9.8497
2025-04-01 11:06 | 3.0191 | 37.1649 | 3.9033 | 10.5701
2025-04-01 11:06 | 0.10667 | 51.528 | 2.6493 | 19.9361
2025-04-01 11:06 | 0 | 48.9965 | 4.4368 | 7.2585
2025-04-01 11:06 | 5.6462 | 63.9873 | 3.7596 | 20.8385
2025-04-01 11:06 | 0 | 19.4516 | 2.8047 | 11.2469
2025-04-01 11:06 | 3.304 | 27.7016 | 7.7943 | 12.3876
2025-04-01 11:06 | 7.9922 | 42.9317 | 1.9725 | 17.4207
2025-04-01 11:06 | 0.51697 | 28.0569 | 2.9031 | 16.2075
2025-04-01 11:06 | 0 | 34.3229 | 3.2504 | 15.3234
2025-04-01 11:06 | 0.78992 | 30.0435 | 2.7438 | 20.4629
2025-04-01 11:06 | 0 | 31.6577 | 5.3084 | 15.0618
2025-04-01 11:06 | 11.7209 | 26.5728 | 1.1023 | 22.0539
2025-04-01 11:06 | 3.674 | 30.5545 | 2.8525 | 15.9271
2025-04-01 11:06 | 0 | 42.2867 | 2.8378 | 12.5492
2025-04-01 11:06 | 0 | 29.5016 | 2.0952 | 15.0459
2025-04-01 11:07 | 8.5719 | 49.851 | 2.2748 | 8.7663
2025-04-01 11:07 | 5.8027 | 34.0542 | 1.9437 | 7.8926
2025-04-01 11:07 | 8.4 | 43.1566 | 2.4359 | 14.3046
2025-04-01 11:07 | 0 | 40.5816 | 3.7841 | 12.9749
2025-04-01 11:07 | 0 | 48.7106 | 3.2562 | 8.7113
2025-04-01 11:07 | 0 | 42.0126 | 5.7829 | 21.4578
2025-04-01 11:07 | 1.1503 | 36.6287 | 3.7276 | 7.1491
2025-04-01 11:07 | 2.1454 | 26.6484 | 4.5561 | 10.2665
2025-04-01 11:07 | 0.77407 | 40.7093 | 1.9008 | 16.534
2025-04-01 11:07 | 2.1918 | 57.2087 | 3.4714 | 11.1955
2025-04-01 11:07 | 7.8648 | 42.4589 | 4.7788 | 5.5215
2025-04-01 11:07 | 0 | 52.4882 | 3.3282 | 17.5624
2025-04-01 11:07 | 9.2701 | 31.9805 | 2.712 | 6.452
2025-04-01 11:07 | 21.9671 | 46.7673 | 2.711 | 6.4549
2025-04-01 11:07 | 0 | 50.3565 | 5.5341 | 14.3971
2025-04-01 11:07 | 0.76043 | 27.4952 | 5.5284 | 17.1983
2025-04-01 11:07 | 2.6295 | 38.0069 | 4.2809 | 14.4803
2025-04-01 11:07 | 0 | 23.2483 | 2.4557 | 12.8379
2025-04-01 11:07 | 6.3922 | 45.0609 | 5.0671 | 11.6198
2025-04-01 11:07 | 4.3988 | 48.15 | 4.8604 | 9.4266
2025-04-01 11:07 | 3.5271 | 43.7386 | 1.9304 | 7.0659
2025-04-01 11:07 | 9.3661 | 36.3537 | 7.3766 | 7.209
2025-04-01 11:07 | 6.1972 | 35.8229 | 4.0457 | 25.5216
2025-04-01 11:07 | 11.8972 | 41.6101 | 2.1776 | 17.1326
2025-04-01 11:07 | 10.1159 | 31.9062 | 2.7757 | 17.7182
2025-04-01 11:07 | 0 | 27.1775 | 2.0501 | 4.0713
2025-04-01 11:07 | 4.6462 | 43.4064 | 2.6982 | 33.4269
2025-04-01 11:07 | 9.0859 | 45.6685 | 3.4862 | 9.7163
2025-04-01 11:07 | 0 | 51.4591 | 6.0537 | 16.1341
2025-04-01 11:07 | 4.6475 | 46.2523 | 3.1103 | 13.7104
2025-04-01 11:07 | 7.4842 | 36.6416 | 1.4031 | 16.6405
2025-04-01 11:07 | 7.617 | 25.1414 | 1.8128 | 22.4709
2025-04-01 11:07 | 5.6277 | 35.2421 | 3.8973 | 13.8792
2025-04-01 11:07 | 5.6989 | 38.0181 | 3.0156 | 11.7599
2025-04-01 11:07 | 0.25149 | 52.8861 | 3.9593 | 4.1999
2025-04-01 11:07 | 11.5533 | 47.107 | 4.1659 | 19.7088
2025-04-01 11:07 | 0 | 34.8707 | 4.8881 | 14.6285
2025-04-01 11:07 | 7.8181 | 27.9554 | 2.5896 | 16.656
2025-04-01 11:07 | 14.4001 | 52.8539 | 2.8523 | 11.938
2025-04-01 11:07 | 5.378 | 58.1267 | 4.8375 | 10.4896
2025-04-01 11:07 | 7.0618 | 44.1454 | 6.1136 | 15.8391
2025-04-01 11:07 | 11.1357 | 45.8199 | 3.1917 | 12.5401
2025-04-01 11:07 | 15.4729 | 34.0837 | 3.135 | 16.2772
2025-04-01 11:07 | 7.2558 | 44.0199 | 2.2323 | 20.6342
2025-04-01 11:07 | 7.8912 | 59.4736 | 1.9233 | 18.7787
2025-04-01 11:07 | 5.0038 | 59.4921 | 2.2636 | 16.2411
2025-04-01 11:07 | 4.6454 | 43.7883 | 3.3966 | 9.2336
2025-04-01 11:07 | 4.9752 | 67.7085 | 2.8252 | 8.3222
2025-04-01 11:07 | 8.5046 | 37.5374 | 3.5242 | 10.4592
2025-04-01 11:07 | 0 | 40.3863 | 5.2473 | 8.9554
2025-04-01 11:07 | 3.4121 | 27.5345 | 3.7617 | 16.5312
2025-04-01 11:07 | 0 | 43.0582 | 2.2974 | 11.9735
2025-04-01 11:07 | 5.4986 | 47.2715 | 1.2175 | 6.5001
2025-04-01 11:07 | 0.36079 | 45.4905 | 2.2768 | 21.7381
2025-04-01 11:07 | 10.0362 | 50.6564 | 4.6697 | 20.7956
2025-04-01 11:07 | 0 | 46.4648 | 6.6724 | 7.8815
2025-04-01 11:07 | 1.6967 | 48.5572 | 6.0365 | 12.8893
2025-04-01 11:07 | 19.385 | 46.4616 | 1.337 | 16.7887
2025-04-01 11:07 | 0.3433 | 33.4413 | 3.9537 | 21.4877
2025-04-01 11:07 | 0 | 35.0157 | 8.8053 | 16.1467
2025-04-01 11:08 | 10.3451 | 40.5263 | 5.8239 | 13.5911
2025-04-01 11:08 | 9.4889 | 38.0965 | 4.5687 | 15.8763
2025-04-01 11:08 | 10.4802 | 52.1034 | 3.1982 | 23.9116
2025-04-01 11:08 | 0 | 66.7756 | 10.2737 | 17.127
2025-04-01 11:08 | 0 | 43.8673 | 10.8546 | 16.9394
2025-04-01 11:08 | 7.9741 | 39.9264 | 5.7169 | 33.8594
2025-04-01 11:08 | 2.1184 | 59.2189 | 5.5618 | 23.3187
2025-04-01 11:08 | 0 | 44.0432 | 4.1448 | 7.7165
2025-04-01 11:08 | 9.1726 | 43.907 | 2.2323 | 4.9256
2025-04-01 11:08 | 9.2533 | 36.4589 | 2.3296 | 22.5501
2025-04-01 11:08 | 9.3947 | 34.2419 | 5.6672 | 17.8109
2025-04-01 11:08 | 6.2693 | 48.5361 | 3.3923 | 10.9362
2025-04-01 11:08 | 4.0765 | 36.7197 | 1.5266 | 12.6977
2025-04-01 11:08 | 4.749 | 47.1747 | 5.4644 | 12.3369
2025-04-01 11:08 | 0.39805 | 39.1727 | 4.538 | 12.4285
2025-04-01 11:08 | 0 | 53.4637 | 3.967 | 6.6401
2025-04-01 11:08 | 8.7405 | 40.9229 | 2.523 | 18.6346
2025-04-01 11:08 | 4.3672 | 33.9541 | 1.6905 | 7.5794
2025-04-01 11:08 | 0 | 37.1152 | 6.8929 | 17.287
2025-04-01 11:08 | 12.7997 | 55.8538 | 2.532 | 14.4155
2025-04-01 11:08 | 0 | 34.7693 | 6.1229 | 18.6236
2025-04-01 11:08 | 0 | 38.9115 | 6.3466 | 4.3341
2025-04-01 11:08 | 1.4037 | 62.2427 | 4.211 | 11.5084
2025-04-01 11:08 | 10.5957 | 47.7469 | 6.0177 | 9.4832
2025-04-01 11:08 | 17.6664 | 44.9874 | 4.4241 | 16.2534
2025-04-01 11:08 | 0.44247 | 45.4307 | 4.6722 | 16.28
2025-04-01 11:08 | 6.1206 | 48.0491 | 3.6667 | 14.2884
2025-04-01 11:08 | 9.305 | 43.2587 | 3.3263 | 9.4825
2025-04-01 11:08 | 0 | 33.9418 | 3.5441 | 4.1612
2025-04-01 11:08 | 10.243 | 18.8531 | 5.2891 | 10.929
2025-04-01 11:08 | 0 | 48.408 | 5.547 | 20.2284
2025-04-01 11:08 | 0 | 64.4396 | 3.81 | 10.2318
2025-04-01 11:08 | 5.7355 | 55.3731 | 5.3101 | 8.7041
2025-04-01 11:08 | 1.0415 | 43.6125 | 3.337 | 11.1484
2025-04-01 11:08 | 7.2132 | 41.3172 | 2.1931 | 1.6474
2025-04-01 11:08 | 0 | 23.6812 | 3.4762 | 15.1488
2025-04-01 11:08 | 0.97575 | 32.1102 | 5.1148 | 7.4748
2025-04-01 11:08 | 3.407 | 46.6846 | 3.7512 | 16.5861
2025-04-01 11:08 | 7.6053 | 33.2391 | 2.6593 | 18.8202
2025-04-01 11:08 | 2.9153 | 26.1249 | 6.0044 | 15.943
2025-04-01 11:08 | 5.5911 | 27.7719 | 3.7544 | 11.1721
2025-04-01 11:08 | 8.8884 | 55.2447 | 2.8549 | 10.4522
2025-04-01 11:08 | 0.99041 | 38.2567 | 2.8753 | 19.4791
2025-04-01 11:08 | 7.9555 | 47.0122 | 1.4792 | 21.632
2025-04-01 11:08 | 7.8081 | 33.1424 | 3.9507 | 8.8187
2025-04-01 11:08 | 10.6343 | 22.5485 | 4.9307 | 7.0098
2025-04-01 11:08 | 2.1247 | 30.1341 | 5.6228 | 14.0795
2025-04-01 11:08 | 2.1182 | 33.7388 | 5.9799 | 12.4367
2025-04-01 11:08 | 2.2773 | 37.0783 | 2.09 | 6.9674
2025-04-01 11:08 | 4.8127 | 30.7004 | 2.2213 | 24.5717
2025-04-01 11:08 | 7.9165 | 44.7985 | 2.9842 | 11.9826
2025-04-01 11:08 | 4.8695 | 68.1658 | 3.6543 | 5.7081
2025-04-01 11:08 | 6.0152 | 34.327 | 2.4177 | 14.6972
2025-04-01 11:08 | 3.8019 | 32.0647 | 4.4445 | 10.3471
2025-04-01 11:08 | 0 | 32.0289 | 5.7105 | 19.1778
2025-04-01 11:08 | 6.4572 | 34.666 | 0.62019 | 21.2109
2025-04-01 11:08 | 1.845 | 36.733 | 2.9159 | 10.1474
2025-04-01 11:08 | 0 | 31.4929 | 2.9801 | 12.3396
2025-04-01 11:08 | 0 | 41.8566 | 3.4693 | 11.3504
2025-04-01 11:08 | 0.48992 | 36.9054 | 3.2597 | 7.96
2025-04-01 11:09 | 6.1722 | 48.6838 | 3.5397 | 7.1658
2025-04-01 11:09 | 6.9445 | 58.9797 | 4.0036 | 16.3982
2025-04-01 11:09 | 1.9492 | 57.4532 | 4.4961 | 14.3227
2025-04-01 11:09 | 5.6039 | 25.9052 | 4.5445 | 15.6952
2025-04-01 11:09 | 12.7548 | 54.1461 | 1.6926 | 9.8471
2025-04-01 11:09 | 0 | 46.0167 | 2.7197 | 25.3212
2025-04-01 11:09 | 7.9762 | 28.2978 | 2.88 | 25.8473
2025-04-01 11:09 | 0 | 50.1107 | 1.7266 | 12.6623
2025-04-01 11:09 | 0 | 53.7153 | 4.6642 | 20.108
2025-04-01 11:09 | 0 | 26.9554 | 2.5899 | 11.1651
2025-04-01 11:09 | 0 | 42.3817 | 4.7819 | 16.6773
2025-04-01 11:09 | 0 | 24.1752 | 2.6833 | 14.8338
2025-04-01 11:09 | 4.0431 | 39.1047 | 7.6297 | 19.8424
2025-04-01 11:09 | 7.6472 | 60.9315 | 5.3783 | 15.0298
2025-04-01 11:09 | 2.0108 | 30.148 | 4.4066 | 14.2058
2025-04-01 11:09 | 4.0286 | 47.1892 | 2.6686 | 12.1694
2025-04-01 11:09 | 0 | 42.9342 | 3.8986 | 10.7054
2025-04-01 11:09 | 12.4602 | 58.5036 | 2.9568 | 8.2988
2025-04-01 11:09 | 0 | 25.2563 | 3.4474 | 17.4124
2025-04-01 11:09 | 5.4338 | 49.0785 | 2.1017 | 23.5054
2025-04-01 11:09 | 5.95 | 36.9094 | 5.911 | 12.6968
2025-04-01 11:09 | 2.5261 | 49.2325 | 6.1326 | 19.6506
2025-04-01 11:09 | 11.4885 | 29.3181 | 5.0741 | 8.2651
2025-04-01 11:09 | 0 | 49.3241 | 6.3119 | 20.9711
2025-04-01 11:09 | 3.2744 | 53.6808 | 5.954 | 10.4574
2025-04-01 11:09 | 2.2862 | 34.2799 | 0.67937 | 14.5202
2025-04-01 11:09 | 0 | 39.1692 | 2.2211 | 5.9336
2025-04-01 11:09 | 0.066955 | 38.5017 | 2.6409 | 11.8651
2025-04-01 11:09 | 3.1533 | 52.8223 | 3.1133 | 5.3414
2025-04-01 11:09 | 4.1043 | 38.3764 | 3.2651 | 17.542
2025-04-01 11:09 | 0 | 31.3106 | 2.5064 | 16.0652
2025-04-01 11:09 | 6.0145 | 32.4721 | 4.7934 | 11.4379
2025-04-01 11:09 | 8.0223 | 24.931 | 2.033 | 8.3676
2025-04-01 11:09 | 0 | 43.6011 | 6.0002 | 18.5795
2025-04-01 11:09 | 2.2787 | 57.1011 | 2.7975 | 24.6246
2025-04-01 11:09 | 2.2743 | 35.5929 | 6.1894 | 16.0319
2025-04-01 11:09 | 2.1941 | 48.8028 | 2.1619 | 9.2322
2025-04-01 11:09 | 3.5852 | 33.3642 | 4.0463 | 13.2817
2025-04-01 11:09 | 1.655 | 43.4099 | 3.6848 | 36.0881
2025-04-01 11:09 | 1.4289 | 42.2587 | 4.5173 | 24.9161
2025-04-01 11:09 | 1.8896 | 34.912 | 2.1872 | 8.5614
2025-04-01 11:09 | 2.0529 | 51.1827 | 3.7523 | 14.3443
2025-04-01 11:09 | 7.6068 | 50.6559 | 5.5877 | 9.9185
2025-04-01 11:09 | 12.4118 | 46.6372 | 3.2221 | 11.9628
2025-04-01 11:09 | 12.7054 | 46.5533 | 6.0532 | 15.3173
2025-04-01 11:09 | 9.77 | 40.4956 | 2.9412 | 17.2282
2025-04-01 11:09 | 9.9681 | 45.2873 | 0.88893 | 22.2317
2025-04-01 11:09 | 2.0238 | 47.0912 | 2.0837 | 14.4125
2025-04-01 11:09 | 3.3602 | 63.901 | 1.2152 | 16.9841
2025-04-01 11:09 | 16.803 | 42.3206 | 4.0615 | 21.7297
2025-04-01 11:09 | 7.3406 | 41.8993 | 2.3242 | 14.1664
2025-04-01 11:09 | 8.2749 | 26.8343 | 5.8309 | 7.5887
2025-04-01 11:09 | 3.4962 | 59.1501 | 2.1479 | 13.9991
2025-04-01 11:09 | 10.5895 | 49.9924 | 1.6624 | 11.4655
2025-04-01 11:09 | 13.0744 | 45.7653 | 2.796 | 8.6615
2025-04-01 11:09 | 6.4938 | 59.2413 | 3.7016 | 23.9397
2025-04-01 11:09 | 10.1793 | 35.9658 | 4.3037 | 12.5617
2025-04-01 11:09 | 0.28472 | 35.8429 | 1.5065 | 20.1388
2025-04-01 11:09 | 0 | 28.8367 | 0.96624 | 16.4931
2025-04-01 11:09 | 4.983 | 32.6885 | 6.9019 | 14.4576
2025-04-01 11:10 | 8.3425 | 49.3528 | 6.1336 | 8.1414
2025-04-01 11:10 | 4.1357 | 46.2315 | 1.9767 | 22.1473
2025-04-01 11:10 | 0 | 37.9175 | 7.2112 | 9.2635
2025-04-01 11:10 | 11.4525 | 48.0255 | 3.1479 | 13.2147
2025-04-01 11:10 | 16.4671 | 40.6841 | 5.728 | 14.9886
2025-04-01 11:10 | 10.2904 | 34.8094 | 5.6942 | 12.3737
2025-04-01 11:10 | 4.8964 | 24.804 | 5.6988 | 28.6908
2025-04-01 11:10 | 0.24648 | 65.7728 | 6.9435 | 14.6305
2025-04-01 11:10 | 1.6631 | 30.7137 | 1.784 | 17.2066
2025-04-01 11:10 | 0 | 54.5303 | 9.0139 | 1.908
2025-04-01 11:10 | 3.4289 | 42.2687 | 3.8112 | 18.342
2025-04-01 11:10 | 7.1849 | 29.9825 | 1.551 | 17.4743
2025-04-01 11:10 | 2.4591 | 38.3664 | 2.3419 | 14.2903
2025-04-01 11:10 | 8.4627 | 46.223 | 1.9862 | 13.5432
2025-04-01 11:10 | 5.0334 | 34.9596 | 2.8188 | 17.2794
2025-04-01 11:10 | 6.5293 | 35.7379 | 2.2063 | 10.3858
2025-04-01 11:10 | 5.3486 | 38.6653 | 5.3477 | 15.0961
2025-04-01 11:10 | 3.5748 | 44.3617 | 3.0601 | 11.5186
2025-04-01 11:10 | 7.3027 | 42.605 | 2.7873 | 14.0933
2025-04-01 11:10 | 9.3638 | 23.0454 | 2.7944 | 24.3478
2025-04-01 11:10 | 9.8365 | 26.502 | 2.8152 | 24.8361
2025-04-01 11:10 | 0 | 38.4616 | 1.8947 | 13.9896
2025-04-01 11:10 | 6.0459 | 41.0664 | 3.6576 | 17.6373
2025-04-01 11:10 | 14.0323 | 26.8062 | 5.0476 | 3.5034
2025-04-01 11:10 | 3.8668 | 40.9046 | 1.6773 | 6.7086
2025-04-01 11:10 | 1.092 | 44.4088 | 3.4211 | 8.1989
2025-04-01 11:10 | 0 | 37.123 | 2.5825 | 11.7944
2025-04-01 11:10 | 16.7373 | 28.9102 | 2.7161 | 9.1219
2025-04-01 11:10 | 9.2251 | 48.188 | 2.4918 | 12.0849
2025-04-01 11:10 | 9.6313 | 43.8791 | 4.4335 | 20.0153
2025-04-01 11:10 | 0.82884 | 29.6341 | 2.9629 | 13.7661
2025-04-01 11:10 | 16.1126 | 29.1451 | 0.77795 | 11.0896
2025-04-01 11:10 | 1.3143 | 36.6798 | 0.30737 | 13.9237
2025-04-01 11:10 | 0 | 61.3743 | 2.7518 | 26.5167
2025-04-01 11:10 | 0 | 43.681 | 3.9337 | 7.5634
2025-04-01 11:10 | 1.9236 | 44.05 | 3.2268 | 28.4281
2025-04-01 11:10 | 2.0118 | 37.1725 | 4.301 | 9.1893
2025-04-01 11:10 | 0.57261 | 46.1427 | 5.0641 | 13.2393
2025-04-01 11:10 | 7.0531 | 43.9625 | 2.7452 | 16.6019
2025-04-01 11:10 | 18.2674 | 30.1906 | 1.348 | 11.6991
2025-04-01 11:10 | 0 | 46.0111 | 5.2756 | 14.2671
2025-04-01 11:10 | 0 | 41.8697 | 4.1497 | 7.1538
2025-04-01 11:10 | 8.0466 | 38.897 | 4.7958 | 9.6751
2025-04-01 11:10 | 0 | 33.0941 | 3.4045 | 10.2748
2025-04-01 11:10 | 1.6487 | 28.6537 | 2.8429 | 5.5985
2025-04-01 11:10 | 4.3608 | 27.2954 | 6.532 | 26.9461
2025-04-01 11:10 | 8.5983 | 32.7979 | 2.3838 | 3.8943
2025-04-01 11:10 | 7.4629 | 26.5155 | 2.7032 | 17.5359
2025-04-01 11:10 | 0 | 43.7051 | 1.5338 | 12.349
2025-04-01 11:10 | 2.0333 | 42.8682 | 5.5128 | 14.5279
2025-04-01 11:10 | 0 | 32.5599 | 5.4549 | 15.9672
2025-04-01 11:10 | 5.575 | 43.8778 | 1.6307 | 10.9938
2025-04-01 11:10 | 2.9353 | 37.51 | 3.4478 | 16.8458
2025-04-01 11:10 | 6.8447 | 50.3749 | 0.96614 | 8.0894
2025-04-01 11:10 | 0 | 33.4964 | 4.2116 | 7.6772
2025-04-01 11:10 | 5.5379 | 32.2246 | 6.0732 | 7.7487
2025-04-01 11:10 | 5.5762 | 41.9599 | 3.8109 | 9.7581
2025-04-01 11:10 | 1.6903 | 24.6492 | 6.2378 | 17.6063
2025-04-01 11:10 | 1.08 | 32.4482 | 7.5086 | 12.7137
2025-04-01 11:10 | 0 | 30.6424 | 4.8651 | 14.5653
2025-04-01 11:11 | 0 | 28.5781 | 2.5604 | 20.3316
2025-04-01 11:11 | 3.7752 | 28.0994 | 2.0239 | 9.3089
2025-04-01 11:11 | 7.8141 | 29.331 | 2.4352 | 4.5519
2025-04-01 11:11 | 0.52395 | 42.9921 | 5.98 | 19.0785
2025-04-01 11:11 | 0 | 36.6162 | 5.0562 | 28.2802
2025-04-01 11:11 | 0 | 54.3747 | 1.2184 | 6.3719
2025-04-01 11:11 | 6.8944 | 41.8611 | 3.42 | 10.6398
2025-04-01 11:11 | 0 | 39.8299 | 2.5288 | 18.5649
2025-04-01 11:11 | 5.584 | 28.0225 | 1.0658 | 13.5969
2025-04-01 11:11 | 0 | 54.645 | 2.1542 | 9.8323
2025-04-01 11:11 | 0 | 33.9643 | 9.3937 | 17.0089
2025-04-01 11:11 | 1.883 | 32.7938 | 6.5337 | 16.3341
2025-04-01 11:11 | 0 | 48.2633 | 4.4825 | 15.4276
2025-04-01 11:11 | 8.1215 | 31.74 | 4.8389 | 8.057
2025-04-01 11:11 | 7.1476 | 43.4753 | 4.4847 | 27.2227
2025-04-01 11:11 | 5.624 | 31.2136 | 4.6819 | 16.9544
2025-04-01 11:11 | 0 | 33.5681 | 4.3117 | 6.167
2025-04-01 11:11 | 0 | 55.3969 | 1.8972 | 21.9098
2025-04-01 11:11 | 0 | 45.2079 | 1.4275 | 29.0292
2025-04-01 11:11 | 0 | 25.723 | 1.9922 | 15.0991
2025-04-01 11:11 | 3.0162 | 52.2138 | 5.5552 | 26.4546
2025-04-01 11:11 | 2.6005 | 34.6998 | 3.7403 | 10.7482
2025-04-01 11:11 | 6.9013 | 46.4697 | 3.9293 | 15.4
2025-04-01 11:11 | 13.0693 | 34.9014 | 5.4329 | 20.4507
2025-04-01 11:11 | 4.5436 | 35.1509 | 5.8703 | 7.0048
2025-04-01 11:11 | 0 | 37.2135 | 3.2392 | 10.6349
2025-04-01 11:11 | 0 | 31.6476 | 10.2904 | 11.1352
2025-04-01 11:11 | 11.2032 | 40.9887 | 1.7099 | 9.9157
2025-04-01 11:11 | 3.8759 | 28.2142 | 5.6354 | 18.9794
2025-04-01 11:11 | 6.4904 | 43.3875 | 10.4475 | 21.0254
2025-04-01 11:11 | 13.7869 | 29.8173 | 3.0211 | 22.1893
2025-04-01 11:11 | 0 | 33.4586 | 4.3836 | 19.3045
2025-04-01 11:11 | 9.0658 | 32.2797 | 3.828 | 1.4979
2025-04-01 11:11 | 14.819 | 35.5677 | 4.6933 | 12.0083
2025-04-01 11:11 | 9.0047 | 61.0589 | 5.1716 | 26.555
2025-04-01 11:11 | 10.3636 | 43.0291 | 6.0556 | 18.9845
2025-04-01 11:11 | 8.1123 | 29.4695 | 2.2729 | 12.8706
2025-04-01 11:11 | 3.977 | 24.7522 | 5.3708 | 10.0414
2025-04-01 11:11 | 8.2818 | 37.2201 | 5.4511 | 25.4056
2025-04-01 11:11 | 4.9853 | 37.013 | 3.161 | 25.7569
2025-04-01 11:11 | 4.069 | 47.7394 | 2.6791 | 16.23
2025-04-01 11:11 | 0 | 29.3776 | 3.1518 | 33.8153
2025-04-01 11:11 | 0 | 34.5755 | 4.171 | 9.7539
2025-04-01 11:11 | 3.2291 | 33.3553 | 5.3725 | 10.377
2025-04-01 11:11 | 0.29199 | 36.5021 | 3.7106 | 25.6101
2025-04-01 11:11 | 1.8534 | 44.267 | 2.8123 | 23.9298
2025-04-01 11:11 | 0 | 35.4074 | 0.9165 | 12.9372
2025-04-01 11:11 | 4.1121 | 30.0737 | 4.7736 | 14.4645
2025-04-01 11:11 | 7.6145 | 46.5755 | 3.6692 | 19.6455
2025-04-01 11:11 | 15.807 | 35.095 | 0.80303 | 18.048
2025-04-01 11:11 | 10.6738 | 39.4398 | 4.31 | 20.0384
2025-04-01 11:11 | 0 | 43.7954 | 5.1948 | 11.0286
2025-04-01 11:11 | 6.2512 | 30.5805 | 3.3355 | 14.7792
2025-04-01 11:11 | 0.59886 | 35.3494 | 2.8426 | 20.4551
2025-04-01 11:11 | 0.56165 | 35.7291 | 6.4282 | 4.8822
2025-04-01 11:11 | 9.6538 | 43.2557 | 4.481 | 6.2487
2025-04-01 11:11 | 4.4409 | 54.2585 | 7.6617 | 11.3805
2025-04-01 11:11 | 4.4745 | 37.656 | 2.5963 | 8.3259
2025-04-01 11:11 | 0 | 44.7743 | 4.4458 | 12.9961
2025-04-01 11:11 | 4.9228 | 37.969 | 1.494 | 7.5775
2025-04-01 11:12 | 0.3309 | 48.5342 | 7.2903 | 19.8649
2025-04-01 11:12 | 0 | 26.8532 | 3.6115 | 8.6205
2025-04-01 11:12 | 7.5832 | 31.7975 | 2.8179 | 25.5285
2025-04-01 11:12 | 1.2578 | 57.052 | 1.799 | 22.8853
2025-04-01 11:12 | 6.5955 | 38.9003 | 2.4697 | 22.358
2025-04-01 11:12 | 8.5137 | 55.6252 | 4.503 | 16.5757
2025-04-01 11:12 | 0 | 27.4388 | 2.2573 | 21.4231
2025-04-01 11:12 | 7.1053 | 60.2604 | 4.5649 | 19.2219
2025-04-01 11:12 | 9.9702 | 35.7101 | 1.4903 | 6.4546
2025-04-01 11:12 | 6.9662 | 37.7577 | 2.2268 | 15.1773
2025-04-01 11:12 | 2.3771 | 34.2606 | 4.8734 | 27.378
2025-04-01 11:12 | 12.9173 | 49.2528 | 5.3457 | 4.6562
2025-04-01 11:12 | 6.1154 | 33.9694 | 1.893 | 8.0241
2025-04-01 11:12 | 4.526 | 45.7185 | 1.771 | 9.1667
2025-04-01 11:12 | 1.2295 | 45.5463 | 4.6645 | 19.6207
2025-04-01 11:12 | 0 | 49.0688 | 1.7437 | 10.1187
2025-04-01 11:12 | 7.9385 | 52.1446 | 2.6445 | 15.346
2025-04-01 11:12 | 7.1261 | 36.091 | 5.4196 | 12.4089
2025-04-01 11:12 | 0.67346 | 29.0481 | 2.9552 | 10.6799
2025-04-01 11:12 | 11.257 | 44.3612 | 3.1462 | 25.8143
2025-04-01 11:12 | 5.3432 | 33.6448 | 3.8512 | 8.4179
2025-04-01 11:12 | 5.8446 | 31.8845 | 3.3163 | 8.6688
2025-04-01 11:12 | 12.8248 | 41.875 | 2.2617 | 12.5128
2025-04-01 11:12 | 12.4648 | 48.8138 | 5.4312 | 29.75
2025-04-01 11:12 | 0.83704 | 49.7509 | 4.7542 | 12.6017
2025-04-01 11:12 | 3.5033 | 26.0166 | 3.0283 | 14.5199
2025-04-01 11:12 | 10.2949 | 40.2147 | 1.3913 | 17.5109
2025-04-01 11:12 | 4.3383 | 34.1703 | 1.5445 | 7.392
2025-04-01 11:12 | 4.5755 | 45.1887 | 4.7758 | 22.4051
2025-04-01 11:12 | 9.5303 | 26.2494 | 5.1182 | 9.0501
2025-04-01 11:12 | 0 | 24.2687 | 2.1579 | 14.6999
2025-04-01 11:12 | 2.3123 | 50.5154 | 1.545 | 7.1816
2025-04-01 11:12 | 0.95478 | 30.2563 | 6.5805 | 26.9961
2025-04-01 11:12 | 3.0225 | 31.4929 | 5.7351 | 21.9608
2025-04-01 11:12 | 3.6424 | 34.1325 | 4.3918 | 16.7749
2025-04-01 11:12 | 2.5435 | 47.9899 | 0.91771 | 15.7075
2025-04-01 11:12 | 10.193 | 33.6311 | 6.9028 | 15.391
2025-04-01 11:12 | 0 | 56.2137 | 3.3965 | 5.8896
2025-04-01 11:12 | 4.6339 | 45.0582 | 4.4619 | 16.2785
2025-04-01 11:12 | 0.55271 | 37.6254 | 1.7035 | 22.8585
2025-04-01 11:12 | 4.0361 | 39.7338 | 1.8859 | 8.483
2025-04-01 11:12 | 4.4881 | 47.6948 | 3.1842 | 10.6134
2025-04-01 11:12 | 4.7718 | 32.3494 | 4.4095 | 20.5256
2025-04-01 11:12 | 0 | 27.9569 | 2.8809 | 7.5663
2025-04-01 11:12 | 0 | 55.5685 | 3.3396 | 14.268
2025-04-01 11:12 | 0.43871 | 33.2869 | 3.1951 | 10.0529
2025-04-01 11:12 | 6.7523 | 36.852 | 1.1312 | 12.9474
2025-04-01 11:12 | 0 | 26.7123 | 2.3834 | 17.6399
2025-04-01 11:12 | 0 | 33.1124 | 2.5397 | 16.7399
2025-04-01 11:12 | 5.2425 | 50.1602 | 4.5348 | 10.8038
2025-04-01 11:12 | 2.0293 | 30.6781 | 3.4089 | 14.5831
2025-04-01 11:12 | 0 | 37.4412 | 2.9113 | 10.8103
2025-04-01 11:12 | 0 | 36.8375 | 5.1551 | 10.9085
2025-04-01 11:12 | 4.4337 | 40.6201 | 8.1682 | 14.9233
2025-04-01 11:12 | 0 | 29.4788 | 5.6715 | 16.5914
2025-04-01 11:12 | 0 | 39.3701 | 5.5225 | 11.1658
2025-04-01 11:12 | 9.2324 | 33.7691 | 1.5929 | 41.1329
2025-04-01 11:12 | 3.3148 | 29.9837 | 5.4145 | 18.6001
2025-04-01 11:12 | 0 | 35.0181 | 1.9812 | 5.0346
2025-04-01 11:12 | 7.1666 | 24.9427 | 3.6368 | 11.6055
2025-04-01 11:13 | 2.83 | 48.3649 | 1.9312 | 27.5916
2025-04-01 11:13 | 3.0587 | 35.6132 | 2.668 | 9.7935
2025-04-01 11:13 | 3.6884 | 47.3993 | 1.8311 | 17.0124
2025-04-01 11:13 | 12.6527 | 33.8407 | 9.1706 | 8.5986
2025-04-01 11:13 | 1.5296 | 31.7175 | 2.9325 | 22.2164
2025-04-01 11:13 | 6.5177 | 33.0154 | 3.6069 | 25.9445
2025-04-01 11:13 | 5.5975 | 40.2773 | 4.8839 | 10.1397
2025-04-01 11:13 | 8.3149 | 33.7904 | 3.3095 | 19.2475
2025-04-01 11:13 | 8.4174 | 42.116 | 0.71078 | 11.9317
2025-04-01 11:13 | 0 | 46.6405 | 1.9992 | 10.52
2025-04-01 11:13 | 9.6243 | 41.2567 | 4.435 | 10.3249
2025-04-01 11:13 | 1.1594 | 31.2993 | 2.4568 | 24.7886
2025-04-01 11:13 | 3.6809 | 33.5678 | 9.1678 | 19.0386
2025-04-01 11:13 | 8.7523 | 47.7109 | 4.8401 | 20.354
2025-04-01 11:13 | 0 | 38.3489 | 3.6822 | 17.3956
2025-04-01 11:13 | 0.82537 | 35.1623 | 2.3429 | 17.1269
2025-04-01 11:13 | 12.4528 | 44.1275 | 3.4399 | 15.0369
2025-04-01 11:13 | 9.371 | 45.2087 | 5.4516 | 23.2769
2025-04-01 11:13 | 2.7596 | 23.3767 | 7.6455 | 11.6871
2025-04-01 11:13 | 2.9099 | 33.351 | 3.2278 | 12.8246
2025-04-01 11:13 | 12.2404 | 41.4213 | 4.9627 | 22.9891
2025-04-01 11:13 | 2.6031 | 32.3481 | 1.9112 | 14.4719
2025-04-01 11:13 | 4.2082 | 36.0751 | 5.6721 | 16.6822
2025-04-01 11:13 | 10.3266 | 47.7742 | 4.2444 | 10.9421
2025-04-01 11:13 | 0.39184 | 31.7844 | 6.0637 | 17.8863
2025-04-01 11:13 | 7.9255 | 54.5477 | 2.5985 | 23.2578
2025-04-01 11:13 | 7.8676 | 55.6734 | 1.082 | 11.0507
2025-04-01 11:13 | 3.6649 | 49.3076 | 3.477 | 13.343
2025-04-01 11:13 | 8.9833 | 41.0115 | 3.1624 | 23.046
2025-04-01 11:13 | 8.6518 | 34.5122 | 4.4659 | 19.4278
2025-04-01 11:13 | 9.0579 | 34.4358 | 4.4036 | 17.3242
2025-04-01 11:13 | 6.439 | 40.0101 | 2.1255 | 18.9999
2025-04-01 11:13 | 12.085 | 42.3723 | 4.9312 | 7.5115
2025-04-01 11:13 | 5.6968 | 52.0021 | 5.8734 | 12.9672
2025-04-01 11:13 | 9.33 | 37.059 | 1.6917 | 11.3096
2025-04-01 11:13 | 1.2897 | 25.2193 | 2.1676 | 20.011
2025-04-01 11:13 | 10.2351 | 34.8992 | 1.8582 | 25.4317
2025-04-01 11:13 | 1.9632 | 36.1151 | 3.9257 | 25.9569
2025-04-01 11:13 | 6.2962 | 19.0836 | 5.2749 | 14.5078
2025-04-01 11:13 | 0 | 46.6841 | 5.7078 | 14.284
2025-04-01 11:13 | 0 | 53.8272 | 3.4784 | 15.6193
2025-04-01 11:13 | 0 | 36.173 | 4.8207 | 8.6589
2025-04-01 11:13 | 0.88925 | 28.6215 | 2.527 | 16.2728
2025-04-01 11:13 | 9.9445 | 56.0733 | 3.0121 | 10.3023
2025-04-01 11:13 | 5.424 | 33.4994 | 4.3582 | 13.5678
2025-04-01 11:13 | 0 | 50.5012 | 4.2604 | 11.7315
2025-04-01 11:13 | 0.45805 | 34.4801 | 9.4155 | 11.8859
2025-04-01 11:13 | 0 | 46.9953 | 3.4336 | 13.8679
2025-04-01 11:13 | 8.4223 | 57.5633 | 3.5233 | 14.4651
2025-04-01 11:13 | 3.5862 | 43.5178 | 7.3151 | 18.0064
2025-04-01 11:13 | 8.2534 | 45.1896 | 4.3608 | 14.1222
2025-04-01 11:13 | 9.8115 | 47.7384 | 5.1907 | 12.3219
2025-04-01 11:13 | 13.813 | 41.9264 | 6.0623 | 10.3804
2025-04-01 11:13 | 0 | 46.0952 | 5.3752 | 19.1217
2025-04-01 11:13 | 4.2122 | 33.4051 | 6.7131 | 12.8968
2025-04-01 11:13 | 0 | 31.5378 | 2.3024 | 10.9644
2025-04-01 11:13 | 0.78314 | 37.743 | 8.7337 | 21.1063
2025-04-01 11:13 | 2.0228 | 51.2833 | 4.5817 | 11.9595
2025-04-01 11:13 | 5.1042 | 36.8329 | 1.2396 | 13.4901
2025-04-01 11:13 | 10.0071 | 32.7331 | 1.4919 | 16.0943
2025-04-01 11:14 | 5.2765 | 43.7408 | 3.3508 | 9.1525
2025-04-01 11:14 | 7.5594 | 33.3571 | 5.4749 | 29.8504
2025-04-01 11:14 | 3.6839 | 31.3334 | 5.3852 | 9.8984
2025-04-01 11:14 | 8.8253 | 40.7186 | 3.9066 | 4.5178
2025-04-01 11:14 | 6.7591 | 31.6451 | 4.4353 | 27.8653
2025-04-01 11:14 | 6.4967 | 25.9189 | 2.6914 | 17.0365
2025-04-01 11:14 | 0.0085085 | 21.3 | 3.4042 | 11.9399
2025-04-01 11:14 | 10.1816 | 28.211 | 3.1559 | 8.3936
2025-04-01 11:14 | 1.6408 | 29.3683 | 1.2759 | 26.3087
2025-04-01 11:14 | 2.2429 | 42.2955 | 1.5561 | 12.6031
2025-04-01 11:14 | 0 | 39.9156 | 3.9871 | 10.349
2025-04-01 11:14 | 9.1331 | 56.966 | 3.0613 | 10.7969
2025-04-01 11:14 | 9.6043 | 45.5644 | 2.6123 | 14.27
2025-04-01 11:14 | 7.3646 | 50.4698 | 2.4655 | 17.0475
2025-04-01 11:14 | 6.4773 | 22.378 | 2.4937 | 9.1144
2025-04-01 11:14 | 0 | 52.0509 | 5.3312 | 5.4817
2025-04-01 11:14 | 4.5673 | 40.0943 | 3.2064 | 13.6155
2025-04-01 11:14 | 0.65827 | 47.6964 | 1.1042 | 5.257
2025-04-01 11:14 | 0 | 36.4149 | 2.0819 | 13.4254
2025-04-01 11:14 | 1.8412 | 32.9992 | 1.774 | 22.9124
2025-04-01 11:14 | 0 | 40.41 | 5.5646 | 23.4733
2025-04-01 11:14 | 10.7759 | 42.8903 | 1.3377 | 11.318
2025-04-01 11:14 | 7.785 | 51.0674 | 6.5432 | 16.3511
2025-04-01 11:14 | 11.1558 | 49.5943 | 3.2594 | 4.1694
2025-04-01 11:14 | 6.868 | 19.2042 | 4.7375 | 12.0596
2025-04-01 11:14 | 5.1358 | 44.0745 | 3.7373 | 23.5702
2025-04-01 11:14 | 13.4729 | 48.953 | 6.151 | 9.2064
2025-04-01 11:14 | 14.5818 | 26.5216 | 3.3637 | 24.677
2025-04-01 11:14 | 5.2176 | 37.2883 | 2.5292 | 20.5768
2025-04-01 11:14 | 5.1362 | 32.6008 | 1.3149 | 17.7433
2025-04-01 11:14 | 4.2087 | 28.8091 | 3.9464 | 5.4008
2025-04-01 11:14 | 3.5117 | 28.3468 | 6.8166 | 12.1497
2025-04-01 11:14 | 1.6767 | 34.4845 | 5.4105 | 31.0396
2025-04-01 11:14 | 6.8851 | 38.9679 | 4.0571 | 17.8294
2025-04-01 11:14 | 0 | 33.9477 | 1.8429 | 10.033
2025-04-01 11:14 | 5.2584 | 40.1731 | 2.762 | 17.2391
2025-04-01 11:14 | 3.3837 | 50.9452 | 2.2533 | 13.0697
2025-04-01 11:14 | 3.5399 | 41.8333 | 1.6852 | 6.5017
2025-04-01 11:14 | 3.3876 | 41.8411 | 9.1638 | 17.1215
2025-04-01 11:14 | 3.7517 | 24.2404 | 3.8266 | 14.1537
2025-04-01 11:14 | 7.0169 | 40.051 | 4.7751 | 12.4761
2025-04-01 11:14 | 0 | 39.8204 | 6.3407 | 10.619
2025-04-01 11:14 | 1.9689 | 37.686 | 1.0644 | 8.009
2025-04-01 11:14 | 14.5972 | 40.5872 | 4.061 | 7.2084
2025-04-01 11:14 | 1.3805 | 38.8037 | 1.2464 | 11.6764
2025-04-01 11:14 | 6.3532 | 40.8959 | 3.0015 | 15.1438
2025-04-01 11:14 | 0 | 32.4311 | 8.3574 | 21.2754
2025-04-01 11:14 | 0 | 38.5695 | 3.2411 | 21.0721
2025-04-01 11:14 | 4.5167 | 24.1307 | 1.9007 | 16.4544
2025-04-01 11:14 | 6.1796 | 49.1684 | 3.9185 | 14.0044
2025-04-01 11:14 | 0 | 32.8077 | 2.0257 | 14.1717
2025-04-01 11:14 | 4.789 | 41.5283 | 1.7004 | 6.7419
2025-04-01 11:14 | 0 | 33.0128 | 1.8545 | 31.2437
2025-04-01 11:14 | 8.2741 | 41.2766 | 4.9418 | 16.6654
2025-04-01 11:14 | 9.2334 | 39.839 | 1.7305 | 12.3576
2025-04-01 11:14 | 12.0296 | 43.7906 | 2.6034 | 16.5994
2025-04-01 11:14 | 12.2808 | 24.7496 | 4.4477 | 7.006
2025-04-01 11:14 | 0 | 46.9076 | 8.7702 | 5.8831
2025-04-01 11:14 | 7.8473 | 16.7386 | 2.5232 | 9.3178
2025-04-01 11:14 | 0 | 37.1165 | 1.0954 | 6.2325
2025-04-01 11:15 | 4.4155 | 38.4641 | 3.6399 | 10.0945
2025-04-01 11:15 | 0.61011 | 31.9772 | 2.0643 | 14.3534
2025-04-01 11:15 | 0 | 33.1348 | 4.2569 | 20.1764
2025-04-01 11:15 | 0 | 22.3666 | 4.1975 | 17.3188
2025-04-01 11:15 | 4.5354 | 37.7911 | 2.163 | 7.6432
2025-04-01 11:15 | 0.67088 | 31.8144 | 2.7531 | 10.5283
2025-04-01 11:15 | 3.5881 | 48.1843 | 3.9826 | 10.641
2025-04-01 11:15 | 5.6646 | 48.7536 | 6.0594 | 15.465
2025-04-01 11:15 | 3.2814 | 36.3229 | 3.4927 | 18.6577
2025-04-01 11:15 | 8.8174 | 43.34 | 2.5239 | 18.0831
2025-04-01 11:15 | 4.419 | 41.4754 | 4.4747 | 10.9669
2025-04-01 11:15 | 0 | 45.4824 | 4.3452 | 13.5059
2025-04-01 11:15 | 5.8888 | 37.3257 | 3.8665 | 7.8756
2025-04-01 11:15 | 6.1967 | 44.2228 | 1.9563 | 22.116
2025-04-01 11:15 | 7.2601 | 20.7917 | 7.0861 | 23.1431
2025-04-01 11:15 | 7.1843 | 28.2034 | 4.0357 | 8.3264
2025-04-01 11:15 | 5.4711 | 44.8224 | 2.6566 | 8.3322
2025-04-01 11:15 | 2.5356 | 24.3434 | 1.2735 | 3.3343
2025-04-01 11:15 | 0.68053 | 29.5971 | 5.0167 | 7.0634
2025-04-01 11:15 | 0 | 28.6381 | 3.1792 | 10.6341
2025-04-01 11:15 | 0 | 32.9682 | 2.878 | 28.6333
2025-04-01 11:15 | 0 | 27.7919 | 4.248 | 5.7126
2025-04-01 11:15 | 0 | 31.726 | 1.4626 | 24.0571
2025-04-01 11:15 | 12.2963 | 40.0446 | 2.9729 | 19.0117
2025-04-01 11:15 | 6.1218 | 39.6927 | 2.5461 | 7.6255
2025-04-01 11:15 | 5.0879 | 40.3933 | 1.517 | 14.632
2025-04-01 11:15 | 1.2366 | 42.7427 | 2.4576 | 13.6585
2025-04-01 11:15 | 2.0158 | 31.7413 | 3.4435 | 11.9357
2025-04-01 11:15 | 4.4941 | 52.8702 | 3.5309 | 7.5347
2025-04-01 11:15 | 7.9662 | 36.2896 | 3.036 | 6.5111
2025-04-01 11:15 | 12.28 | 27.0586 | 1.7306 | 24.4208
2025-04-01 11:15 | 6.0444 | 38.3398 | 4.9963 | 11.5461
2025-04-01 11:15 | 11.154 | 36.5936 | 2.6538 | 17.4397
2025-04-01 11:15 | 0 | 31.312 | 5.9447 | 9.8015
2025-04-01 11:15 | 7.6568 | 59.4466 | 4.5518 | 9.5198
2025-04-01 11:15 | 5.4568 | 29.5237 | 3.8568 | 18.0726
2025-04-01 11:15 | 0 | 35.2494 | 1.7651 | 17.5842
2025-04-01 11:15 | 7.3049 | 46.4735 | 2.5254 | 12.2597
2025-04-01 11:15 | 7.4264 | 52.8748 | 1.2145 | 8.3519
2025-04-01 11:15 | 0 | 26.266 | 6.1069 | 10.4164
2025-04-01 11:15 | 11.0826 | 46.0282 | 2.7812 | 4.4348
2025-04-01 11:15 | 0 | 47.3897 | 6.0774 | 8.2202
2025-04-01 11:15 | 19.6278 | 31.8226 | 2.5695 | 7.4872
2025-04-01 11:15 | 5.5301 | 27.7584 | 4.107 | 8.6025
2025-04-01 11:15 | 0 | 32.7531 | 3.9644 | 15.4671
2025-04-01 11:15 | 1.4275 | 32.1488 | 1.7338 | 20.3799
2025-04-01 11:15 | 20.573 | 23.9302 | 2.122 | 21.7162
2025-04-01 11:15 | 7.0304 | 36.4326 | 3.7828 | 19.1184
2025-04-01 11:15 | 0 | 46.6732 | 2.3185 | 9.8938
2025-04-01 11:15 | 9.2246 | 39.3255 | 4.5606 | 6.4089
2025-04-01 11:15 | 9.318 | 32.2135 | 3.5236 | 10.394
2025-04-01 11:15 | 5.2264 | 36.4954 | 2.8274 | 17.5272
2025-04-01 11:15 | 0 | 39.9219 | 4.0655 | 13.7094
2025-04-01 11:15 | 6.9407 | 30.629 | 7.0003 | 12.3033
2025-04-01 11:15 | 11.1703 | 35.7497 | 3.4637 | 5.7383
2025-04-01 11:15 | 3.9494 | 41.5658 | 2.6426 | 17.4888
2025-04-01 11:15 | 6.1534 | 37.916 | 4.3055 | 10.6176
2025-04-01 11:15 | 0.4088 | 38.937 | 4.796 | 30.301
2025-04-01 11:15 | 6.2795 | 35.2663 | 4.3739 | 19.0556
2025-04-01 11:15 | 0 | 33.02 | 2.1583 | 10.3197
2025-04-01 11:16 | 0.069909 | 34.8509 | 2.2203 | 20.5341
2025-04-01 11:16 | 0.62358 | 43.1008 | 5.8688 | 17.5584
2025-04-01 11:16 | 0 | 34.3902 | 1.9895 | 6.1917
2025-04-01 11:16 | 0 | 50.4611 | 3.2317 | 6.5502
2025-04-01 11:16 | 5.0036 | 45.1123 | 2.4666 | 13.8019
2025-04-01 11:16 | 10.7399 | 48.9399 | 8.0918 | 20.9963
2025-04-01 11:16 | 5.8314 | 30.4898 | 3.8393 | 13.1992
2025-04-01 11:16 | 0 | 42.0492 | 4.0996 | 12.9484
2025-04-01 11:16 | 0 | 39.167 | 1.3614 | 11.8322
2025-04-01 11:16 | 5.9342 | 26.1762 | 1.1264 | 18.4192
2025-04-01 11:16 | 6.3641 | 37.8837 | 3.6137 | 8.9798
2025-04-01 11:16 | 1.1553 | 23.5444 | 3.6988 | 12.6815
2025-04-01 11:16 | 5.4451 | 24.8729 | 3.4023 | 25.5541
2025-04-01 11:16 | 4.4921 | 36.4753 | 3.8243 | 13.4145
2025-04-01 11:16 | 8.0116 | 38.5973 | 3.5327 | 15.2815
2025-04-01 11:16 | 4.2114 | 40.4688 | 2.5091 | 9.8879
2025-04-01 11:16 | 9.6689 | 51.161 | 3.4913 | 12.3665
2025-04-01 11:16 | 0 | 43.4401 | 3.4721 | 12.6467
2025-04-01 11:16 | 4.8322 | 29.9719 | 4.9386 | 7.7501
2025-04-01 11:16 | 5.0951 | 28.1689 | 3.2379 | 14.388
2025-04-01 11:16 | 8.5168 | 46.6483 | 2.8036 | 5.9415
2025-04-01 11:16 | 10.5129 | 59.6784 | 1.2474 | 13.4848
2025-04-01 11:16 | 8.8736 | 61.8495 | 5.5221 | 26.3947
2025-04-01 11:16 | 5.1847 | 46.4037 | 4.9264 | 10.6925
2025-04-01 11:16 | 3.785 | 30.4511 | 1.628 | 12.5603
2025-04-01 11:16 | 15.7864 | 38.3639 | 6.6179 | 19.0312
2025-04-01 11:16 | 0.65682 | 42.2661 | 6.5069 | 27.8116
2025-04-01 11:16 | 2.4395 | 23.4756 | 5.3384 | 11.7326
2025-04-01 11:16 | 11.1989 | 37.4897 | 2.1721 | 12.8994
2025-04-01 11:16 | 0 | 37.6101 | 4.0432 | 26.8175
2025-04-01 11:16 | 0 | 39.7204 | 5.9986 | 19.7842
2025-04-01 11:16 | 0 | 40.4696 | 5.1889 | 6.2172
2025-04-01 11:16 | 3.7636 | 28.7213 | 2.8254 | 20.49
2025-04-01 11:16 | 1.8835 | 56.3682 | 1.8093 | 11.5287
2025-04-01 11:16 | 0 | 34.4223 | 5.0887 | 11.2778
2025-04-01 11:16 | 3.4855 | 17.9585 | 4.3012 | 22.3618
2025-04-01 11:16 | 6.0117 | 46.9512 | 2.5834 | 12.193
2025-04-01 11:16 | 1.745 | 32.6122 | 6.9462 | 11.3681
2025-04-01 11:16 | 1.5621 | 34.7264 | 4.3692 | 8.1627
2025-04-01 11:16 | 0 | 31.9895 | 2.1576 | 26.4714
2025-04-01 11:16 | 4.0019 | 19.0333 | 5.0357 | 15.0458
2025-04-01 11:16 | 0 | 35.862 | 2.6575 | 16.3339
2025-04-01 11:16 | 1.5899 | 27.1549 | 3.6046 | 28.7766
2025-04-01 11:16 | 6.2243 | 43.5697 | 7.612 | 20.8297
2025-04-01 11:16 | 12.1173 | 50.6788 | 4.6085 | 14.9743
2025-04-01 11:16 | 4.7014 | 55.9819 | 6.9031 | 6.4715
2025-04-01 11:16 | 7.1099 | 34.9322 | 7.3412 | 31.9809
2025-04-01 11:16 | 2.8807 | 43.3871 | 6.1425 | 9.2919
2025-04-01 11:16 | 3.5834 | 40.8861 | 5.2651 | 7.6221
2025-04-01 11:16 | 0.95684 | 34.0897 | 0.99616 | 19.4096
2025-04-01 11:16 | 2.196 | 39.4335 | 4.1251 | 16.0351
2025-04-01 11:16 | 0 | 32.1789 | 5.4361 | 5.7101
2025-04-01 11:16 | 9.8014 | 50.3466 | 2.95 | 8.4584
2025-04-01 11:16 | 0 | 32.0364 | 3.1709 | 11.5986
2025-04-01 11:16 | 6.8364 | 41.3178 | 4.5737 | 21.3614
2025-04-01 11:16 | 2.4804 | 29.2955 | 1.8609 | 6.2644
2025-04-01 11:16 | 3.3904 | 25.6285 | 5.7472 | 11.9684
2025-04-01 11:16 | 1.2361 | 35.5802 | 2.5578 | 11.1278
2025-04-01 11:16 | 4.1414 | 31.6418 | 4.4536 | 21.3204
2025-04-01 11:16 | 2.958 | 54.1342 | 3.0258 | 9.3906
2025-04-01 11:17 | 9.5264 | 26.4335 | 3.7472 | 9.8798
2025-04-01 11:17 | 5.7902 | 46.2888 | 2.6224 | 9.0043
2025-04-01 11:17 | 0.47868 | 32.8179 | 3.752 | 12.3607
2025-04-01 11:17 | 3.9694 | 39.5574 | 5.0737 | 8.7643
2025-04-01 11:17 | 5.562 | 31.3164 | 1.8136 | 13.6644
2025-04-01 11:17 | 0.9288 | 36.4634 | 2.6892 | 12.6751
2025-04-01 11:17 | 5.4796 | 32.5218 | 0.95035 | 17.5562
2025-04-01 11:17 | 0 | 35.8197 | 1.7768 | 12.1084
2025-04-01 11:17 | 0 | 42.3561 | 1.2093 | 17.5512
2025-04-01 11:17 | 5.983 | 48.34 | 5.4791 | 14.4764
2025-04-01 11:17 | 0 | 61.3998 | 2.7535 | 17.5128
2025-04-01 11:17 | 3.438 | 43.3119 | 4.4244 | 14.2492
2025-04-01 11:17 | 3.3308 | 37.4146 | 5.155 | 6.4069
2025-04-01 11:17 | 0 | 29.611 | 0.79682 | 17.4515
2025-04-01 11:17 | 6.1147 | 54.0178 | 2.6186 | 18.249
2025-04-01 11:17 | 6.8569 | 38.266 | 4.8378 | 17.3736
2025-04-01 11:17 | 1.5181 | 42.2815 | 1.51 | 19.9838
2025-04-01 11:17 | 0 | 43.5073 | 4.3235 | 12.741
2025-04-01 11:17 | 5.0744 | 29.4997 | 3.5916 | 12.7465
2025-04-01 11:17 | 1.9661 | 32.5196 | 3.3845 | 13.9712
2025-04-01 11:17 | 8.2657 | 28.0021 | 4.9158 | 13.1628
2025-04-01 11:17 | 6.3613 | 37.1535 | 1.6917 | 24.6547
2025-04-01 11:17 | 9.4922 | 39.951 | 4.922 | 10.6664
2025-04-01 11:17 | 11.1525 | 26.4746 | 1.3547 | 14.6976
2025-04-01 11:17 | 6.0414 | 33.7241 | 4.8977 | 17.4905
2025-04-01 11:17 | 0 | 33.2894 | 4.3887 | 16.9631
2025-04-01 11:17 | 2.9925 | 53.3078 | 3.7776 | 21.2637
2025-04-01 11:17 | 8.9623 | 36.1188 | 2.086 | 5.1947
2025-04-01 11:17 | 2.9777 | 40.6959 | 3.2085 | 7.4587
2025-04-01 11:17 | 1.1284 | 42.6993 | 3.1003 | 11.2342
2025-04-01 11:17 | 0.92037 | 32.0518 | 3.0252 | 25.9261
2025-04-01 11:17 | 0 | 35.2134 | 3.114 | 11.0797
2025-04-01 11:17 | 0 | 32.912 | 4.1388 | 15.5181
2025-04-01 11:17 | 1.8679 | 37.2004 | 2.9671 | 17.4551
2025-04-01 11:17 | 0.86271 | 40.4814 | 2.3799 | 20.7942
2025-04-01 11:17 | 8.378 | 37.3652 | 4.831 | 17.5454
2025-04-01 11:17 | 20.523 | 39.2051 | 2.7696 | 11.2893
2025-04-01 11:17 | 7.7316 | 46.9295 | 3.4943 | 12.4137
2025-04-01 11:17 | 4.5817 | 44.838 | 1.2505 | 32.1797
2025-04-01 11:17 | 4.0713 | 33.1212 | 3.6068 | 17.3662
2025-04-01 11:17 | 9.4951 | 34.534 | 5.0949 | 11.0368
2025-04-01 11:17 | 0 | 34.5139 | 3.1612 | 11.5902
2025-04-01 11:17 | 7.5895 | 29.5247 | 1.1546 | 21.2434
2025-04-01 11:17 | 1.7676 | 41.879 | 1.6368 | 15.6687
2025-04-01 11:17 | 7.9631 | 41.9943 | 3.509 | 16.6638
2025-04-01 11:17 | 2.8588 | 36.5807 | 1.2565 | 8.2539
2025-04-01 11:17 | 0 | 24.9034 | 2.1307 | 17.9439
2025-04-01 11:17 | 2.2078 | 29.8633 | 5.3778 | 21.855
2025-04-01 11:17 | 7.2994 | 36.9345 | 3.2643 | 18.56
2025-04-01 11:17 | 7.1109 | 41.4288 | 1.5244 | 23.2904
2025-04-01 11:17 | 4.9769 | 31.1699 | 2.3423 | 15.4982
2025-04-01 11:17 | 0 | 31.3001 | 2.8132 | 18.7691
2025-04-01 11:17 | 0 | 48.1669 | 7.7923 | 11.6088
2025-04-01 11:17 | 1.3772 | 40.8436 | 7.0316 | 10.6328
2025-04-01 11:17 | 4.0916 | 37.2403 | 3.4666 | 20.6611
2025-04-01 11:17 | 3.3135 | 27.7866 | 4.2325 | 12.1825
2025-04-01 11:17 | 0.071563 | 52.7084 | 3.8268 | 13.1636
2025-04-01 11:17 | 5.4237 | 30.6301 | 3.6074 | 24.1156
2025-04-01 11:17 | 0 | 27.3592 | 4.1083 | 3.6518
2025-04-01 11:17 | 11.2565 | 28.7189 | 5.042 | 20.9626
2025-04-01 11:18 | 8.0498 | 40.4367 | 4.6383 | 18.6655
2025-04-01 11:18 | 11.1436 | 24.3407 | 5.5311 | 14.2724
2025-04-01 11:18 | 10.3421 | 40.8467 | 2.8359 | 6.2045
2025-04-01 11:18 | 9.4407 | 28.7507 | 2.324 | 19.6259
2025-04-01 11:18 | 0.94251 | 33.4611 | 3.5427 | 7.7698
2025-04-01 11:18 | 0 | 40.495 | 7.6712 | 15.1011
2025-04-01 11:18 | 13.3806 | 32.7672 | 1.5514 | 16.3859
2025-04-01 11:18 | 6.4598 | 36.487 | 5.8937 | 28.3705
2025-04-01 11:18 | 2.5724 | 38.7059 | 2.7332 | 4.3477
2025-04-01 11:18 | 0 | 27.6822 | 4.7997 | 18.5249
2025-04-01 11:18 | 8.6354 | 35.8447 | 2.0644 | 16.8785
2025-04-01 11:18 | 20.4986 | 37.9948 | 5.0275 | 26.4141
2025-04-01 11:18 | 4.9516 | 42.5739 | 3.2723 | 15.9888
2025-04-01 11:18 | 6.3403 | 44.9124 | 4.7722 | 15.8625
2025-04-01 11:18 | 7.1683 | 40.4844 | 2.9723 | 20.7201
2025-04-01 11:18 | 0.79853 | 31.8926 | 2.6815 | 20.6445
2025-04-01 11:18 | 2.5723 | 53.013 | 3.8662 | 14.1462
2025-04-01 11:18 | 9.9606 | 43.3001 | 5.9718 | 15.7811
2025-04-01 11:18 | 5.6803 | 31.703 | 6.7249 | 12.6465
2025-04-01 11:18 | 10.0818 | 34.1756 | 3.8757 | 10.2042
2025-04-01 11:18 | 6.9445 | 36.3204 | 1.7484 | 8.9274
2025-04-01 11:18 | 3.1748 | 38.7367 | 7.8792 | 13.816
2025-04-01 11:18 | 0.85865 | 39.4001 | 2.1904 | 14.703
2025-04-01 11:18 | 1.6174 | 29.6643 | 4.4033 | 21.5669
2025-04-01 11:18 | 0 | 47.2975 | 9.4313 | 15.8078
2025-04-01 11:18 | 1.5588 | 34.3256 | 3.495 | 12.2013
2025-04-01 11:18 | 0 | 35.7412 | 5.3727 | 6.7898
2025-04-01 11:18 | 0 | 24.6249 | 9.449 | 14.2358
2025-04-01 11:18 | 4.4706 | 35.5185 | 2.8749 | 23.233
2025-04-01 11:18 | 21.0763 | 20.8791 | 2.7339 | 10.1702
2025-04-01 11:18 | 12.0235 | 18.776 | 2.5019 | 8.9734
2025-04-01 11:18 | 17.1351 | 26.0226 | 0.34585 | 23.4071
2025-04-01 11:18 | 5.7096 | 32.222 | 1.8027 | 6.1059
2025-04-01 11:18 | 12.2074 | 27.5465 | 2.9037 | 19.6851
2025-04-01 11:18 | 1.3405 | 40.1336 | 2.8237 | 13.3899
2025-04-01 11:18 | 0.7297 | 28.5693 | 2.7464 | 15.4463
2025-04-01 11:18 | 3.6327 | 36.7885 | 1.8835 | 7.5866
2025-04-01 11:18 | 6.3083 | 36.216 | 3.4347 | 18.9123
2025-04-01 11:18 | 6.4518 | 25.6658 | 1.3674 | 9.3034
2025-04-01 11:18 | 3.588 | 25.7166 | 3.3347 | 20.9881
2025-04-01 11:18 | 0 | 25.4279 | 3.5439 | 15.1238
2025-04-01 11:18 | 0 | 33.6047 | 5.1565 | 14.4218
2025-04-01 11:18 | 1.3492 | 40.3931 | 1.0079 | 19.7328
2025-04-01 11:18 | 6.5389 | 35.8617 | 3.2752 | 14.2675
2025-04-01 11:18 | 0 | 34.6331 | 2.9268 | 20.9591
2025-04-01 11:18 | 2.715 | 26.0175 | 3.2059 | 20.4561
2025-04-01 11:18 | 1.6026 | 23.2414 | 2.4833 | 13.2339
2025-04-01 11:18 | 0 | 34.9666 | 5.1262 | 15.7243
2025-04-01 11:18 | 3.971 | 33.0997 | 0.56559 | 8.5163
2025-04-01 11:18 | 9.4659 | 27.2025 | 8.3896 | 25.9866
2025-04-01 11:18 | 8.9091 | 31.8374 | 2.7668 | 20.9977
2025-04-01 11:18 | 0 | 30.7232 | 3.64 | 9.2927
2025-04-01 11:18 | 5.9251 | 33.4615 | 2.7197 | 28.8191
2025-04-01 11:18 | 17.0598 | 21.6241 | 4.0023 | 18.6597
2025-04-01 11:18 | 5.2845 | 35.9074 | 3.4581 | 9.3937
2025-04-01 11:18 | 0 | 48.979 | 3.772 | 6.4019
2025-04-01 11:18 | 0 | 36.4046 | 2.9284 | 25.4119
2025-04-01 11:18 | 9.0726 | 32.6222 | 0.71119 | 7.2679
2025-04-01 11:18 | 9.5679 | 41.3845 | 2.2932 | 12.3477
2025-04-01 11:18 | 7.8794 | 28.6911 | 1.9606 | 16.8442
2025-04-01 11:19 | 0 | 36.142 | 3.5443 | 31.9404
2025-04-01 11:19 | 3.2055 | 34.1109 | 4.1086 | 14.734
2025-04-01 11:19 | 0 | 55.6667 | 3.0354 | 19.6373
2025-04-01 11:19 | 4.0314 | 24.5742 | 2.8067 | 18.576
2025-04-01 11:19 | 2.6856 | 25.72 | 3.0708 | 6.0445
2025-04-01 11:19 | 1.6349 | 31.5443 | 3.5826 | 18.7066
2025-04-01 11:19 | 3.9079 | 46.7165 | 2.9039 | 14.6243
2025-04-01 11:19 | 3.7162 | 47.9962 | 6.5053 | 8.182
2025-04-01 11:19 | 5.1632 | 50.6498 | 2.7104 | 13.9872
2025-04-01 11:19 | 0.75111 | 29.8763 | 3.8901 | 6.9907
2025-04-01 11:19 | 9.3726 | 45.6685 | 1.6754 | 13.9575
2025-04-01 11:19 | 6.5582 | 30.4914 | 11.0461 | 9.0826
2025-04-01 11:19 | 0.081287 | 31.2831 | 3.5237 | 11.3433
2025-04-01 11:19 | 8.8387 | 30.9577 | 3.4329 | 6.8998
2025-04-01 11:19 | 10.5855 | 30.2783 | 3.9735 | 1.3405
2025-04-01 11:19 | 3.0947 | 33.9536 | 4.8132 | 27.3846
2025-04-01 11:19 | 6.5906 | 33.6603 | 3.1625 | 20.3867
2025-04-01 11:19 | 0.31134 | 27.9827 | 4.1081 | 5.5925
2025-04-01 11:19 | 2.0913 | 52.927 | 1.6791 | 5.5658
2025-04-01 11:19 | 0.46307 | 29.6382 | 5.2855 | 22.5596
2025-04-01 11:19 | 0 | 42.6837 | 5.7539 | 6.3673
2025-04-01 11:19 | 11.0249 | 39.221 | 5.3434 | 21.8048
2025-04-01 11:19 | 0 | 43.6462 | 3.2459 | 19.123
2025-04-01 11:19 | 6.841 | 41.8387 | 0.81245 | 14.7405
2025-04-01 11:19 | 3.0453 | 26.7413 | 1.5529 | 9.2173
2025-04-01 11:19 | 11.7704 | 40.7284 | 4.958 | 5.4121
2025-04-01 11:19 | 0 | 47.922 | 5.0127 | 11.2513
2025-04-01 11:19 | 1.867 | 24.6396 | 4.1743 | 16.3287
2025-04-01 11:19 | 5.4126 | 40.1038 | 2.0745 | 13.0435
2025-04-01 11:19 | 0 | 37.9155 | 0.54522 | 11.5081
2025-04-01 11:19 | 0 | 46.0286 | 3.9629 | 12.5195
2025-04-01 11:19 | 0 | 41.4213 | 1.9232 | 22.9604
2025-04-01 11:19 | 14.8239 | 27.7931 | 5.8034 | 13.4547
2025-04-01 11:19 | 0 | 31.7069 | 3.2085 | 29.0539
2025-04-01 11:19 | 0 | 25.1499 | 4.846 | 12.0274
2025-04-01 11:19 | 4.0691 | 36.5614 | 6.2395 | 9.2569
2025-04-01 11:19 | 7.2323 | 42.6487 | 3.4986 | 9.0926
2025-04-01 11:19 | 3.9236 | 41.8913 | 3.6089 | 18.1744
2025-04-01 11:19 | 3.7586 | 33.6201 | 5.1986 | 22.7007
2025-04-01 11:19 | 5.1004 | 25.1247 | 8.5801 | 12.1357
2025-04-01 11:19 | 6.6125 | 37.3254 | 5.6547 | 13.5735
2025-04-01 11:19 | 0 | 41.1352 | 1.4925 | 11.1069
2025-04-01 11:19 | 11.4451 | 41.5886 | 4.4405 | 6.863
2025-04-01 11:19 | 5.804 | 27.2489 | 5.5564 | 15.1403
2025-04-01 11:19 | 0 | 39.5613 | 3.1021 | 12.738
2025-04-01 11:19 | 7.3573 | 27.4474 | 7.0949 | 22.5811
2025-04-01 11:19 | 0 | 49.5339 | 7.0319 | 10.4973
2025-04-01 11:19 | 4.4622 | 32.5069 | 7.0479 | 20.6074
2025-04-01 11:19 | 4.4242 | 39.2783 | 2.4858 | 17.8856
2025-04-01 11:19 | 10.5842 | 57.3549 | 1.79 | 18.0574
2025-04-01 11:19 | 1.3416 | 59.653 | 4.4262 | 19.3016
2025-04-01 11:19 | 0 | 33.5453 | 5.6709 | 5.6036
2025-04-01 11:19 | 13.0865 | 26.9434 | 2.3471 | 18.929
2025-04-01 11:19 | 2.4577 | 39.1408 | 5.1991 | 16.092
2025-04-01 11:19 | 3.0714 | 29.0985 | 1.601 | 21.018
2025-04-01 11:19 | 5.8363 | 23.9099 | 3.686 | 12.9902
2025-04-01 11:19 | 4.3319 | 29.1977 | 3.575 | 13.9675
2025-04-01 11:19 | 0.51472 | 39.1087 | 4.2297 | 15.0413
2025-04-01 11:19 | 11.1154 | 55.5133 | 1.8099 | 13.723
2025-04-01 11:19 | 0 | 47.6628 | 1.8871 | 15.1607
2025-04-01 11:20 | 2.1592 | 32.9387 | 1.8341 | 13.6408
2025-04-01 11:20 | 19.3413 | 23.5575 | 0.60458 | 15.8866
2025-04-01 11:20 | 1.6053 | 40.2939 | 4.1672 | 25.098
2025-04-01 11:20 | 6.1452 | 29.1575 | 3.0636 | 6.3251
2025-04-01 11:20 | 6.6738 | 42.9846 | 2.7692 | 17.0483
2025-04-01 11:20 | 1.8454 | 35.7957 | 2.3374 | 28.282
2025-04-01 11:20 | 4.1813 | 35.5602 | 6.4766 | 13.4157
2025-04-01 11:20 | 12.8301 | 37.2343 | 3.4345 | 23.1887
2025-04-01 11:20 | 0 | 24.7362 | 1.2844 | 20.8563
2025-04-01 11:20 | 8.8713 | 38.3695 | 4.9222 | 5.1611
2025-04-01 11:20 | 1.2917 | 37.9948 | 4.7943 | 16.1759
2025-04-01 11:20 | 6.5346 | 44.8061 | 5.9917 | 31.3733
2025-04-01 11:20 | 4.0669 | 29.3577 | 4.0299 | 21.313
2025-04-01 11:20 | 3.6104 | 24.6613 | 4.6858 | 12.6805
2025-04-01 11:20 | 4.1187 | 45.1947 | 3.5955 | 18.9323
2025-04-01 11:20 | 3.2127 | 32.0571 | 3.818 | 22.304
2025-04-01 11:20 | 4.7007 | 39.7409 | 4.5227 | 11.8893
2025-04-01 11:20 | 1.3093 | 23.9946 | 3.3817 | 8.4592
2025-04-01 11:20 | 0 | 27.1753 | 1.9536 | 14.9567
2025-04-01 11:20 | 0 | 41.5943 | 3.556 | 17.6056
2025-04-01 11:20 | 12.8942 | 34.2512 | 4.6094 | 15.0906
2025-04-01 11:20 | 5.8122 | 31.1078 | 5.3282 | 18.3152
2025-04-01 11:20 | 0.75598 | 25.5386 | 0.93988 | 15.7457
2025-04-01 11:20 | 0 | 37.9913 | 1.2477 | 12.5377
2025-04-01 11:20 | 9.5221 | 38.8329 | 4.0002 | 15.4251
2025-04-01 11:20 | 7.7718 | 47.9089 | 6.173 | 12.9175
2025-04-01 11:20 | 1.9807 | 34.6318 | 1.4394 | 13.5232
2025-04-01 11:20 | 0.5739 | 26.6414 | 10.4496 | 19.5903
2025-04-01 11:20 | 3.1061 | 31.8999 | 5.0859 | 23.6948
2025-04-01 11:20 | 5.1795 | 48.0612 | 5.8851 | 8.9278
2025-04-01 11:20 | 6.1391 | 29.0268 | 5.8823 | 9.5034
2025-04-01 11:20 | 2.6139 | 40.6368 | 6.0471 | 10.5406
2025-04-01 11:20 | 8.4355 | 30.6236 | 2.5502 | 26.6593
2025-04-01 11:20 | 2.8323 | 35.424 | 2.0963 | 11.5072
2025-04-01 11:20 | 0 | 38.5501 | 4.8505 | 12.2271
2025-04-01 11:20 | 0 | 27.2925 | 2.0123 | 7.6054
2025-04-01 11:20 | 0 | 35.9173 | 2.5684 | 12.8233
2025-04-01 11:20 | 0 | 21.7171 | 4.5003 | 20.3217
2025-04-01 11:20 | 0.46127 | 35.0876 | 2.6121 | 12.7229
2025-04-01 11:20 | 10.8581 | 36.0307 | 4.7739 | 11.2706
2025-04-01 11:20 | 0 | 44.7157 | 8.153 | 15.6803
2025-04-01 11:20 | 0 | 31.6384 | 3.3551 | 15.1459
2025-04-01 11:20 | 0.74148 | 40.2002 | 1.8038 | 21.5403
2025-04-01 11:20 | 0.19813 | 32.7767 | 0.52839 | 8.4913
2025-04-01 11:20 | 16.1512 | 29.4396 | 3.6981 | 14.17
2025-04-01 11:20 | 2.147 | 34.6798 | 3.4722 | 5.0739
2025-04-01 11:20 | 6.3626 | 44.596 | 5.015 | 10.7708
2025-04-01 11:20 | 0 | 34.9552 | 1.325 | 20.311
2025-04-01 11:20 | 14.2012 | 44.5395 | 2.1318 | 8.1973
2025-04-01 11:20 | 0 | 53.758 | 2.04 | 22.8172
2025-04-01 11:20 | 3.9575 | 25.8648 | 5.0022 | 20.5358
2025-04-01 11:20 | 7.4818 | 38.7578 | 3.6779 | 22.914
2025-04-01 11:20 | 0 | 43.2798 | 5.911 | 14.2446
2025-04-01 11:20 | 0.2484 | 27.1432 | 3.3183 | 17.4608
2025-04-01 11:20 | 2.2236 | 39.0882 | 1.678 | 15.738
2025-04-01 11:20 | 5.4896 | 29.948 | 3.592 | 8.4763
2025-04-01 11:20 | 2.9981 | 36.0813 | 1.7972 | 10.743
2025-04-01 11:20 | 3.0317 | 41.3438 | 6.8246 | 18.0816
2025-04-01 11:20 | 6.3421 | 26.5916 | 2.8312 | 19.3182
2025-04-01 11:20 | 6.6046 | 42.7451 | 2.5876 | 10.8655
2025-04-01 11:21 | 0 | 26.3491 | 3.8099 | 16.3423
2025-04-01 11:21 | 0.14876 | 34.0863 | 5.8249 | 17.0509
2025-04-01 11:21 | 0 | 35.9562 | 2.6279 | 25.1129
2025-04-01 11:21 | 0 | 32.6973 | 3.828 | 21.8209
2025-04-01 11:21 | 0 | 34.867 | 4.1189 | 13.0599
2025-04-01 11:21 | 9.3856 | 26.1052 | 3.2982 | 14.6017
2025-04-01 11:21 | 9.3371 | 28.8513 | 2.831 | 6.1132
2025-04-01 11:21 | 1.9224 | 40.1059 | 3.0584 | 17.3177
2025-04-01 11:21 | 9.372 | 26.809 | 3.7148 | 23.0498
2025-04-01 11:21 | 13.9399 | 36.7626 | 0.88085 | 16.9723
2025-04-01 11:21 | 0 | 40.8296 | 4.635 | 12.2745
2025-04-01 11:21 | 0 | 33.7573 | 4.2175 | 8.2288
2025-04-01 11:21 | 0 | 36.7913 | 4.7295 | 22.4028
2025-04-01 11:21 | 0 | 35.4348 | 3.2137 | 23.0685
2025-04-01 11:21 | 7.3984 | 26.0627 | 6.6481 | 25.1407
2025-04-01 11:21 | 10.4927 | 32.8407 | 2.0029 | 14.1772
2025-04-01 11:21 | 10.6021 | 19.581 | 3.3715 | 16.4369
2025-04-01 11:21 | 9.3138 | 33.0106 | 4.5738 | 9.5064
2025-04-01 11:21 | 0 | 30.747 | 2.7091 | 14.1445
2025-04-01 11:21 | 1.5301 | 25.8707 | 3.4623 | 15.6133
2025-04-01 11:21 | 0.11729 | 44.6355 | 5.2474 | 10.5612
2025-04-01 11:21 | 0 | 37.8385 | 10.6723 | 16.0027
2025-04-01 11:21 | 4.191 | 35.3556 | 2.8758 | 17.3498
2025-04-01 11:21 | 0 | 23.8019 | 4.0191 | 21.2582
2025-04-01 11:21 | 3.8461 | 47.4756 | 2.5695 | 6.3385
2025-04-01 11:21 | 0.1934 | 33.8986 | 2.6499 | 9.6189
2025-04-01 11:21 | 4.7198 | 49.4519 | 5.6669 | 7.2462
2025-04-01 11:21 | 4.8719 | 41.5798 | 8.4964 | 5.8799
2025-04-01 11:21 | 0 | 37.5332 | 3.0469 | 6.7996
2025-04-01 11:21 | 0 | 38.2395 | 2.714 | 14.8415
2025-04-01 11:21 | 0 | 37.1439 | 3.4304 | 9.4255
2025-04-01 11:21 | 3.5526 | 37.9734 | 7.8348 | 10.8785
2025-04-01 11:21 | 6.1857 | 35.7614 | 4.7693 | 8.0753
2025-04-01 11:21 | 5.042 | 32.8987 | 6.1261 | 16.1573
2025-04-01 11:21 | 6.6942 | 29.5507 | 7.9097 | 17.6055
2025-04-01 11:21 | 11.6104 | 30.4852 | 2.378 | 16.8527
2025-04-01 11:21 | 1.9176 | 42.5573 | 2.1095 | 12.8778
2025-04-01 11:21 | 0 | 37.1702 | 3.3129 | 15.499
2025-04-01 11:21 | 0 | 20.8244 | 5.6271 | 10.2994
2025-04-01 11:21 | 7.5966 | 39.0426 | 4.8327 | 15.1448
2025-04-01 11:21 | 0.45213 | 36.4601 | 2.5438 | 9.9275
2025-04-01 11:21 | 9.2543 | 51.7904 | 4.6 | 14.2813
2025-04-01 11:21 | 3.9121 | 41.3649 | 4.1459 | 13.3397
2025-04-01 11:21 | 2.6036 | 48.3176 | 2.9313 | 19.0837
2025-04-01 11:21 | 15.7794 | 39.4059 | 4.8644 | 13.9172
2025-04-01 11:21 | 7.9387 | 48.1248 | 3.7899 | 9.3805
2025-04-01 11:21 | 0 | 43.6234 | 3.8307 | 10.8222
2025-04-01 11:21 | 15.9326 | 26.8633 | 5.5191 | 11.6659
2025-04-01 11:21 | 17.6312 | 42.2359 | 5.7944 | 15.7129
2025-04-01 11:21 | 4.6556 | 26.4676 | 4.2609 | 7.1724
2025-04-01 11:21 | 1.3079 | 50.9901 | 4.3076 | 21.9002
2025-04-01 11:21 | 0 | 37.6504 | 2.4603 | 13.5146
2025-04-01 11:21 | 0 | 46.3815 | 2.7595 | 10.5042
2025-04-01 11:21 | 6.3064 | 31.9411 | 4.8099 | 10.7985
2025-04-01 11:21 | 5.6837 | 30.1549 | 5.8267 | 8.7189
2025-04-01 11:21 | 0 | 22.7135 | 4.1161 | 9.5904
2025-04-01 11:21 | 3.6518 | 30.3514 | 3.3868 | 20.4154
2025-04-01 11:21 | 0 | 24.9679 | 2.5364 | 24.3682
2025-04-01 11:21 | 5.8092 | 57.7784 | 4.7209 | 5.7346
2025-04-01 11:21 | 4.6756 | 29.9321 | 1.0235 | 4.1412
2025-04-01 11:22 | 0 | 20.6333 | 8.5772 | 8.2717
2025-04-01 11:22 | 5.2266 | 35.6662 | 3.0021 | 21.3607
2025-04-01 11:22 | 0 | 40.0784 | 1.7088 | 14.8811
2025-04-01 11:22 | 4.342 | 43.2843 | 3.5756 | 18.4094
2025-04-01 11:22 | 4.5952 | 34.7316 | 3.1961 | 13.9864
2025-04-01 11:22 | 4.8007 | 35.036 | 3.3012 | 14.8441
2025-04-01 11:22 | 3.9273 | 51.8375 | 1.1665 | 21.3115
2025-04-01 11:22 | 5.1986 | 39.1681 | 1.519 | 14.9233
2025-04-01 11:22 | 6.8098 | 55.9839 | 4.6773 | 12.5215
2025-04-01 11:22 | 0 | 33.9996 | 2.2894 | 17.4385
2025-04-01 11:22 | 0.48274 | 51.9744 | 5.6858 | 12.5333
2025-04-01 11:22 | 7.6027 | 30.9151 | 0.93004 | 19.3834
2025-04-01 11:22 | 12.3186 | 46.1048 | 5.8576 | 19.8682
2025-04-01 11:22 | 1.3212 | 30.0842 | 3.9464 | 6.2693
2025-04-01 11:22 | 10.6971 | 42.2356 | 5.3082 | 13.6888
2025-04-01 11:22 | 1.7206 | 31.704 | 4.7941 | 17.6576
2025-04-01 11:22 | 4.1474 | 23.9431 | 7.2421 | 14.9393
2025-04-01 11:22 | 7.3936 | 28.7111 | 0.5727 | 9.3791
2025-04-01 11:22 | 10.3838 | 28.4456 | 4.5676 | 21.8845
2025-04-01 11:22 | 0 | 24.7316 | 1.1948 | 12.2027
2025-04-01 11:22 | 0 | 25.902 | 3.9412 | 17.2621
2025-04-01 11:22 | 8.8056 | 24.7736 | 7.3612 | 24.5019
2025-04-01 11:22 | 3.1829 | 35.7182 | 3.0252 | 18.0037
2025-04-01 11:22 | 5.7884 | 19.3018 | 5.9551 | 22.4642
2025-04-01 11:22 | 0.42438 | 40.5246 | 2.3586 | 15.0184
2025-04-01 11:22 | 0 | 20.9098 | 5.1813 | 15.2476
2025-04-01 11:22 | 2.7643 | 33.7578 | 4.5661 | 17.2502
2025-04-01 11:22 | 0.71865 | 35.0307 | 2.7535 | 7.626
2025-04-01 11:22 | 0 | 47.1053 | 5.5117 | 17.1817
2025-04-01 11:22 | 0 | 31.091 | 2.658 | 11.6743
2025-04-01 11:22 | 0 | 30.0019 | 0.38859 | 13.7649
2025-04-01 11:22 | 4.4931 | 38.8728 | 2.2737 | 17.7807
2025-04-01 11:22 | 4.2419 | 30.1799 | 2.0069 | 21.3261
2025-04-01 11:22 | 5.1521 | 26.8244 | 4.104 | 15.8603
2025-04-01 11:22 | 0 | 38.4405 | 3.3185 | 15.705
2025-04-01 11:22 | 5.9707 | 33.9 | 1.0244 | 15.3688
2025-04-01 11:22 | 1.1706 | 32.7229 | 6.1261 | 17.9835
2025-04-01 11:22 | 0.64131 | 27.0294 | 3.5484 | 8.7385
2025-04-01 11:22 | 5.2255 | 29.8655 | 4.2806 | 15.2318
2025-04-01 11:22 | 3.1348 | 33.7472 | 9.1442 | 6.2374
2025-04-01 11:22 | 0 | 36.9757 | 3.5307 | 15.001
2025-04-01 11:22 | 0 | 34.3034 | 4.2823 | 15.0919
2025-04-01 11:22 | 2.3798 | 22.4823 | 2.7764 | 23.7421
2025-04-01 11:22 | 18.3028 | 29.3611 | 5.4965 | 14.6979
2025-04-01 11:22 | 2.9422 | 42.4956 | 2.6194 | 11.3662
2025-04-01 11:22 | 1.2091 | 45.1289 | 3.1829 | 9.4331
2025-04-01 11:22 | 2.0998 | 40.8693 | 2.4973 | 20.9278
2025-04-01 11:22 | 3.2726 | 21.0324 | 3.5868 | 8.8803
2025-04-01 11:22 | 1.9956 | 51.4775 | 3.5935 | 15.4913
2025-04-01 11:22 | 4.142 | 39.7072 | 2.9085 | 20.4103
2025-04-01 11:22 | 8.9364 | 32.2044 | 0.53187 | 15.8958
2025-04-01 11:22 | 12.9558 | 33.8738 | 5.9359 | 11.7292
2025-04-01 11:22 | 2.742 | 39.1692 | 7.9978 | 11.1489
2025-04-01 11:22 | 4.2878 | 34.6683 | 5.3429 | 15.914
2025-04-01 11:22 | 0.43384 | 35.6867 | 1.6639 | 16.707
2025-04-01 11:22 | 0 | 35.5558 | 7.5189 | 2.828
2025-04-01 11:22 | 7.0638 | 27.2174 | 4.547 | 16.3666
2025-04-01 11:22 | 1.003 | 34.6492 | 3.2411 | 23.2901
2025-04-01 11:22 | 0 | 33.1866 | 1.0857 | 10.2931
2025-04-01 11:22 | 4.2706 | 26.4959 | 4.151 | 10.504
2025-04-01 11:23 | 0 | 17.1413 | 5.6604 | 22.291
2025-04-01 11:23 | 0 | 43.2812 | 6.2503 | 8.1174
2025-04-01 11:23 | 2.8449 | 55.5116 | 3.0468 | 5.8928
2025-04-01 11:23 | 0 | 26.6981 | 0.70755 | 15.8776
2025-04-01 11:23 | 4.9874 | 28.0136 | 1.4771 | 8.7291
2025-04-01 11:23 | 5.6511 | 18.4398 | 6.5558 | 25.691
2025-04-01 11:23 | 7.0454 | 41.6313 | 3.1101 | 18.3723
2025-04-01 11:23 | 9.6265 | 34.0608 | 0.69351 | 7.1297
2025-04-01 11:23 | 9.302 | 44.279 | 1.6965 | 26.1921
2025-04-01 11:23 | 11.6489 | 41.6385 | 3.3932 | 12.041
2025-04-01 11:23 | 0 | 28.1354 | 2.0823 | 14.605
2025-04-01 11:23 | 12.0005 | 53.8264 | 3.7889 | 15.5492
2025-04-01 11:23 | 0 | 27.5328 | 2.9977 | 9.4538
2025-04-01 11:23 | 3.0786 | 61.9111 | 1.6508 | 10.8025
2025-04-01 11:23 | 3.1056 | 50.499 | 5.1625 | 13.0226
2025-04-01 11:23 | 12.1148 | 38.5383 | 3.0897 | 11.6154
2025-04-01 11:23 | 9.5267 | 27.9588 | 2.1686 | 19.3398
2025-04-01 11:23 | 2.8222 | 37.7419 | 5.0039 | 11.2971
2025-04-01 11:23 | 5.6938 | 31.2927 | 3.6338 | 16.3985
2025-04-01 11:23 | 4.7132 | 38.018 | 2.929 | 19.9363
2025-04-01 11:23 | 0.84694 | 42.3987 | 1.9616 | 17.5581
2025-04-01 11:23 | 3.8664 | 35.0557 | 3.9975 | 11.1705
2025-04-01 11:23 | 0 | 59.6676 | 3.0327 | 12.7968
2025-04-01 11:23 | 6.4652 | 32.0733 | 5.6454 | 12.3577
2025-04-01 11:23 | 0 | 27.8036 | 2.0197 | 20.1811
2025-04-01 11:23 | 3.5159 | 18.8244 | 2.3928 | 17.5955
2025-04-01 11:23 | 4.521 | 30.7393 | 1.8891 | 10.4415
2025-04-01 11:23 | 4.5755 | 29.8957 | 3.5569 | 14.4604
2025-04-01 11:23 | 0 | 40.7853 | 2.8245 | 4.3406
2025-04-01 11:23 | 6.7565 | 28.6099 | 4.7377 | 15.4817
2025-04-01 11:23 | 1.8575 | 49.3588 | 6.4013 | 22.4541
2025-04-01 11:23 | 0 | 57.3949 | 3.7358 | 31.3431
2025-04-01 11:23 | 6.3458 | 43.9929 | 1.1532 | 22.2678
2025-04-01 11:23 | 9.326 | 52.298 | 3.9458 | 27.322
2025-04-01 11:23 | 0 | 30.3301 | 2.3288 | 23.7685
2025-04-01 11:23 | 7.8577 | 50.2896 | 2.3571 | 18.3798
2025-04-01 11:23 | 3.0156 | 35.3008 | 1.7556 | 14.845
2025-04-01 11:23 | 0 | 27.4223 | 4.4695 | 16.657
2025-04-01 11:23 | 1.6578 | 23.0778 | 6.1129 | 15.9803
2025-04-01 11:23 | 0 | 36.8192 | 3.0179 | 22.8144
2025-04-01 11:23 | 0.60959 | 43.6467 | 4.9714 | 9.5383
2025-04-01 11:23 | 1.5009 | 39.4602 | 4.106 | 28.0726
2025-04-01 11:23 | 10.0789 | 33.6566 | 3.367 | 8.4845
2025-04-01 11:23 | 0 | 39.6044 | 3.3197 | 13.5171
2025-04-01 11:23 | 5.7961 | 46.5419 | 9.2919 | 9.8594
2025-04-01 11:23 | 2.5463 | 22.209 | 3.2027 | 16.1495
2025-04-01 11:23 | 0.25343 | 27.453 | 2.0264 | 7.9017
2025-04-01 11:23 | 1.3606 | 40.3307 | 3.0597 | 13.5613
2025-04-01 11:23 | 3.1001 | 41.0917 | 2.7347 | 14.2293
2025-04-01 11:23 | 9.6767 | 39.1424 | 5.4467 | 18.2145
2025-04-01 11:23 | 7.7422 | 28.1696 | 3.5486 | 11.3567
2025-04-01 11:23 | 0.88308 | 35.595 | 8.3839 | 5.0014
2025-04-01 11:23 | 1.5212 | 36.254 | 4.351 | 8.6864
2025-04-01 11:23 | 1.9131 | 58.0427 | 3.0147 | 9.0309
2025-04-01 11:23 | 9.1762 | 33.0392 | 6.711 | 7.747
2025-04-01 11:23 | 4.7018 | 36.071 | 3.6513 | 8.068
2025-04-01 11:23 | 1.989 | 50.235 | 3.7283 | 19.7554
2025-04-01 11:23 | 2.2263 | 36.9397 | 3.0815 | 14.7203
2025-04-01 11:23 | 4.8086 | 40.9411 | 1.4171 | 7.323
2025-04-01 11:23 | 0 | 34.7333 | 0.94429 | 23.4161
2025-04-01 11:24 | 0 | 43.38 | 5.227 | 8.9869
2025-04-01 11:24 | 0 | 48.7743 | 2.6589 | 13.4288
2025-04-01 11:24 | 3.9217 | 36.7253 | 6.5437 | 26.96
2025-04-01 11:24 | 2.5684 | 52.0747 | 2.3168 | 5.0166
2025-04-01 11:24 | 6.0789 | 43.8843 | 1.8321 | 11.622
2025-04-01 11:24 | 0 | 24.5247 | 3.7947 | 15.5231
2025-04-01 11:24 | 5.3403 | 33.5998 | 1.5334 | 16.0694
2025-04-01 11:24 | 0 | 28.9307 | 5.7861 | 17.7463
2025-04-01 11:24 | 0 | 24.3109 | 6.0797 | 13.0735
2025-04-01 11:24 | 7.5087 | 31.2618 | 3.8422 | 11.5558
2025-04-01 11:24 | 0 | 39.6258 | 3.166 | 14.3232
2025-04-01 11:24 | 2.1323 | 51.8441 | 3.1238 | 8.0644
2025-04-01 11:24 | 0 | 36.6944 | 4.2147 | 16.2694
2025-04-01 11:24 | 0 | 57.21 | 3.1681 | 13.0053
2025-04-01 11:24 | 9.4296 | 47.6041 | 4.1861 | 11.696
2025-04-01 11:24 | 1.7069 | 36.0646 | 7.3435 | 7.7484
2025-04-01 11:24 | 8.2957 | 33.5406 | 4.6325 | 7.0123
2025-04-01 11:24 | 3.3926 | 35.695 | 2.0079 | 6.7907
2025-04-01 11:24 | 0 | 51.5328 | 2.3178 | 21.2952
2025-04-01 11:24 | 6.7819 | 25.4532 | 3.871 | 18.3495
2025-04-01 11:24 | 0 | 34.1687 | 3.4536 | 18.5455
2025-04-01 11:24 | 1.8712 | 48.3081 | 2.3275 | 9.9001
2025-04-01 11:24 | 9.8853 | 34.9652 | 3.7149 | 14.3577
2025-04-01 11:24 | 2.8091 | 35.0621 | 5.1779 | 5.5731
2025-04-01 11:24 | 0 | 40.9239 | 5.0056 | 19.291
2025-04-01 11:24 | 12.5272 | 25.4851 | 3.6664 | 7.0792
2025-04-01 11:24 | 0.26595 | 29.9351 | 3.7049 | 15.1023
2025-04-01 11:24 | 5.4712 | 33.5458 | 3.6621 | 9.4793
2025-04-01 11:24 | 4.9386 | 26.7709 | 2.79 | 17.4388
2025-04-01 11:24 | 5.9075 | 38.1099 | 2.9636 | 15.1941
2025-04-01 11:24 | 3.2347 | 24.1174 | 1.3756 | 14.5363
2025-04-01 11:24 | 11.1907 | 29.9814 | 2.4148 | 22.0856
2025-04-01 11:24 | 6.9283 | 37.4688 | 3.8475 | 10.6264
2025-04-01 11:24 | 3.496 | 35.9802 | 5.1467 | 26.8391
2025-04-01 11:24 | 0 | 25.3064 | 2.1875 | 9.5471
2025-04-01 11:24 | 4.4051 | 37.1077 | 5.3344 | 19.9502
2025-04-01 11:24 | 0 | 43.9556 | 3.0487 | 18.2715
2025-04-01 11:24 | 0 | 28.5664 | 3.1784 | 17.4951
2025-04-01 11:24 | 2.1941 | 38.308 | 4.4375 | 13.2063
2025-04-01 11:24 | 13.9611 | 35.6521 | 4.5001 | 9.2332
2025-04-01 11:24 | 0 | 34.5422 | 6.7267 | 7.1074
2025-04-01 11:24 | 1.1853 | 41.4967 | 2.2857 | 4.3797
2025-04-01 11:24 | 0 | 37.6448 | 1.8691 | 13.8375
2025-04-01 11:24 | 8.8497 | 33.9265 | 5.2509 | 8.1798
2025-04-01 11:24 | 9.0133 | 49.6626 | 2.5538 | 7.8772
2025-04-01 11:24 | 0 | 25.5941 | 1.8933 | 12.9357
2025-04-01 11:24 | 1.6316 | 31.9863 | 4.2431 | 19.6696
2025-04-01 11:24 | 0.3252 | 27.5187 | 4.2421 | 22.2596
2025-04-01 11:24 | 3.4795 | 33.2682 | 6.0602 | 14.029
2025-04-01 11:24 | 4.1666 | 30.8631 | 2.4701 | 9.4851
2025-04-01 11:24 | 6.0631 | 40.4216 | 5.4569 | 6.8915
2025-04-01 11:24 | 0.90112 | 39.9985 | 2.9901 | 13.5687
2025-04-01 11:24 | 11.5597 | 22.6141 | 4.8312 | 8.4584
2025-04-01 11:24 | 8.4608 | 43.4847 | 1.0885 | 15.3834
2025-04-01 11:24 | 4.102 | 31.7172 | 1.9383 | 21.2552
2025-04-01 11:24 | 8.9356 | 29.486 | 2.8487 | 8.4597
2025-04-01 11:24 | 6.2306 | 30.8519 | 3.2481 | 17.313
2025-04-01 11:24 | 0.028817 | 38.1127 | 1.1734 | 16.6735
2025-04-01 11:24 | 3.9412 | 36.7286 | 3.6397 | 20.4857
2025-04-01 11:24 | 2.5196 | 17.8076 | 2.6342 | 15.1376
2025-04-01 11:25 | 1.6709 | 44.0498 | 3.0911 | 24.3897
2025-04-01 11:25 | 0 | 52.4122 | 7.2217 | 8.6654
2025-04-01 11:25 | 2.2447 | 42.2336 | 3.4981 | 8.5111
2025-04-01 11:25 | 0 | 24.7677 | 5.4034 | 19.6486
2025-04-01 11:25 | 11.0885 | 49.9067 | 4.2896 | 25.5739
2025-04-01 11:25 | 6.3642 | 49.243 | 3.1294 | 13.1122
2025-04-01 11:25 | 2.8038 | 22.8088 | 5.9545 | 29.4465
2025-04-01 11:25 | 4.3968 | 27.4274 | 5.3454 | 18.5654
2025-04-01 11:25 | 0 | 31.2703 | 3.8289 | 14.9333
2025-04-01 11:25 | 6.1932 | 36.896 | 2.1738 | 25.8291
2025-04-01 11:25 | 1.6953 | 27.8322 | 4.0843 | 14.1789
2025-04-01 11:25 | 0 | 28.9659 | 5.0258 | 18.0558
2025-04-01 11:25 | 1.7022 | 48.3413 | 3.8513 | 12.9595
2025-04-01 11:25 | 3.1022 | 20.7177 | 4.5945 | 11.3619
2025-04-01 11:25 | 0 | 33.0152 | 2.9263 | 27.5617
2025-04-01 11:25 | 0.96949 | 37.0577 | 1.8641 | 9.144
2025-04-01 11:25 | 7.8761 | 26.916 | 3.5187 | 1.8536
2025-04-01 11:25 | 1.9668 | 30.958 | 1.8472 | 8.7046
2025-04-01 11:25 | 6.0916 | 44.4191 | 1.5956 | 14.9363
2025-04-01 11:25 | 0 | 35.8461 | 5.5467 | 10.9945
2025-04-01 11:25 | 0 | 50.7525 | 3.5012 | 22.3877
2025-04-01 11:25 | 0 | 23.0771 | 2.6823 | 2.5105
2025-04-01 11:25 | 8.6206 | 17.0811 | 5.2167 | 18.2621
2025-04-01 11:25 | 0 | 44.333 | 3.4044 | 10.0236
2025-04-01 11:25 | 0 | 33.0295 | 3.0035 | 7.8609
2025-04-01 11:25 | 0 | 17.783 | 1.5572 | 21.3239
2025-04-01 11:25 | 0 | 23.5079 | 0.33864 | 5.7668
2025-04-01 11:25 | 5.9247 | 43.7437 | 1.1813 | 12.1241
2025-04-01 11:25 | 2.1981 | 49.6304 | 3.434 | 12.7832
2025-04-01 11:25 | 6.6473 | 33.8213 | 3.5414 | 8.2818
2025-04-01 11:25 | 5.6633 | 39.0079 | 2.4046 | 17.1527
2025-04-01 11:25 | 0 | 40.2167 | 2.3277 | 29.0703
2025-04-01 11:25 | 0 | 41.6324 | 5.3897 | 7.8927
2025-04-01 11:25 | 7.0135 | 33.6525 | 3.1549 | 12.013
2025-04-01 11:25 | 12.1568 | 51.5537 | 2.6087 | 17.6395
2025-04-01 11:25 | 13.0921 | 23.1721 | 2.4155 | 27.6396
2025-04-01 11:25 | 0.67178 | 49.5364 | 11.0584 | 13.575
2025-04-01 11:25 | 0 | 50.6469 | 3.8062 | 8.7606
2025-04-01 11:25 | 7.4532 | 29.069 | 2.8635 | 22.6942
2025-04-01 11:25 | 10.7455 | 22.6545 | 5.9151 | 18.4432
2025-04-01 11:25 | 0 | 39.6744 | 5.8928 | 7.8369
2025-04-01 11:25 | 6.9981 | 45.587 | 0.948 | 18.284
2025-04-01 11:25 | 14.1049 | 28.5321 | 2.0382 | 13.6494
2025-04-01 11:25 | 5.9842 | 24.0468 | 5.0955 | 12.9594
2025-04-01 11:25 | 6.903 | 29.1922 | 6.6551 | 13.6571
2025-04-01 11:25 | 1.4498 | 19.6252 | 4.4745 | 5.029
2025-04-01 11:25 | 18.18 | 29.6734 | 5.7223 | 26.4601
2025-04-01 11:25 | 0.42624 | 35.1429 | 2.8224 | 21.428
2025-04-01 11:25 | 3.2659 | 34.8844 | 5.0516 | 6.9851
2025-04-01 11:25 | 0 | 34.2774 | 3.3603 | 9.615
2025-04-01 11:25 | 7.2539 | 25.7075 | 2.3745 | 12.2901
2025-04-01 11:25 | 4.3937 | 24.4302 | 7.1288 | 24.4955
2025-04-01 11:25 | 12.2271 | 27.448 | 4.4229 | 7.8157
2025-04-01 11:25 | 1.0576 | 44.4123 | 4.4846 | 13.8988
2025-04-01 11:25 | 0 | 33.1663 | 4.8533 | 6.3288
2025-04-01 11:25 | 7.2719 | 33.9252 | 3.7302 | 6.9926
2025-04-01 11:25 | 7.1463 | 35.8163 | 5.909 | 9.0846
2025-04-01 11:25 | 1.7105 | 47.9587 | 0.58289 | 29.3771
2025-04-01 11:25 | 13.146 | 17.7101 | 3.1428 | 14.9564
2025-04-01 11:25 | 1.3965 | 35.6622 | 2.9556 | 11.3664
2025-04-01 11:26 | 6.4422 | 40.5135 | 4.3709 | 19.7067
2025-04-01 11:26 | 1.5268 | 27.6997 | 2.7723 | 11.8113
2025-04-01 11:26 | 0 | 31.0845 | 3.6283 | 8.2725
2025-04-01 11:26 | 0 | 24.3581 | 5.5284 | 21.8167
2025-04-01 11:26 | 0 | 38.3385 | 2.6269 | 28.4694
2025-04-01 11:26 | 5.3377 | 53.1788 | 4.0844 | 13.9284
2025-04-01 11:26 | 3.4849 | 44.2743 | 6.9353 | 27.0546
2025-04-01 11:26 | 0.56622 | 20.4036 | 1.5468 | 18.6483
2025-04-01 11:26 | 12.8995 | 48.9491 | 4.3708 | 15.356
2025-04-01 11:26 | 6.6541 | 41.97 | 2.881 | 20.1785
2025-04-01 11:26 | 0 | 22.8625 | 3.2352 | 11.611
2025-04-01 11:26 | 0 | 31.2145 | 3.3656 | 15.9743
2025-04-01 11:26 | 4.7296 | 31.5463 | 2.6809 | 14.3701
2025-04-01 11:26 | 12.0889 | 45.9887 | 2.6062 | 9.4671
2025-04-01 11:26 | 1.3181 | 46.0878 | 5.3153 | 13.0089
2025-04-01 11:26 | 0 | 39.3241 | 4.1423 | 9.0893
2025-04-01 11:26 | 12.951 | 34.3539 | 4.3378 | 13.386
2025-04-01 11:26 | 5.4314 | 39.8327 | 1.0998 | 6.916
2025-04-01 11:26 | 3.0908 | 49.9185 | 2.1618 | 6.542
2025-04-01 11:26 | 4.8295 | 21.1649 | 4.6979 | 13.3731
2025-04-01 11:26 | 1.6022 | 34.3405 | 5.2104 | 18.9256
2025-04-01 11:26 | 2.8803 | 24.4337 | 2.5788 | 17.4394
2025-04-01 11:26 | 2.7091 | 17.805 | 4.1895 | 19.7468
2025-04-01 11:26 | 4.3736 | 48.6115 | 6.9916 | 2.1501
2025-04-01 11:26 | 1.308 | 28.7286 | 3.6423 | 14.6369
2025-04-01 11:26 | 2.7117 | 32.1389 | 4.194 | 14.8952
2025-04-01 11:26 | 8.1215 | 33.994 | 2.1832 | 19.5986
2025-04-01 11:26 | 0 | 24.4214 | 2.4135 | 7.8821
2025-04-01 11:26 | 0 | 31.5758 | 3.1862 | 12.556
2025-04-01 11:26 | 4.2347 | 26.5415 | 2.545 | 8.0162
2025-04-01 11:26 | 0 | 44.9024 | 1.8113 | 25.4487
2025-04-01 11:26 | 0 | 22.4771 | 5.5681 | 13.2644
2025-04-01 11:26 | 0 | 24.1972 | 4.6722 | 12.5653
2025-04-01 11:26 | 0 | 34.7095 | 1.4962 | 9.8689
2025-04-01 11:26 | 5.5831 | 47.7722 | 5.9606 | 10.853
2025-04-01 11:26 | 0 | 35.1006 | 2.8779 | 11.1237
2025-04-01 11:26 | 1.0896 | 30.8437 | 2.3143 | 10.5054
2025-04-01 11:26 | 3.9831 | 24.8399 | 3.9807 | 9.7253
2025-04-01 11:26 | 16.5945 | 28.2221 | 1.6263 | 5.1126
2025-04-01 11:26 | 10.2524 | 41.0554 | 5.5693 | 13.3442
2025-04-01 11:26 | 5.6738 | 35.3385 | 3.3658 | 21.3916
2025-04-01 11:26 | 2.6065 | 33.3879 | 8.7673 | 22.0888
2025-04-01 11:26 | 1.6442 | 31.9982 | 3.1898 | 19.5521
2025-04-01 11:26 | 7.4147 | 24.3967 | 3.5943 | 6.8213
2025-04-01 11:26 | 0 | 34.7052 | 2.7106 | 22.5637
2025-04-01 11:26 | 4.073 | 40.7769 | 4.0894 | 11.6366
2025-04-01 11:26 | 7.4773 | 25.0459 | 2.6981 | 4.8039
2025-04-01 11:26 | 0 | 34.5907 | 5.103 | 15.3121
2025-04-01 11:26 | 1.6052 | 42.7659 | 2.39 | 7.6846
2025-04-01 11:26 | 0.47674 | 40.2041 | 5.5951 | 16.467
2025-04-01 11:26 | 0 | 59.4538 | 4.2708 | 5.7215
2025-04-01 11:26 | 5.6911 | 25.4951 | 3.7688 | 12.062
2025-04-01 11:26 | 10.3017 | 32.9103 | 3.1008 | 21.4497
2025-04-01 11:26 | 16.4126 | 30.4888 | 4.6625 | 23.7204
2025-04-01 11:26 | 6.0282 | 40.8931 | 1.2161 | 13.2563
2025-04-01 11:26 | 1.5039 | 38.6152 | 5.3815 | 8.2246
2025-04-01 11:26 | 0 | 24.3611 | 4.373 | 14.667
2025-04-01 11:26 | 5.5828 | 28.9089 | 3.2972 | 8.9739
2025-04-01 11:26 | 9.0191 | 33.6622 | 2.7797 | 10.3908
2025-04-01 11:26 | 4.2067 | 46.471 | 5.1074 | 18.032
2025-04-01 11:27 | 2.0075 | 32.434 | 3.8873 | 10.08
2025-04-01 11:27 | 0 | 40.7472 | 3.0649 | 13.0763
2025-04-01 11:27 | 1.5104 | 49.3763 | 1.977 | 14.4409
2025-04-01 11:27 | 0.74051 | 39.3607 | 0.92841 | 20.3094
2025-04-01 11:27 | 3.2186 | 47.8857 | 1.0116 | 10.8081
2025-04-01 11:27 | 0 | 33.6571 | 3.4779 | 12.511
2025-04-01 11:27 | 13.1819 | 27.623 | 3.3766 | 10.9505
2025-04-01 11:27 | 2.4667 | 46.5213 | 1.8888 | 16.9793
2025-04-01 11:27 | 0 | 36.0425 | 3.9368 | 5.4649
2025-04-01 11:27 | 9.7188 | 31.6727 | 3.0933 | 19.9807
2025-04-01 11:27 | 3.8117 | 51.1824 | 2.8781 | 13.6167
2025-04-01 11:27 | 11.9184 | 28.8345 | 2.4139 | 8.3933
2025-04-01 11:27 | 2.2836 | 28.0792 | 2.5079 | 16.1745
2025-04-01 11:27 | 6.3535 | 28.8854 | 2.582 | 8.1712
2025-04-01 11:27 | 3.3361 | 44.7524 | 3.3604 | 12.8349
2025-04-01 11:27 | 0 | 24.9463 | 1.5305 | 13.0013
2025-04-01 11:27 | 0 | 24.8097 | 1.3934 | 21.1373
2025-04-01 11:27 | 0 | 35.5311 | 3.8781 | 9.5507
2025-04-01 11:27 | 0 | 40.5816 | 3.3914 | 19.4888
2025-04-01 11:27 | 0.0011225 | 25.3396 | 3.891 | 18.333
2025-04-01 11:27 | 0 | 39.2054 | 2.2496 | 12.9134
2025-04-01 11:27 | 8.2923 | 23.5915 | 1.2689 | 20.5021
2025-04-01 11:27 | 1.4537 | 40.3817 | 8.0135 | 5.508
2025-04-01 11:27 | 0 | 39.0742 | 4.7198 | 4.5386
2025-04-01 11:27 | 6.1432 | 35.3192 | 3.1416 | 6.2967
2025-04-01 11:27 | 8.8252 | 36.3281 | 1.2279 | 15.1454
2025-04-01 11:27 | 0.30081 | 33.5574 | 0.78033 | 18.8537
2025-04-01 11:27 | 0.67301 | 34.7249 | 3.2385 | 12.9825
2025-04-01 11:27 | 2.9716 | 46.9228 | 3.2335 | 7.9574
2025-04-01 11:27 | 2.8996 | 37.2415 | 4.5788 | 15.2256
2025-04-01 11:27 | 0 | 28.726 | 4.8023 | 12.6419
2025-04-01 11:27 | 0 | 32.331 | 1.7745 | 21.271
2025-04-01 11:27 | 2.1415 | 42.0649 | 3.4866 | 18.1307
2025-04-01 11:27 | 6.4499 | 23.5062 | 5.6869 | 16.0182
2025-04-01 11:27 | 6.2749 | 38.5051 | 1.9701 | 8.0242
2025-04-01 11:27 | 9.7908 | 28.4099 | 6.3388 | 11.9984
2025-04-01 11:27 | 7.4838 | 32.8415 | 3.3926 | 25.6998
2025-04-01 11:27 | 8.1957 | 35.067 | 3.8745 | 16.4719
2025-04-01 11:27 | 5.1986 | 28.3302 | 5.598 | 26.1398
2025-04-01 11:27 | 15.6622 | 33.4294 | 1.3036 | 14.2411
2025-04-01 11:27 | 8.4613 | 38.5343 | 4.7626 | 8.1919
2025-04-01 11:27 | 3.5502 | 43.3211 | 5.4231 | 12.0009
2025-04-01 11:27 | 9.2104 | 39.5909 | 4.6911 | 8.549
2025-04-01 11:27 | 6.9916 | 35.3958 | 6.9465 | 23.5924
2025-04-01 11:27 | 15.2425 | 17.3738 | 0.64881 | 19.3059
2025-04-01 11:27 | 8.5634 | 39.5924 | 2.2144 | 8.8495
2025-04-01 11:27 | 5.5107 | 26.1554 | 7.8109 | 7.0696
2025-04-01 11:27 | 0 | 28.1156 | 4.6575 | 14.5688
2025-04-01 11:27 | 0 | 30.6649 | 1.4751 | 11.528
2025-04-01 11:27 | 13.2207 | 47.7272 | 3.526 | 6.793
2025-04-01 11:27 | 2.1643 | 30.2715 | 6.3094 | 36.3416
2025-04-01 11:27 | 4.9007 | 25.5095 | 2.3852 | 6.9212
2025-04-01 11:27 | 4.2979 | 37.6609 | 4.188 | 9.9778
2025-04-01 11:27 | 1.3698 | 36.3313 | 1.0408 | 19.6931
2025-04-01 11:27 | 0 | 39.2902 | 3.575 | 15.2296
2025-04-01 11:27 | 0 | 27.0096 | 1.9616 | 17.3787
2025-04-01 11:27 | 11.2283 | 32.647 | 2.6115 | 12.0602
2025-04-01 11:27 | 9.0273 | 31.3542 | 5.3607 | 18.3658
2025-04-01 11:27 | 0 | 29.3887 | 4.4769 | 15.9711
2025-04-01 11:27 | 6.4321 | 38.0308 | 1.0195 | 8.0375
2025-04-01 11:28 | 13.4201 | 42.0104 | 5.5057 | 13.7296
2025-04-01 11:28 | 4.6012 | 50.2644 | 4.0554 | 18.5944
2025-04-01 11:28 | 0 | 51.1015 | 6.0524 | 11.8921
2025-04-01 11:28 | 9.2625 | 33.2121 | 3.541 | 20.348
2025-04-01 11:28 | 10.6326 | 18.9206 | 2.3446 | 9.1865
2025-04-01 11:28 | 10.5074 | 27.1178 | 4.1509 | 15.756
2025-04-01 11:28 | 0 | 31.2426 | 2.7227 | 14.9697
2025-04-01 11:28 | 8.58 | 31.3103 | 2.9051 | 33.1401
2025-04-01 11:28 | 8.5471 | 46.544 | 3.9149 | 12.189
2025-04-01 11:28 | 9.5788 | 38.2312 | 2.7897 | 12.8603
2025-04-01 11:28 | 11.0026 | 45.4078 | 6.4476 | 9.4921
2025-04-01 11:28 | 12.2466 | 45.9521 | 2.5663 | 16.1552
2025-04-01 11:28 | 7.77 | 29.9908 | 1.2327 | 17.5473
2025-04-01 11:28 | 8.5394 | 27.6343 | 2.9122 | 16.3075
2025-04-01 11:28 | 8.1823 | 26.9834 | 7.0803 | 7.0421
2025-04-01 11:28 | 16.7309 | 47.9683 | 7.414 | 12.1316
2025-04-01 11:28 | 0.11152 | 35.8655 | 2.7007 | 13.0253
2025-04-01 11:28 | 9.4882 | 20.2553 | 2.7673 | 16.4128
2025-04-01 11:28 | 7.7994 | 38.4333 | 1.343 | 11.038
2025-04-01 11:28 | 0 | 36.4903 | 4.5286 | 19.1864
2025-04-01 11:28 | 0 | 49.0454 | 0.86746 | 6.2498
2025-04-01 11:28 | 12.9955 | 25.9649 | 3.8824 | 17.0098
2025-04-01 11:28 | 7.9772 | 30.3536 | 3.1809 | 16.6028
2025-04-01 11:28 | 3.6342 | 39.6419 | 4.3289 | 9.9721
2025-04-01 11:28 | 5.2516 | 38.1689 | 2.9596 | 7.9459
2025-04-01 11:28 | 1.7976 | 34.0089 | 5.2472 | 17.4954
2025-04-01 11:28 | 8.1943 | 42.972 | 3.6075 | 16.4561
2025-04-01 11:28 | 7.4133 | 41.9813 | 3.4832 | 9.4487
2025-04-01 11:28 | 8.3832 | 46.6561 | 1.8363 | 9.3429
2025-04-01 11:28 | 14.4267 | 46.6255 | 3.3717 | 18.3704
2025-04-01 11:28 | 5.9126 | 37.241 | 2.754 | 7.8356
2025-04-01 11:28 | 12.2925 | 31.1203 | 3.5762 | 9.2706
2025-04-01 11:28 | 10.6597 | 28.8057 | 4.2246 | 9.4596
2025-04-01 11:28 | 0 | 41.8115 | 2.0249 | 12.2587
2025-04-01 11:28 | 0 | 33.7426 | 4.2764 | 12.0303
2025-04-01 11:28 | 3.1951 | 40.4697 | 3.3441 | 9.4206
2025-04-01 11:28 | 0 | 26.4193 | 3.2257 | 17.8984
2025-04-01 11:28 | 0.31132 | 30.4977 | 1.253 | 6.9115
2025-04-01 11:28 | 0 | 41.2087 | 2.4897 | 9.1273
2025-04-01 11:28 | 6.6809 | 34.2166 | 1.5444 | 17.5093
2025-04-01 11:28 | 0 | 31.5313 | 3.2777 | 5.6646
2025-04-01 11:28 | 9.6783 | 22.96 | 2.9117 | 10.1104
2025-04-01 11:28 | 0.7404 | 34.058 | 2.4006 | 14.7057
2025-04-01 11:28 | 0 | 25.9921 | 4.3812 | 21.9206
2025-04-01 11:28 | 2.6419 | 15.6088 | 4.9113 | 9.8965
2025-04-01 11:28 | 0 | 44.1585 | 3.1132 | 16.435
2025-04-01 11:28 | 9.6479 | 18.7202 | 1.985 | 8.6599
2025-04-01 11:28 | 0 | 32.2399 | 4.2302 | 19.9947
2025-04-01 11:28 | 0 | 36.296 | 2.6805 | 18.3107
2025-04-01 11:28 | 0.11507 | 25.2665 | 2.803 | 21.3943
2025-04-01 11:28 | 3.1012 | 37.1423 | 2.8048 | 17.757
2025-04-01 11:28 | 3.6487 | 29.5207 | 2.9075 | 21.3789
2025-04-01 11:28 | 0.38224 | 24.2345 | 2.4306 | 8.3914
2025-04-01 11:28 | 0.63592 | 29.7823 | 1.3055 | 12.4729
2025-04-01 11:28 | 7.5619 | 33.5043 | 1.8316 | 12.2453
2025-04-01 11:28 | 10.9011 | 21.7651 | 5.4162 | 11.3576
2025-04-01 11:28 | 0 | 35.8796 | 4.275 | 11.0737
2025-04-01 11:28 | 14.1546 | 43.1968 | 5.147 | 12.5991
2025-04-01 11:28 | 9.1622 | 31.6267 | 4.2948 | 8.5763
2025-04-01 11:28 | 1.1913 | 36.1692 | 1.8469 | 13.7801
2025-04-01 11:29 | 2.4045 | 38.3408 | 4.8411 | 13.3144
2025-04-01 11:29 | 2.4168 | 28.0707 | 5.2985 | 11.2506
2025-04-01 11:29 | 0 | 26.4306 | 6.3691 | 15.0539
2025-04-01 11:29 | 2.7186 | 34.0385 | 2.733 | 13.0416
2025-04-01 11:29 | 0 | 33.8289 | 2.0928 | 10.0578
2025-04-01 11:29 | 0 | 25.4043 | 1.6095 | 17.722
2025-04-01 11:29 | 0 | 24.0203 | 1.2294 | 19.09
2025-04-01 11:29 | 4.1556 | 31.1539 | 4.8124 | 17.1396
2025-04-01 11:29 | 7.8569 | 52.5338 | 4.5867 | 11.6069
2025-04-01 11:29 | 6.6023 | 24.1277 | 6.2248 | 16.0387
2025-04-01 11:29 | 6.1572 | 15.6585 | 4.8302 | 21.1628
2025-04-01 11:29 | 9.5497 | 21.9953 | 5.5674 | 11.6492
2025-04-01 11:29 | 0 | 53.7494 | 6.0092 | 12.879
2025-04-01 11:29 | 6.3604 | 31.8914 | 1.798 | 15.5569
2025-04-01 11:29 | 6.1336 | 37.0146 | 1.8604 | 10.6866
2025-04-01 11:29 | 0 | 36.4326 | 0.69505 | 13.787
2025-04-01 11:29 | 0 | 37.3123 | 0.92113 | 8.0943
2025-04-01 11:29 | 3.4925 | 27.7496 | 2.3808 | 14.1569
2025-04-01 11:29 | 2.0393 | 23.8178 | 3.0862 | 24.0641
2025-04-01 11:29 | 7.7472 | 31.0313 | 3.2634 | 20.8679
2025-04-01 11:29 | 0 | 22.1834 | 0.40854 | 7.5258
2025-04-01 11:29 | 9.4701 | 23.0296 | 3.5947 | 16.7566
2025-04-01 11:29 | 4.564 | 34.8921 | 2.0142 | 19.5338
2025-04-01 11:29 | 0 | 30.6628 | 5.6286 | 17.1795
2025-04-01 11:29 | 2.7054 | 32.5617 | 3.3745 | 13.2662
2025-04-01 11:29 | 3.3597 | 40.3476 | 2.3467 | 16.6422
2025-04-01 11:29 | 1.9864 | 26.7266 | 0.83765 | 15.9308
2025-04-01 11:29 | 2.534 | 51.0435 | 0.19159 | 15.9792
2025-04-01 11:29 | 2.4208 | 34.2378 | 5.7547 | 13.8522
2025-04-01 11:29 | 4.0418 | 38.4744 | 2.3793 | 14.4419
2025-04-01 11:29 | 0 | 31.011 | 1.6822 | 14.1155
2025-04-01 11:29 | 1.0001 | 50.962 | 1.8762 | 31.1825
2025-04-01 11:29 | 8.763 | 27.362 | 3.8319 | 19.763
2025-04-01 11:29 | 2.3929 | 42.2587 | 2.7883 | 16.4634
2025-04-01 11:29 | 11.9486 | 30.0291 | 2.6926 | 6.8107
2025-04-01 11:29 | 4.3716 | 45.4852 | 2.6378 | 16.9244
2025-04-01 11:29 | 10.2181 | 39.4003 | 2.7833 | 2.4029
2025-04-01 11:29 | 0 | 46.3778 | 4.0591 | 4.9143
2025-04-01 11:29 | 4.7539 | 27.1037 | 2.7519 | 17.5315
2025-04-01 11:29 | 5.6002 | 33.4175 | 1.0522 | 14.5412
2025-04-01 11:29 | 1.9814 | 40.499 | 4.8964 | 16.3357
2025-04-01 11:29 | 7.2311 | 31.9471 | 6.1426 | 9.3465
2025-04-01 11:29 | 8.6976 | 30.6729 | 3.5192 | 12.6384
2025-04-01 11:29 | 3.1082 | 39.8231 | 2.9105 | 16.1874
2025-04-01 11:29 | 5.3173 | 21.6737 | 2.0355 | 11.7433
2025-04-01 11:29 | 0.62693 | 33.5993 | 1.4132 | 17.2973
2025-04-01 11:29 | 5.6644 | 41.8167 | 1.5011 | 14.0398
2025-04-01 11:29 | 2.2023 | 31.6693 | 3.763 | 17.2334
2025-04-01 11:29 | 1.1788 | 27.4627 | 4.7621 | 7.2745
2025-04-01 11:29 | 0.77549 | 31.0176 | 4.4297 | 5.3172
2025-04-01 11:29 | 8.5681 | 25.9832 | 5.6809 | 16.6743
2025-04-01 11:29 | 11.3044 | 26.1919 | 2.3715 | 15.9316
2025-04-01 11:29 | 9.5046 | 28.6952 | 2.8874 | 18.613
2025-04-01 11:29 | 11.3253 | 26.9143 | 3.9848 | 4.9596
2025-04-01 11:29 | 7.7313 | 33.9722 | 2.0536 | 16.9768
2025-04-01 11:29 | 1.8105 | 28.0075 | 4.463 | 24.0315
2025-04-01 11:29 | 1.1474 | 49.0246 | 1.9333 | 17.371
2025-04-01 11:29 | 0 | 32.2531 | 2.8826 | 17.2289
2025-04-01 11:29 | 9.2644 | 37.1572 | 3.1932 | 12.3353
2025-04-01 11:29 | 2.5979 | 25.6426 | 5.6982 | 7.6168
2025-04-01 11:30 | 0 | 35.7014 | 2.3603 | 6.1094
2025-04-01 11:30 | 12.7209 | 23.2363 | 3.9894 | 19.9327
2025-04-01 11:30 | 3.2071 | 27.5882 | 3.1195 | 16.5283
2025-04-01 11:30 | 6.5781 | 33.1347 | 3.9667 | 12.2548
2025-04-01 11:30 | 8.2092 | 51.5462 | 2.4498 | 21.0651
2025-04-01 11:30 | 1.0611 | 33.4992 | 2.6541 | 4.7097
2025-04-01 11:30 | 0.12993 | 40.2588 | 2.1603 | 7.4425
2025-04-01 11:30 | 3.72 | 23.9881 | 3.6073 | 14.026
2025-04-01 11:30 | 4.4113 | 17.6459 | 3.1129 | 11.0124
2025-04-01 11:30 | 0.33108 | 38.98 | 5.0505 | 18.7633
2025-04-01 11:30 | 9.6765 | 26.0916 | 1.4369 | 10.1273
2025-04-01 11:30 | 5.8073 | 20.4271 | 4.5507 | 10.4584
2025-04-01 11:30 | 6.09 | 42.2177 | 3.2838 | 27.0176
2025-04-01 11:30 | 0 | 47.0195 | 0.53811 | 11.9953
2025-04-01 11:30 | 0 | 30.6698 | 3.1034 | 17.0068
2025-04-01 11:30 | 2.1191 | 42.8582 | 3.4188 | 7.1624
2025-04-01 11:30 | 8.9386 | 42.3642 | 1.1614 | 9.0628
2025-04-01 11:30 | 0.59968 | 41.7107 | 3.802 | 7.3447
2025-04-01 11:30 | 10.6763 | 32.9765 | 1.8334 | 4.3967
2025-04-01 11:30 | 7.6871 | 38.482 | 2.0758 | 12.461
2025-04-01 11:30 | 0.44467 | 39.0289 | 1.6894 | 10.6597
2025-04-01 11:30 | 11.7732 | 41.6408 | 4.3413 | 8.5064
2025-04-01 11:30 | 2.7848 | 24.6736 | 3.8821 | 12.1789
2025-04-01 11:30 | 7.7273 | 22.1978 | 2.7359 | 8.2238
2025-04-01 11:30 | 8.8867 | 31.8955 | 0.87083 | 20.2502
2025-04-01 11:30 | 9.3945 | 27.9473 | 2.9947 | 17.3134
2025-04-01 11:30 | 0 | 28.8601 | 6.8389 | 16.7508
2025-04-01 11:30 | 7.9104 | 26.9846 | 1.9218 | 11.0629
2025-04-01 11:30 | 4.9031 | 25.7746 | 2.0447 | 17.2067
2025-04-01 11:30 | 3.0107 | 46.0319 | 7.5647 | 16.8948
2025-04-01 11:30 | 2.8554 | 31.1054 | 5.5113 | 12.3964
2025-04-01 11:30 | 1.294 | 20.368 | 2.6548 | 12.202
2025-04-01 11:30 | 3.2124 | 42.61 | 2.187 | 6.5148
2025-04-01 11:30 | 14.8617 | 29.3333 | 0.40607 | 15.3757
2025-04-01 11:30 | 13.4495 | 25.9839 | 9.9241 | 12.102
2025-04-01 11:30 | 10.8934 | 50.7381 | 3.0852 | 18.8101
2025-04-01 11:30 | 5.9472 | 40.3919 | 4.2539 | 8.5326
2025-04-01 11:30 | 11.8312 | 40.8457 | 1.6715 | 12.1982
2025-04-01 11:30 | 3.0998 | 44.9726 | 4.5638 | 18.8451
2025-04-01 11:30 | 0 | 45.699 | 5.6106 | 17.7488
2025-04-01 11:30 | 0.26353 | 42.9693 | 3.1021 | 20.926
2025-04-01 11:30 | 8.474 | 33.9909 | 2.4362 | 15.7188
2025-04-01 11:30 | 8.4866 | 24.7334 | 5.9216 | 16.7356
2025-04-01 11:30 | 0 | 44.4953 | 2.8287 | 6.9343
2025-04-01 11:30 | 8.7004 | 28.9628 | 1.6217 | 14.8297
2025-04-01 11:30 | 1.0371 | 31.7968 | 0.47328 | 15.3568
2025-04-01 11:30 | 0 | 25.5989 | 5.7354 | 9.6451
2025-04-01 11:30 | 0 | 37.0341 | 6.3884 | 29.0659
2025-04-01 11:30 | 0 | 39.2597 | 1.9611 | 8.2913
2025-04-01 11:30 | 0 | 25.9323 | 3.5544 | 6.7004
2025-04-01 11:30 | 0 | 23.4576 | 6.8462 | 13.4489
2025-04-01 11:30 | 0 | 40.0161 | 0.59073 | 1.7022
2025-04-01 11:30 | 0 | 38.796 | 2.5363 | 10.1285
2025-04-01 11:30 | 2.7258 | 22.8403 | 5.6349 | 26.9554
2025-04-01 11:30 | 7.7775 | 19.2379 | 7.15 | 14.6217
2025-04-01 11:30 | 0 | 26.1738 | 5.2155 | 13.8294
2025-04-01 11:30 | 0 | 31.1943 | 5.7301 | 10.8474
2025-04-01 11:30 | 0 | 42.5133 | 3.6133 | 7.1364
2025-04-01 11:30 | 0 | 32.4352 | 4.931 | 6.1584
2025-04-01 11:30 | 7.8159 | 30.8389 | 5.1723 | 14.3755
2025-04-01 11:31 | 9.2992 | 30.3465 | 3.7427 | 16.7393
2025-04-01 11:31 | 4.3688 | 37.9301 | 2.5813 | 17.8029
2025-04-01 11:31 | 0 | 18.3596 | 1.4238 | 12.6253
2025-04-01 11:31 | 0 | 36.9692 | 2.7344 | 13.0125
2025-04-01 11:31 | 0.98294 | 36.263 | 7.0042 | 4.4585
2025-04-01 11:31 | 0.41332 | 43.548 | 1.2688 | 12.6295
2025-04-01 11:31 | 0 | 20.7008 | 2.2221 | 16.6645
2025-04-01 11:31 | 8.0862 | 37.1021 | 3.3748 | 8.1492
2025-04-01 11:31 | 0 | 39.2353 | 2.0874 | 12.2314
2025-04-01 11:31 | 2.3161 | 29.8508 | 4.2458 | 12.8527
2025-04-01 11:31 | 4.2351 | 38.3458 | 4.848 | 8.9115
2025-04-01 11:31 | 1.4772 | 24.1035 | 5.1726 | 16.2375
2025-04-01 11:31 | 0 | 30.1061 | 2.4688 | 9.8954
2025-04-01 11:31 | 0.1711 | 38.5817 | 0.29405 | 16.0753
2025-04-01 11:31 | 6.468 | 30.9213 | 2.0897 | 14.3018
2025-04-01 11:31 | 1.8333 | 44.5629 | 9.6881 | 10.3778
2025-04-01 11:31 | 2.8415 | 47.7328 | 4.766 | 14.6685
2025-04-01 11:31 | 0 | 45.4784 | 4.5597 | 7.9112
2025-04-01 11:31 | 0 | 18.4597 | 1.4434 | 7.3387
2025-04-01 11:31 | 0 | 45.9742 | 5.4994 | 13.338
2025-04-01 11:31 | 0 | 30.8862 | 4.1733 | 19.0247
2025-04-01 11:31 | 11.3474 | 36.6472 | 3.9448 | 12.8448
2025-04-01 11:31 | 0.96399 | 30.1489 | 4.3829 | 15.6661
2025-04-01 11:31 | 0 | 45.4547 | 5.5847 | 9.631
2025-04-01 11:31 | 9.8114 | 40.1942 | 3.896 | 12.5084
2025-04-01 11:31 | 4.978 | 31.1528 | 7.5715 | 15.1674
2025-04-01 11:31 | 0.11393 | 26.5552 | 2.7694 | 27.3002
2025-04-01 11:31 | 3.6306 | 32.4408 | 1.2246 | 20.0556
2025-04-01 11:31 | 3.6057 | 39.1665 | 2.6334 | 16.6745
2025-04-01 11:31 | 1.1377 | 34.6538 | 3.1462 | 21.5899
2025-04-01 11:31 | 0 | 22.0065 | 2.839 | 16.0865
2025-04-01 11:31 | 0 | 14.1336 | 4.7195 | 2.0155
2025-04-01 11:31 | 4.2688 | 35.7895 | 1.9312 | 17.232
2025-04-01 11:31 | 0 | 29.5449 | 2.7019 | 9.4973
2025-04-01 11:31 | 10.4397 | 30.5159 | 1.7059 | 8.92
2025-04-01 11:31 | 6.823 | 30.3062 | 3.1679 | 21.1431
2025-04-01 11:31 | 3.1008 | 31.9423 | 0.55296 | 5.0978
2025-04-01 11:31 | 0 | 20.3301 | 2.4525 | 12.272
2025-04-01 11:31 | 0.71073 | 30.5115 | 5.7086 | 16.0096
2025-04-01 11:31 | 0 | 32.1015 | 6.8709 | 27.8029
2025-04-01 11:31 | 2.3474 | 50.3859 | 1.9076 | 8.8502
2025-04-01 11:31 | 7.0717 | 33.5422 | 3.1407 | 11.7085
2025-04-01 11:31 | 7.601 | 27.4663 | 4.7081 | 10.2213
2025-04-01 11:31 | 0 | 44.6296 | 1.4612 | 12.8971
2025-04-01 11:31 | 9.2289 | 24.1775 | 4.7366 | 10.3449
2025-04-01 11:31 | 4.4359 | 26.7343 | 2.9558 | 14.0897
2025-04-01 11:31 | 8.1683 | 27.8154 | 5.1273 | 7.8535
2025-04-01 11:31 | 0 | 28.4841 | 3.915 | 5.4481
2025-04-01 11:31 | 0 | 26.1147 | 1.7215 | 10.8047
2025-04-01 11:31 | 4.0355 | 36.1452 | 3.477 | 20.6209
2025-04-01 11:31 | 12.362 | 24.0258 | 7.495 | 9.4862
2025-04-01 11:31 | 0 | 53.9637 | 4.7341 | 27.3518
2025-04-01 11:31 | 0 | 47.0511 | 1.4257 | 13.2651
2025-04-01 11:31 | 11.034 | 41.2887 | 3.6913 | 28.4029
2025-04-01 11:31 | 4.9898 | 32.7239 | 3.2113 | 16.9041
2025-04-01 11:31 | 16.2765 | 42.7407 | 3.169 | 13.0982
2025-04-01 11:31 | 7.7247 | 21.0323 | 3.7967 | 11.5104
2025-04-01 11:31 | 3.4726 | 37.5567 | 5.0299 | 15.5711
2025-04-01 11:31 | 0 | 22.4443 | 6 | 14.814
2025-04-01 11:31 | 2.0054 | 36.7695 | 2.4319 | 12.8843
2025-04-01 11:32 | 12.1902 | 34.621 | 1.5085 | 20.9713
2025-04-01 11:32 | 0.044187 | 27.9066 | 5.7215 | 15.4287
2025-04-01 11:32 | 5.0949 | 47.8534 | 0.83023 | 11.3379
2025-04-01 11:32 | 11.6761 | 33.2068 | 5.052 | 16.1598
2025-04-01 11:32 | 4.6755 | 40.6491 | 2.5199 | 14.9805
2025-04-01 11:32 | 0 | 38.766 | 4.4128 | 14.8023
2025-04-01 11:32 | 3.134 | 40.0104 | 2.5043 | 5.7551
2025-04-01 11:32 | 9.7596 | 37.2237 | 6.1059 | 10.8829
2025-04-01 11:32 | 6.3905 | 40.229 | 1.7904 | 19.9425
2025-04-01 11:32 | 12.8766 | 33.901 | 2.1144 | 20.3288
2025-04-01 11:32 | 2.8961 | 29.058 | 3.1786 | 9.3875
2025-04-01 11:32 | 0 | 35.5325 | 4.2506 | 10.8212
2025-04-01 11:32 | 12.2144 | 34.3495 | 7.5204 | 13.2412
2025-04-01 11:32 | 13.2744 | 48.3029 | 5.9309 | 10.778
2025-04-01 11:32 | 0 | 14.0391 | 3.896 | 16.9418
2025-04-01 11:32 | 1.8791 | 24.4143 | 2.8285 | 22.0628
2025-04-01 11:32 | 2.0352 | 44.4844 | 1.5048 | 6.7187
2025-04-01 11:32 | 0.62834 | 27.8553 | 4.7473 | 13.0102
2025-04-01 11:32 | 12.5824 | 21.3771 | 3.0341 | 11.8953
2025-04-01 11:32 | 5.4079 | 30.4181 | 4.4073 | 12.6036
2025-04-01 11:32 | 2.7087 | 45.5897 | 1.5571 | 9.6385
2025-04-01 11:32 | 7.0631 | 31.8437 | 3.5687 | 17.1625
2025-04-01 11:32 | 16.9866 | 32.9515 | 3.4711 | 24.9307
2025-04-01 11:32 | 3.9146 | 37.0569 | 1.7066 | 13.1567
2025-04-01 11:32 | 0 | 31.7356 | 6.3975 | 6.7061
2025-04-01 11:32 | 1.3005 | 41.7696 | 4.3889 | 18.8899
2025-04-01 11:32 | 4.3145 | 37.8533 | 1.755 | 14.2232
2025-04-01 11:32 | 0 | 31.9869 | 3.4121 | 9.6693
2025-04-01 11:32 | 10.5095 | 41.3257 | 5.515 | 10.503
2025-04-01 11:32 | 7.9031 | 25.245 | 4.3652 | 10.028
2025-04-01 11:32 | 0 | 22.3728 | 2.4572 | 7.9915
2025-04-01 11:32 | 0.79474 | 35.8829 | 5.1543 | 16.2988
2025-04-01 11:32 | 0 | 28.8108 | 4.122 | 5.7192
2025-04-01 11:32 | 1.8275 | 17.7039 | 5.6719 | 8.1189
2025-04-01 11:32 | 2.4395 | 19.3091 | 3.8497 | 12.3142
2025-04-01 11:32 | 4.9034 | 30.8086 | 3.5109 | 11.5784
2025-04-01 11:32 | 7.0886 | 32.2725 | 6.0018 | 8.1304
2025-04-01 11:32 | 1.5557 | 23.0294 | 3.3532 | 17.5771
2025-04-01 11:32 | 5.5227 | 24.8199 | 6.1826 | 20.0908
2025-04-01 11:32 | 0 | 22.6034 | 1.3725 | 10.9348
2025-04-01 11:32 | 0 | 26.2477 | 2.3945 | 16.1014
2025-04-01 11:32 | 3.2589 | 24.3373 | 0.8257 | 9.8518
2025-04-01 11:32 | 0 | 40.3744 | 2.0384 | 17.5234
2025-04-01 11:32 | 2.4884 | 35.9904 | 2.1275 | 21.1367
2025-04-01 11:32 | 1.7257 | 27.2482 | 4.4199 | 23.1783
2025-04-01 11:32 | 0.96052 | 32.5716 | 3.2321 | 13.3764
2025-04-01 11:32 | 6.0799 | 41.8146 | 5.007 | 13.5843
2025-04-01 11:32 | 0 | 19.0481 | 0.9787 | 10.2244
2025-04-01 11:32 | 5.3255 | 37.4467 | 5.1951 | 7.948
2025-04-01 11:32 | 13.3294 | 43.1753 | 2.0486 | 12.2752
2025-04-01 11:32 | 2.6278 | 19.5921 | 0.40371 | 15.7877
2025-04-01 11:32 | 8.2749 | 23.9964 | 2.9778 | 15.5782
2025-04-01 11:32 | 13.986 | 31.4608 | 2.5327 | 5.3749
2025-04-01 11:32 | 5.4047 | 29.2826 | 2.814 | 20.369
2025-04-01 11:32 | 0 | 28.2118 | 3.3932 | 6.5605
2025-04-01 11:32 | 3.4861 | 26.8938 | 2.1955 | 18.2816
2025-04-01 11:32 | 3.3764 | 27.4674 | 4.5514 | 13.8096
2025-04-01 11:32 | 0.11055 | 22.3338 | 3.2949 | 11.031
2025-04-01 11:32 | 15.1939 | 23.9922 | 1.7822 | 13.8961
2025-04-01 11:32 | 3.1237 | 32.2833 | 4.5067 | 23.0373
2025-04-01 11:33 | 0 | 23.4274 | 3.5666 | 12.2184
2025-04-01 11:33 | 6.2003 | 26.5297 | 3.9591 | 6.1138
2025-04-01 11:33 | 8.9821 | 22.8859 | 1.8748 | 10.2333
2025-04-01 11:33 | 9.8346 | 34.4806 | 2.9273 | 8.3589
2025-04-01 11:33 | 5.4333 | 22.2704 | 5.0276 | 26.2079
2025-04-01 11:33 | 3.5919 | 25.4067 | 1.4969 | 13.67
2025-04-01 11:33 | 4.9269 | 23.529 | 3.1991 | 10.7972
2025-04-01 11:33 | 4.8985 | 28.6142 | 4.065 | 21.8792
2025-04-01 11:33 | 0.88922 | 17.1536 | 4.3357 | 15.8491
2025-04-01 11:33 | 0 | 25.5302 | 1.1852 | 8.007
2025-04-01 11:33 | 7.5771 | 40.3939 | 3.7424 | 28.3455
2025-04-01 11:33 | 2.7141 | 19.1741 | 4.5281 | 14.6466
2025-04-01 11:33 | 3.3396 | 37.6658 | 4.6691 | 6.0391
2025-04-01 11:33 | 3.3923 | 35.7597 | 3.4621 | 14.5839
2025-04-01 11:33 | 2.9118 | 21.0814 | 4.9241 | 32.4938
2025-04-01 11:33 | 2.7078 | 43.2009 | 3.4635 | 4.3758
2025-04-01 11:33 | 0 | 34.771 | 6.818 | 7.0759
2025-04-01 11:33 | 0.43502 | 31.17 | 6.1631 | 18.136
2025-04-01 11:33 | 1.2021 | 19.5843 | 2.5073 | 17.2034
2025-04-01 11:33 | 10.6218 | 35.9007 | 6.4269 | 13.5809
2025-04-01 11:33 | 0 | 18.6507 | 3.0937 | 7.3707
2025-04-01 11:33 | 3.2934 | 19.925 | 3.0032 | 15.3369
2025-04-01 11:33 | 2.1273 | 31.7906 | 1.9732 | 15.6032
2025-04-01 11:33 | 0 | 23.0819 | 3.9938 | 12.2194
2025-04-01 11:33 | 2.163 | 30.1874 | 5.5995 | 13.3486
2025-04-01 11:33 | 1.3446 | 26.0843 | 2.1152 | 20.4703
2025-04-01 11:33 | 0 | 27.3541 | 4.1569 | 17.1422
2025-04-01 11:33 | 7.769 | 42.4575 | 6.8801 | 11.4005
2025-04-01 11:33 | 11.1432 | 38.8843 | 2.3699 | 20.0838
2025-04-01 11:33 | 6.9883 | 25.9553 | 2.8405 | 18.9139
2025-04-01 11:33 | 3.3813 | 24.3746 | 2.5134 | 8.8437
2025-04-01 11:33 | 0 | 21.2219 | 1.2195 | 14.8672
2025-04-01 11:33 | 2.5473 | 35.6416 | 5.7229 | 11.6318
2025-04-01 11:33 | 0 | 38.4576 | 5.1369 | 14.8522
2025-04-01 11:33 | 7.4876 | 25.8905 | 1.2957 | 16.5891
2025-04-01 11:33 | 8.1563 | 29.2936 | 3.1746 | 18.8714
2025-04-01 11:33 | 3.7323 | 16.2403 | 3.8097 | 15.2623
2025-04-01 11:33 | 0 | 28.2118 | 3.7495 | 13.2083
2025-04-01 11:33 | 3.8557 | 27.0269 | 6.647 | 11.6883
2025-04-01 11:33 | 3.8095 | 26.4517 | 2.0649 | 17.5814
2025-04-01 11:33 | 6.4506 | 20.0905 | 1.7941 | 13.8323
2025-04-01 11:33 | 8.6207 | 42.2146 | 2.3401 | 10.6587
2025-04-01 11:33 | 0.010989 | 35.3395 | 3.9889 | 7.2915
2025-04-01 11:33 | 0.92039 | 36.4462 | 6.2148 | 15.1912
2025-04-01 11:33 | 5.2837 | 30.8144 | 2.8389 | 10.6044
2025-04-01 11:33 | 6.7175 | 34.3607 | 4.4339 | 12.4229
2025-04-01 11:33 | 5.2282 | 30.089 | 3.9772 | 8.8763
2025-04-01 11:33 | 20.5896 | 26.0738 | 4.1321 | 14.8502
2025-04-01 11:33 | 0 | 27.6904 | 5.3831 | 20.4125
2025-04-01 11:33 | 0 | 40.6957 | 6.701 | 9.2006
2025-04-01 11:33 | 0.41542 | 37.0389 | 4.0909 | 9.9716
2025-04-01 11:33 | 12.661 | 29.2786 | 4.5328 | 12.5353
2025-04-01 11:33 | 8.9456 | 30.6332 | 1.9876 | 13.2456
2025-04-01 11:33 | 3.6654 | 28.6238 | 3.3416 | 8.7907
2025-04-01 11:33 | 12.1344 | 20.2663 | 2.4012 | 17.7583
2025-04-01 11:33 | 4.3117 | 24.4053 | 2.1374 | 13.1242
2025-04-01 11:33 | 9.3388 | 30.081 | 3.0728 | 11.5318
2025-04-01 11:33 | 0 | 33.5274 | 5.8182 | 6.1632
2025-04-01 11:33 | 3.3019 | 34.8678 | 2.4738 | 11.1306
2025-04-01 11:33 | 5.6674 | 32.9085 | 2.9481 | 17.7868
2025-04-01 11:34 | 1.8578 | 38.6691 | 3.7629 | 7.0799
2025-04-01 11:34 | 4.4456 | 37.3793 | 4.2414 | 11.2406
2025-04-01 11:34 | 0 | 26.6904 | 3.4359 | 13.3821
2025-04-01 11:34 | 0.99037 | 35.6245 | 2.5643 | 16.8347
2025-04-01 11:34 | 9.4195 | 40.1538 | 4.0061 | 13.5892
2025-04-01 11:34 | 0 | 25.6368 | 2.0625 | 13.4694
2025-04-01 11:34 | 6.3297 | 32.1342 | 1.3225 | 17.4685
2025-04-01 11:34 | 4.977 | 18.9253 | 3.0657 | 18.7037
2025-04-01 11:34 | 0 | 25.8217 | 3.9935 | 15.1096
2025-04-01 11:34 | 0 | 35.4772 | 1.9912 | 18.6318
2025-04-01 11:34 | 2.1218 | 35.6989 | 7.5911 | 8.8522
2025-04-01 11:34 | 9.8399 | 30.6716 | 2.7941 | 20.8724
2025-04-01 11:34 | 0 | 29.9727 | 3.677 | 14.2208
2025-04-01 11:34 | 0 | 24.1175 | 3.1994 | 15.1942
2025-04-01 11:34 | 3.053 | 44.1825 | 4.6356 | 12.8115
2025-04-01 11:34 | 4.244 | 36.1297 | 3.76 | 20.9717
2025-04-01 11:34 | 4.0364 | 25.7274 | 5.0109 | 10.5009
2025-04-01 11:34 | 11.8445 | 33.0882 | 2.9176 | 11.0688
2025-04-01 11:34 | 0.58532 | 32.1447 | 7.3851 | 13.3521
2025-04-01 11:34 | 0.24653 | 27.4608 | 2.3412 | 5.5557
2025-04-01 11:34 | 0 | 23.5771 | 2.6251 | 23.2778
2025-04-01 11:34 | 3.1017 | 33.074 | 3.6461 | 10.7818
2025-04-01 11:34 | 15.5209 | 20.5358 | 4.3618 | 11.4936
2025-04-01 11:34 | 1.8415 | 20.8646 | 3.0236 | 21.2647
2025-04-01 11:34 | 10.0759 | 28.7521 | 2.4157 | 19.9537
2025-04-01 11:34 | 3.2842 | 37.6554 | 4.0371 | 12.6655
2025-04-01 11:34 | 0 | 34.4736 | 4.8305 | 10.386
2025-04-01 11:34 | 7.2273 | 42.5128 | 3.9645 | 10.1943
2025-04-01 11:34 | 0 | 29.3 | 1.9196 | 15.656
2025-04-01 11:34 | 0 | 38.4512 | 7.7568 | 18.1516
2025-04-01 11:34 | 5.6818 | 27.9864 | 5.1737 | 24.7335
2025-04-01 11:34 | 9.3701 | 29.147 | 3.2974 | 12.0357
2025-04-01 11:34 | 3.9864 | 27.2351 | 1.7276 | 7.2019
2025-04-01 11:34 | 0 | 29.1905 | 5.0948 | 6.4421
2025-04-01 11:34 | 3.0037 | 28.5268 | 2.3677 | 17.3516
2025-04-01 11:34 | 10.6774 | 36.8935 | 3.0897 | 10.6532
2025-04-01 11:34 | 7.0205 | 23.5685 | 1.9264 | 18.1827
2025-04-01 11:34 | 0 | 33.9692 | 1.7778 | 5.5016
2025-04-01 11:34 | 7.8304 | 25.0074 | 2.4781 | 31.4719
2025-04-01 11:34 | 3.2055 | 29.8086 | 5.023 | 12.5724
2025-04-01 11:34 | 10.4227 | 28.9429 | 3.8901 | 19.9797
2025-04-01 11:34 | 2.8407 | 25.2157 | 0.89009 | 17.1087
2025-04-01 11:34 | 2.7946 | 36.6182 | 6.7744 | 15.5774
2025-04-01 11:34 | 2.0624 | 34.1684 | 2.6996 | 17.3565
2025-04-01 11:34 | 7.2958 | 28.8193 | 3.8868 | 8.8297
2025-04-01 11:34 | 0.31222 | 19.5986 | 5.6923 | 6.814
2025-04-01 11:34 | 6.7741 | 36.2534 | 2.2966 | 3.2367
2025-04-01 11:34 | 10.7739 | 25.6254 | 4.1171 | 10.3919
2025-04-01 11:34 | 0 | 35.4752 | 4.6122 | 12.0125
2025-04-01 11:34 | 9.9241 | 26.7511 | 7.1652 | 13.5386
2025-04-01 11:34 | 3.2008 | 33.5793 | 4.958 | 14.5909
2025-04-01 11:34 | 5.7877 | 24.3165 | 2.3887 | 32.7712
2025-04-01 11:34 | 0 | 21.2957 | 5.9364 | 14.3714
2025-04-01 11:34 | 3.7101 | 37.0865 | 2.9228 | 27.8112
2025-04-01 11:34 | 7.3542 | 11.0216 | 2.9723 | 15.9914
2025-04-01 11:34 | 0 | 30.9523 | 4.1172 | 11.2424
2025-04-01 11:34 | 3.2084 | 26.1842 | 2.3619 | 22.1385
2025-04-01 11:34 | 10.4578 | 36.0129 | 2.968 | 8.5465
2025-04-01 11:34 | 1.2432 | 20.8764 | 4.1111 | 13.8214
2025-04-01 11:34 | 15.8012 | 42.859 | 2.0368 | 12.7235
2025-04-01 11:35 | 20.2745 | 34.3737 | 4.6979 | 16.4243
2025-04-01 11:35 | 6.9001 | 28.762 | 4.8614 | 10.1449
2025-04-01 11:35 | 6.8162 | 30.5526 | 4.1946 | 28.4486
2025-04-01 11:35 | 0 | 45.0157 | 4.7314 | 18.958
2025-04-01 11:35 | 12.0639 | 30.8574 | 4.6624 | 21.2005
2025-04-01 11:35 | 2.7463 | 29.3008 | 3.4778 | 9.9588
2025-04-01 11:35 | 0.42521 | 25.9848 | 1.4618 | 17.903
2025-04-01 11:35 | 0.084871 | 40.8514 | 2.767 | 7.4758
2025-04-01 11:35 | 0 | 24.3863 | 4.9667 | 7.6227
2025-04-01 11:35 | 0 | 48.7785 | 4.013 | 6.5702
2025-04-01 11:35 | 7.8203 | 30.4628 | 1.3148 | 13.6142
2025-04-01 11:35 | 3.4237 | 47.4086 | 2.5657 | 14.8493
2025-04-01 11:35 | 0 | 27.9972 | 1.5106 | 5.4404
2025-04-01 11:35 | 6.6375 | 29.2992 | 4.1607 | 10.2888
2025-04-01 11:35 | 2.7727 | 28.1664 | 2.9825 | 9.6665
2025-04-01 11:35 | 0.88695 | 23.9935 | 3.2347 | 11.3707
2025-04-01 11:35 | 5.9401 | 31.9953 | 4.2982 | 14.6695
2025-04-01 11:35 | 1.6005 | 21.945 | 5.7822 | 7.7573
2025-04-01 11:35 | 1.4925 | 24.1046 | 2.5577 | 5.8294
2025-04-01 11:35 | 3.7014 | 32.0584 | 2.5298 | 8.405
2025-04-01 11:35 | 6.4679 | 22.569 | 4.6592 | 11.1219
2025-04-01 11:35 | 2.2718 | 33.6691 | 0.32765 | 4.99
2025-04-01 11:35 | 4.9867 | 25.2801 | 1.6152 | 7.6644
2025-04-01 11:35 | 6.7519 | 29.6414 | 3.304 | 3.3517
2025-04-01 11:35 | 0.68482 | 24.237 | 1.5293 | 6.2437
2025-04-01 11:35 | 0 | 24.8901 | 3.5783 | 16.6649
2025-04-01 11:35 | 0 | 19.7628 | 6.2766 | 17.9844
2025-04-01 11:35 | 8.602 | 22.4202 | 1.4554 | 8.7958
2025-04-01 11:35 | 5.3859 | 40.2624 | 6.605 | 12.6488
2025-04-01 11:35 | 6.5403 | 19.7674 | 0.70415 | 12.8306
2025-04-01 11:35 | 2.9898 | 23.0099 | 2.0185 | 13.8862
2025-04-01 11:35 | 10.241 | 28.9541 | 1.3371 | 12.4442
2025-04-01 11:35 | 1.3962 | 26.7667 | 4.1833 | 15.3897
2025-04-01 11:35 | 7.3355 | 28.8434 | 4.4127 | 4.4208
2025-04-01 11:35 | 0 | 31.8026 | 4.7121 | 17.6612
2025-04-01 11:35 | 0 | 45.4716 | 1.8142 | 5.1379
2025-04-01 11:35 | 0 | 24.9429 | 5.2924 | 14.6564
2025-04-01 11:35 | 1.7635 | 29.011 | 6.7473 | 8.4165
2025-04-01 11:35 | 0 | 36.8954 | 4.0816 | 13.4893
2025-04-01 11:35 | 0.90293 | 26.884 | 5.1569 | 8.37
2025-04-01 11:35 | 4.6379 | 23.7063 | 6.2532 | 5.4277
2025-04-01 11:35 | 0 | 27.8983 | 1.0901 | 15.4028
2025-04-01 11:35 | 9.8882 | 34.9332 | 2.5944 | 21.076
2025-04-01 11:35 | 0.25927 | 37.3972 | 2.1775 | 17.6734
2025-04-01 11:35 | 0.95867 | 29.791 | 7.4825 | 12.4594
2025-04-01 11:35 | 3.9997 | 40.2786 | 1.9357 | 31.3518
2025-04-01 11:35 | 1.5224 | 37.6385 | 0.90025 | 5.9094
2025-04-01 11:35 | 3.3793 | 19.0068 | 5.2882 | 18.8707
2025-04-01 11:35 | 0 | 20.6112 | 3.6011 | 16.4883
2025-04-01 11:35 | 1.1226 | 31.0729 | 3.759 | 19.3941
2025-04-01 11:35 | 13.6202 | 50.88 | 2.7132 | 5.7846

## I-CIMS LODs+LOQs
[6 empty] | Calculations | [7 empty] | Results
| 4:2 FTOH |  |  | 6:2 FTOH |  |  |  | 8:2 FTOH |  |  | 10:2 FTOH |  |  | Mean BG Signal (cps) | LOD (pptv) | LOQ (pptv)
Cf | 3.2033 | cps/pptv | Cf | 8.7278 | cps/pptv |  | Cf | 8.2272 | cps/pptv | Cf | 9.0228 | cps/pptv | 4:2 FTOH | 4.5111 | 4.5491 | 33.8178
[X] LOD | 4.5491 | pptv | [X] LOD | 3.5175 | pptv |  | [X] LOD | 1.6854 | pptv | [X] LOD | 2.3505 | pptv | 6:2 FTOH | 37.0118 | 3.5175 | 17.1304
[X] LOQ | 33.8178 | pptv | [X] LOQ | 17.1304 | pptv |  | [X] LOQ | 13.0065 | pptv | [X] LOQ | 13.6683 | pptv | 8:2 FTOH | 3.7489 | 1.6854 | 13.0065
t | 1 | sec | t | 1 | sec |  | t | 1 | sec | t | 1 | sec | 10:2 FTOH | 14.3834 | 2.3505 | 13.6683
B | 4.5111 | cps | B | 37.0118 | cps |  | B | 3.7489 | cps | B | 14.3834 | cps
LOD | 3 | Set to 3 | LOD | 3 | Set to 3 |  | LOD | 3 | Set to 3 | LOD | 3 | Set to 3
LOQ | 10 | Set to 10 | LOQ | 10 | Set to 10 |  | LOQ | 10 | Set to 10 | LOQ | 10 | Set to 10

## Unit Conversions
Avg. Flow Rate (mL/min) | Sampling Time (min) | Avg. Vol. Sampled (mL) |  | MDL (ng on tube) | MDL (µg/m3) | MDL (pptv) | RL (ng on tube) | RL (µg/m3) | RL (pptv)
102.3842 | 25 | 2559.6042 | 4:2 FTOH | 0.094896 | 0.037075 | 3.4324 | 0.28469 | 0.11122 | 10.2973
|  |  | 6:2 FTOH | 0.088081 | 0.034412 | 2.3108 | 0.26424 | 0.10324 | 6.9325
Volume of 1 mole gas @ STP (L): | 24.45 |  | 8:2 FTOH | 0.13954 | 0.054516 | 2.8719 | 0.41862 | 0.16355 | 8.6157
|  |  | 10:2 FTOH | 0.26435 | 0.10328 | 4.4761 | 0.79304 | 0.30983 | 13.4284
| Mol wt (g/mol) |  |  | LOD (pptv) | LOD (µg/m3) | LOQ (pptv) | LOQ (µg/m3)
4:2 FTOH | 264.09 |  | 4:2 FTOH | 4.5491 | 0.049136 | 33.8178 | 0.36527
6:2 FTOH | 364.1 |  | 6:2 FTOH | 3.5175 | 0.052382 | 17.1304 | 0.2551
8:2 FTOH | 464.12 |  | 8:2 FTOH | 1.6854 | 0.031994 | 13.0065 | 0.24689
10:2 FTOH | 564.13 |  | 10:2 FTOH | 2.3505 | 0.054232 | 13.6683 | 0.31537