## Figure 1
Figure 1. On-line solid phase extraction (SPE) liquid chromatography/mass spectrometry (LC/MS) chromatograms of m/z122 illustrating background Trizma intensities under various SPE/mobile phase conditions using a 5-mL injection of deionized water. (A) and (B) are preserved with 2.0 g/L 2-CA, 7.75 g/L Trizma and 200 mg/L ascorbic acid while (C) has no preservatives.
No Metadata associated with this Figure

## Figure 2
Figure 2. On-line solid phase extraction (SPE) liquid chromatography/mass spectrometry (LC/MS) chromatograms (overlaid MS/MS segments) of a calibration standard with Method 543 analytes at concentration levels of 20-80 ng/L. Peak 1 is methomyl-13C2,15N, 2 is 3-hydroxycarbofuran, 3 is fenamiphos sulfoxide, 4 is fenamiphos sulfone, 5 is carbofuran-13C6, 6 is fenamiphos, 7 is tebuconazole, 8 is tebufenozide, 9 is bensulide-d14, 10 is bensulide
No Metadata associated with this Figure

## Table 1
Table 1. Chemicals evaluated for inclusion in EPA Method 543, retention times and MS/MS parameters
Analyte | Retention Time (min) | Precursor Ion (m/z) | Product Ion (m/z) | Cone Voltage (volts) | Collision Energy (volts) | IS used
CCL 3 chemicals and degradates
3-hydroxycarbofuran | 5.91 | 238 | 181 | 20 | 10 | 1
fenamiphos sulfoxide | 6.31 | 320 | 233 | 32 | 26 | 2
fenamiphos sulfone | 6.7 | 336 | 266 | 28 | 22 | 2
fenamiphos | 7.59 | 304 | 217 | 28 | 24 | 3
tebuconazole | 7.75 | 308 | 70 | 26 | 20 | 3
tebufenozide | 8.06 | 353 | 133 | 12 | 22 | 3
bensulide | 8.27 | 398 | 314 | 20 | 12 | 3
Internal standards
methomyl-13C2,15N (IS#1) | 5.58 | 166 | 91 | 15 | 10 | -
carbofuran-13C6 (IS#2) | 6.82 | 228 | 171 | 20 | 10 | -
bensulide-d14 (IS#3) | 8.25 | 412 | 316 | 22 | 12 | -
No Metadata associated with this table

## Table 2
Table 2. Mean recovery and precision (n=4) of method analytes in preserved tap water (TOC=3 mg/L) using a 5-mL and 2-mL on-line SPE injection.
|  | 5-mL injection volume |  | 2-mL injection volume
| Spike Level ng/L | Mean Rec % | RSD % | Mean Rec % | RSD %
3-hydroxycarbofuran | 20 | 58.2 | 5.9 | 91.1 | 3.5
fenamiphos sulfoxide | 20 | 80.6 | 2.6 | 102 | 2
fenamiphos sulfone | 20 | 93.8 | 2.3 | 104 | 2
fenamiphos | 8 | 99.6 | 2.1 | 91.5 | 1.5
tebuconazole | 8 | 109 | 2.3 | 95.8 | 2.3
tebufenozide | 8 | 112 | 4.4 | 89.7 | 1.9
bensulide | 20 | 103 | 2.9 | 93.7 | 6.6
methomyl-13C2,15N |  | NDa | ND | ND | ND
carbofuran-13C6 | 32 | 81.3 | 0.9 | 98.1 | 1.3
bensulide-d14 | 80 | 96.7 | 2.5 | 101 | 2
a Methomyl-13C2,15N was not added as an IS until after this study.
Metadata for Table 2
2 mL injection into 5-mL loop
MS Method: MRM_05_29_13 |  |  | Elution = 20 mM NH4OAc/ 100% ACN
Quan Method: OLSPE_07_03_13 |  |  | Condition = 100% ACN
LC Method: CCL3_05_29_13_AB |  |  | 90:10 initial mobile phase 20 mM NH4OAc:ACN
Cal File: OLSPE_07_24_13_c
Quan File: OLSPE_07_24_13_c |  |  | Load with 20 mM NH4OAc at 2mL/min
Wash =20 mM NH4OAc at 2mL/min
Tap water LFSMs ng/L
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Mean | % Recovery | %RSD | Matrix Blank ng/L | LRB ng/L | LFB % Recovery
3-hydroxycarbofuran | 20 | 17.8933 | 17.6503 | 19.1209 | 18.1896 | 18.2135 | 91.0676 | 3.5351 | 0 | 0 | 98.9
fenamiphos sulfoxide | 20 | 19.7636 | 20.5533 | 20.7024 | 20.3615 | 20.3452 | 101.726 | 2.0254 | 0 | 0 | 99.2
fenamiphos sulfone | 20 | 20.782 | 20.7871 | 21.1717 | 20.1862 | 20.7317 | 103.6587 | 1.9628 | 0 | 0 | 100.8
fenamiphos | 8 | 7.3819 | 7.385 | 7.1538 | 7.3737 | 7.3236 | 91.545 | 1.5471 | 0 | 0 | 93.5
tebuconazole | 8 | 7.7991 | 7.5178 | 7.8368 | 7.5076 | 7.6653 | 95.8166 | 2.3085 | 0.4775 | 0 | 99.8
tebufenozide | 8 | 7.3604 | 7.1644 | 7.1238 | 7.048 | 7.1741 | 89.6769 | 1.8568 | 0 | 0 | 93.3
bensulide | 20 | 19.5419 | 19.6989 | 18.6996 | 17.0132 | 18.7384 | 93.692 | 6.5693 | 0 | 0 | 105.7
carbofuran-13C6 | 16 | 15.8095 | 15.8859 | 15.4267 | 15.6535 | 15.6939 | 98.0869 | 1.2916 | 99.1 | 89.8 | 102.3
tebuconazole-d6 | 16 | 16.3537 | 16.7613 | 17.0311 | 16.7306 | 16.7192 | 104.4948 | 1.6661 | 105.4 | 91.8 | 107.6
bensulide-d14 | 80 | 78.7648 | 79.8437 | 81.6039 | 82.358 | 80.6426 | 100.8033 | 2.0289 | 102.9 | 81 | 104.7

## Table 3
Table 3. Mean recovery and precision (n=5) of method analytes in deionized water and challenging drinking water matrices preserved with 2.0 g/L 2-CA, 7.75 g/L Trizma and 200 mg/L ascorbic acid using a 2-mL injection volume
Analyte | Spike Level ng/L | DI water |  | Tap water 1a |  | Tap water 2b |  | Tap water 3c
|  | Mean Recovery % | RSD % | Mean Recovery % | RSD % | Mean Recovery % | RSD % |  | Mean Recovery % | RSD %
Mid-Level Fortification
3-hydroxycarbofuran | 50 | 96.8 | 9.1 | 98.2 | 4.7 | 109 | 4 |  | 90.1 | 5.5
fenamiphos sulfoxide | 50 | 96.3 | 1.7 | 95.3 | 3.2 | 96.6 | 3 |  | 100 | 4
fenamiphos sulfone | 50 | 94.6 | 5.5 | 99.5 | 3.8 | 99.4 | 2.9 |  | 101 | 2.2
fenamiphos | 20 | 100 | 4.5 | 99.2 | 2.1 | 96 | 4.8 |  | 107 | 2.7
tebuconazole | 20 | 99.9 | 3.6 | 99.4 | 6.9 | 90.9 | 4.8 |  | 86.4 | 2.6
tebufenozide | 20 | 95.6 | 5.3 | 105 | 1.1 | 99.4 | 2.5 |  | 98.1 | 2.5
bensulide | 50 | 92 | 2 | 106 | 4.7 | 101 | 4 |  | 99.8 | 2
methomyl-13C2,15N | 80 | 97.2 | 5.6 | 112 | 4.3 | 96.7 | 3 |  | 106 | 7.1
carbofuran-13C6 | 32 | 100 | 2.7 | 103 | 1.9 | 94.2 | 2.6 |  | 98.7 | 5.1
bensulide-d14 | 80 | 96.1 | 4.3 | 103 | 2.2 | 99.2 | 3.3 |  | 96 | 2.9
Low-Level Fortification
3-hydroxycarbofuran | 5 | 108 | 7.2 | 108 | 14 | 111 | 3.5 |  | 98.2 | 8.5
fenamiphos sulfoxide | 5 | 108 | 4.9 | 95.4 | 2.1 | 102 | 4 |  | 106 | 3.3
fenamiphos sulfone | 5 | 108 | 5.4 | 94.1 | 6.7 | 95 | 1.8 |  | 97.6 | 3.9
fenamiphos | 2 | 102 | 2.6 | 91.9 | 3.8 | 97.2 | 6.7 |  | 97.9 | 5
tebuconazole | 2 | 104 | 3.8 | 87.1 | 6.4 | 103 | 9.8 |  | 98.1 | 3.3
tebufenozide | 2 | 103 | 6.7 | 97.5 | 5.1 | 105 | 8.3 |  | 105 | 9.1
bensulide | 5 | 107 | 11 | 89.8 | 3.8 | 88.4 | 11 |  | 100 | 5.3
methomyl-13C2,15N | 80 | 98.2 | 1.4 | 98.3 | 4.3 | 97.2 | 2.4 |  | 102 | 4.5
carbofuran-13C6 | 32 | 96.1 | 3.9 | 97.4 | 2.2 | 94.3 | 1.5 |  | 99.5 | 2.1
bensulide-d14 | 80 | 95.7 | 3.5 | 98.3 | 1.7 | 91.7 | 2.2 |  | 94.6 | 3.2
a Tap water from a surface water source; TOC = 3.20 mg/L and hardness = 68.5 mg/L as calcium carbonate
b Tap water from a surface water source; TOC = 2.44 mg/L and hardness = 137 mg/L as calcium carbonate
c Tap water from a ground water source; TOC = 0.71 mg/L and hardness = 342 mg/L as calcium carbonate
Metadata for Table 3
LFBs and tap waters with 7.75 g/L trizma, 2 g/L 2-CA, 200 mg/L ascorbic
2 mL injection into 5-mL loop
Elution =20 mM NH4OAc in water/ 100% ACN
Condition = 100% ACN
90:10 initial mobile phase 20 MM NH4OAc : ACN
Load & Wash with 20 mM NH4OAc at 2mL/min
LFBs (DI Water), ng/L
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Rep #5 | Mean | % Recovery | %RSD |  | LRB ng/L | CCC cal 6 % Rec | QCS %Rec | CCC cal 3 % Rec
3-hydroxycarbofuran | 50 | 48.66 | 54.8171 | 43.5599 | 45.0943 | 49.964 | 48.4191 | 96.8382 | 9.1227 |  | 0 | 100.7 | 107.5 | 105
fenamiphos sulfoxide | 50 | 48.8044 | 48.8582 | 48.1479 | 46.9001 | 47.9615 | 48.1344 | 96.2689 | 1.6511 |  | 0 | 97.5 | 113.1 | 97.9
fenamiphos sulfone | 50 | 46.2861 | 51.9334 | 46.1906 | 46.083 | 45.9735 | 47.2933 | 94.5866 | 5.4903 |  | 0 | 96.8 | 124.8 | 95
fenamiphos | 20 | 20.5803 | 20.7906 | 18.918 | 19.1537 | 20.6173 | 20.012 | 100.0599 | 4.4897 |  | 0 | 102.5 | 118.3 | 96.1
tebuconazole | 20 | 20.4312 | 20.7587 | 18.8527 | 19.8353 | 20.063 | 19.9882 | 99.941 | 3.6321 |  | 0 | 98.9 | 118 | 100.6
tebufenozide | 20 | 19.9794 | 19.9441 | 17.7597 | 18.3014 | 19.5818 | 19.1133 | 95.5664 | 5.3301 |  | 0 | 96.5 | 108.5 | 88
bensulide | 50 | 46.5315 | 46.9507 | 45.0055 | 46.5198 | 45.1064 | 46.0228 | 92.0455 | 1.956 |  | 0 | 96 | 111.6 | 93.7
methomyl-13C2, 15N | 80 | 79.0356 | 72.9202 | 82.0482 | 81.3353 | 73.3429 | 77.7364 | 97.1706 | 5.5973 |  | 92.6 | 94.1 | 90.9 | 98
carbofuran-13C6 | 32 | 31.4884 | 31.0959 | 32.3351 | 33.3536 | 32.2807 | 32.1107 | 100.3461 | 2.7154 |  | 102.9 | 97.6 | 95.1 | 100.9
bensulide-d14 | 80 | 72.4759 | 75.8097 | 80.8507 | 79.4601 | 75.7765 | 76.8746 | 96.0932 | 4.3233 |  | 99.9 | 90.6 | 86.7 | 93.2
Tap water 1, ng/L
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Rep #5 | Mean | % REC | %RSD |  | MB ng/L | CCC cal 5 % Rec | LRB ng/L
3-hydroxycarbofuran | 50 | 48.0175 | 49.9017 | 52.358 | 46.1048 | 49.0803 | 49.0925 | 98.1849 | 4.7115 |  | 0 | 106.6 | 0
fenamiphos sulfoxide | 50 | 49.966 | 47.9008 | 47.2943 | 45.7541 | 47.289 | 47.6408 | 95.2817 | 3.1968 |  | 0 | 102.2 | 0
fenamiphos sulfone | 50 | 52.8437 | 49.5504 | 49.894 | 48.3526 | 48.1892 | 49.766 | 99.5319 | 3.7619 |  | 0 | 100.2 | 0
fenamiphos | 20 | 20.4921 | 19.9298 | 19.56 | 19.7119 | 19.4569 | 19.8301 | 99.1507 | 2.0707 |  | 0 | 96.5 | 0
tebuconazole | 20 | 21.8806 | 18.5804 | 20.0722 | 20.2836 | 18.5655 | 19.8765 | 99.3823 | 6.9435 |  | 0 | 95 | 0
tebufenozide | 20 | 21.1016 | 20.7696 | 20.899 | 21.1604 | 20.6385 | 20.9138 | 104.569 | 1.0501 |  | 0 | 100.6 | 0
bensulide | 50 | 54.1356 | 50.2548 | 50.4125 | 56.1317 | 53.0025 | 52.7874 | 105.5749 | 4.7457 |  | 0 | 99.2 | 0
methomyl-13C2, 15N | 80 | 90.0341 | 86.5887 | 84.225 | 94.0199 | 91.0143 | 89.1764 | 111.4705 | 4.2989 |  | 101.5 | 94.7 | 93.4
carbofuran-13C6 | 32 | 32.27 | 33.3552 | 32.591 | 33.7683 | 33.5119 | 33.0993 | 103.4353 | 1.9282 |  | 98.9 | 95.7 | 96.2
bensulide-d14 | 80 | 79.4981 | 84.2088 | 83.6061 | 83.0537 | 82.8837 | 82.6501 | 103.3126 | 2.2224 |  | 92.7 | 100.9 | 94.7
Tap water 2, ng/L
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Rep #5 | Mean | % REC | %RSD |  | MB ng/L | CCC cal 5 % Rec | LRB ng/L
3-hydroxycarbofuran | 50 | 54.2157 | 53.3898 | 53.3484 | 54.001 | 58.6149 | 54.714 | 109.4279 | 4.0449 |  | 0 | 108.9 | 0
fenamiphos sulfoxide | 50 | 47.7695 | 46.4095 | 48.9253 | 50.3398 | 48.053 | 48.2994 | 96.5988 | 3.0125 |  | 0 | 101.8 | 0
fenamiphos sulfone | 50 | 48.6749 | 50.5452 | 48.7852 | 48.6296 | 51.868 | 49.7006 | 99.4011 | 2.924 |  | 0 | 102.2 | 0
fenamiphos | 20 | 19.9443 | 19.1026 | 17.9197 | 20.2221 | 18.8149 | 19.2007 | 96.0036 | 4.7986 |  | 0 | 96.9 | 0
tebuconazole | 20 | 19.4277 | 17.7059 | 17.0968 | 18.4666 | 18.2088 | 18.1812 | 90.9058 | 4.7932 |  | 0 | 96.5 | 0
tebufenozide | 20 | 20.4611 | 19.913 | 19.1484 | 20.1483 | 19.7333 | 19.8808 | 99.4042 | 2.4737 |  | 0 | 102.9 | 0
bensulide | 50 | 52.3245 | 51.0102 | 46.9337 | 51.061 | 50.2179 | 50.3095 | 100.619 | 4.0396 |  | 0 | 101.9 | 0
methomyl-13C2, 15N | 80 | 78.0847 | 76.4471 | 80.8524 | 76.638 | 74.6012 | 77.3247 | 96.6558 | 3.0109 |  | 95 | 95.1 | 100.8
carbofuran-13C6 | 32 | 30.9812 | 30.3893 | 30.5574 | 28.971 | 29.8888 | 30.1575 | 94.2423 | 2.5536 |  | 95.5 | 91.9 | 97.3
bensulide-d14 | 80 | 76.904 | 80.2214 | 83.15 | 77.0221 | 79.6583 | 79.3912 | 99.239 | 3.2536 |  | 95.4 | 91 | 96.8
Tap water 3, ng/L
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Rep #5 | Mean | % REC | %RSD |  | MB ng/L | CCC cal 6 % Rec
3-hydroxycarbofuran | 50 | 47.6484 | 42.1544 | 47.5939 | 43.477 | 44.4481 | 45.0644 | 90.1287 | 5.4854 |  | 0 | 81.9
fenamiphos sulfoxide | 50 | 51.7616 | 52.1054 | 50.6396 | 49.1105 | 47.3223 | 50.1879 | 100.3758 | 3.9513 |  | 0 | 107.4
fenamiphos sulfone | 50 | 51.4972 | 50.6568 | 51.5817 | 50.1166 | 48.9417 | 50.5588 | 101.1176 | 2.1547 |  | 0 | 107.5
fenamiphos | 20 | 20.5417 | 20.8729 | 21.8839 | 21.4242 | 21.8053 | 21.3056 | 106.5279 | 2.7459 |  | 0 | 116
tebuconazole | 20 | 17.9019 | 16.7768 | 17.0526 | 17.5867 | 17.0637 | 17.2763 | 86.3817 | 2.6407 |  | 0 | 106
tebufenozide | 20 | 19.0724 | 19.6682 | 20.3266 | 19.263 | 19.8094 | 19.6279 | 98.1396 | 2.5017 |  | 0 | 106.8
bensulide | 50 | 50.1635 | 50.8049 | 48.1751 | 50.3128 | 50.0239 | 49.896 | 99.7921 | 2.0165 |  | 0 | 104.8
methomyl-13C2, 15N | 80 | 82.049 | 87.7093 | 77.0937 | 93.1839 | 85.6511 | 85.1374 | 106.4218 | 7.0929 |  | 112 | 109
carbofuran-13C6 | 32 | 29.7449 | 30.5509 | 31.1743 | 33.6027 | 32.9217 | 31.5989 | 98.7466 | 5.1227 |  | 102.2 | 91.2
bensulide-d14 | 80 | 74.275 | 74.7594 | 77.3962 | 79.6675 | 77.7946 | 76.7785 | 95.9732 | 2.92 |  | 95 | 79.8

## Table 4
Table 4. Aqueous sample holding time data (n=5) for chlorinated tap water samplesa fortified with method analytesb and preservedc
Analyte |  | Fortified Conc. (ng/L) | Day 0 |  | Day 7 |  | Day 14 |  | Day 22 |  | Day 28
|  |  | Mean %Recovery | % RSD | Mean %Recovery | % RSD | Mean %Recovery | % RSD | Mean %Recovery | % RSD | Mean %Recovery | % RSD
3-hydroxycarbofuran |  | 50 | 115 | 3.8 | 114 | 3.8 | 112 | 5 | 100 | 7.3 | 94.1 | 3.9
fenamiphos sulfoxide |  | 50 | 108 | 3.1 | 107 | 2.6 | 106 | 2.9 | 110 | 4 | 110 | 2.9
fenamiphos sulfone |  | 50 | 109 | 3.6 | 107 | 0.9 | 104 | 3.2 | 109 | 4.9 | 110 | 3.2
fenamiphos |  | 20 | 109 | 1.8 | 104 | 2.9 | 107 | 2.2 | 105 | 6.9 | 99.3 | 3.4
tebuconazole |  | 20 | 96.2 | 4.1 | 110 | 4.7 | 112 | 4.7 | 102 | 5.3 | 108 | 2.2
tebufenozide |  | 20 | 109 | 4.4 | 116 | 2.1 | 113 | 0.9 | 112 | 4.9 | 105 | 3.2
bensulide |  | 50 | 105 | 2.2 | 105 | 3.8 | 106 | 3.8 | 102 | 8 | 103 | 3.2
methomyl-13C2,15N |  | 80 | 97.1 | 2.9 | 96.4 | 3 | 99.1 | 2.2 | 97.9 | 5.1 | 106 | 3.5
carbofuran-13C6 |  | 32 | 97.1 | 2.6 | 97.9 | 2.2 | 102 | 1.7 | 97.8 | 5 | 99.7 | 1.5
bensulide-d14 |  | 80 | 85.9 | 2.6 | 101 | 4.7 | 97.8 | 1.6 | 94.7 | 5 | 100 | 1.3
a | TOC = 0.72 mg/L and hardness = 154 mg/L as calcium carbonate
b | Internal standards were not added to samples until the day of extraction
c | Preserved with 2.0 g/L 2-CA, 7.75 g/L Trizma and 200 mg/L ascorbic acid
Metadata for Table 4
Day 0 LFMS, ng/L
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Rep #5 | Mean | % REC | %RSD |  | MB ng/L | CCC cal 5 % Recovery | LRB ng/L
3-hydroxycarbofuran | 50 | 56.5548 | 58.8518 | 53.7562 | 58.2878 | 58.7846 | 57.247 | 114.494 | 3.7757 |  | 0 | 109.4 | 0
fenamiphos sulfoxide | 50 | 54.1546 | 54.8551 | 51.2157 | 55.4933 | 53.4291 | 53.8296 | 107.6591 | 3.0692 |  | 0 | 98.4 | 0
fenamiphos sulfone | 50 | 54.1433 | 54.2224 | 51.8937 | 57.1685 | 55.5314 | 54.5919 | 109.1837 | 3.5637 |  | 0 | 98.3 | 0
fenamiphos | 20 | 21.633 | 22.164 | 21.7555 | 21.9602 | 21.12 | 21.7265 | 108.6327 | 1.8172 |  | 0 | 100.4 | 0
tebuconazole | 20 | 19.3025 | 20.4266 | 19.0912 | 19.0873 | 18.2548 | 19.2325 | 96.1624 | 4.0509 |  | 0 | 102.8 | 0
tebufenozide | 20 | 23.1393 | 21.9457 | 20.7512 | 22.298 | 21.046 | 21.836 | 109.1802 | 4.4194 |  | 0 | 109.5 | 0
bensulide | 50 | 52.2638 | 50.9041 | 53.1862 | 53.7233 | 51.5413 | 52.3237 | 104.6475 | 2.207 |  | 0 | 102.1 | 0
methomyl-13C2, 15N | 80 | 76.6678 | 78.1577 | 81.4527 | 76.2853 | 75.9465 | 77.702 | 97.1275 | 2.9088 |  | 94.3 | 96.3 | 92.2
carbofuran-13C6 | 32 | 31.0088 | 30.9945 | 32.1843 | 29.9352 | 31.1737 | 31.0593 | 97.0603 | 2.5702 |  | 91.7 | 96.8 | 98.7
tebuconazole-d6 | 32 | 27.5654 | 27.4667 | 26.3396 | 27.7685 | 28.2624 | 27.4805 | 85.8766 | 2.5753 |  | 94.4 | 91.6 | 85.1
bensulide-d14 | 80 | 76.6716 | 78.9041 | 80.1938 | 77.5295 | 79.7068 | 78.6012 | 98.2515 | 1.8787 |  | 93.9 | 97.1 | 93.4
QC criteria
|  | Day 7 LFMS, ng/L | [9 empty] | 70-130%
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Rep #5 | Mean | % Recovery | %RSD |  | MB ng/L | CCC cal 5 % Recovery | LRB ng/L
3-hydroxycarbofuran | 50 | 57.1089 | 58.1936 | 59.0143 | 53.4257 | 58.0908 | 57.1666 | 114.3333 | 3.8445 |  | 0 | 100 | 0
fenamiphos sulfoxide | 50 | 51.5332 | 53.7767 | 52.2983 | 53.6842 | 55.043 | 53.2671 | 106.5342 | 2.5762 |  | 0 | 97.3 | 0
fenamiphos sulfone | 50 | 53.7779 | 53.5656 | 52.7658 | 54.1242 | 53.7012 | 53.5869 | 107.1739 | 0.9391 |  | 0 | 96.8 | 0
fenamiphos | 20 | 21.3401 | 20.3879 | 21.2909 | 19.9763 | 20.4896 | 20.697 | 103.4848 | 2.883 |  | 0 | 104.9 | 0
tebuconazole | 20 | 22.7373 | 21.4258 | 22.2975 | 20.4822 | 23.0463 | 21.9978 | 109.9891 | 4.7461 |  | 0 | 114.7 | 0
tebufenozide | 20 | 22.7728 | 22.978 | 23.1666 | 23.0973 | 24.0553 | 23.214 | 116.07 | 2.1258 |  | 0 | 109.7 | 0
bensulide | 50 | 53.1356 | 51.8786 | 50.6494 | 50.7169 | 55.4083 | 52.3578 | 104.7155 | 3.7901 |  | 0 | 97.2 | 0
methomyl-13C2, 15N | 80 | 77.9843 | 74.7355 | 77.5815 | 80.208 | 74.9625 | 77.0943 | 96.3679 | 2.9602 |  | 101.2 | 100.6 | 97.4
carbofuran-13C6 | 32 | 32.522 | 30.8875 | 31.0058 | 31.2693 | 31.0179 | 31.3405 | 97.939 | 2.1537 |  | 103.7 | 99.8 | 91.8
bensulide-d14 | 80 | 79.8612 | 79.3593 | 73.9908 | 79.1351 | 75.2218 | 77.5136 | 96.8921 | 3.4862 |  | 98.8 | 91.7 | 86.9
Day 14 LFMS, ng/L
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Rep #5 | Mean | % Recovery | %RSD |  | MB ng/L | CCC cal 5 % Recovery | LRB ng/L
3-hydroxycarbofuran | 50 | 56.8456 | 57.2631 | 54.4523 | 58.5804 | 51.5401 | 55.7363 | 111.4727 | 4.9872 |  | 0 | 116.6 | 0
fenamiphos sulfoxide | 50 | 51.6673 | 52.992 | 55.787 | 53.1115 | 52.4192 | 53.1954 | 106.3908 | 2.9278 |  | 0 | 93.6 | 0
fenamiphos sulfone | 50 | 49.9394 | 54.1114 | 53.2324 | 51.0874 | 52.3754 | 52.1492 | 104.2983 | 3.1923 |  | 0 | 95.7 | 0
fenamiphos | 20 | 21.3637 | 22.17 | 21.3174 | 20.8604 | 21.2443 | 21.3912 | 106.9558 | 2.2372 |  | 0 | 100 | 0
tebuconazole | 20 | 22.4375 | 23.8155 | 21.3487 | 21.6359 | 21.3868 | 22.1249 | 110.6244 | 4.7092 |  | 0 | 103.9 | 0
tebufenozide | 20 | 22.736 | 22.5693 | 22.1962 | 22.5161 | 22.6476 | 22.533 | 112.6651 | 0.91326 |  | 0 | 103 | 0
bensulide | 50 | 51.554 | 56.5264 | 53.2578 | 51.8521 | 52.1537 | 53.0688 | 106.1376 | 3.8391 |  | 0 | 104.7 | 0
methomyl-13C2, 15N | 80 | 80.8875 | 81.3526 | 78.0355 | 78.8138 | 77.4697 | 79.3118 | 99.1398 | 2.1764 |  | 98 | 100.2 | 96
carbofuran-13C6 | 32 | 32.8798 | 32.3357 | 31.6471 | 33.0798 | 32.4361 | 32.4757 | 101.4865 | 1.7117 |  | 94.8 | 106.8 | 94.4
tebuconazole-d6 | 32 | 31.271 | 33.3935 | 31.1472 | 31.4381 | 30.3302 | 31.516 | 98.4875 | 3.5944 |  | 87 | 100.5 | 96.7
bensulide-d14 | 80 | 77.9936 | 76.3546 | 78.1397 | 79.599 | 79.0491 | 78.2272 | 97.784 | 1.582 |  | 91.1 | 99.9 | 100.9
Day 21 LFMS, ng/L
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Rep #5 | Mean | % Recovery | %RSD |  | MB ng/L | LRB ng/L | CCC cal 5 % Recovery
3-hydroxycarbofuran | 50 | 47.4196 | 50.7706 | 46.5286 | 55.8074 | 49.548 | 50.0148 | 100.0296 | 7.2939 |  | 0 | 0 | 99.3
fenamiphos sulfoxide | 50 | 54.1349 | 53.7876 | 54.0126 | 58.7783 | 53.6602 | 54.8747 | 109.7494 | 3.991 |  | 0 | 0 | 99.2
fenamiphos sulfone | 50 | 56.1154 | 52.7463 | 52.4473 | 58.4966 | 52.9808 | 54.5573 | 109.1146 | 4.8636 |  | 0 | 0 | 100.6
fenamiphos | 20 | 19.8392 | 20.5152 | 20.3564 | 23.5417 | 20.9868 | 21.0479 | 105.2394 | 6.9037 |  | 0 | 0 | 95.2
tebuconazole | 20 | 20.3395 | 19.2426 | 20.0165 | 22.1812 | 20.2545 | 20.4068 | 102.0341 | 5.3027 |  | 0 | 0 | 93.3
tebufenozide | 20 | 22.5586 | 21.7913 | 21.5677 | 24.2136 | 21.7431 | 22.3749 | 111.8744 | 4.8989 |  | 0 | 0.072697 | 99
bensulide | 50 | 49.3206 | 48.3945 | 49.4557 | 58.3946 | 50.1633 | 51.1457 | 102.2915 | 8.0179 |  | 0 | 0 | 98.4
methomyl-13C2, 15N | 80 | 79.8459 | 76.7743 | 81.5242 | 72.0535 | 81.2336 | 78.2863 | 97.8579 | 5.0588 |  | 98.1 | 89.8 | 98.6
carbofuran-13C6 | 32 | 31.5978 | 32.2257 | 32.307 | 28.5631 | 31.7463 | 31.288 | 97.7749 | 4.964 |  | 99.4 | 90.4 | 98.4
tebuconazole-d6 | 32 | 30.8712 | 29.6544 | 31.6907 | 27.9768 | 31.2772 | 30.294 | 94.6689 | 4.9595 |  | 91.6 | 83.8 | 93.8
bensulide-d14 | 80 | 79.9602 | 79.9533 | 80.5024 | 70.2091 | 78.6811 | 77.8612 | 97.3265 | 5.5608 |  | 95.8 | 99.1 | 102.6
Day 28 LFMS, ng/L
| Spike Level ng/L | Rep #1 | Rep #2 | Rep #3 | Rep #4 | Rep #5 | Mean | % Recovery | %RSD |  | MB ng/L | LRB ng/L | CCC cal 5 % Recovery
3-hydroxycarbofuran | 50 | 45.4945 | 45.5871 | 46.4469 | 49.7989 | 47.8845 | 47.0424 | 94.0848 | 3.8581 |  | 0 | 0 | 100.2
fenamiphos sulfoxide | 50 | 52.9227 | 57.3593 | 54.6216 | 54.364 | 54.9524 | 54.844 | 109.688 | 2.9273 |  | 0 | 0 | 98.8
fenamiphos sulfone | 50 | 51.8526 | 55.362 | 56.3634 | 54.6922 | 55.8418 | 54.8224 | 109.6448 | 3.2299 |  | 0 | 0 | 99.3
fenamiphos | 20 | 19.5295 | 20.1122 | 18.8607 | 20.6111 | 20.2308 | 19.8688 | 99.3442 | 3.4437 |  | 0 | 0 | 99.4
tebuconazole | 20 | 21.328 | 22.0977 | 21.6701 | 20.905 | 21.8508 | 21.5703 | 107.8517 | 2.1599 |  | 0 | 0 | 99.8
tebufenozide | 20 | 21.5246 | 21.222 | 19.858 | 21.3967 | 21.2615 | 21.0525 | 105.2627 | 3.2222 |  | 0 | 0 | 102.2
bensulide | 50 | 52.2697 | 49.6942 | 49.3254 | 52.739 | 52.4862 | 51.3029 | 102.6058 | 3.217 |  | 0 | 0 | 101.1
methomyl-13C2, 15N | 80 | 84.9246 | 83.4107 | 84.5815 | 80.8467 | 88.933 | 84.5393 | 105.6741 | 3.4671 |  | 103.1 | 99.2 | 95.2
carbofuran-13C6 | 32 | 32.6258 | 31.2729 | 31.8781 | 31.8029 | 31.9925 | 31.9144 | 99.7326 | 1.5169 |  | 98.7 | 94 | 98.5
bensulide-d14 | 80 | 80.1839 | 79.2763 | 81.9068 | 79.3678 | 79.8656 | 80.1201 | 100.1501 | 1.3296 |  | 97.8 | 98.8 | 97.5

## Table 5
Table 5. Recovery and precision of final method analytes obtained in multi-laboratory verification studies
% Mean Recovery (% RSD)
| Lab 1a |  | Lab 2b
Analyte | DI water (n=4) | Tap water (n=4) | DI water (n=7) | Tap water (n=7)
3-hydroxycarbofuran | 101 (11) | 104 (3.9) | 107 (1.6) | 93.1 (4.8)
fenamiphos sulfoxide | 98.8 (4.9) | 107 (3.8) | 104 (1.3) | 101 (3.5)
fenamiphos sulfone | 93.3 (5.8) | 99.4 (7.3) | 100 (0.8) | 102 (0.9)
fenamiphos | 102 (4.5) | 108 (2.6) | 101 (2.0) | 104 (2.8)
tebuconazole | 113 (4.9) | 112 (6.6) | 87.7 (4.2) | 101 (6.5)
tebufenozide | 116 (8.4) | 122 (7.0) | 97.0 (1.4) | 100 (5.1)
bensulide | 99.6 (2.6) | 104 (3.3) | 102 (2.7) | 101 (1.2)
methomyl-13C2,15N | 95.6 (13) | 96.8 (11) | 99.2 (2.2) | 102 (3.1)
carbofuran-13C6 | 102 (4.4) | 96.2 (3.5) | 102 (2.1) | 103 (3.4)
bensulide-d14 | 88.5 (4.2) | 82.6 (4.2) | 104 (2.5) | 105 (3.6)
a Fortified at 10-25 ng/L; treated drinking water from a ground water source
b Fortified at 2.5-6.5 ng/L; treated drinking water from a surface water source
Metadata for Table 5
Lab 2
LFSM (tap water)
Sample |  |  | 3-Hydroxycarbofuran | Fenamiphos sulfoxide | Fenamiphos sulfone | Fenamiphos | Tebuconazole | Tebufenozide | Bensulide | Methomyl-13C2-15N | Carbofuran-13C6 | Bensulide-d14
Name | Type | Level | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy
LFSM 1 | QC | 9 | 91.1886 | 101.6447 | 101.3126 | 101.9882 | 109.7798 | 108.3641 | 100.9512 | 102.1853 | 106.0134 | 97.1047
LFSM 2 | QC | 9 | 93.8414 | 96.4232 | 100.212 | 103.7583 | 97.3947 | 101.6174 | 99.7717 | 106.6727 | 105.7117 | 103.0067
LFSM 3 | QC | 9 | 102.1573 | 96.5766 | 103.2466 | 99.1323 | 107.3343 | 106.1741 | 102.6791 | 104.9058 | 108.1835 | 107.7867
LFSM 6 | QC | 9 | 87.521 | 106.0394 | 101.7663 | 106.446 | 90.9795 | 96.4924 | 102.5168 | 100.2408 | 100.6988 | 108.4499
LFSM 7 | QC | 9 | 93.1805 | 98.6064 | 101.5552 | 107.1477 | 99.5433 | 97.7087 | 100.864 | 99.5048 | 98.4623 | 105.7395
LFSM 8 | QC | 9 | 91.3568 | 102.9297 | 101.7071 | 105.0757 | 96.936 | 96.7986 | 99.4471 | 99.5599 | 101.5762 | 104.5506
LFSM 9 | QC | 9 | 92.6194 | 101.0904 | 101.6048 | 101.6991 | 103.5641 | 95.8348 | 100.9868 | 98.2499 | 101.0069 | 106.2385
|  | Expected Conc. (ng/L) | 6.25 | 6.25 | 6.25 | 2.5 | 2.5 | 2.5 | 6.25 | 12.8 | 5.12 | 12.8
|  | Average | 93.1236 | 100.4729 | 101.6292 | 103.6068 | 100.7902 | 100.4272 | 101.031 | 101.617 | 103.0933 | 104.6967
|  | Std Dev | 4.484 | 3.5081 | 0.89028 | 2.8595 | 6.5259 | 5.0756 | 1.2284 | 3.1258 | 3.5385 | 3.8213
|  | %RSD | 4.815% | 3.492% | 0.876% | 2.76% | 6.475% | 5.054% | 1.216% | 3.076% | 3.432% | 3.65%
LFB
Sample |  |  | 3-Hydroxycarbofuran | Fenamiphos sulfoxide | Fenamiphos sulfone | Fenamiphos | Tebuconazole | Tebufenozide | Bensulide | Methomyl-13C2-15N | Carbofuran-13C6 | Bensulide-d14
Name | Type | Level | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy | Accuracy
LFB 2 | QC | 9 | 107.6384 | 103.9098 | 101.311 | 101.6786 | 93.9155 | 99.0086 | 104.4644 | 100.0019 | 102.1612 | 102.3938
LFB 3 | QC | 9 | 104.7755 | 104.08 | 99.8623 | 97.8618 | 84.7325 | 94.8063 | 98.4665 | 101.0228 | 104.5772 | 107.8601
LFB 4 | QC | 9 | 104.8519 | 103.157 | 98.9905 | 98.9592 | 86.1445 | 95.7922 | 97.8641 | 101.7441 | 104.3001 | 107.1361
LFB 5 | QC | 9 | 108.1335 | 107.0171 | 101.046 | 103.964 | 90.2268 | 96.9648 | 103.2042 | 95.2178 | 98.4535 | 100.8713
LFB 6 | QC | 9 | 108.7244 | 103.9021 | 100.4437 | 100.5552 | 84.6687 | 96.6861 | 102.1181 | 99.0927 | 101.7315 | 105.1515
LFB 7 | QC | 9 | 107.927 | 103.152 | 100.1179 | 101.4375 | 84.367 | 96.2328 | 99.9574 | 99.0618 | 102.0963 | 105.2117
LFB 8 | QC | 9 | 108.6597 | 105.3345 | 101.1005 | 100.2293 | 89.8746 | 97.2628 | 104.3474 | 98.0913 | 100.7264 | 102.219
|  | Expected Conc. (ng/L) | 6.25 | 6.25 | 6.25 | 2.5 | 2.5 | 2.5 | 6.25 | 12.8 | 5.12 | 12.8
|  | Average | 107.2443 | 104.3647 | 100.4103 | 100.6694 | 87.7042 | 96.6791 | 101.4889 | 99.176 | 102.0066 | 104.4062
|  | Std Dev | 1.7042 | 1.3788 | 0.8261 | 1.9781 | 3.6803 | 1.3129 | 2.7381 | 2.1429 | 2.0912 | 2.6431
|  | %RSD | 1.589% | 1.321% | 0.8227% | 1.965% | 4.196% | 1.358% | 2.698% | 2.161% | 2.05% | 2.532%
Lab 1 - LFBs (DI water)
Lab 1 - LFSMs (tap water)

## Table 6
Table 6. Comparison of multi-laboratory LCMRLs to the HRLs of method analytes
Analyte | HRLa | LCMRL ng/L
| ng/L | EPA Lab | Lab 1 | Lab 2
3-hydroxycarbofuran | 420 | 1.7 | 1.9 | 2.1
bensulide | 35000 | 1.2 | 1.9 | 0.33
fenamiphos | 700 | 1.4 | 4.3 | 0.25
fenamiphos sulfone | NAb | 0.27 | 0.66 | 0.072
fenamiphos sulfoxide | NA | 1.3 | 3.2 | 0.21
tebuconazole | 210000 | 0.47 | 0.83 | 0.13
tebufenozide | 126000 | 1.2 | 0.82 | 0.73
a HRLs obtained from the Office of Water’s Contaminant Information Sheets for the Final CCL 3 Chemicals.16
b Not available.
Metadata for Table 6
EPA LCMRL Dataset
Analyte | 3-hydroxycarbofuran | fenamiphos sulfoxide | fenamiphos sulfone | fenamiphos | tebuconazole | tebufenozide | bensulide
Spike concentration, ng/L | 1 | 1 | 1 | 0.4 | 0.4 | 0.4 | 1
Rep 1 | 1.2981 | 0.9219 | 1.3255 | 0.40088 | 0.55008 | 0.54263 | 1.0149
Rep 2 | 1.1843 | 1.1526 | 1.1248 | 0.44007 | 0.50467 | 0.40091 | 1.1172
Rep 3 | 0.9536 | 1.1635 | 1.0964 | 0.48127 | 0.51811 | 0.47245 | 1.205
Rep 4 | 1.5889 | 1.0379 | 1.0413 | 0.374 | 0.45233 | 0.39799 | 0.99512
Spike concentration, ng/L | 2 | 2 | 2 | 0.8 | 0.8 | 0.8 | 2
Rep 1 | 1.8639 | 2.2343 | 2.2129 | 0.81036 | 0.83723 | 0.7154 | 2.2302
Rep 2 | 2.2369 | 2.0576 | 2.1723 | 0.80358 | 0.94724 | 0.85052 | 1.9107
Rep 3 | 1.7229 | 2.2152 | 2.1611 | 0.84244 | 0.73838 | 0.82166 | 1.9346
Rep 4 | 1.861 | 2.1069 | 2.137 | 0.8366 | 0.94984 | 0.77486 | 1.8829
Spike concentration, ng/L | 5 | 5 | 5 | 2 | 2 | 2 | 5
Rep 1 | 4.3889 | 5.2001 | 5.3663 | 2.2509 | 2.0507 | 2.3355 | 4.7198
Rep 2 | 5.0134 | 4.9765 | 5.239 | 2.1181 | 1.8969 | 2.0066 | 5.172
Rep 3 | 4.7025 | 4.2582 | 4.829 | 1.7291 | 1.7161 | 2.0332 | 4.0135
Rep 4 | 4.5408 | 5.1428 | 5.1594 | 2.1549 | 1.9347 | 2.0108 | 5.2392
| 4.2008 | 4.2099 | 4.8098 | 1.8424 | 1.7792 | 1.7549 | 4.2427
Spike concentration, ng/L | 7.5 | 7.5 | 7.5 | 3 | 3 | 3 | 7.5
Rep 1 | 7.8074 | 7.5984 | 7.6646 | 3.1384 | 3.4354 | 3.0068 | 7.5808
Rep 2 | 6.8502 | 7.193 | 7.5503 | 2.8119 | 2.9054 | 2.8748 | 7.5498
Rep 3 | 6.4288 | 7.6178 | 7.5428 | 3.1613 | 2.7334 | 3.0572 | 7.7478
Rep 4 | 7.7356 | 7.4367 | 7.6826 | 3.2398 | 3.1587 | 2.9819 | 8.2626
Spike concentration, ng/L | 10 | 10 | 10 | 4 | 4 | 4 | 10
Rep 1 | 9.3155 | 9.2039 | 9.944 | 3.7644 | 3.5558 | 3.6384 | 10.0735
Rep 2 | 10.133 | 9.525 | 10.0263 | 4.11 | 4.8251 | 4.2233 | 10.1619
Rep 3 | 8.674 | 10.3233 | 10.2843 | 4.303 | 4.0409 | 4.0407 | 10.5602
Rep 4 | 10.2118 | 9.902 | 9.7616 | 3.9285 | 3.6032 | 3.7817 | 9.6516
Spike concentration, ng/L | 20 | 20 | 20 | 8 | 8 | 8 | 20
Rep 1 | 20.2432 | 22.2864 | 21.4921 | 8.3264 | 8.2358 | 8.5817 | 20.1879
Rep 2 | 20.2544 | 21.3992 | 20.915 | 8.6563 | 9.7242 | 8.9132 | 22.3728
Rep 3 | 18.3502 | 20.5054 | 21.1612 | 8.3708 | 8.2099 | 7.9721 | 19.0349
Rep 4 | 18.1936 | 21.4271 | 19.8846 | 8.9677 | 8.7934 | 8.4836 | 21.0501
Spike concentration, ng/L | 15 | 15 | 15 | 6 | 6 | 6 | 15
Rep 1 | 14.0096 | 13.9157 | 15.2545 | 5.9009 | 5.8959 | 5.7926 | 15.5442
Rep 2 | 14.977 | 15.82 | 16.1369 | 6.6851 | 6.731 | 6.6858 | 17.7462
Rep 3 | 15.8875 | 16.6744 | 17.337 | 6.5683 | 6.3453 | 6.5215 | 16.6273
Rep 4 | 14.2723 | 14.4908 | 15.0675 | 7.5653 | 6.1156 | 6.6204 | 13.9983
Lab 1 LCMRL Dataset
Sample |  |  | 3-Hydroxycarbofuran Results |  |  |  | Fenamiphos sulfoxide Results |  |  |  | Fenamiphos sulfone Results |  |  |  | Fenamiphos Results |  |  |  | Tebuconazole Results |  |  |  | Tebufenozide Results |  |  |  | Bensulide Results
Name | Type | Level | RT | Area | Final Conc. | Accuracy | RT | Area | Final Conc. | Accuracy | RT | Area | Final Conc. | Accuracy | RT | Area | Final Conc. | Accuracy | RT | Area | Final Conc. | Accuracy | RT | Area | Final Conc. | Accuracy | RT | Area | Final Conc. | Accuracy
started 140122
H2O Blank | DoubleBlank | [9 empty] | 8.1737 | 83.3746 | 28.1962 |  | 9.0575 | 46.517 | 16.4129 |  | 9.2568 | 230.0968 | 7.052 |  | 9.5219 | 749.7497 | 7.5674 |  | 9.8115 | 247.7244 | 13.624
Cal 3 | Cal | 3 |  |  |  |  | 7.6872 | 61.6611 | 0.04639 | 95.0029 | 8.2211 | 93.8392 | 0.042099 | 86.2161 | 8.9514 | 0 | 0 | 0 | 9.2568 | 92.2771 | 0.076569 | 392.056 | 9.5773 | 86.0811 | 0.01613 | 82.5891 | 9.8344 | 50.0988 | 0.040599 | 83.1439
Cal 3 | Cal | 3 | 7.4651 | 0 | 0 | 0 | 7.675 | 61.4511 | 0.048921 | 100.1855 | 8.2327 | 114.1973 | 0.057744 | 118.2557 | 8.8317 | 0 | 0 | 0 | 9.2802 | 57.551 | 0.067039 | 343.2593 | 9.5346 | 75.6777 | 0.017576 | 89.9964 | 9.8059 | 56.0298 | 0.058344 | 119.4831
Cal 4 | Cal | 4 |  |  |  |  | 7.6911 | 265.1356 | 0.21278 | 108.9471 | 8.2211 | 308.6962 | 0.17926 | 91.7822 | 9.059 | 103.0416 | 0.075042 | 96.0474 | 9.2308 | 61.6609 | 0.073004 | 93.4387 | 9.5503 | 407.1963 | 0.095008 | 121.6018 | 9.8039 | 161.7565 | 0.19286 | 98.745
Cal 5 | Cal | 5 | 7.43 | 0 | 0 | 0 | 7.6955 | 1972.0327 | 0.75623 | 96.7909 | 8.1972 | 2924.8131 | 0.78992 | 101.1038 | 9.0806 | 1156.9718 | 0.328 | 104.9613 | 9.2685 | 1059.4576 | 0.33591 | 107.4897 | 9.5366 | 3289.8281 | 0.33348 | 106.7128 | 9.8089 | 1462.2496 | 0.77941 | 99.7585
Cal 6 | Cal | 6 | 7.4074 | 412.3673 | 3.5351 | 113.1227 | 7.69 | 7934.0651 | 3.0813 | 98.6012 | 8.1986 | 11880.9077 | 3.238 | 103.6175 | 9.0723 | 4465.2934 | 1.2375 | 98.9985 | 9.2568 | 4258.4727 | 1.2417 | 99.3321 | 9.5277 | 12550.1741 | 1.2477 | 99.8146 | 9.8033 | 5820.3917 | 3.0737 | 98.36
Cal 7 | Cal | 7 | 7.4048 | 1643.836 | 12.0983 | 96.7865 | 7.6895 | 33287.9476 | 12.5614 | 100.4909 | 8.2027 | 45287.1855 | 12.3725 | 98.9797 | 9.0731 | 17844.8633 | 4.9995 | 99.9894 | 9.2568 | 16947.5442 | 4.986 | 99.7196 | 9.5311 | 50207.3214 | 4.9621 | 99.2416 | 9.8052 | 23639.8098 | 12.5668 | 100.5343
Cal 8 | Cal | 8 | 7.3897 | 6753.5722 | 50.0704 | 100.1409 | 7.6878 | 142644.8087 | 49.9907 | 99.9814 | 8.1993 | 205285.1283 | 50.0208 | 100.0416 | 9.0734 | 74847.6877 | 20.0006 | 100.0032 | 9.2685 | 67518.9965 | 20.0042 | 100.0208 | 9.5308 | 229202.8237 | 20.008 | 100.04 | 9.8087 | 102179.8208 | 49.9874 | 99.9748
LRB 1 | Blank |  | 7.5828 | 0 | 0 |  | 7.6745 | 19.0963 | 0.0083904 |  | 8.1973 | 76.0624 | 0.012495 |  | 9.0955 | 55.804 | 0.028219 |  | 9.2451 | 175.4561 | 0.083556 |  | 9.5346 | 197.821 | 0.021909 |  | 9.7949 | 92.7576 | 0.046368
CC 6 | CC | 6 | 7.4046 | 424.4429 | 4.1461 | 132.6766 | 7.6858 | 7764.8629 | 3.5557 | 113.7828 | 8.2012 | 11162.028 | 3.1752 | 101.6063 | 9.0689 | 4181.9616 | 1.2099 | 96.7883 | 9.2685 | 3869.1057 | 1.1575 | 92.6039 | 9.5265 | 12362.5088 | 1.2589 | 100.7102 | 9.8055 | 5464.9512 | 2.9566 | 94.6102
LFB 1 | QC | 9 | 7.385 | 139.9399 | 1.8755 | 119.9924 | 7.6863 | 4034.4808 | 1.9194 | 122.8044 | 8.2029 | 5495.3864 | 1.6336 | 104.519 | 9.078 | 2128.9088 | 0.64904 | 103.846 | 9.2568 | 1918.8974 | 0.61541 | 98.4653 | 9.5277 | 5958.3353 | 0.63787 | 102.0598 | 9.8127 | 2665.5389 | 1.5092 | 96.5608
LFB 2 | QC | 9 | 7.3716 | 185.4905 | 2.0828 | 133.2554 | 7.6867 | 3885.2625 | 1.6437 | 105.1639 | 8.1987 | 5442.7255 | 1.5359 | 98.2639 | 9.0793 | 2109.2323 | 0.61123 | 97.7964 | 9.2568 | 1928.0496 | 0.57321 | 91.7132 | 9.5316 | 5837.8704 | 0.57739 | 92.3817 | 9.8039 | 2676.0198 | 1.3988 | 89.4953
LFB 3 | QC | 9 | 7.3833 | 203.1675 | 2.2093 | 141.3496 | 7.6842 | 3709.5851 | 1.565 | 100.1307 | 8.1974 | 5330.688 | 1.5229 | 97.4362 | 9.0705 | 2068.9701 | 0.6071 | 97.1365 | 9.2568 | 2020.7304 | 0.63344 | 101.3509 | 9.5236 | 5947.4372 | 0.62339 | 99.743 | 9.8093 | 2659.4754 | 1.4741 | 94.3109
LFB 4 | QC | 9 | 7.3833 | 189.0033 | 2.0277 | 129.7338 | 7.687 | 3859.3287 | 1.5392 | 98.4769 | 8.1953 | 5524.9512 | 1.555 | 99.4895 | 9.0725 | 2088.1816 | 0.60371 | 96.5936 | 9.2451 | 2143.5026 | 0.64523 | 103.2375 | 9.5272 | 6196.6215 | 0.6243 | 99.8886 | 9.8062 | 2799.3582 | 1.4914 | 95.4206
LFB 5 | QC | 9 | 7.428 | 220.6597 | 2.1888 | 140.0391 | 7.6845 | 3981.8795 | 1.526 | 97.6318 | 8.1985 | 5607.9566 | 1.5223 | 97.3955 | 9.0731 | 2030.5802 | 0.56702 | 90.7229 | 9.2685 | 1999.1816 | 0.60615 | 96.9843 | 9.5254 | 5914.8427 | 0.59839 | 95.7421 | 9.8064 | 2841.5301 | 1.5198 | 97.2359
LFB 6 | QC | 9 | 7.3616 | 190.1892 | 2.0336 | 130.1083 | 7.6793 | 3797.7923 | 1.5121 | 96.7417 | 8.1886 | 5458.7002 | 1.5156 | 96.9662 | 9.065 | 1970.3262 | 0.56285 | 90.0555 | 9.2568 | 1980.4062 | 0.61789 | 98.862 | 9.5252 | 6216.3206 | 0.6475 | 103.6003 | 9.8025 | 2648.3875 | 1.4591 | 93.3515
LFB 7 | QC | 9 | 7.4067 | 193.5845 | 2.015 | 128.9156 | 7.6787 | 4204.5496 | 1.6207 | 103.6919 | 8.1982 | 5957.2372 | 1.6009 | 102.425 | 9.0724 | 2135.7738 | 0.58984 | 94.3749 | 9.2568 | 2177.5652 | 0.61816 | 98.9059 | 9.5268 | 6334.6978 | 0.60065 | 96.1047 | 9.8043 | 2888.6864 | 1.448 | 92.6452
LFB 8 | QC | 9 | 7.3482 | 161.5698 | 1.8069 | 115.6071 | 7.6828 | 4133.8652 | 1.6063 | 102.7694 | 8.1911 | 5928.7151 | 1.6068 | 102.7999 | 9.0704 | 2239.9957 | 0.62312 | 99.6997 | 9.2568 | 2335.8714 | 0.6868 | 109.888 | 9.5234 | 6179.5054 | 0.61009 | 97.6148 | 9.8002 | 2846.0858 | 1.4857 | 95.0535
LRB 2 | Blank |  | 7.5594 | 0 | 0 |  | 7.827 | 0 | 0 |  | 7.9523 | 0 | 0 | [5 empty] | 9.2308 | 43.1979 | 0.045526 |  | 9.923 | 0 | 0.0017783 |  | 9.8128 | 0 | 0
LRB 3 | Blank | [17 empty] | 9.2074 | 57.2484 | 0.049376 |  | 9.923 | 0 | 0.0017783
LRB 4 | Blank | [9 empty] | 8.5705 | 8.15 | 0 |  | 8.8599 | 29.7289 | 0.020412 |  | 9.3738 | 0 | 0 |  | 9.9347 | 0 | 0.0017783 |  | 9.9067 | 0 | 0
Cal 4 | CC | 4 |  |  |  |  | 7.6865 | 451.6391 | 0.17525 | 89.7282 | 8.1895 | 700.5822 | 0.18086 | 92.6021 | 9.0929 | 245.3935 | 0.0787 | 100.73 | 9.2451 | 253.4654 | 0.10276 | 131.522 | 9.5165 | 811.8415 | 0.080688 | 103.2735 | 9.7935 | 419.8606 | 0.21287 | 108.9929
Cal 9 | CC | 10 | 7.3754 | 14407.7515 | 92.0646 | 92.0646 | 7.6805 | 295008.7335 | 82.6067 | 82.6067 | 8.1921 | 434581.1702 | 94.9797 | 94.9797 | 9.0683 | 160147.5248 | 40.9814 | 102.4534 | 9.2568 | 138342.3096 | 49.2109 | 123.0271 | 9.5232 | 476465.6636 | 40.1369 | 100.3422 | 9.803 | 217366.7937 | 106.9571 | 106.9571
7.5-1 | QC | 11 | 7.3859 | 8078.5386 | 22.0129 | 88.0515 | 7.6861 | 181994.9204 | 24.9101 | 99.6402 | 8.1984 | 284637.5953 | 27.1789 | 108.7155 | 9.0722 | 105410.7114 | 10.6083 | 106.0828 | 9.2568 | 118695.1315 | 13.0907 | 130.9066 | 9.5286 | 434719.8312 | 14.883 | 148.8301 | 9.8099 | 141173.1449 | 26.9974 | 107.9894
7.5-2 | QC | 11 | 7.3899 | 8025.4606 | 21.7285 | 86.9142 | 7.6847 | 182305.1558 | 24.7884 | 99.1537 | 8.1964 | 271633.3602 | 26.3453 | 105.381 | 9.0709 | 98459.7453 | 10.0545 | 100.5454 | 9.2568 | 101256.7761 | 10.9971 | 109.9713 | 9.5292 | 400891.4605 | 13.7017 | 137.0168 | 9.8096 | 136649.6211 | 25.9577 | 103.8309
7.5-3 | QC | 11 | 7.3869 | 7539.9016 | 21.6195 | 86.478 | 7.687 | 176962.6088 | 25.4027 | 101.6108 | 8.199 | 264141.8289 | 26.8872 | 107.5487 | 9.0729 | 101018.9152 | 10.8237 | 108.2367 | 9.2568 | 100928.6657 | 11.1111 | 111.1113 | 9.5299 | 392699.4449 | 13.606 | 136.0601 | 9.8064 | 137504.7952 | 26.4529 | 105.8117
7.5-4 | QC | 11 | 7.3698 | 7940.3313 | 24.447 | 97.7878 | 7.6853 | 177858.2956 | 27.3606 | 109.4425 | 8.1944 | 262959.3631 | 26.9221 | 107.6885 | 9.0691 | 96041.873 | 10.362 | 103.6198 | 9.2568 | 94265.0185 | 10.7151 | 107.1514 | 9.5258 | 374723.3145 | 13.4363 | 134.3629 | 9.8057 | 133309.8699 | 26.5162 | 106.0647
7-1 | QC | 12 | 7.3924 | 3850.5725 | 11.707 | 93.656 | 7.683 | 81652.9973 | 12.6885 | 101.5079 | 8.196 | 120989.6507 | 13.1735 | 105.3877 | 9.0704 | 43698.5823 | 4.8915 | 97.8307 | 9.2568 | 44485.248 | 5.0367 | 100.7344 | 9.5287 | 180995.6101 | 6.815 | 136.2999 | 9.8063 | 61784.7396 | 12.6376 | 101.101
CC 7 | CC | 7 | 7.4056 | 3836.0286 | 13.4585 | 107.6679 | 7.6874 | 80515.041 | 14.4608 | 115.6866 | 8.1979 | 117991.2088 | 13.09 | 104.7197 | 9.0716 | 41787.3101 | 4.7666 | 95.3315 | 9.2568 | 39098.1585 | 4.5346 | 90.6918 | 9.5305 | 154068.8761 | 5.9761 | 119.5226 | 9.8054 | 55744.5482 | 11.7041 | 93.633
7-2 | QC | 12 | 7.3965 | 3600.1734 | 11.6458 | 93.1663 | 7.6879 | 79366.9125 | 13.0907 | 104.7257 | 8.1966 | 114283.8509 | 13.2739 | 106.1908 | 9.0717 | 41412.6616 | 4.9457 | 98.9144 | 9.2568 | 43287.3526 | 4.9923 | 99.8455 | 9.5284 | 158982.4752 | 6.1209 | 122.4179 | 9.8049 | 58159.6868 | 12.1249 | 96.9992
7-3 | QC | 12 | 7.3981 | 3386.8831 | 10.7277 | 85.8216 | 7.6889 | 77056.6089 | 12.4054 | 99.2432 | 8.1973 | 111325.0479 | 13.4772 | 107.8178 | 9.073 | 39728.0508 | 4.948 | 98.9607 | 9.2568 | 42152.6172 | 5.1803 | 103.6055 | 9.5286 | 161932.6333 | 6.6225 | 132.4494 | 9.8048 | 58614.5054 | 13.0039 | 104.0312
7-4 | QC | 12 | 7.3741 | 3664.4764 | 10.8501 | 86.8009 | 7.6896 | 74944.1664 | 11.3485 | 90.7879 | 8.1981 | 110248.4275 | 12.7257 | 101.8054 | 9.073 | 40102.9458 | 4.7548 | 95.0956 | 9.2568 | 42787.3703 | 5.0728 | 101.456 | 9.5317 | 163332.3647 | 6.4517 | 129.0345 | 9.8062 | 60367.8926 | 12.925 | 103.3997
6.5-1 | QC | 13 | 7.3957 | 2005.2928 | 6.1085 | 97.7355 | 7.6905 | 38185.2611 | 5.7582 | 92.1307 | 8.1957 | 54473.661 | 6.1773 | 98.8363 | 9.0738 | 20045.7019 | 2.3119 | 92.4753 | 9.2568 | 22088.8932 | 2.5852 | 103.4097 | 9.5279 | 79180.5281 | 3.1531 | 126.1232 | 9.8031 | 29012.2361 | 6.1905 | 99.048
6.5-2 | QC | 13 | 7.4009 | 1907.4383 | 5.8158 | 93.0528 | 7.6892 | 38016.2172 | 5.7002 | 91.2036 | 8.2004 | 54993.8597 | 6.152 | 98.4321 | 9.0733 | 19982.0786 | 2.2737 | 90.9461 | 9.2685 | 22442.9813 | 2.5061 | 100.2423 | 9.5285 | 79964.3011 | 3.0396 | 121.5856 | 9.8098 | 29960.5375 | 6.0991 | 97.5853
CC 6 | CC | 6 | 7.395 | 803.0234 | 2.9754 | 95.2118 | 7.6911 | 17495.3089 | 2.8018 | 89.6577 | 8.2005 | 24381.5247 | 2.9335 | 93.8725 | 9.0816 | 8989.4258 | 1.1008 | 88.0678 | 9.2568 | 9700.2116 | 1.1972 | 95.7774 | 9.5341 | 34873.9759 | 1.4643 | 117.1431 | 9.8082 | 12599.7193 | 2.8146 | 90.0678
6.5-3 | QC | 13 | 7.3937 | 2018.2083 | 6.306 | 100.8962 | 7.6929 | 38974.1261 | 6.0454 | 96.7268 | 8.2022 | 55636.503 | 6.4786 | 103.6581 | 9.0797 | 20385.8048 | 2.4151 | 96.6044 | 9.2685 | 21414.0191 | 2.5213 | 100.8521 | 9.5396 | 79312.1727 | 3.1766 | 127.0633 | 9.8156 | 29362.0748 | 6.3016 | 100.8249
6.5-4 | QC | 13 | 7.3775 | 1707.0076 | 5.9945 | 95.9113 | 7.6921 | 39248.7562 | 6.7702 | 108.3231 | 8.2007 | 55712.8203 | 6.6109 | 105.775 | 9.0773 | 20664.1737 | 2.4947 | 99.788 | 9.2685 | 20831.9099 | 2.3983 | 95.9328 | 9.5328 | 74930.367 | 2.9377 | 117.5098 | 9.8121 | 28320.5004 | 5.9437 | 95.0996
6-1 | QC | 14 | 7.4066 | 1043.6812 | 4.0099 | 128.3158 | 7.6881 | 18455.9418 | 3.3047 | 105.7512 | 8.2004 | 27404.916 | 3.1181 | 99.7802 | 9.0769 | 10027.1522 | 1.1608 | 92.8634 | 9.2685 | 10078.5251 | 1.1814 | 94.5114 | 9.533 | 38363.6984 | 1.5287 | 122.2999 | 9.8104 | 13242.662 | 2.8087 | 89.8775
6-2 | QC | 14 | 7.3992 | 1053.1341 | 4.13 | 132.1594 | 7.689 | 19030.6586 | 3.4967 | 111.8955 | 8.1985 | 27535.0021 | 3.2526 | 104.084 | 9.0737 | 9395.4621 | 1.1299 | 90.3944 | 9.2685 | 9797.8737 | 1.1573 | 92.5852 | 9.5362 | 36832.0239 | 1.4787 | 118.2955 | 9.8141 | 12926.6089 | 2.7614 | 88.3634
6-3 | QC | 14 | 7.3865 | 933.9537 | 3.6045 | 115.3431 | 7.6885 | 18681.3017 | 3.2787 | 104.9199 | 8.2026 | 26635.3349 | 3.0796 | 98.547 | 9.0722 | 9384.0119 | 1.1046 | 88.3654 | 9.2568 | 10536.7579 | 1.2673 | 101.3851 | 9.5296 | 35853.6792 | 1.4689 | 117.5142 | 9.8146 | 12095.3653 | 2.6368 | 84.3765
CC 7 | CC | 7 | 7.3676 | 4012.3265 | 12.5202 | 100.1619 | 7.6848 | 80634.7863 | 12.9117 | 103.2936 | 8.1981 | 119378.8106 | 13.2019 | 105.6152 | 9.0728 | 41750.5515 | 4.749 | 94.9797 | 9.2568 | 43273.6332 | 5.2787 | 105.5747 | 9.5328 | 170526.617 | 6.91 | 138.2002 | 9.8079 | 56911.0581 | 12.5353 | 100.2822
6-4 | QC | 14 | 7.4128 | 1003.7458 | 3.5617 | 113.9751 | 7.6882 | 18595.0485 | 2.9961 | 95.8767 | 8.1997 | 27945.4665 | 3.1732 | 101.5409 | 9.0737 | 9758.0921 | 1.1279 | 90.2319 | 9.2568 | 11048.334 | 1.3203 | 105.6238 | 9.5283 | 38249.3008 | 1.5575 | 124.6002 | 9.8089 | 13211.4896 | 2.8637 | 91.6392
3.5-1 | QC | 18 | 7.6316 | 0 | 0 | 0 | 7.6516 | 407.8234 | 0.070802 | 72.4981 | 8.1775 | 832.2476 | 0.095314 | 97.598 | 9.0657 | 230.0379 | 0.041096 | 105.2131 | 9.2308 | 610.6112 | 0.10905 | 279.1833 | 9.511 | 1198.9097 | 0.054524 | 139.5899 | 9.7814 | 430.8047 | 0.096868 | 99.1886
3.5-2 | QC | 18 | 7.5965 | 0 | 0 | 0 | 7.6804 | 439.9325 | 0.074897 | 76.6911 | 8.186 | 967.3461 | 0.11051 | 113.1534 | 9.0578 | 243.8374 | 0.042376 | 108.4905 | 9.2074 | 491.1202 | 0.093541 | 239.4801 | 9.5658 | 1042.706 | 0.047176 | 120.7772 | 9.8035 | 439.861 | 0.097912 | 100.2581
CC 6 | CC | 6 | 7.3989 | 884.5213 | 3.3309 | 106.5904 | 7.6732 | 17848.4179 | 2.9992 | 95.9745 | 8.185 | 27626.4319 | 3.4388 | 110.0432 | 9.0538 | 8454.9809 | 1.0727 | 85.8131 | 9.2451 | 13145.2587 | 1.7219 | 137.7512 | 9.5129 | 36302.812 | 1.6269 | 130.1536 | 9.7948 | 11842.4856 | 2.8264 | 90.4457
3.5-3 | QC | 18 |  |  |  |  | 7.6634 | 389.9184 | 0.064523 | 66.0691 | 8.169 | 869.7373 | 0.09584 | 98.1369 | 9.0578 | 159.2355 | 0.031453 | 80.5247 | 9.2568 | 507.3331 | 0.094333 | 241.507 | 9.5052 | 996.6821 | 0.044336 | 113.5064 | 9.7919 | 492.3479 | 0.10784 | 110.4288
3.5-4 | QC | 18 | 7.6316 | 0 | 0 | 0 | 7.678 | 557.2914 | 0.091621 | 93.8165 | 8.189 | 913.322 | 0.10314 | 105.6077 | 9.0711 | 167.8512 | 0.032864 | 84.1364 | 9.2451 | 392.4027 | 0.081852 | 209.5554 | 9.5346 | 983.5122 | 0.045014 | 115.2425 | 9.7808 | 332.2 | 0.073781 | 75.5489
CC 7 | CC | 7 | 7.349 | 4035.9157 | 11.5619 | 92.4949 | 7.669 | 79253.3069 | 11.6528 | 93.2222 | 8.1819 | 125947.0925 | 13.531 | 108.2482 | 9.0566 | 43254.7424 | 4.7837 | 95.6738 | 9.2451 | 60220.5243 | 7.2157 | 144.3143 | 9.5118 | 175918.4484 | 6.9607 | 139.214 | 9.792 | 55997.5472 | 12.0527 | 96.4217
started 140130
Cal 3 | Cal | 3 | 7.4301 | 218.0958 | 0.039724 | 81.3515 | 7.7079 | 236.1267 | 0.050859 | 104.155 | 8.1948 | 542.3816 | 0.054066 | 110.7227 | 9.0581 | 131.8688 | 0.021702 | 111.1206 | 9.2454 | 447.3778 | 0.02087 | 106.8603 | 9.5114 | 933.6898 | 0.024871 | 127.35 | 9.7896 | 335.1873 | 0.051243 | 104.9413
Cal 4 | Cal | 4 | 7.3483 | 1059.978 | 0.23566 | 120.6613 | 7.6643 | 1183.88 | 0.18767 | 96.0875 | 8.1858 | 2260.0603 | 0.18924 | 96.8913 | 9.0552 | 703.2811 | 0.074103 | 94.846 | 9.2454 | 1398.3401 | 0.081823 | 104.7262 | 9.5167 | 3078.0294 | 0.072605 | 92.929 | 9.7897 | 1135.7072 | 0.20017 | 102.4873
Cal 5 | Cal | 5 | 7.3537 | 3295.6349 | 0.75945 | 97.2096 | 7.6747 | 5332.9389 | 0.7908 | 101.2221 | 8.1857 | 9146.7032 | 0.73277 | 93.794 | 9.061 | 3146.4134 | 0.29766 | 95.2506 | 9.2454 | 4692.3357 | 0.29112 | 93.158 | 9.5147 | 11732.0504 | 0.26344 | 84.2996 | 9.8003 | 3987.3924 | 0.72445 | 92.7302
Cal 6 | Cal | 6 | 7.3557 | 13638.501 | 3.1625 | 101.2009 | 7.671 | 21099.2605 | 3.0628 | 98.0091 | 8.1867 | 38541.3797 | 3.0432 | 97.3818 | 9.059 | 13283.1588 | 1.2218 | 97.7401 | 9.2454 | 18146.6652 | 1.1601 | 92.805 | 9.5161 | 51476.3572 | 1.1502 | 92.0122 | 9.7997 | 16748.7255 | 3.0982 | 99.1411
Cal 7 | Cal | 7 | 7.3557 | 55424.0346 | 12.4441 | 99.553 | 7.674 | 89996.8858 | 12.5685 | 100.5482 | 8.1869 | 165500.389 | 12.6583 | 101.266 | 9.0582 | 56749.7116 | 5.0546 | 101.0921 | 9.2454 | 73578.441 | 5.1279 | 102.5577 | 9.5149 | 216355.0538 | 5.1774 | 103.5477 | 9.7979 | 63201.6509 | 12.5915 | 100.7324
Cal 8 | Cal | 8 | 7.3575 | 248155.7939 | 50.0096 | 100.0193 | 7.6727 | 397444.8613 | 49.989 | 99.9779 | 8.187 | 732278.602 | 49.9714 | 99.9428 | 9.0582 | 249074.4852 | 19.9899 | 99.9497 | 9.2454 | 320464.8837 | 19.9774 | 99.8872 | 9.515 | 962060.9455 | 19.9692 | 99.846 | 9.7998 | 290981.2851 | 49.9835 | 99.9669
LRB 1 | Blank |  | 7.5704 | 0 | 0 |  | 7.6396 | 361.2966 | 0.068101 |  | 8.1694 | 724.5889 | 0.066146 |  | 9.0452 | 247.7756 | 0.031818 |  | 9.2454 | 658.512 | 0.034015 |  | 9.5181 | 1246.7587 | 0.03141 |  | 9.8255 | 354.6595 | 0.052774
CC 6 | CC | 6 | 7.3595 | 13950.3542 | 3.3301 | 106.5617 | 7.6734 | 20965.9744 | 3.133 | 100.2565 | 8.1871 | 37738.9015 | 2.955 | 94.5612 | 9.0576 | 13233.6167 | 1.2069 | 96.5527 | 9.2454 | 16745.1899 | 1.046 | 83.677 | 9.5155 | 49306.7799 | 1.0776 | 86.204 | 9.7963 | 16626.5266 | 3.0065 | 96.2073
LFB 1 | QC | 9 | 7.3525 | 28695.39 | 6.9624 | 111.3985 | 7.6698 | 44919.2361 | 6.7931 | 108.6901 | 8.186 | 82530.452 | 6.5564 | 104.903 | 9.0557 | 28646.2663 | 2.6471 | 105.8848 | 9.2454 | 35404.3589 | 2.3276 | 93.102 | 9.5159 | 112896.0493 | 2.5677 | 102.7092 | 9.798 | 34234.9151 | 6.4707 | 103.5317
LFB 2 | QC | 9 | 7.3533 | 28944.5308 | 6.7274 | 107.6384 | 7.6714 | 44814.1928 | 6.4944 | 103.9098 | 8.1859 | 81948.3502 | 6.3319 | 101.311 | 9.0565 | 28279.6043 | 2.542 | 101.6786 | 9.2454 | 36212.3345 | 2.3479 | 93.9155 | 9.5136 | 110253.2525 | 2.4752 | 99.0086 | 9.7994 | 35028.9873 | 6.529 | 104.4644
LFB 3 | QC | 9 | 7.3593 | 28446.8099 | 6.5485 | 104.7755 | 7.673 | 45347.6239 | 6.505 | 104.08 | 8.1889 | 82665.553 | 6.2414 | 99.8623 | 9.0585 | 27842.6235 | 2.4465 | 97.8618 | 9.2454 | 34386.1647 | 2.1183 | 84.7325 | 9.5145 | 111111.1071 | 2.3702 | 94.8063 | 9.7991 | 34741.3666 | 6.1542 | 98.4665
LFB 4 | QC | 9 | 7.3525 | 28671.2062 | 6.5532 | 104.8519 | 7.6701 | 45257.0075 | 6.4473 | 103.157 | 8.1851 | 81714.098 | 6.1869 | 98.9905 | 9.0563 | 28085.8276 | 2.474 | 98.9592 | 9.2454 | 34728.881 | 2.1536 | 86.1445 | 9.5134 | 111536.34 | 2.3948 | 95.7922 | 9.7964 | 34293.1944 | 6.1165 | 97.8641
LFB 5 | QC | 9 | 7.3536 | 27689.2142 | 6.7583 | 108.1335 | 7.6706 | 43976.8227 | 6.6886 | 107.0171 | 8.1858 | 78763.9374 | 6.3154 | 101.046 | 9.0564 | 27877.1394 | 2.5991 | 103.964 | 9.2454 | 34260.5585 | 2.2557 | 90.2268 | 9.5137 | 106326.0634 | 2.4241 | 96.9648 | 9.7959 | 34083.768 | 6.4503 | 103.2042
CC 7 | CC | 7 | 7.3531 | 56603.7242 | 13.0938 | 104.7506 | 7.6689 | 91693.4289 | 13.1939 | 105.5514 | 8.1856 | 163699.0736 | 12.4274 | 99.4193 | 9.0575 | 49789.7774 | 4.4162 | 88.3246 | 9.2454 | 67225.8579 | 4.3195 | 86.3907 | 9.5126 | 217967.5911 | 4.8029 | 96.0581 | 9.7942 | 69763.198 | 12.7533 | 102.0262
LFB 6 | QC | 9 | 7.3505 | 28976.7708 | 6.7953 | 108.7244 | 7.6677 | 44403.4101 | 6.4939 | 103.9021 | 8.1799 | 80892.641 | 6.2777 | 100.4437 | 9.0546 | 27843.9206 | 2.5139 | 100.5552 | 9.2454 | 33497.218 | 2.1167 | 84.6687 | 9.5108 | 110512.6455 | 2.4172 | 96.6861 | 9.7942 | 35148.9504 | 6.3824 | 102.1181
LFB 7 | QC | 9 | 7.3538 | 28750.9305 | 6.7454 | 107.927 | 7.6697 | 44061.7126 | 6.447 | 103.152 | 8.1837 | 80914.7051 | 6.2574 | 100.1179 | 9.0547 | 28193.4103 | 2.5359 | 101.4375 | 9.2454 | 33396.0842 | 2.1092 | 84.367 | 9.514 | 110046.9345 | 2.4058 | 96.2328 | 9.7976 | 34411.0007 | 6.2473 | 99.9574
LFB 8 | QC | 9 | 7.3594 | 28666.5271 | 6.7912 | 108.6597 | 7.6701 | 44574.9189 | 6.5834 | 105.3345 | 8.1837 | 80626.5221 | 6.3188 | 101.1005 | 9.0569 | 27477.8729 | 2.5057 | 100.2293 | 9.2454 | 34581.5987 | 2.2469 | 89.8746 | 9.5129 | 108084.5491 | 2.4316 | 97.2628 | 9.7966 | 34929.2804 | 6.5217 | 104.3474
CC 6 | CC | 6 | 7.3524 | 13373.2484 | 3.1349 | 100.3181 | 7.6706 | 20956.7819 | 3.075 | 98.4011 | 8.1818 | 36942.9555 | 2.9032 | 92.9038 | 9.0566 | 13385.0023 | 1.2247 | 97.9744 | 9.2454 | 15753.7888 | 1.0121 | 80.9706 | 9.5121 | 49740.1325 | 1.1177 | 89.4196 | 9.7907 | 17152.4188 | 3.1895 | 102.0642
4-1 | QC | 4 | 7.3518 | 846.3653 | 0.20139 | 103.112 | 7.6707 | 1006.0074 | 0.17356 | 88.8642 | 8.1845 | 2057.6281 | 0.18435 | 94.3898 | 9.0581 | 701.5742 | 0.078442 | 100.3989 | 9.231 | 1276.8627 | 0.080114 | 102.5396 | 9.5141 | 2867.0548 | 0.072685 | 93.0309 | 9.7779 | 761.6621 | 0.13979 | 71.5736
4-2 | QC | 4 | 7.36 | 740.8993 | 0.16296 | 83.437 | 7.663 | 1256.8357 | 0.20128 | 103.0586 | 8.1766 | 2003.5742 | 0.1768 | 90.5225 | 9.0452 | 774.5048 | 0.084397 | 108.0217 | 9.231 | 1176.6035 | 0.072998 | 93.4311 | 9.5127 | 2487.3127 | 0.063358 | 81.0926 | 9.7824 | 871.906 | 0.1623 | 83.0996
CC 7 | CC | 7 | 7.3501 | 54260.1465 | 12.374 | 98.9921 | 7.6718 | 83040.3382 | 11.8059 | 94.4472 | 8.1819 | 146318.731 | 12.0212 | 96.1693 | 9.0565 | 51018.6116 | 4.8781 | 97.5629 | 9.2454 | 59208.8292 | 3.9856 | 79.7129 | 9.5115 | 202191.1619 | 4.6647 | 93.2936 | 9.7902 | 69348.853 | 13.238 | 105.9038
4-3 | QC | 4 | 7.3483 | 975.9603 | 0.21428 | 109.7103 | 7.663 | 1231.2033 | 0.19342 | 99.0325 | 8.1788 | 2165.6732 | 0.18687 | 95.6776 | 9.0493 | 702.0068 | 0.075943 | 97.2014 | 9.2571 | 1290.9025 | 0.079217 | 101.3911 | 9.511 | 3142.3069 | 0.077837 | 99.6249 | 9.8011 | 971.8045 | 0.17868 | 91.4877
4-4 | QC | 4 | 7.3717 | 891.109 | 0.1958 | 100.2531 | 7.6818 | 1014.8214 | 0.16294 | 83.4252 | 8.18 | 2047.4839 | 0.17693 | 90.5898 | 9.0581 | 529.8915 | 0.059388 | 76.0119 | 9.231 | 1108.4819 | 0.067419 | 86.2908 | 9.512 | 2707.1563 | 0.0681 | 87.1628 | 9.8013 | 888.9729 | 0.16403 | 83.9822
4.5-1 | QC | 10 | 7.3453 | 1731.6376 | 0.404 | 103.4302 | 7.6747 | 2322.4045 | 0.36247 | 92.7982 | 8.1796 | 3960.3586 | 0.33878 | 86.7321 | 9.0544 | 1384.3626 | 0.14274 | 91.3238 | 9.231 | 2360.7371 | 0.15753 | 100.7869 | 9.5114 | 5603.6853 | 0.14038 | 89.8154 | 9.7876 | 1840.6525 | 0.3623 | 92.7542
4.5-2 | QC | 10 | 7.3478 | 1837.7555 | 0.42156 | 107.9264 | 7.6729 | 2412.4532 | 0.36938 | 94.5677 | 8.1791 | 4213.3524 | 0.35119 | 89.9111 | 9.0577 | 1437.8613 | 0.14447 | 92.4306 | 9.2454 | 2254.5319 | 0.15282 | 97.7762 | 9.5085 | 5665.0414 | 0.14436 | 92.36 | 9.7875 | 1803.334 | 0.36119 | 92.4712
4.5-3 | QC | 10 | 7.3483 | 1864.4598 | 0.42898 | 109.8257 | 7.666 | 2414.8433 | 0.3706 | 94.88 | 8.1841 | 4070.7392 | 0.33657 | 86.1674 | 9.059 | 1315.2638 | 0.13177 | 84.3048 | 9.2454 | 2263.741 | 0.14974 | 95.8059 | 9.5089 | 5573.3884 | 0.13887 | 88.8479 | 9.7896 | 1879.0581 | 0.36803 | 94.221
CC 6 | CC | 6 | 7.3558 | 13363.1603 | 3.3041 | 105.7322 | 7.6678 | 20325.5832 | 3.1457 | 100.6639 | 8.1827 | 34545.4222 | 2.9448 | 94.2352 | 9.053 | 12174.1097 | 1.2087 | 96.6935 | 9.2454 | 14618.4535 | 1.0447 | 83.5766 | 9.5115 | 47464.2301 | 1.1855 | 94.8394 | 9.7933 | 15853.2709 | 3.2782 | 104.9016
4.5-4 | QC | 10 | 7.3559 | 1704.3401 | 0.41129 | 105.2975 | 7.6665 | 2392.8222 | 0.38511 | 98.5957 | 8.1838 | 4290.3222 | 0.36732 | 94.039 | 9.0495 | 1620.3456 | 0.16601 | 106.2117 | 9.2454 | 1832.951 | 0.12686 | 81.1674 | 9.5114 | 5844.7841 | 0.15384 | 98.4238 | 9.792 | 1912.1705 | 0.39753 | 101.7752
5-1 | QC | 5 | 7.3629 | 3483.5854 | 0.80315 | 102.803 | 7.6747 | 4888.6138 | 0.72592 | 92.9174 | 8.182 | 8416.5478 | 0.6669 | 85.3633 | 9.0581 | 3119.0657 | 0.29168 | 93.3388 | 9.2454 | 3549.349 | 0.22907 | 73.3018 | 9.511 | 10917.8923 | 0.25734 | 82.3494 | 9.7917 | 3666.1952 | 0.69852 | 89.4105
5-2 | QC | 5 | 7.3483 | 3528.9531 | 0.82656 | 105.8003 | 7.6711 | 4783.9414 | 0.72146 | 92.3473 | 8.1837 | 8759.6977 | 0.68167 | 87.2534 | 9.0506 | 2857.9514 | 0.26346 | 84.306 | 9.2454 | 3402.9141 | 0.22108 | 70.7455 | 9.5079 | 11230.8839 | 0.26677 | 85.3673 | 9.7989 | 3630.7866 | 0.6974 | 89.2671
5-3 | QC | 5 | 7.3483 | 3620.3924 | 0.8161 | 104.4604 | 7.6685 | 5179.9684 | 0.7513 | 96.1661 | 8.1811 | 8958.6722 | 0.68864 | 88.1465 | 9.0548 | 2976.2621 | 0.27079 | 86.6532 | 9.2454 | 3714.3129 | 0.24474 | 78.3175 | 9.5106 | 11116.2181 | 0.2667 | 85.3427 | 9.7907 | 3781.9082 | 0.73445 | 94.0102
5-4 | QC | 5 | 7.3563 | 3363.4417 | 0.76903 | 98.4358 | 7.6654 | 4888.7755 | 0.72056 | 92.2316 | 8.1823 | 8690.6323 | 0.68753 | 88.0041 | 9.0547 | 2860.7999 | 0.26797 | 85.751 | 9.2454 | 3859.8037 | 0.25451 | 81.4448 | 9.5102 | 11547.6459 | 0.27673 | 88.5521 | 9.7923 | 3741.7671 | 0.72595 | 92.9215
CC 7 | CC | 7 | 7.3553 | 59157.1906 | 13.5861 | 108.6885 | 7.6655 | 90148.0217 | 12.9107 | 103.2856 | 8.1827 | 159108.1305 | 12.1618 | 97.2946 | 9.0529 | 55378.1407 | 4.9267 | 98.5346 | 9.2454 | 56839.3671 | 3.8095 | 76.1891 | 9.5098 | 213093.1598 | 4.8819 | 97.638 | 9.7935 | 65011.7736 | 12.3792 | 99.0338
8-1 | QC | 8 | 7.3529 | 251901.5777 | 47.6407 | 95.2813 | 7.6701 | 397007.8827 | 46.9417 | 93.8834 | 8.1831 | 708706.2177 | 47.0749 | 94.1498 | 9.0556 | 241913.2685 | 18.8792 | 94.3958 | 9.2454 | 257122.9281 | 17.9585 | 89.7925 | 9.512 | 1051356.8449 | 23.5898 | 117.949 | 9.7935 | 281752.9997 | 53.155 | 106.31
8-2 | QC | 8 | 7.355 | 252037.7362 | 49.3511 | 98.7021 | 7.6734 | 388440.5541 | 47.6909 | 95.3819 | 8.1852 | 693586.2083 | 46.9422 | 93.8844 | 9.0555 | 238768.0777 | 18.9709 | 94.8547 | 9.2454 | 269065.2421 | 16.9886 | 84.9429 | 9.5137 | 1000209.1511 | 20.6068 | 103.034 | 9.7929 | 298861.7842 | 51.0275 | 102.0549
8-3 | QC | 8 | 7.3512 | 253290.3374 | 49.7772 | 99.5543 | 7.6701 | 384648.3988 | 47.4854 | 94.9707 | 8.1826 | 675936.0135 | 46.464 | 92.928 | 9.0543 | 235115.439 | 18.9533 | 94.7667 | 9.2454 | 269771.67 | 17.0399 | 85.1996 | 9.5117 | 970441.61 | 20.081 | 100.405 | 9.793 | 293342.905 | 50.2358 | 100.4716
8-4 | QC | 8 | 7.3525 | 247985.1979 | 49.2197 | 98.4394 | 7.6686 | 383178.4837 | 47.6706 | 95.3413 | 8.1802 | 671527.6882 | 46.2564 | 92.5129 | 9.0544 | 231315.1945 | 18.7013 | 93.5066 | 9.2454 | 264430.19 | 16.6695 | 83.3476 | 9.5113 | 952320.7395 | 19.6836 | 98.4182 | 9.7915 | 293128.8406 | 50.0318 | 100.0636
5.5-1 | QC | 11 | 7.3388 | 10140.0391 | 2.2616 | 144.7442 | 7.6666 | 12210.8703 | 1.7167 | 109.8661 | 8.1828 | 17958.549 | 1.487 | 95.1661 | 9.0501 | 6495.7504 | 0.62868 | 100.5884 | 9.2454 | 8537.2572 | 0.63341 | 101.3456 | 9.5129 | 25536.8593 | 0.66859 | 106.974 | 9.7935 | 8187.365 | 1.7663 | 113.0402
5.5-2 | QC | 11 | 7.3477 | 9750.5189 | 2.0822 | 133.26 | 7.6682 | 11469.5108 | 1.546 | 98.9448 | 8.1838 | 18035.5551 | 1.4077 | 90.0911 | 9.0542 | 6551.2151 | 0.59791 | 95.6651 | 9.2454 | 8154.1263 | 0.59386 | 95.0172 | 9.5098 | 24854.8189 | 0.63955 | 102.3275 | 9.7936 | 8125.1486 | 1.7219 | 110.2004
5.5-3 | QC | 11 | 7.3451 | 10243.276 | 2.1975 | 140.6416 | 7.6679 | 11905.3741 | 1.6113 | 103.1209 | 8.1796 | 17533.1659 | 1.4028 | 89.777 | 9.0537 | 6146.0645 | 0.57538 | 92.0602 | 9.2454 | 7832.7382 | 0.53209 | 85.134 | 9.5108 | 24954.4968 | 0.60023 | 96.0366 | 9.7911 | 8027.6842 | 1.5889 | 101.6866
5.5-4 | QC | 11 | 7.3396 | 9982.3228 | 2.0741 | 132.7419 | 7.6681 | 11493.815 | 1.508 | 96.5098 | 8.1818 | 17616.6367 | 1.3632 | 87.2438 | 9.054 | 6374.5611 | 0.577 | 92.32 | 9.231 | 7687.6205 | 0.5049 | 80.7841 | 9.5112 | 25313.3559 | 0.58926 | 94.2812 | 9.7933 | 7521.2605 | 1.4397 | 92.1405
CC 6 | CC | 6 | 7.3544 | 13393.481 | 3.3307 | 106.5828 | 7.6677 | 19633.6535 | 3.0576 | 97.8436 | 8.1787 | 34117.5697 | 2.9625 | 94.7994 | 9.0523 | 10955.9871 | 1.1094 | 88.7511 | 9.2454 | 14057.5983 | 1.0456 | 83.6454 | 9.5075 | 48808.2265 | 1.2677 | 101.418 | 9.7904 | 15944.072 | 3.4303 | 109.7682
Lab 2 LCMRL Dataset